27482 | 1220648-2024-10369 | 19174114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174114&pc=OZD | 04/19/2024 | 06/21/2024 |  | 04/24/2024 |  |  | 04/19/2024 | 04/24/2024 |  |  | 09/30/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444572 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 38-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding. the bleeding was treated with a sandbag and fem-stop. heparin was withheld and the patient was transfused with two units of packed red blood cells and platelets. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27483 | 1220648-2024-10370 | 19174251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174251&pc=OZD | 04/19/2024 | 06/21/2024 |  | 04/24/2024 |  |  | 04/19/2024 | 04/24/2024 |  |  | 04/30/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353500 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 76-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding. the bleeding was treated with pressure and perclose sutures. due to continued oozing, the impella was weaned and explanted. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
27484 | 1220648-2024-10371 | 19174315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174315&pc=OZD | 04/13/2022 | 04/24/2024 |  | 04/24/2024 |  |  | 04/13/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187159 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the doctor could not get distal pulses. the impella was removed, and an intra-aortic balloon pump was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27485 | 1220648-2024-10372 | 19174395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174395&pc=OZD | 04/19/2024 | 06/03/2024 |  | 04/24/2024 |  |  | 04/19/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451902 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. however, the impella cp device was removed due to persistent slow bleeding around the insertion site. there was no change in hemodynamics after removal of the impella cp device, and the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6, g3, and h10 have been revised from the initial submission.
27486 | 1220648-2024-10373 | 19174533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174533&pc=OZD | 10/23/2022 | 04/15/2025 |  | 04/24/2024 |  |  | 10/23/2022 | 04/24/2024 |  |  | 04/30/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216236 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, it was found that the placement signal was flat. the pressure bag clamp was open with pressure over 250mmhg, and the physician aspirated the placement line with no success. however, the patient was deemed to have sufficient cardiac contractility, so the device was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27487 | 1220648-2024-10374 | 19174609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174609&pc=OZD | 04/19/2024 | 06/10/2024 |  | 04/24/2024 |  |  | 04/19/2024 | 04/24/2024 |  |  | 09/30/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, low diastolic flow was noted, and the decision was made to wean the impella pump which was eventually successfully removed. there was no report of adverse effects to the patient. | [5 empty] | The investigation into the access site bleeding ¿ minor and the low pump flow issues has been completed since the original report was filed. the device was returned for investigation. visual inspection found no issues on the pump. the root cause of access site bleeding was use related. the root cause of the low flow was most likely a patient condition as no issue was found on the pump. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27488 | 1220648-2024-10375 | 19174833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174833&pc=PYX | 04/19/2024 | 07/15/2024 |  | 04/24/2024 |  |  | 04/19/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025462727 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 48-year-old male with cardiomyopathy was implanted with an impella rp flex device for mechanical circulatory support. the patient needed the support due to needing a left ventricular assist device placement and a tricuspid valve ring repair. the patient developed oozing at the access site. the medical staff set a drive line and chest tube drainage. the medical staff then administered two units of red blood cells, one unit of fresh frozen plasma, and one unit of platelets. the medical staff switched the purge from heparin to bicarb due to the chest tube output. the patient had a washout and a moderate amount of clots were removed from the patient's chest. the patient remains on support. | [5 empty] | The investigation into access site bleeding, optical signal issue, thrombus and vascular damage-peripheral vessel has been completed. the product and data logs were returned for evaluation. impella rp flex pump was returned in cut open state. visually inspected the pump and found biomaterial ingested over inflow cage of pump. borescope view was obtained and observed the entire cannula to be blocked with significant biomaterial. sensor resistance tests could not be performed since pump returned in cut state. no other abnormalities were found. data logs were reviewed, and several dp sensor placement signal not reliable alarms observed. the placement signal drift trend was observed. the pump flow calculations were disabled during this time. no optical signal issue found. the cause of the access site bleeding issue was most likely patient condition related due to bleeding at multiple sites. as the necessary information was not provided, the root cause of the thrombus was not determined. the cause of the placement signal issue was due to sensor drift per return product investigation and log analysis. the cause of the vascular damage issue was not determined due to insufficient clinical information. additional instructions for use impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)". | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
27489 | 1220648-2024-10376 | 19174957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19174957&pc=OZD | 04/19/2024 | 06/04/2024 |  | 04/24/2024 |  |  | 04/19/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463953 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 54-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that the team had to troubleshoot the pump during the support for purge flow and purge pressure issues. the team replaced the purge cassette but, that alone was not enough, and the team made the choice to add tissue plasminogen activator for the purge. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported high purge pressure has been completed. no product was returned. data logs show that 263 hours into the case there is a 7 minute rise in purge pressure triggering hpp and purge system blocked alarms. this remains elevated for just under 12 hours and then resolves over a span of about 2 hours. the clinical states that tpa was added to the purge which then resolved the issue. the root cause of the high purge pressure was most likely biomaterial. the instructions for use for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 impact code 4644 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27490 | 1220648-2024-10377 | 19175133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19175133&pc=OZD | 04/20/2024 | 06/11/2024 |  | 04/24/2024 |  |  | 04/20/2024 | 04/24/2024 |  |  | 06/30/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385783 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | White | Malfunction | No | Yes | Migration or Expulsion of Device | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 64-year-old female with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during transfer to a tertiary center, the transfer center discovered that the impella device was not appropriately sutured. the impella device migrated out of position. the impella device was explanted, and the patient remains on extra-corporeal membrane oxygenation support. | [5 empty] | The investigation for the reported positioning issue has been completed. data logs showed there were no issues for the pump during support. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. added- d6a/d6b | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27491 | 1220648-2024-10378 | 19175194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19175194&pc=OZD | 04/21/2024 | 06/12/2024 |  | 04/24/2024 |  |  | 04/21/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATIC IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATIC IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815943 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/21/2024 |  |  |  |  | The user facility reported a 56-year-old male was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there was an automatic impella controller (aic) console failure alarm. the patient's pressure dropped. the console was restarted, but the failure occurred two more times. the aic was replaced. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) failure alarm has been completed. the console was returned for evaluation and the failure mode was unable to be reproduced. no hardware issue was found with the console. upon force shutdown, the console rebooted automatically, as happened at the customer site. data logs were reviewed finding several unexpected shutdowns and reboots. the root cause of intermittent unexpected controller shutdown was not determined due to the inability to reproduce. d2-updated common device name in accordance with updated procedures, sections d6 and g3 have been revised from the initial submission.
27492 | 1220648-2024-10379 | 19175374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19175374&pc=OZD | 04/22/2024 | 07/25/2024 |  | 04/24/2024 |  |  | 04/22/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATIC IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATIC IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145435 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was an automatic impella controller (aic) console failure alarm. the aic was replaced. no patient harm was reported. | [5 empty] | The investigation for the console (aic) yellow alarm has been completed. data logs were reviewed which confirmed the reported failure mode given there were multiple controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that signals received from optical bench ( snr, contrast, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. the console's front optical connector was inspected before cleaning and after cleaning. there was a persistent imperfection within the core of the front optical connector. the consoles fringe pattern revealed no observable abnormality. the root cause of the controller error is due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. d2-corrected common device name f6/f8 should have been left blank in the initial report. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
27493 | 1220648-2024-10380 | 19175767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19175767&pc=OZD | 01/29/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 01/29/2022 | 04/24/2024 |  |  | 05/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was repositioned to 3.6 and the groin was redressed with minimal ooze. additionally, the patient received 2 units of packed red blood cells due to a steep drop in hemoglobin. later, the patient experienced moderate hematoma in the right groin. the patient remained on support and was sent to unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27494 | 1220648-2024-10381 | 19177095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19177095&pc=OZD | 10/31/2023 | 09/17/2025 |  | 04/24/2024 |  |  | 03/26/2024 | 04/24/2024 |  |  | 07/31/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374296 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/26/2024 |  |  |  |  | The complaint reported a 35-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. abiomed¿s long-term outcome and quality indicator impella registry notified us of an allegation of pump stop with the impella device. the impella suddenly stopped working. patient was doing well and went for emergency impella replacement. additionally, while on support, the patient had a gastrointestinal (gi) bleed. heparin was withheld and purge was changed to sodium bicarb. the patient also had runs of ventricular tachycardia (vt). echo was ordered to confirm positioning and the staff monitored. | The customer reported that the device lost placement signal. no further information provided. |  |  |  |  | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the pump stop issue was an open electrical line, based data log review. the cause of open electric line is unknown. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the issue vascular damage was not determined since product was not returned and limited clinical information was provided. as the necessary information and returned device were not provided, the root cause of the vt/vf was not determined. potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. data logs were provided for the case run after 10/24/2023. they show that the placement signal was at 3000, with a reported hard failure trend on all 5s optical parameters. at the end of case run, a sudden pump stop was reported with pump stop alarm 115, which indicates an open electric line. the cause of the pump stop issue was an open electrical line, based data log review. the cause of open electric line is unknown. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. g1 updated with correct mdr reporting contact email. f6 and f8 should have been left blank. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-10381.
27495 | 1220648-2024-10260 | 19162348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162348&pc=OZD | 04/16/2024 | 05/17/2024 |  | 04/23/2024 |  |  | 04/16/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 63-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, battery failure alarms occurred on the automated impella controller (aic). the end user reported the aic was always plugged into the red outlet and smelled hot. the aic was swapped, and no further issues or alarms occurred. the patient remains stable. | [5 empty] | The investigation for the reported console battery alarm issue has been completed. the console was returned for evaluation . battery failure issues were unable to be reproduced. during examination of the inside of the console, it was discovered that there was a blown capacitor @ c175 on the pbm board. this blown capacitor is likely what the nurse smelled when they said the console smelled ¿hot¿. the logs were reviewed which confirmed the reported battery failure alarms. there were numerous battery alarms including: battery failure (transport connection blown/missing) , battery level low (50%), and battery failure (low voltage). the root cause of this issue was a blown capacitor found @ c175 on the pbm board but the cause of the blown capacitor is unknown. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27496 | 1220648-2024-10261 | 19162518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162518&pc=OZD | 04/17/2022 | 05/21/2025 |  | 04/23/2024 |  |  | 04/17/2022 | 04/23/2024 |  |  | 09/30/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135330 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced a suction event that was resolved by giving one unit of packed red blood cells. it was also reported while the pump was being repositioned with transesophageal echocardiography, the pump became caught in the cords of the mitral valve, and the impella device came across valve. the medical team could not advance into the left ventricle, so the impella cp was removed and exchanged for a new impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27497 | 1220648-2024-10262 | 19162573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162573&pc=OZD | 04/18/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/18/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183191 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ectopy while the impella was running with higher flows. the ectopy resolved with the p-level lowered to p-4. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27498 | 1220648-2024-10263 | 19162631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162631&pc=OZD | 05/02/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 05/02/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184800 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was significant bleeding from the access site. heparin was withheld, manual pressure was applied to achieve hemostasis, and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27499 | 1220648-2024-10264 | 19162656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162656&pc=OZD | 05/02/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 05/02/2022 | 04/23/2024 |  |  | 08/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134848 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Swelling/ Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on impella support post-procedure, the patient's groin became swollen. the patient also had difficulty keeping their right leg immobile, contributing to the insertion site swelling. two percloses, manual pressure for five minutes, and a fem stop were applied to achieve hemostasis. the device was explanted shortly after this event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27500 | 1220648-2024-10265 | 19162679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162679&pc=OZD | 04/16/2024 | 05/16/2024 |  | 04/23/2024 |  |  | 04/16/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported that the automated impella controller (aic) purge disc appeared loose inside its slot causing a "purge disc not detected" alarm. the aic was exchanged. there was no report of harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the reported purge disc not detected issue could not be reproduced. there was no issue found during the case. the device functioned/operated to specification. in accordance with updated procedures, section d1 has been revised from the initial submission. | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27501 | 1220648-2024-10267 | 19162711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162711&pc=OZD | 04/20/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/20/2022 | 04/23/2024 |  |  | 01/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187618 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing at the groin site. the patient received one unit of blood to help resolve the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27502 | 1220648-2024-10266 | 19162715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162715&pc=OZD | 09/28/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 09/28/2022 | 04/23/2024 |  |  | 06/30/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 90-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported due to complication suffered during the tandem support pci (percutaneous coronary intervention), the left coronary system was never revascularized. while on support it was noted that the previously deployed perclose sutures were tightened but the physician was not pleased with the result so a third perclose was added and a manual compression was held for 10 minutes. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27503 | 1220648-2024-10268 | 19162787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162787&pc=OZD | 03/22/2024 | 04/23/2024 |  | 04/23/2024 |  |  | 03/22/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1360581 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) did not start properly and displayed "system check failed - replace console immediately" error messages. the issue was unable to be resolved with a hard reset. therefore, the console was exchanged. there was no reported patient harm. | [5 empty] | The investigation has been completed. the product was returned for investigation. the console was tested and the failure mode was reproduced so when the console was turned on, it immediately output a system self check fail. the super cap on the power battery manager (pbm) circuit board was confirmed to have corroded the pbm communication trace next to it. during data analysis, the automated impella controller (aic) logs showed pbm1 communication errors. the cause of the aic issue was contamination of a communication line on the pbm.
27504 | 1220648-2024-10269 | 19162813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162813&pc=OZD | 05/03/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 05/03/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173114 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. both perclose's failed during closure. the team was able to get control of the bleeding and manual pressure was held. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27505 | 19162841 | 19162841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162841&pc=OZD | 03/19/2024 | 04/11/2024 | 04/17/2024 | 04/23/2024 | 04/17/2024 | 04/23/2024 |  | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | No answer provided |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 3 MO |  | Male |  |  |  | Injury | Yes | No | [7 empty] | Low Blood Pressure/ Hypotension | Perforation | [11 empty] | Hospitalization | Required Intervention | [23 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | This patient is a (b)(6) year-old male with past medical history significant for severe aortic stenosis undergoing transcatheter aortic valve replacement (tavr) evaluation and presented to the hospital for elective left heart catheterization (lhc). impella placed prior to attempted percutaneous coronary intervention (pci) to distal left main (lm). patient became hypotensive with concern for perforation of left ventricle (lv) and therefore patient had emergent pericardiocentesis. cardiothoracic surgery (cts) then performed a pericardial window. patient was stabilized and transferred to cardiovascular intensive care unit (cvicu).
27506 | 1220648-2024-10270 | 19162851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162851&pc=OZD | 04/20/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 04/20/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172836 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported a 62-year-old patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma at the impella access site. femstop was placed on the left femoral artery and heparin was withheld. it was also reported that the patient experienced a loss of doppler pulses in the left leg. an external femoral to femoral distal perfusion catheter was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27507 | 1220648-2024-10271 | 19162896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162896&pc=OZD | 05/08/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 05/08/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174265 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The complainant reported the 69-year-old female patient was implanted with an impella cp support due to a st-segment elevation myocardial infarction and complicated cardiogenic shock. it was reported that the impella site was bleeding overnight,. a suture was placed as intervention. the impella was redressed and re-angled. it was reported that the oozing was controlled. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27508 | 1220648-2024-10272 | 19162920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162920&pc=OZD | 05/09/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 05/09/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182520 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported the 52-year-old male patient was placed onto impella cp support due to st-elevated myocardial infarction. while on support, the patient experienced access site bleeding that was resolved by stopping heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27509 | 1220648-2024-10273 | 19162950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19162950&pc=OZD | 04/22/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/22/2022 | 04/23/2024 |  |  | 02/29/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195558 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/22/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient¿s morning labs were partially hemolyzed, as their urine was red and <30ml/hr. acute hemolysis was later noted. information regarding resolution was not provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27510 | 1220648-2024-10274 | 19163015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163015&pc=PYX | 03/30/2024 | 05/16/2024 |  | 04/23/2024 |  |  | 04/01/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025461549 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Death | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 34-year-old female undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that an implanted impella rp flex pump experienced a prolonged period of low flows during patient support. it was noted that the pump was in good positioning during this period of time. there was no noted harm to the patient as a result of the low flows. | [5 empty] | The investigation for the low pump flow has been completed. the pump and data logs were not returned for investigation. the root cause of the low pump flow was not determined since product and data logs were not returned. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27511 | 1220648-2024-10275 | 19163056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163056&pc=OZD | 09/29/2022 | 03/28/2025 |  | 04/23/2024 |  |  | 09/29/2022 | 04/23/2024 |  |  | 06/30/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported an 81-year-old female on impella cp for mechanical circulatory support. it was reported the patient was placed onto impella cp due to cardiogenic shock accompanied by severe multiple vessel disease and coronary clot. while on support, the patient experienced hemolysis and limb ischemia with no specified resolutions. additionally, the patient experienced access site bleeding, which was resolved by redressing, angle matching, and knee immobilization. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27512 | 1220648-2024-10278 | 19163166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163166&pc=OZD | 09/29/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 09/29/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172997 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported an 84-year-old patient needing mechanical circulatory support. it was reported that there was access site bleeding. the physician wanted to leave the peel-away sheath in but during suturing, the sheath nicked which caused light bleeding. this was treated by removing the peel-away sheath, angle matching and placing aggressive forward-tension sutures. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27513 | 1220648-2024-10277 | 19163172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163172&pc=OZD | 04/06/2024 | 05/29/2024 |  | 04/23/2024 |  |  | 04/06/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467789 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 71-year-old male with acute myocardial infarction and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient exhibited access site bleeding. the patient was febrile and the access site was warm and red. the medical staff withheld anticoagulants and administered two units of platelets and antibiotics due to a possible infection. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation for the reported access site bleed and sepsis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and sepsis could not be determined. the instructions for use referenced in the initial report is from ifu for the impella 5.5 with smartassist for use during cardiogenic shock , section: potential adverse events (united states), which now includes statement "(including ventricular perforation).¿ added- d6a/d6b
27514 | 1220648-2024-10279 | 19163222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163222&pc=OZD | 04/20/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/20/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177143 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/23/2022 |  |  |  |  | The user facility reported an 84-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella for left ventricular unloading. in preparing the patient for explant, a clot was observed on the tip of the peel away sheath. a fluid bolus of 2000 cc¿s was provided to the patient along with intravenous (iv) calcium. it is stated that the patient was placed on hypothermic protocol and continued to experience arrhythmias. a cardioversion was performed but was stated to be unsuccessful as the patient remains in ventricular fibrillation. a cut down of the femoral artery was performed. the medical professional successfully removed the impella device, clots, and used two (2) preclose sutures and surgically closed the femoral artery access site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ this report is being filed as part of a retrospective review of historical records.
27515 | 1220648-2024-10280 | 19163243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163243&pc=OZD | 02/04/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 02/04/2022 | 04/23/2024 |  |  | 01/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039014 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the staff noted suspected hemolysis and hematuria and escalated the patient from the impella cp to the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27516 | 1220648-2024-10281 | 19163259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163259&pc=OZD | 05/10/2022 | 05/22/2025 |  | 04/23/2024 |  |  | 05/10/2022 | 04/23/2024 |  |  | 03/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191935 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported that the 63-year-old patient was on the impella cp and issues were encountered during positioning as the pump had slipped into the aorta. weaning was already initiated, at time of report. impella was then explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27517 | 1220648-2024-10282 | 19163390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163390&pc=OZD | 04/10/2024 | 06/06/2024 |  | 04/23/2024 |  |  | 04/11/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025473457 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a patient of unknown age and gender was implanted unsuccessfully with an impella cp device for mechanical circulatory support. the impella device had been opened but was unused because the medical staff did not confirm that the aortic valve was a mechanical valve. the impella device was not used on the patient, but the introducer was used. no patient harm was reported. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. the impella was not used with the patient however the introducer that is a part of the impella kit was used with the patient as the aortic valve was a mechanical valve as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27518 | 1220648-2024-10283 | 19163483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163483&pc=OZD | 04/11/2024 | 05/20/2024 |  | 04/23/2024 |  |  | 04/11/2024 | 04/23/2024 |  |  | 04/30/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218533 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported an 88-year-old male with acute myocardial infarction and cardiogenic shock was implanted unsuccessfully with an impella cp device for mechanical circulatory support. the patient was admitted for chest pain and shortness of breath and also has a history of cerebral vascular accident, diabetes, and sepsis. an introducer malfunction was noticed when the impella device was attempted to be implanted into the femoral artery. the clear peel away lock got stuck on the catheter but was eventually cut off successfully. the impella device was then implanted successfully. support lasted for 90 hours, and the impella device was weaned and explanted after a successful percutaneous coronary intervention. | [5 empty] | The investigation for the introducer damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues introducer damage could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27519 | 1220648-2024-10284 | 19163603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163603&pc=OZD | 04/13/2024 | 06/26/2024 |  | 04/23/2024 |  |  | 04/13/2024 | 04/23/2024 |  |  | 11/30/2025 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding. dressing site visualized and redressed, insertion angle maintained and adequate. one unit of packed red blood cells was provided. additionally, while on support, the patient had a gastrointestinal bleed and heparin was withheld. | It was further reported that that sometime after support was completed, care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the access site bleed and gastrointestinal bleed/vascular issue has been completed. the cp pump was not returned. patient was bleeding at impella access site and hemostasis achieved after dressing change/insertion angle adjustment. the cause of the access site bleeding issue was most likely use related per clinical review. the cause of the gi bleed/vascular damage issue was not determined since product was not returned and unknown relation to impella based on clinical information. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states) added- b5 mso review, h6 clinical code 4476, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27520 | 1220648-2024-10285 | 19163658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163658&pc=OZD | 02/05/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 02/05/2022 | 04/23/2024 |  |  | 11/30/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161857 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced no distal pulse in the left dorsalis pedis pulse, posterior tibial pulse. the foot was cool but not mottled. an echo was performed which showed impella at 3cm and pointing posterior toward the mitral valve. reposition attempted, but unable to get the device to sit in the mid ventricular space. patient received 2 units of blood for suspected gi bleed. the patient was weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27521 | 1220648-2024-10286 | 19163714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163714&pc=OZD | 10/03/2022 | 05/12/2025 |  | 04/23/2024 |  |  | 10/03/2022 | 04/23/2024 |  |  | 05/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2023227302 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The complainant reported a 72 year old male patient on impella cp for mechanical circulatory support. while on support the patient had slight bleeding at the impella groin side. systemic heparin was reduced to resolve the bleeding. the impella was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27522 | 1220648-2024-10287 | 19163753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163753&pc=OZD | 04/15/2024 | 06/20/2024 |  | 04/23/2024 |  |  | 04/15/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447110 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Improper Flow or Infusion | Positioning Problem |  |  |  | Hematoma | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The user facility reported an 80-year-old male patient in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma and heparin was paused. manual pressure was held as treatment. additionally, the patient had ventricular fibrillation (vf) on support. it was uncertain if the vf was caused by the pump's position. furthermore, the pump had low pump flows and positioning issues. there were persistent suction alarms and p-3. an echocardiogram confirmed the pump was shallow. the decision was made to explant the impella. the patient was reported to be stable. | [5 empty] | The investigation has been completed. the impella and data logs were not returned for review. the suction alarms was most likely related to positioning issue. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. based on clinical description, the root cause of positioning issue was most likely use related. as the necessary information was not provided, the root cause of the ventricular fibrillation was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 codes 4639, 3009 and 2954 were reported incorrectly on the previously submitted report in accordance with updated procedures. the correct codes have been added to medical device problem codes. .h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
27523 | 1220648-2024-10288 | 19163912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163912&pc=OZD | 02/05/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 02/05/2022 | 04/23/2024 |  |  | 11/30/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Ischemia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. initially the medical staff did not want to place the impella due to the inability to see if left ventricular thrombus was present. the decision was made to place an intra-aortic balloon pump (iabp) instead. the impella was exchanged for the iabp via the right groin without issue. after repositioning, it was noted that the access site was bleeding. pressure dressing added to the impella site did not alleviate the issue as the initial dressing was observed to be saturated with blood once in the operating room. also, the patient¿s right foot was cold, and the medical staff was unable to find a pulse with doppler. the patient¿s labs were ¿grossly¿ hemolyzing. the impella was measuring 4.3 centimeters from the aortic valve according to echocardiogram. the device was repositioned at bedside and flows increased. the patient¿s dressing was changed, and the right groin access site stopped bleeding. the patient was successfully weaned, and the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. correction to initial mfr 1220648-2024-10288: section h6, health effect - clinical code, the code for 2135 perforation of vessels should have been omitted as there was no allegation of perforated vessels.
27524 | 1220648-2024-10289 | 19163965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19163965&pc=OZD | 04/15/2024 | 06/07/2024 |  | 04/23/2024 |  |  | 04/15/2024 | 04/23/2024 |  |  | 10/31/2025 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418121 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The user facility reported a 63-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the team mistakenly inserted the 23 french sheath and not the 14 french sheath meant for impella cp placement. the limb became ischemic due to the placement. the team placed a distal perfusion catheter for flow to the limb. the patient remains on support. | [5 empty] | The investigation into the access site bleeding ¿ minor and the limb ischemia - reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding is a use issue, as the physician used a 23fr sheath in place of the 14fr. there is no evidence to indicate the packaging contained incorrect components. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
27525 | 1220648-2024-10290 | 19164195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164195&pc=OZD | 04/15/2022 | 05/30/2025 |  | 04/23/2024 |  |  | 04/15/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during support there was loss of pulsatility during multiple inflations. after the peel away sheath was removed and the repositioning sheath advanced into the arteriotomy, bleeding was noted at the impella site. the impella was then successfully weaned and explanted, without issues, 50 minutes of manual pressure was held, hemostasis achieved, and the patient was transferred to the icu in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. correction to initial mfr 1220648-2024-10290: section h6, health effect - clinical code for limb ischemia should be omitted as the patient did not experience limb ischemia during impella support.
27526 | 1220648-2024-10291 | 19164277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164277&pc=OZD | 02/28/2024 | 06/20/2024 |  | 04/23/2024 |  |  | 04/16/2024 | 04/23/2024 |  |  | 03/31/2025 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329130 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | A medwatch report filed by the family of a 72-year-old female patient in acute myocardial infarction/cardiogenic shock was reported to have been implanted with an impella cp device for mechanical circulatory support. the family and physician had conversations and the patient was made do not resuscitate/ do not intubate (dnr/dni) and was transferred to the tertiary center for further care. after 20 hours of support, the patient expired with care withdrawn on the following morning. the family months later filed the medwatch to the food and drug admistration (fda), which the fda shared with abiomed. the medwatch mentioned limb ischemia as a patient harm caused by the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 death and date of death was inadvertently left off the previously submitted report and was added. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has been completed. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27527 | 1220648-2024-10292 | 19164319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164319&pc=OZD | 04/17/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/17/2022 | 04/23/2024 |  |  | 10/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced limb ischemia due to blood flow occlusion in the superficial femoral artery, there was no distal pulse audible on the doppler. systemic heparin purge was added. the impella was pulled back to 4.18 cm from 5.13 from inlet to the aortic valve after hemolysis was noted. it was reported the patient also had an elevated white blood count. the impella was explanted due to sepsis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ hemolysis ifu quotes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions:including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
27528 | 1220648-2024-10293 | 19164433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164433&pc=OZD | 04/17/2024 | 06/04/2024 |  | 04/23/2024 |  |  | 04/17/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was an observation made of a rapidly expanding hematoma. the hematoma prompted explant of the impella device after only 15 hours of support, and placement of a covered stent in the subclavian artery. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. per data analysis, issue was classified as vascular damage. the root cause of vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27529 | 1220648-2024-10294 | 19164519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164519&pc=OZD | 04/17/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/17/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162332 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella had a positioning issue that was solved, however, it is unclear how it was resolved. also, a bronchoscopy revealed a severe thrombus in the left bronchus. the physician added extra-corporeal membrane oxygenation due to the unclear pulmonary distress. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27530 | 1220648-2024-10295 | 19164623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164623&pc=OZD | 04/17/2024 | 06/04/2024 |  | 04/23/2024 |  |  | 04/17/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 57-year-old female with cardiomyopathy was implanted unsuccessfully with an impella cp device for mechanical circulatory support. the medical staff found what appeared to be bubbles of air in the left ventricle and aorta and decided to remove the impella device. the patient was treated chemically and remains stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the embolism issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the embolism issue was not determined since product was not returned, insufficient clinical information provided, and data logs were not provided. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27531 | 1220648-2024-10296 | 19164765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164765&pc=OZD | 04/17/2024 | 06/04/2024 |  | 04/23/2024 |  |  | 04/17/2024 | 04/23/2024 |  |  | 09/30/2025 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445680 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that when the medical doctor went to move the repo sheath there was difficulty. the physician used two heave hemostats to break it off. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the difficult removal has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues difficult removal could not be determined. added- d6a/d6b.
27532 | 1220648-2024-10297 | 19164766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164766&pc=OZD | 04/17/2022 | 05/23/2025 |  | 04/23/2024 |  |  | 04/17/2022 | 04/23/2024 |  |  | 11/30/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163645 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, groin oozing was noted and was resolved with sutures, pressure, and redressing. additionally, the pump experienced intermittent suction alarms that were resolved with giving volume. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27533 | 1220648-2024-10298 | 19164841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164841&pc=OZD | 10/26/2022 | 06/02/2025 |  | 04/23/2024 |  |  | 10/26/2022 | 04/23/2024 |  |  | 07/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240462 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The complainant reported the 83 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced access site bleeding. the impella cp was switched for another impella cp. cvc and flotrac applied to control bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27534 | 1220648-2024-10299 | 19164871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164871&pc=OZD | 04/21/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/21/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170781 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 71-year-old patient needing mechanical circulatory support. it was reported that while on impella support the physician noted profuse bleeding around the repo sheath, and a near instant hematoma required the pump be explanted as hemostasis could not be achieved. contralateral access was obtained and an 8mm x 4cm balloon was used for tamponade. two perclose devices were deployed, and successful hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27535 | 1220648-2024-10300 | 19164904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164904&pc=PYX | 04/17/2024 | 10/06/2025 |  | 04/23/2024 |  |  | 04/18/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479524 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 44-year-old male postoperative cardiothoracic surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that multiple complications were encountered during impella rp support. following implant, it was noted that oozing was seen at the access site. a mattress suture was subsequently placed to achieve hemostasis. roughly four days into support, the pump began to experience decreased flows. it was believed that a clot was present in the cannula and causing a decrease in flow rate. as a result of the noted issue, the decision was made to remove the pump and place the patient on protek duo. there was no patient harm. | The investigation uncovered placement signal issues during support. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the low pump flow and access site bleed has been completed. the impella device was returned for evaluation. upon functional testing, the pump passed electrical testing. putting the pump in a bucket of water for extended period of time showed the sensor to drift +10mmhg. pressurized flow loop reproduced the dp sensor drift. the root cause of the low pump flow was likely due to placement signal issues from dp sensor drift. according to clinical details, the patient was reported to be oozing from the access site; adding a second mattress suture resolved the bleeding. the root cause of the access site bleeding major issue was not determined due to a lack of clinical details. the instructions for use referenced in the initial report is for the impella rp flex with smartassist instructions for use & clinical reference manual in the following sections: section: using the automated impella controller with the impella rp flex smartassist system catheter. section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | B5 updated to include placement signal issue description. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. g1 mdr reporting contact email and reporting contact fax number updated. f6 and f8 should have been left blank in the initial report. d10 concomitant medical products and therapy dates updated.
27536 | 1220648-2024-10301 | 19164908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19164908&pc=OZD | 05/10/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 05/10/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188808 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/10/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, ischemic lower extremity was noted with repositioning of the sheath in place. the impella cp was escalated to an impella 5.0 and relocated to axillary site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27537 | 1220648-2024-10302 | 19165046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165046&pc=OZD | 04/21/2022 | 05/20/2025 |  | 04/23/2024 |  |  | 04/21/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178855 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump advanced under the papillary muscle. during an attempt to reposition pump, the pump came back entirely into aorta. despite multiple attempts to advance the pump and use of a buddy wire, placement of the pump back across the aortic valve was unsuccessful. the pump was reinserted and pulled back into the descending aorta; the device was kept at p1 for the remainder of the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27538 | 1220648-2024-10303 | 19165119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165119&pc=OZD | 04/18/2024 | 06/05/2024 |  | 04/23/2024 |  |  | 04/18/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450083 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 77-year-old female post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during attempted delivery of an impella 5.5 pump, the catheter appeared to bend/break where it meets the motor housing. as a result of the noted damage, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation for the reported product damage has been completed. impella 5.5 pump was returned and visually inspected. it was observed that the catheter was pinched/clamp and kinked closer to the motor housing. the catheter was damaged due to improper technique likely occurring during insertion. there were no purge issues, pump stop, or other abnormalities found relative to the functioning of the 5.5 pump. the cause of the delivery issue was use related due to improper technique per return product investigation and clinical review. added-d9 device return date added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27539 | 1220648-2024-10304 | 19165141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165141&pc=OZD | 04/21/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/21/2022 | 04/23/2024 |  |  | 10/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support the patient experienced lost pulses in left lower extremity. the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27540 | 1220648-2024-10305 | 19165177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165177&pc=OZD | 04/26/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 04/26/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189353 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2022 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on impella support, patient developed hematuria, as both foley and urine were reported to be tinged. heparin was placed on hold and impella repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27541 | 1220648-2024-10306 | 19165227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165227&pc=OZD | 04/17/2024 | 06/21/2024 |  | 04/23/2024 |  |  | 04/18/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447391 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 57-year-old male with acute myocardial infarction and cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma and was treated with one unit of packed red blood cells, withholding of heparin, and manual pressure. the patient also lost a pulse in their limb. the patient was transferred to the intensive care unit and remains under medical support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into access site bleeding ¿ major and limb ischemia - reversible issues has been completed. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.2 common device name revised as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27542 | 1220648-2024-10307 | 19165229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165229&pc=OZD | 04/27/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 04/27/2022 | 04/23/2024 |  |  | 11/30/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168591 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a 61-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support. while on support, patient had ventricular tachycardia (vt) occasionally occurred and was accompanied with suction. suction disappeared when vt was solved by crm. also suggested physician to add volume to the patient. blood pressure dropped severely once during vt and was solved by defibrillation. in addition, a clot was found at left leg. physician suspected the clot was formed due to vt, iliac stenosis and vessel tortuosity. patient was sent to outpatient therapy for thrombectomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27543 | 1220648-2024-10308 | 19165444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165444&pc=OZD | 04/30/2022 | 04/01/2025 |  | 04/23/2024 |  |  | 04/30/2022 | 04/23/2024 |  |  | 02/29/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2022 |  |  |  |  | The user facility reported an 85-year-old male on impella cp for mechanical circulatory support. it was reported that following the impella cp implant, the patient had an access site bleed. it was noted that the patient experienced a possible femoral bleed due to sitting up and flexing their right leg. an echocardiogram (echo) was completed, it was reported that patient was do not resuscitate (dnr) and all actions were stopped. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
27544 | 1220648-2024-10309 | 19165448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165448&pc=OZD | 04/18/2024 | 06/10/2024 |  | 04/23/2024 |  |  | 04/18/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448595 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the team inserted the impella cp via the short 14 french sheath. the team observed the 14 french to have malfunctioned. the orange cap split from the sheath. it came off and bleeding started through the diaphragm. the short 14 french was replaced by the longer length 14 french sheath. support was initiated and the team completed the percutaneous coronary intervention, weaned, and explanted the impella cp. | [5 empty] | The investigation into the introducer damage ¿ mechanical and the access site bleeding ¿ minor issues has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding issue was most likely a damaged introducer hub, and the cause of the damaged introducer is not determined since the product was returned for investigation. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27545 | 1220648-2024-10310 | 19165869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165869&pc=OZD | 04/17/2024 | 05/29/2024 |  | 04/23/2024 |  |  | 04/17/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the display of the automated impella controller (aic) was not operational. the console was replaced to resolve the issue. no patient harm was reported. | [5 empty] | The investigation into the optical signal issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the optical data loss was due to a software issue. d9. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
27546 | 1220648-2024-10311 | 19165942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19165942&pc=OZD | 04/19/2024 | 06/27/2024 |  | 04/23/2024 |  |  | 04/19/2024 | 04/23/2024 |  |  |  |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441892 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 20-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella device failed on day nine of support, followed by a pump stop and motor current high alarm. the impella device was explanted and the patient remained on support. | It was further reported that the patient expired sometime after explant. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. impella 5.5 was returned for evaluation. upon visual inspection, a large biomaterial was observed on the impeller blade but not covering the purge gap. the pump was able to purge successfully in a bucket of water. pump was attempted to run multiple times with a pump stop occurring each time. data logs were reviewed and right logs at end of case show saturated flows with sporadic changes in motor current and motor speed with multiple motor current spike occurring before pump stop occurs. additionally, purge pressure was high at time of pump stop. clinical details noted that patient was receiving lot's of transfusions due to blood loss unrelated to the impella and pump was attempted to be repositioned when pump stop occurred. the root cause of the pump stop is due to biomaterial ingestion. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in the following sections: section: direct aortic insertion. section: using the automated impella controller with the impella catheter. section: impella stopped. added- b5 mso review, d9 device return date added d6a/d6b. f6/f8-should have been left blank in the initial report.
27547 | 1220648-2024-10312 | 19168572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19168572&pc=OZD | 10/27/2022 | 04/23/2024 |  | 04/23/2024 |  |  | 10/27/2022 | 04/23/2024 |  |  | 06/30/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023232042 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | A complainant reported the 82 year old female patient was on impella cp support due to lmt infarction. while on support, there was bleeding from the sg insertion site. the impella was not working well, but it was managed by adding volume with blood transfusion. the patient was managed without administering heparin to the purge solution. however, the purge flow rate gradually decreased possibly forming a blood clot. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27548 | 1220648-2024-10313 | 19169247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169247&pc=OZD | 10/31/2022 | 03/27/2025 |  | 04/23/2024 |  |  | 10/31/2022 | 04/23/2024 |  |  | 07/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238252 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient encountered ventricular tachycardia (vt) during support. the patient encountered recurrent vt episodes and there was suspicion that it was linked to the pump's pigtail. the decision was made to explant the pump and escalate to a 5.5 pump without the pigtail. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27549 | 1220648-2024-10314 | 19169385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169385&pc=OZD | 11/07/2022 | 03/27/2025 |  | 04/23/2024 |  |  | 11/07/2022 | 04/23/2024 |  |  | 07/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023240890 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced bleeding at all access areas requiring several units of blood products to be administered. due to the patient's condition, the impella device was explanted. in addition, it was reported that the patient underwent vascular surgery to repair the groin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27550 | 1220648-2024-10315 | 19169466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169466&pc=OZD | 11/12/2022 | 03/27/2025 |  | 04/23/2024 |  |  | 11/12/2022 | 04/23/2024 |  |  | 07/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2023236769 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unclear Information | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was removed and replaced with a new pump. no information was provided to determine why the impella device was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27551 | 1220648-2024-10316 | 19169501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169501&pc=OZD | 11/20/2022 | 03/27/2025 |  | 04/23/2024 |  |  | 11/20/2022 | 04/23/2024 |  |  | 03/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023204928 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had delivery issues. a second implant failed due to the patient's vascular anatomy as the device was unable to pass through the vasculature. extracorporeal membrane oxygenation (ecmo) support was utilized. no other information was available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27552 | 1220648-2024-10317 | 19169550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169550&pc=OZD | 04/22/2022 | 04/04/2025 |  | 04/23/2024 |  |  | 04/22/2022 | 04/23/2024 |  |  | 09/30/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141658 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was observed that the impella access site had a basketball size hematoma with a suspected retroperitoneal (rp) bleed. it was reported that the condition was managed by administering blood products, holding manual pressure and ultimately removal of the impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27553 | 1220648-2024-10318 | 19169585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169585&pc=OZD | 04/22/2022 | 05/20/2025 |  | 04/23/2024 |  |  | 04/22/2022 | 04/23/2024 |  |  | 10/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2022 |  |  |  |  | The user facility reported a 74 year old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the patient developed hemolysis and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27554 | 1220648-2024-10319 | 19169607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169607&pc=OZD | 04/23/2022 | 05/19/2025 |  | 04/23/2024 |  |  | 04/23/2022 | 04/23/2024 |  |  | 11/30/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2022 |  |  |  |  | The complainant reported a 48 year old male patient was implanted with an impella cp for mechanical circulatory support due to st-segment elevation myocardial infarction (stemi). while on support, a downtrend in platelets and hemolysis was noted. one unit of packed red blood cells (prbc) was required every other day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27555 | 1220648-2024-10320 | 19169687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169687&pc=OZD | 04/23/2022 | 05/14/2025 |  | 04/23/2024 |  |  | 04/25/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 2023187800 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a 74 year old male patient was implanted with an impella cp for mechanical circulatory support due to a st-segment elevation myocardial infarction (stemi). while on support thrombocytopenia was noted. one unit of packed red blood cells was transfused. frequent suction events were also noted. a bedside pericardiocentesis was performed and physician repositioned the catheter. position verified with bedside transthoracic echocardiogram; however, suction alarms persisted. lastly, the impella pigtail was noted to be underneath the papillary muscle. the doctor was able to reposition the impella at bedside. the pump was later explanted and the patient escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27556 | 1220648-2024-10321 | 19169697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169697&pc=OZD | 04/25/2022 | 04/15/2025 |  | 04/23/2024 |  |  | 04/25/2022 | 04/23/2024 |  |  | 12/31/2023 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023163171 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced leg ischemia. the decision was made to switch to extracorporeal membrane oxygenation (ecmo) and the impella 5.0. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27557 | 1220648-2024-10322 | 19169704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169704&pc=OZD | 06/20/2023 | 12/19/2024 |  | 04/23/2024 |  |  | 03/25/2024 | 04/23/2024 |  |  | 04/30/2025 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333069 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional |  |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | Abiomed received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), a 58- year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was noted that worsening of cardiogenic shock and worsening of renal failure occurred with a possible relationship to the impella 5.5 used on this patient. patient was hospitalised to address the issue. | It was noted that worsening of cardiogenic shock and worsening of renal failure occurred with a possible relationship to the impella 5.5 used on this patient. the patient wanted do not resuscitate/do not intubate (dnr/dni) and comfort care. care was withdrawn and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation vf/vt/af/at and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-10322. h6 health effect - impact code 1802 death.
27558 | 1220648-2024-10323 | 19169706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169706&pc=OZD | 04/25/2022 | 04/07/2025 |  | 04/23/2024 |  |  | 04/25/2022 | 04/23/2024 |  |  | 03/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023191372 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had moderate malperfusion of the right leg. the medical team was not concerned as the pump was planned to be removed. the patient was reported to be stable with no further issues. patient is planned to be weaned off support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27559 | 1220648-2024-10324 | 19169708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169708&pc=OZD | 04/26/2022 | 04/04/2025 |  | 04/23/2024 |  |  | 04/26/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184869 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was no systemic heparin at this time due to bleeding at impella cp site. moreover, hemolysis was noted due to hematuria in urine. patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27560 | 1220648-2024-10325 | 19169718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19169718&pc=OZD | 04/26/2022 | 03/26/2025 |  | 04/23/2024 |  |  | 04/26/2022 | 04/23/2024 |  |  | 01/31/2024 |  | 04/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181346 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, leg malperfusion was noted. the impella was weaned and the patient was transferred to the cath lab for explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27561 | 1220648-2024-10167 | 19151325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151325&pc=OZD | 04/12/2024 | 05/28/2024 |  | 04/22/2024 |  |  | 04/12/2024 | 04/22/2024 |  |  | 09/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407697 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 78-year-old male in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that following implant, the impella cp pump became stuck in the papillary muscle. attempts to reposition the pump were unsuccessful and the decision was made to replace the device. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the pump was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27562 | 1220648-2024-10168 | 19151432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151432&pc=OZD | 04/27/2022 | 05/20/2025 |  | 04/22/2024 |  |  | 04/27/2022 | 04/22/2024 |  |  | 02/29/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193987 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Insufficient Flow or Under Infusion | [5 empty] | Hemolysis | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a 75-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to non-st segment elevation myocardial infarction (nstemi) and mitral valve disease. while on support, the patient had bloody urine and laboratory draws were hemolyzed. the patient started having suction alarms despite being given 300 ml fluid bolus and reducing the impella to p2. the impella position was verified with echocardiography and was found to be correctly positioned free floating in left ventricle and there was no interference with valve leaflets. systemic heparin was removed from the purge solution. it was also reported that there was a vessel dissection, and the vessel was stented to treat. the medical team decided to explant the impella early due to hemolysis, suction alarms, and the vessel dissection. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine. hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure. patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27563 | 1220648-2024-10169 | 19151509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151509&pc=OZD | 04/27/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 04/27/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support a kink was observed in the right femoral artery sheath which caused a small hematoma that required no intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27564 | 1220648-2024-10170 | 19151606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151606&pc=OZD | 04/27/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 04/27/2022 | 04/22/2024 |  |  | 12/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172013 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a large hematoma in the right chest below the repositioning sheath, and the medical team decided to explant the impella cp as a result of the hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27565 | 1220648-2024-10171 | 19151690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151690&pc=OZD | 05/04/2022 | 03/31/2025 |  | 04/22/2024 |  |  | 05/04/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023174758 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, physician suspected limb ischemia. high dose catecholamines was needed. physician discussed options and decided to wait for stabilization and observe the leg. no further intervention was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27566 | 1220648-2024-10172 | 19151698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151698&pc=OZD | 04/01/2024 | 05/17/2024 |  | 04/22/2024 |  |  | 04/01/2024 | 04/22/2024 |  |  | 09/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025445536 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. continuous suction alarms with low flows were encountered during impella device support. repositioning was attempted; however, the issue remained. as a result of the persistent low flows, the decision was made to explant the device. there was no report of patient harm. | [5 empty] | The investigation into the low pump flow issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ d6a, d6b. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27567 | 1220648-2024-10173 | 19151701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151701&pc=OZD | 05/04/2022 | 03/31/2025 |  | 04/22/2024 |  |  | 05/04/2022 | 04/22/2024 |  |  | 11/30/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169559 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was still experiencing surgical bleeding. it was noted that the patient was receiving platelets and one unit of blood, and there was to be a switch to bicarbonate in purge. it was reported that this complication was resolved by surgical cut-down and removal of impella device. hemostasis was achieved without complication via sutured arteriotomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27568 | 1220648-2024-10175 | 19151751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151751&pc=OZD | 04/03/2024 | 06/07/2024 |  | 04/22/2024 |  |  | 04/03/2024 | 04/22/2024 |  |  | 10/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024417016 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 80-year-old male patient on impella cp support was bleeding from the nose and the puncture site. pressure bandage and additional suturing were unsuccessful in controlling the bleed. therefore, the decision was made to explant the pump to resolve the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical information provided, the root cause of the access site bleeding issue was most likely patient condition as the bleeding was from multiple sites. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27569 | 1220648-2024-10174 | 19151755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151755&pc=OZD | 05/04/2022 | 03/31/2025 |  | 04/22/2024 |  |  | 05/04/2022 | 04/22/2024 |  |  | 02/29/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186455 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a 68-year-old female was implanted with an impella cp for mechanical circulatory support. the patient experienced limb ischemia. antegrade perfusion was not possible. the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27570 | 1220648-2024-10176 | 19151784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151784&pc=OZD | 04/04/2024 | 06/06/2024 |  | 04/22/2024 |  |  | 04/04/2024 | 04/22/2024 |  |  | 09/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025465232 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella 5.5 support for 65-year-old male patient, the purge pressure was high. there were no kinks or external influences noted. there was no reported patient harm. | [5 empty] | The investigation for the reported high purge pressure has been completed. the pump was not returned for investigation. the data log shows a 2 to 3-hour gradual rise in purge pressure with a reported high purge pressure alarm. this purge pressure trend is consistent until the end of the case, with slight saturated pump flows and then followed by a small dip in pump flow at the end. no motor current spike observed but flow changes event found chained to high purge pressure events. the cause of the high purge pressure issue was most likely biomaterial, based on data log review. the instructions for use referenced in the initial report impella 5.5 with smartassist in section "using the automated impella controller with the impella catheter". added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
27571 | 1220648-2024-10177 | 19151806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151806&pc=OZD | 04/12/2024 | 05/29/2024 |  | 04/22/2024 |  |  | 04/12/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456614 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 66 year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient has been reported to have multiple cardiac and systemic comorbidities. on the fourth day of support, the patient developed signs of limb ischemia in the right leg. when the medical staff consulted vascular surgery, the patient was considered a poor surgical risk due to multiple organ failure, continuous renal replacement therapy, and coagulopathy. due to not tolerating flat bed rest, the patient started pulling several lines, including the impella device¿s, so the medical staff decided to explant the impella device. the patient remained under medical supervision following the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27572 | 1220648-2024-10178 | 19151836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151836&pc=OZD | 05/05/2022 | 09/26/2025 |  | 04/22/2024 |  |  | 05/05/2022 | 04/22/2024 |  |  | 02/28/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022046178 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2022 |  |  |  |  | The user facility reported a 59-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that bleeding was noted at the groin access site. it was also observed that the device was in the papillary muscle close to the mitral valve. the medical professional repositioned the impella cp numerous times with no success. the decision was made to replace the device. explant was performed and a new impella cp was placed without further complications. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b5 bleeding was erroneously included on the initial manufacturer device report number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting address was inadvertently omitted on the initial manufacturer device report numbers. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 health code clinical code 1888 was erroneously entered on the initial manufacturer device report number.
27573 | 1220648-2024-10179 | 19151848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151848&pc=OZD | 04/28/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 04/28/2022 | 04/22/2024 |  |  | 10/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147722 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, patient developed hemolysis. no information related to the resolution of hemolysis was provided. additionally, patient developed infection/sepsis. as a result, impella pump was removed after hemodynamic stabilization. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27574 | 1220648-2024-10180 | 19151859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151859&pc=OZD | 03/04/2024 | 06/21/2024 |  | 04/22/2024 |  |  | 04/05/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443712 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during placement procedure for patient, the impella cp was unable to be placed as priming was unable to be completed. two consoles were used and the purge cassette was exchanged but the issue did not resolve. a new impella was used without issue. there was no reported patient harm. | [5 empty] | The investigation for the reported priming issue has been completed. the pump and cassette were returned for investigation. no physical evidence of damage was noted on the purge cassette. the purge cassette piston was not moving during the priming test. the purge was coming out after cutting the purge line next to the purge disk. further investigation of the purge line revealed a dextrose blockage, which was solved after soaking the purge in hot water for less than ten minutes. the cause of the priming issue was a dextrose buildup. d.6a revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
27575 | 1220648-2024-10181 | 19151900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151900&pc=OZD | 04/05/2024 | 05/21/2024 |  | 04/22/2024 |  |  | 04/05/2024 | 04/22/2024 |  |  | 07/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432785 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old female undergoing high risk percutaneous coronary intervention was planned for an impella device placement for mechanical circulatory support. upon insertion of the impella pump, the patient coded, pump placement was aborted, and decision was made by the physician to place the patient on extracorporeal membrane oxygenation support. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related as the patient coded upon insertion and they were not able to pass the impella through the vasculature. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27576 | 1220648-2024-10183 | 19151901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151901&pc=OZD | 04/02/2024 | 06/10/2024 |  | 04/22/2024 |  |  | 04/02/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458020 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the pump was dislocated when the peel away sheath was pulled away. the impella was unable to be repositioned. therefore, a new impella was implanted. there were no kinks on either the cannula or catheter. the patient was noted to have been transferred to another hospital, but this was not a result of the positioning issues. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical information provided, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27577 | 1220648-2024-10184 | 19151945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151945&pc=OZD | 04/29/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 04/29/2022 | 04/22/2024 |  |  | 10/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a small perforation of the right femoral artery (rfa) due to tortuous anatomy and tissue obstruction above the access site. a covered stent and/or balloon tamponade was put in place to resolve the perforation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27578 | 1220648-2024-10185 | 19151954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19151954&pc=OZD | 04/06/2024 | 06/25/2024 |  | 04/22/2024 |  |  | 04/06/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025447597 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 52-year-old male patient with an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during impella 5.5 support, the nurse described a purge leak at the level of the purge filter. there was no low purge pressure alarm, but purge pressure was less than 330 mmhg. a purge filter and reservoir bypass procedure was completed to resolve the issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak - pump issue has been completed since the original report was submitted. the device was not returned for investigation. as per the clinical details, a purge leak was reported from the purge filter without any purge-related alarms being seen. the purge sidearm was bypassed due to the purge leak issue. the cause of the purge leak issue was most likely a leak from the sidearm, but the exact location of the leak was unknown. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ d.6a revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 code 481 was reported incorrectly on the previously submitted report. a new code has been added to component codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27579 | 1220648-2024-10186 | 19152009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152009&pc=OZD | 04/08/2024 | 06/13/2024 |  | 04/22/2024 |  |  | 04/08/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025452922 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, there was a problem with the purge pressure being high. there were no kinks observed on the purge line. the purge cassette was not replaced. the purge heparin concentration was 25 iu/ml. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that no purge line kink was observed, and purge cassette was changed and did not resolve issue. tpa was added to purge and was able to resolve high purge pressure alarms. the root cause of the high purge pressure was most likely biomaterial. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27580 | 1220648-2024-10187 | 19152027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152027&pc=OZD | 04/04/2024 | 06/10/2024 |  | 04/22/2024 |  |  | 04/04/2024 | 04/22/2024 |  |  | 10/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417175 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported when the impella cp was connected, the automated impella controller (aic) displayed pump not detected/recognized error. the console was exchanged but the issue did not resolve. the pump was therefore explanted and a new one was placed. the issue resolved and there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the pump not detected issue has been completed since the original report was filed. the device was returned for investigation. upon product inspection, it was observed that the patient cable was able to freely rotate and translate relative to the kink protector. the optical and the data lines were broken. the root cause of the pump not detected issue was determined to be an electrical opening due to discontinuity in the wiring.
27581 | 1220648-2024-10188 | 19152052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152052&pc=OZD | 05/18/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 05/18/2022 | 04/22/2024 |  |  | 11/30/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported the 18-year-old male was on impella cp support. during support there was an alarm of placement signal low. impella was repositioned and eventually became stable. additionally, it was reported that the patient experienced access site oozing. heparin was withheld. it was also reported that the patient had pulmonary congestion with clots. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27582 | 1220648-2024-10189 | 19152064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152064&pc=OZD | 04/09/2024 | 06/05/2024 |  | 04/22/2024 |  |  | 04/09/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477915 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 66-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during ogps, the staff noted a lower skin temperature on the left impella leg. this was immediately highlighted to the perfusionist/physician and the rn. the patient was in critical cardiogenic shock (cgs) situation and physicians decided to plan angio ultrasound. the patient was escalated to impella 5.5 and the femoral impella cp was explanted. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27583 | 1220648-2024-10190 | 19152095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152095&pc=OZD | 05/19/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 05/19/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. hemolysis was noted while the patient was on support. the family withdrew care. no further information was provided regarding medical or surgical intervention. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27584 | 1220648-2024-10191 | 19152098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152098&pc=OZD | 04/30/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 04/30/2022 | 04/22/2024 |  |  | 09/30/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132605 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a central venous catheter was then inserted into the internal jugular vein before returning the patient to their room, but it was reported that this resulted in perforation of an artery. physicians decided to remove the impella at that time in order to achieve hemostasis. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27585 | 1220648-2024-10192 | 19152105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152105&pc=OZD | 04/10/2024 | 06/03/2024 |  | 04/22/2024 |  |  | 04/10/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263727 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | Upon review by abiomed's medical safety officer, the patient expired on support over a week later. there is not enough information to associate the use of the device with the patient's outcome. | The complainant reported that during implant procedure for impella 5.5, prior to any patient contact, the automated impella controller (aic) (ic4734) displayed 'battery error'. the aic was therefore replaced. there was no reported patient harm. |  |  |  |  | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. the logs were reviewed and the battery failure issue was confirmed. there was an instance of battery failure alarm (transport connection blown/missing) and battery comm. failure alarms (smbus comm loss or sda short) found from the reported occurrence date. results of running batttest on telnet revealed that ¿fu¿ was missing on both batteries. the console was tested in collaboration with field service. battery failure alarm issue was able to be reproduced. the issue was determined to be caused by depleted batteries. the console was last disconnected from the alternating current (ac) power about 11 months ago and was never plugged back to the ac power since. these depleted batteries could not be charged back up because they were locked out due to low cell voltage ( less than 2.3v). the root cause of this issue was not routinely charging the aic. d.4 revised serial number and unique identifier (udi) # as previously entered incorrectly. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 codes 2199 and 2900 were reported incorrectly on the previously submitted report. new codes were add to health effect - impact code and medical device problem code. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | H6 revised type of investigation codes as they were inadvertently omitted from the previously submitted report.
27586 | 1220648-2024-10193 | 19152120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152120&pc=OZD | 04/05/2024 | 06/11/2024 |  | 04/22/2024 |  |  | 04/05/2024 | 04/22/2024 |  |  | 09/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399687 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for a male patient, the automated impella controller (aic) displayed a "purge pressure high" alarm followed by a "purge system blocked" message. the issue resolved by itself without any manipulation. however, soon after there was a "purge pressure low" alarm. a drop was noted outside of the pressure reservoir. there were no kinks observed. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of high purge pressure was not determined as no product was returned and insufficient clinical information was provided. in accordance with updated procedures, sections d6 have been revised from the initial submission.
27587 | 1220648-2024-10199 | 19152522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152522&pc=OZD | 05/22/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 05/22/2022 | 04/22/2024 |  |  | 02/29/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199997 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced hemolysis while on support. heparin was withheld and one unit of packed red blood cells was administered. the patient continued on impella cp support with stable hemodynamics. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27588 | 1220648-2024-10200 | 19152567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152567&pc=OZD | 04/06/2024 | 06/11/2024 |  | 04/22/2024 |  |  | 04/06/2024 | 04/22/2024 |  |  | 03/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205169 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a low purge pressure and high purge flow alarm. a leakage was found on the yellow lure-lock. it is not entirely clear which side was the problem. the clinical staff disconnected the luer connections with the aid of a clamp. this damaged the luer connection. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak - pump issue has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that leak was coming from the yellow luer connection but unclear on what side was causing the leak and no damage was observed. the cause of the purge leak was sidearm yellow luer damage.
27589 | 1220648-2024-10201 | 19152592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152592&pc=OZD | 04/12/2024 | 05/29/2024 |  | 04/22/2024 |  |  | 04/12/2024 | 04/22/2024 |  |  | 10/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024424959 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 45 year-old male patient with acute myocardial infarction and cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on support due to becoming hypertensive after a balloon angioplasty. while on support, the patient's dressing became saturated. medical staff found hematoma and oozing and used manual pressure to mitigate. around the same time, the patient lost distal pulses in the right leg. the impella device was then explanted successfully. the patient remains on medical support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of access site bleeding was determined to be anticoagulation management because of the high act values of the patient. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27590 | 1220648-2024-10202 | 19152687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152687&pc=OZD | 05/25/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 05/25/2022 | 04/22/2024 |  |  | 01/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022048136 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the there was no systemic heparin at this time due to bleeding at the impella cp site. the patient received 1 unit of packed red blood cells. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27591 | 1220648-2024-10203 | 19152693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152693&pc=OZD | 04/14/2024 | 06/03/2024 |  | 04/22/2024 |  |  | 04/14/2024 | 04/22/2024 |  |  | 04/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024330408 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Failure to Deliver | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella 5.5 implant that was unsuccessful. the impella was not able to be passed though the axillary artery. an impella cp was placed instead. additionally, a hematoma developed over the access site. blood products were provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the delivery issues and the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. the root cause of the delivery issue was most likely patient condition related, the impella was not able to be pass though the axillary artery as per the clinical description. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27592 | 1220648-2024-10204 | 19152718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152718&pc=OZD | 03/29/2024 | 04/22/2024 |  | 04/22/2024 |  |  | 04/08/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213121 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during impella cp support for patient, the automated impella controller (aic) did not recognize the purge disc. therefore, setup was unable to be completed. the purge cassette was exchanged and the console was reset but the issue did not resolve. the pump was unable to be used as there was no backup controller available. there was no reported patient harm. | [5 empty] | The investigation has been completed. the product was returned for investigation. the console was tested. the console was missing the purge flag. during data analysis, the console repeatedly was unable to detect purge discs after the case started. the cause of the issue was determined to be a missing purge flag.
27593 | 1220648-2024-10205 | 19152774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152774&pc=OZD | 04/10/2024 | 09/18/2024 |  | 04/22/2024 |  |  | 04/15/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1390619 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Use of Device Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that while on impella cp support a patient developed limb ischemia. a distal perfusion catheter and fasciotomy was performed to resolve the limb ischemia. | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that while on impella cp support the automated impella controller (aic) froze and gave a red alarm. the aic was replaced and the issue resolve. no patient harm was reported. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | Manufacturer device report 1220648-2024-10205 was incorrectly submitted as a complaint against the impella cp pump, the correct product is automated impella controller, please see below corrections: b1 corrected to product problem. b2 should have been left blank . b5 event narrative updated to reflect correct description. d1 brand name. d2 common device name. d4 model number, catalog number, serial number, lot number, unique identifier (udi), and the expiration date should have been blank. f6/f8-should have been left blank. h1 type of reportable event change to malfunction. h5 changed to no. h6 code 1942, 4641, 4625, 2993, 3038 wert reported incorrectly. new codes were added to health effect - clinical code, health effect - impact code, medical device problem code, and component code. h8 usage of device changed to reuse. h10 instructions for use should have not been included.
27594 | 1220648-2024-10206 | 19152810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152810&pc=OZD | 05/01/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 05/24/2022 | 04/22/2024 |  |  | 07/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114223 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | [9 empty] | Other | Death | Required Intervention | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced major access site bleeding requiring blood transfusion. the patient also had a cerebral vascular accident (cva), limb ischemia and sepsis. the patient expired while on support; however, there were no allegations towards the impella as the cause of death. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that on (b)(6) 2022 the patient experienced serious bleeding at the impella access site and required blood products as a result of the blood loss on (b)(6) 2022 it was reported that the patient experienced an ischemia stroke and limb ischemia during impella cp support. on (b)(6) 2022 the patient was noted to have sepsis while on impella cp support. the impella was successfully explanted on (b)(6) 2022, and the following day on (b)(6) 2022 the patient experienced a sudden cardiac death which was noted to be related to impella use. no further information was provided. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. the following has been corrected from the initial mfr 1220648-2024-10206: section b3: date of event has been corrected. section b2: checked death and added date of death. section b5: description of event has been corrected. section h1: corrected type of reportable event to death. section h6 health effect - impact code: added death 1802 code.
27595 | 1220648-2024-10207 | 19152817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152817&pc=OZD | 04/15/2024 | 06/07/2024 |  | 04/22/2024 |  |  | 04/15/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025441950 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported that while on support the patient developed an access site infection. antibiotics given daily and infection is improving. the patient remains on support and is reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen. | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in section: table 3.2 impella catheter components. added- d6a/d6b.
27596 | 1220648-2024-10208 | 19152897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152897&pc=OZD | 05/27/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 05/27/2022 | 04/22/2024 |  |  | 12/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173625 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient became agitated and displaced the impella from the left ventricle into the aorta. the physician was advised to remove slack, but physician declined to remove the slack. the pump then advanced too deep into the left ventricle and had suction issues. the p level on the impella was decreased significantly to resolve suction issues and care was withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27597 | 1220648-2024-10209 | 19152914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152914&pc=OZD | 04/16/2024 | 05/15/2024 |  | 04/22/2024 |  |  | 04/16/2024 | 04/22/2024 |  |  | 10/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024421370 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 76-year-old male patient of unknown race and ethnicity noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that access site bleeding occurred during implantation of an impella cp. a covered stent and balloon tamponade were attempted, and blood products were given. the impella was explanted in the procedural area. the patient outcome is unknown. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding could not be determined since the product was returned and limited clinical details were provided. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
27598 | 1220648-2024-10210 | 19152998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19152998&pc=OZD | 11/15/2022 | 05/08/2024 |  | 04/22/2024 |  |  | 11/15/2022 | 04/22/2024 |  |  | 06/30/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234945 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, they experienced some bruising and a hematoma at the impella cp insertion site. two perclose were applied. angioseal and manual pressure was held to resolve the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since the hemostasis was achieved by manual pressure along with 2x perclose sutures. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. corrected information provided in d6a, d6b, h6, h10 and h11.
27599 | 1220648-2024-10211 | 19153006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153006&pc=OZD | 04/16/2024 | 06/04/2024 |  | 04/22/2024 |  |  | 04/16/2024 | 04/22/2024 |  |  | 11/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423430 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported an 81 year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during the setup of the impella device, catheter fluid was not observed exiting the outlet portion. a new catheter was used to complete the setup. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the priming issue has been completed. the purge cassette was not returned. impella cp pump was visually inspected, and no purge gap blockage observed. pump pigtail was found to be slightly bent but was not detached leading to any failure during support. test purge cassette was connected to aic and primed with purge fluid by connecting to the complaint pump and purge fluid was passing at the purge gap near pump housing without any issues. purge pressures and flows were found to be normal. pump was connected to console, started at auto and pump was running with normal flows of 3.8 l/min. no abnormalities were found with the pump. data logs were reviewed and no purge cassette failure, air in purge alarm were found. purge system open was found at end of first console shutdown and purge pressure low alarm was observed after switching console but resolved within 30 seconds. no priming issue related alarms were found and issue did not reproduce with return product. there was no issue found during the case per return product investigation and log review. the device functioned/operated to specification. added- d9 device return date d4-updated model number. added- d6a/d6b.
27600 | 1220648-2024-10212 | 19153021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153021&pc=OZD | 01/10/2024 | 07/17/2024 |  | 04/22/2024 |  |  | 01/10/2024 | 04/22/2024 |  |  | 09/30/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278538 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/10/2024 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on impella cp support and while on support, there was access site bleeding. the doctor tightened the preclose, held pressure, and placed a fem stop. patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27601 | 1220648-2024-10213 | 19153036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153036&pc=OZD | 04/17/2024 | 09/03/2024 |  | 04/22/2024 |  |  | 04/17/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1675422 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported when preparing for support, the automated impella controller (aic) had a purge disc/pc malfunction. the aic was removed from service. no patient harm or involvement has been reported. | [5 empty] | The investigation of aic - purge disc/flag issues has been completed. the logs revealed that during the purge start operation, the console was unable to consistently detect the presence of the purge disc and continuously looped through the purge start sequence. during the console testing the purge disc retainer was confirmed to be loose, which caused the purge pressure sensor to inconsistently measure purge pressure and the purge flag to inconsistently detect the presence of the purge disc. the root cause of the issue was a defective component. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-10213: d2 common device name d4 model number e2 health professional inadvertently was selected yes f6/f8-should have been left blank g1 reporting contact facility name g1 reporting contact fax number missing | The automated impella controller was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
27602 | 1220648-2024-10214 | 19153056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153056&pc=OZD | 04/10/2024 | 06/13/2024 |  | 04/22/2024 |  |  | 04/10/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464361 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support for a 38-year-old female patient, the pump was at p0 and the pump would not start again. attempts to restart the pump were made but were unsuccessful. therefore, the pump was explanted, and a new pump was placed. there was no reported patient harm. the device will not be returned as it was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump stop has been completed. no product was returned for evaluation. data logs were reviewed and right data log at the end of case show a motor current spike with a simultaneous speed deviation in motor speed. additionally, after motor current spike, pump flow steps up and become saturated for about 45 seconds before alarm #13 (¿impella stopped ¿ motor current high¿) triggered and pump stop occurred. additionally, purge flow begins to drop with an increase in purge pressure after motor current spike. clinical details noted that multiple attempts were made to restart pump after pump stop occurred but clinical team was not able to restart pump. the root cause of the pump stop is most likely due to biomaterial ingestion. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in the using the automated impella controller with the impella catheter, impella stopped, and the troubleshooting the purge system sections. added- d6a/d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27603 | 1220648-2024-10215 | 19153078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153078&pc=OZD | 04/17/2024 | 05/20/2024 |  | 04/22/2024 |  |  | 04/17/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080532 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 31-year-old male with cardiomyopathy was planned for an impella device for mechanical circulatory support. the team attempted to start a new case with the automated impella controller (aic). when the white cable was plugged into the aic, and it did not recognize the device. the aic kept prompting to make the black-to-black connection. the aic was replaced, and the same white cable was plugged into a new aic without any issues. there was no report of patient harm. | [5 empty] | The investigation for the console (aic) pump issue has been completed. the console was returned and was tested to attempt to recreate the failure mode but this was unsuccessful. the console passed all testing during the investigation leading to the cause of this failure mode to be most likely due to a rarely occurring epm crystal failure on the impellatronic board. data logs were reviewed finding that in step 3, the epm is power cycled after the console does not see a pump connected for 20 seconds and then the console misidentifies pump 492325 asking the user in step 3 to "connect cable: connect grey connectors screen". this is the incorrect message for a 5.5 pump. the pump is moved over to (b)(6) and the failure mode does not reproduce. the pump detection issue was most likely due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. added- d9 device return date d4-updated model number | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27604 | 1220648-2024-10217 | 19153140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153140&pc=OZD | 05/29/2022 | 05/23/2025 |  | 04/22/2024 |  |  | 05/29/2022 | 04/22/2024 |  |  | 02/29/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported a patient on impella cp support due to cardiogenic shock. while on support, a fever spike, and oozing at the axillary site was noted. the patient is overall well tee¿d up for transplant, ongoing bacteremia remains a concern. impella cp was explanted and replaced with a impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27605 | 1220648-2024-10216 | 19153145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153145&pc=OZD | 12/12/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 12/12/2022 | 04/22/2024 |  |  | 07/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253586 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The complainant reported the 73-year-old female patient was implanted with an impella cp and the patient experienced a massive stroke, confirmed with head ct. the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records.
27606 | 1220648-2024-10219 | 19153163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153163&pc=OZD | 12/12/2022 | 08/05/2025 |  | 04/22/2024 |  |  | 12/12/2022 | 04/22/2024 |  |  | 10/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272395 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding at the groin access site. it was reported that the site bleed was managed with compression by suture. the patient was weaned from the impella, the device was explanted, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the manufacturer device report 1220648-2024-10219 in accordance with updated procedures f6/f8-should have been left blank on the initial manufacturer device report 1220648-2024-10219 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10219 in accordance with updated procedures. h6 component code revised from the initial manufacturer device report 1220648-2024-10219 in accordance with updated procedures.
27607 | 1220648-2024-10221 | 19153198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153198&pc=OZD | 04/17/2024 | 06/10/2024 |  | 04/22/2024 |  |  | 04/17/2024 | 04/22/2024 |  |  | 06/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024395161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 59-year-old male patient of unknown race in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient experienced low pump flow and hemolysis. the impella cp was explanted due to hemolysis and shortly after the explant the patient arrested and a new impella cp was placed. the patient remains on support and is in critical condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the low pump flow and the hemolysis issues has been completed since the original report was filed. the device was returned for investigation. the returned pump was visually inspected and big chunk of biomaterial was observed around impella gap. the cause of both saturated flow and hemolysis issues are biomaterial ingestion. d6a, d6b, d9.
27608 | 1220648-2024-10218 | 19153263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153263&pc=OZD | 04/17/2024 | 06/03/2024 |  | 04/22/2024 |  |  | 04/17/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485057 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported an 87-year-old female patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that upon explanting the pump, the patient experienced a pericardial effusion. the patient was tapped and over 800 cc was removed. the patient became unstable and under echo appeared that there was a clot in the apex. a perforation was found in the lateral wall of the left ventricle. the patient expired in the operating room. the relationship between the impella and the cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve. | The investigation into the ventricle perforation and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the ventricle perforation was not determined since product was not returned and there is a lack of clinical details. as the necessary information was not provided, the root cause of the thrombus was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6, g1 and h10 have been revised from the initial submission.
27609 | 1220648-2024-10220 | 19153292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153292&pc=OZD | 04/10/2024 | 06/07/2024 |  | 04/22/2024 |  |  | 04/10/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263726 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during implant procedure for impella 5.5, prior to any patient contact, following prior automated impella controller (aic) replacement, it was noted that the side arm of the holder where the purge disk goes in, in the current aic was broken off. the aic was once again replaced. there was no reported patient harm. | There is no association between impella use and the outcome. |  |  |  |  | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. the cause of the issue was a broken purge retainer. b5, d9.
27610 | 1220648-2024-10222 | 19153307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153307&pc=OZD | 12/12/2022 | 04/02/2025 |  | 04/22/2024 |  |  | 12/12/2022 | 04/22/2024 |  |  | 08/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135458 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the impella had low purge pressure and there was leaking. when the solution was switched from sodium bicarbonate to heparin, the tubing was to be changed. after the change, there were multiple purge system open and purge pressure low alarms. a purge bypass was performed, and the purge pressure and flow normalized. echo was performed and showed that the impella was deep in the left ventricle but there were no placement alarms and motor current was pulsatile. the team did not want to reposition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a added implant date. section d6b explant date is unknown. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27611 | 1220648-2024-10223 | 19153320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153320&pc=OZD | 04/17/2024 | 07/11/2024 |  | 04/22/2024 |  |  | 04/17/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025459789 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 36-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. after inserting the impella pump and starting support, there was a decrease in purge flow and increase in purge pressure. there was also an alarm. the purge cassette was replaced, but there was no improvement. it was noted as there was a high risk of the impella pump stopping, it replaced with a new impella pump, which worked without any problems. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | Additional information received indicated the impella 5.5 device was received from the customer for evaluation. according to clinical, impella 5.5 inserted surgically via right axillary artery. d5w with sodium bicarbonate used in purge solution. soon after insertion and starting support, clinical team noted a decrease in purge flow and an increase in purge pressure. purge cassette was replaced and did not resolve issue. decision was made to remove pump and replace with a new impella 5.5 that worked without issue. data logs were not received for evaluation/analysis. however, the impella 5.5 pump was received, and inspection of the returned pump showed no evidence of biomaterial, or a kinked purge line was noted. purge gap was clear, and no biomaterial was present in pump cannula or on impeller blade. high purge pressure reproduction testing was performed. pump was placed in bucket and purged for approximately 10 minutes and high purge pressure did not reproduce. pump was then turned on to p-9 run for 5 minutes and high purge pressure alarm was triggered. the pump catheter was then cut below motor housing and purge fluid was able to flow through the catheter up to where it was cut. the motor housing was attempted to purge again with a needle and no purge was able to pass through. the pump cannula was cut off and impeller blade was pulled gently to widen the purge gap to check for biomaterial. no biomaterial was found within the purge gap. impeller blade was cut off to view purge gap fully from motor housing and no abnormalities were noted. a motor teardown was performed, and no abnormalities were observed on magnet or within motor. once magnet was removed from motor, purge was able to flow through motor housing. product history review revealed the pump passed all post-sterile inspection checks. clinical details noted that high purge pressure was observed soon after pump was started. purge cassette was changed in attempt to resolve issue without success. the root cause of the high purge pressure cannot be determined as location of purge blockage was unable to be found and no data logs were returned. correction was made to the following fields(s)/section(s): d4: model number corrected. d6a: implant date populated. d6b: explant date populated.
27612 | 1220648-2024-10224 | 19153376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153376&pc=OZD | 05/01/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 05/01/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. post placement, a moderate hematoma was identified, and the groin site was ecchymotic and saturated with blood. one (1) unit of blood was transfused, and femstop, manual pressure was applied to the hematoma and dressing was changed to resolve the issue. it was also reported that a penumbra device was utilized to successfully remove a thrombus from the right femoral artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27613 | 1220648-2024-10225 | 19153426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153426&pc=OZD | 06/09/2022 | 05/20/2025 |  | 04/22/2024 |  |  | 06/09/2022 | 04/22/2024 |  |  | 12/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the device was repositioned overnight, and the patient¿s hemoglobin was elevated. the pump speed was reduced to level p7. the patient was increasingly tachycardic with a heart rate reaching 160 beats per minute. the impella cp pump speed level was decreased, and the patient was successfully weaned to extubation. the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27614 | 1220648-2024-10226 | 19153429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153429&pc=OZD | 04/17/2024 | 06/13/2024 |  | 04/22/2024 |  |  | 04/17/2024 | 04/22/2024 |  |  | 11/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428976 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 76 year-old male patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. the right limb had worse discoloration than the left limb. the patient was placed on extracorporeal membrane oxygenation to resolve limb ischemia. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27615 | 1220648-2024-10227 | 19153618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153618&pc=OZD | 04/16/2024 | 06/24/2024 |  | 04/22/2024 |  |  | 04/16/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025470111 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 74-year-old male with cardiomyopathy was implanted with an impella cp device and placed on veno-arterial extracorporeal membrane oxygenation for mechanical circulatory support. during support, there were purge alarms for purge flow. the purge cassette was not replaced despite alarms for purge flow of low 20's. impella pump support was maintained, and there was no report of patient harm. | [5 empty] | The investigation into the purge leak ¿ cassette has been completed since the original report was submitted. the device was not returned for investigation. as per clinical, there were no purge alarms or purge leak observed. there was no problem found with the pump relative to the increased purge flow event per clinical. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27616 | 1220648-2024-10228 | 19153636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153636&pc=OZD | 05/01/2022 | 05/22/2025 |  | 04/22/2024 |  |  | 05/01/2022 | 04/22/2024 |  |  | 11/30/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157644 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The complainant reported the 59 year old male patient was presented to the cardiac catheterization laboratory for the primary indication of acute myocardial infarction and cardiogenic shock. while on the impella support, the staff noted a positioning issue where the device was positioned in the papillary muscle, close to mitral valve. the physician attempted to reposition several times; however, the attempts were not successful. no further information provided on additional interventions for impella placement issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27617 | 1220648-2024-10229 | 19153737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153737&pc=OZD | 04/12/2024 | 09/26/2024 |  | 04/22/2024 |  |  | 04/12/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456648 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the physician experienced difficulty advancing the impella cp cannula into the patient's left ventricle due to tortuous anatomy. the implant procedure was aborted. while in the procedural area the patient arrested and, despite several rounds of cardiopulmonary resuscitation, the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-10229 was provided in sections b2, b5, g3, and h1, and h6. | The investigation into the delivery issues and the unspecified (allegation) ¿ death has been completed since the original report was submitted. the device was not returned for investigation. the cause of the delivery issue was patient condition related due to pump unable to advance with patient tortuous vessel condition. the root cause of the unspecified death allegation was not determined. d6a, d6b.
27618 | 1220648-2024-10230 | 19153965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19153965&pc=OZD | 04/11/2024 | 05/15/2024 |  | 04/22/2024 |  |  | 04/12/2024 | 04/22/2024 |  |  | 08/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443679 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | Material Too Soft/Flexible | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that there were introducer issues. the staff was serial dilating the right femoral artery up to 14 french x 13 centimeters peel away sheath and the doctor claimed there was damage to the dilator. upon assessing the dilator, no cracks were found. the doctor used a 14 french x 25 centimeters peel away sheath and dilator without issue. the doctor felt the introducer was flimsy and felt there was possible kinking and cracking of the dilator. | [5 empty] | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the introducer damage issue was not determined since product/images were not returned. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27619 | 1220648-2024-10231 | 19154156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154156&pc=OZD | 04/13/2024 | 04/14/2025 |  | 04/22/2024 |  |  | 04/13/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436365 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support post cardiac surgery when an intra-aortic balloon pump was not sufficient. the pump was explanted after four days when there was an issue with the optical signal and low flows. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue and the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the optical signal issue and the low pump flow could not be determined as no data logs and/or product were returned. in accordance with updated procedures, section d6 has been revised from the initial submission. | The investigation of low pump flow and optical signal issue has been completed. the returned product was severed from the patient cable, therefore no reproduction testing or 5s parameter verification could be conducted. optical signal issue: upon review of the data logs, it is apparent that the console is the potential cause of the issue. low pump flow: the returned pump was severed so reproduction testing could not be performed. the data logs show no motor current spikes and there was no biomaterial in the purge gap or on the impeller. there was no damage to the impeller. no patient condition related issues were noted. the root cause of the low pump flow was not determined as insufficient product required for reproduction testing was returned d9 date device returned to manufacturer added. d6a and d6b revised on manufacturer device report 1220648-2024-10231 in accordance with updated procedures. f6/f8-should have been left blank on manufacturer device report 1220648-2024-10231. g1 reporting contact email revised on manufacturer device report 1220648-2024-10231 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10231.
27620 | 1220648-2024-10232 | 19154159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154159&pc=OZD | 05/01/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 05/01/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173019 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced impella access site bleeding, heparin induced thrombocytopenia, and hemolysis. these issues were treated by giving the patient argatroban, discontinuing heparin from the purge fluid, and monitoring any purge trends for access site bleeding and hemolysis. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27621 | 1220648-2024-10233 | 19154228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154228&pc=OZD | 04/13/2024 | 06/05/2024 |  | 04/22/2024 |  |  | 04/13/2024 | 04/22/2024 |  |  | 10/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414427 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported that starting the impella cp procedure, the sheath was attempted but would not advance into the femoral artery. angiogram showed severe peripheral vascular disease. further attempts to place the sheath and impella were aborted. the impella cp was not used, just the 14 french sheath. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the difficulty inserting the introducer issue was most likely patient condition related, sheath attempted but would not advance into femoral artery due to severe pvd. the instructions for use for the impella cp with smartassist system in section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27622 | 1220648-2024-10234 | 19154354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154354&pc=OZD | 08/16/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 08/16/2022 | 04/22/2024 |  |  | 12/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163701 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced bleeding requiring blood transfusions. in addition, there was a hematoma at the puncture site that formed when the impella support was surgically removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27623 | 1220648-2024-10235 | 19154384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154384&pc=OZD | 11/22/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 11/22/2022 | 04/22/2024 |  |  | 07/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256333 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2022 |  |  |  |  | The complainant reported a patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient experienced a gastrointestinal bleed. systemic heparin was withheld, and the patient was weaned off of impella support. after removal, a thrombus was found to have formed on the pigtail of the device, no patient intervention required. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27624 | 1220648-2024-10236 | 19154408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154408&pc=OZD | 04/14/2024 | 05/29/2024 |  | 04/22/2024 |  |  | 04/14/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454963 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 71-year-old male patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed sepsis. the medical staff thought the sepsis was possibly urosepsis. the medical staff weaned the patient off the impella device and then explanted it. the patient remains under medical support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the infection/sepsis issue has been completed since the original report was filed. the pump was returned for investigation. there were light traces of foreign material seen in the outflow cage and around an impeller. data logs were returned and no motor current spike, positioning issue alarms were seen. because evidence related to reported issue has not been provided, the root cause of the infection cannot be determined. it was reported that the pump p-level was decreased which resolved the hemolysis. therefore, per clinical information and product review, the root cause of the hemolysis issue was biomaterial. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27625 | 1220648-2024-10239 | 19154672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154672&pc=OZD | 04/15/2024 | 07/23/2024 |  | 04/22/2024 |  |  | 04/15/2024 | 04/22/2024 |  |  | 06/30/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359737 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Tachycardia | Thromboembolism | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/15/2024 |  |  |  |  | It was further reported that after the impella 5.5 device was removed, the patient developed a significant hematoma in right upper arm in bicep space with affected muscle. a fasciotomy was performed and there was good doppler signal at the radial artery and a palpable brachial artery pulse. it was also reported that the patient's right lower extremity had discoloration and skin changes consistent with diffuse mottling. computed tomography angiography (cta) imaging was performed and found an occlusive thromboembolism noted in the right superficial femoral artery. the patient was taken to operating room for a thrombectomy and fasciotomy. the patient recovered without further sequelae. | The user facility reported a 70-year-old male patient with acute myocardial infarction and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. following defibrillation for an unrelated patient condition, the implanted impella 5.5 pump was unable to provide support levels about p4 without encountering suction coinciding with a flat motor current. it was noted that the patient had existing left arterial clots which may have dislodged during defibrillation. as a result of the persistent issue, the decision was made to explant the pump. an intra aortic balloon pump was subsequently placed for ongoing support. the patient survived the event. |  |  |  |  | The instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ added-h6 clinical code 2654 and 1884, h6 impact code 4625, h6 med dev prob code 2993, h6 component code 925 b1-selected adverse event b2-selected required intervention to prevent permanent impairment/damage f6/f8 should have been removed in the initial report #1220648-2024-10239 h1-changed to serious injury h6-investigation findings in supplemental report 1220648-2024-10239-01 changed to 114. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported access site bleeding and ventricular tachycardia has been completed. without the device nor additional clinical details, the root cause of the access site bleeding and ventricular tachycardia cannot be determined since. the root cause of the investigation for the low pump flow has been updated. the cause of the low pump flow issue most likely biomaterial ingestion as the clinical details confirmed they imaged and visible clots clinging to inflow cage and tip of impella. added- h6 clinical code changed from 4582 to 2095 in the initial report #1220648-2024-10239. added- h6 impact code 4647 in the initial report #1220648-2024-10239. added- h6 component code 925 (was erroneously left off of supplemental report 1220648-2024-10239 #2). | The investigation for the reported low pump flow has been completed. data logs were reviewed finding no motor current spikes nor positioning alarms which indicate a flow issue. many suction alarms occurred. the cause of the low pump flow issue cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. added- d6a/d6b.
27626 | 1220648-2024-10241 | 19154881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154881&pc=OZD | 04/19/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 08/23/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient encountered sepsis. it was also noted that various catheters have been inserted, and the source of infection is unknown. additional information and event resolution is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27627 | 1220648-2024-10242 | 19154898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154898&pc=OZD | 04/06/2024 | 05/30/2024 |  | 04/22/2024 |  |  | 04/15/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 45-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. approximately five minutes after the impella pump was pulled, the patient experienced a stroke and was taken to interventional radiology for stroke management. at bedside, the impella pump was pulled by interventional cardiologist, and a clot was noted on the catheter. within minutes, the patient began to experience slurred speech and focal deficits and was taken to interventional radiology. computed tomography scan showed ischemic acute left m2 segment lesion. it was noted that the patient had been systemically heparinized for several days leading up to stroke and echocardiograms leading up to the event showed no evidence of a clot. the staff reported the patient had no residual deficits and was doing well. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. the instructions for use for the impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states) as referenced in the initial report now includes statement "cardiac or vascular injury (including ventricular perforation).¿ it also states in section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ in accordance with updated procedures, have been revised from the initial submission.
27628 | 1220648-2024-10243 | 19154993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19154993&pc=OZD | 04/15/2024 | 06/13/2024 |  | 04/22/2024 |  |  | 04/15/2024 | 04/22/2024 |  |  | 08/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443590 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/15/2024 |  |  |  |  | The user facility reported a 55-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella cp placement, the patient had no distal pulses in the impella side. the physician is aware of the issue and is assessing the patient frequently. the patient has complicated peripheral artery disease (pad) and no interventions are planned at this time. the patient was reported as stable and remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the reported limb ischemia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | Instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.10 additional manufacturer narrative instructions for use were reported incorrectly on the previously submitted report and have been revised in accordance with updated procedures.
27629 | 1220648-2024-10244 | 19155335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19155335&pc=OZD | 09/28/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 09/28/2022 | 04/22/2024 |  |  | 05/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231096 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced oozing at the groin site and the following day there was not much oozing. the patient remained on cp support for one additional day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27630 | 1220648-2024-10237 | 19155479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19155479&pc=OZD | 04/14/2024 | 06/03/2024 |  | 04/22/2024 |  |  | 04/14/2024 | 04/22/2024 |  |  | 10/31/2025 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445624 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the pump was not priming properly. the staff suspected a purge issue and elected to replace the impella cp with a new device. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the priming issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the priming issue cannot be determined as no product was returned and no abnormalities were observed from field logs.
27631 | 1220648-2024-10238 | 19155670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19155670&pc=OZD | 08/19/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 08/19/2022 | 04/22/2024 |  |  | 03/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206300 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the device began showing no pulsatility in the placement signal, and bubbles were noted at the impella outlet. due to risk for air embolism, the physician made the decision to escalate to impella 5.5, and the impella cp was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27632 | 1220648-2024-10240 | 19155724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19155724&pc=OZD | 05/02/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 05/02/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183500 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Difficult to Insert | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon initial impella insertion, the wire backed across the valve, which required the impella to be explanted to regain left ventricle (lv) access. once lv was re-accessed, impella was inserted in normal fashion without issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27633 | 1220648-2024-10245 | 19156200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19156200&pc=OZD | 10/03/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 10/03/2022 | 04/22/2024 |  |  | 06/30/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236867 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s right foot was noted to have petechia and splotchy. pulse could be found via doppler but not nearly as strong as the pre-device. decision was made to leave alone and monitor possible rash on right foot. all catheters were removed and a 5 french sheath was sewn in to be pulled in day patient area. palpable pulse felt post runoff but diminished in comparison to left foot. impella slowly weaned down over 15 minutes and patient tolerated weaning. impella was then turned off and removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27634 | 1220648-2024-10246 | 19156322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19156322&pc=OZD | 04/15/2024 | 06/21/2024 |  | 04/22/2024 |  |  | 04/15/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449992 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, a purge system blocked alarm message displayed intermittently. motor current was monitored, and tissue plasminogen activator (tpa) on purge was started. the first bag of tpa was still running with no change in purge pressure or purge flow and intermittent purge system blocked alarm. p-level was decreased to p-6. the patient remained on impella pump support, and was noted to be stable. | [5 empty] | The investigation for the high purge pressure has been completed. the product was not returned for investigation. data logs show that 18 hours into the case purge pressure begins to rise over a period of around 10 hours. then there is a sudden rise in purge pressure over 11 minutes with drop in purge flow. this is after increasing p-level. there are no motor current (mc) spikes. the purge pressure and purge flow do not recover. the root cause of the high purge pressure issue was most likely gradual biomaterial buildup. the instructions for use for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6-med dev prob code 1491. added- d6a/d6b h6-impact code 4644. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27635 | 1220648-2024-10247 | 19157141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19157141&pc=OZD | 11/09/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 11/09/2022 | 04/22/2024 |  |  | 09/30/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023261059 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/09/2022 |  |  |  |  | The complainant reported the 65 year old male patient needing mechanical circulatory support. while on impella support there was a small hematoma that was noted. no resolution provided for hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27636 | 1220648-2024-10248 | 19157268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19157268&pc=OZD | 04/16/2024 | 06/06/2024 |  | 04/22/2024 |  |  | 04/16/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464497 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Activation, Positioning or Separation Problem | [5 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 69-year-old female with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed multiple runs of sustained ventricular tachycardia. after an echocardiogram was performed, the medical staff discovered that the impella device was placed in the sub-valvular structures of the mitral valve and/or papillary muscle. attempts to reposition the impella device were unsuccessful, and the device was ultimately replaced. the patient survived the event and remained under medical support. | [5 empty] | The investigation has been completed. the product and data were not returned for review. based on the clinical description, the root cause of positioning issue was most likely use related. as the necessary information was not provided, the root cause of the ventricular tachycardia was not determined. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 medical device problem code was revised in accordance with updated procedures. new code has been added to medical device problem codes. code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27637 | 1220648-2024-10249 | 19158183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19158183&pc=OZD | 04/16/2024 | 06/03/2024 |  | 04/22/2024 |  |  | 04/16/2024 | 04/22/2024 |  |  |  |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466355 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male undergoing high risk percutaneous coronary intervention was planned for an impella cp device placement for mechanical circulatory support. the team utilized the single access method with companion 7 french sheath. however, the 7 french sheath failed and kinked. another device was successfully used. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the introducer damage has been completed. the returned 7fr lp sheath was visually inspected and tip damage was observed. data log review not applicable. if the needle stick in the 14fr introducer is very far out towards the edge of the 14fr valve, the user is most likely to run into the sharp edge of the inner geometry of the peel away hub when doing single access. the cause of the introducer damage was most likely user related since the needle stick was likely too lateral.
27638 | 1220648-2024-10250 | 19158196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19158196&pc=OZD | 11/19/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 11/19/2022 | 04/22/2024 |  |  | 08/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260096 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/19/2022 |  |  |  |  | The complainant reported the 86 year old male patient needing mechanical circulatory support. it was reported that pedal pulses were absent. device was removed in the cath lab. perclose was performed prior to implantation. once the pump was removed perclose was tightened, and manual pressure was applied to the right groin. hemostasis was obtained post manual pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27639 | 1220648-2024-10251 | 19158513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19158513&pc=OZD | 02/05/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 02/05/2022 | 04/22/2024 |  |  | 10/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158792 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 02/05/2022 |  |  |  |  | The user facility reported a 67-year-old patient needing impella cp support due to a progressive decline. it was reported that there was oozing from the insertion site around 14 and venous sheath. the site was redressed but the oozing continued. the impella was slowly weaned and successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27640 | 1220648-2024-10252 | 19159394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19159394&pc=OZD | 02/05/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 02/05/2022 | 04/22/2024 |  |  | 10/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156241 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 02/05/2022 |  |  |  |  | The user facility reported a 64-year-old patient needing mechanical circulatory support. it was reported that the patient on impella cp had ventricular tachycardia (vt). the impella was inserted, the plan was to wean drips accordingly, the patient then went into vt twice. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27641 | 1220648-2024-10253 | 19159954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19159954&pc=OZD | 04/16/2022 | 09/30/2025 |  | 04/22/2024 |  |  | 04/16/2022 | 04/22/2024 |  |  | 03/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064264 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Suction Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/16/2022 |  |  |  |  | The user facility reported a 65-year-old patient needing mechanical circulatory support. it was reported that they encountered positioning issues. the pump got entangled in the mitral valve apparatus and had been unable to free it so the pump was left there on p2. they were then able to reposition the pump but encountered suction when having it on p4. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report in accordance with updated procedures. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 medical device problem code 2170 was erroneously entered on the initial manufacturer device report number.
27642 | 1220648-2024-10254 | 19160009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19160009&pc=OZD | 04/17/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 04/17/2022 | 04/22/2024 |  |  | 12/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171231 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The user facility reported a 69-year-old male on impella cp for mechanical circulatory support. it was reported that the patient had bleeding at the sheath. the impella cp was inserted with ease and the peel away sheath was removed. repositioning advanced and the patient began bleeding around repo sheath. angle match and forward suturing helped but did not resolve the issue. systemic heparin was started through the side port to keep patent. and femostop was placed as the groin site. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27643 | 1220648-2024-10255 | 19160047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19160047&pc=OZD | 04/19/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 04/19/2022 | 04/22/2024 |  |  | 02/28/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022045431 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/19/2022 |  |  |  |  | The user facility reported a 45-year-old patient on impella cp support due to stemi (st-segment elevation myocardial infarction). while on support, access site bleeding was noted. angle match dressing was placed with forward tension sutures. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27644 | 1220648-2024-10257 | 19160178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19160178&pc=OZD | 04/20/2022 | 05/13/2025 |  | 04/22/2024 |  |  | 04/20/2022 | 04/22/2024 |  |  | 07/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022109222 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was noted that the impella was withdrawn due to suspected leak by the cross clamp. following coronary artery bypass graft times four, and mechanical mitral valve placement, the impella was unable to re-advance into the left ventricle and was explanted in the operating room. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27645 | 1220648-2024-10256 | 19160241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19160241&pc=OZD | 04/20/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 04/20/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188330 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on impella cp support the patient developed bleeding at the impella cp insertion site. the impella was removed post procedure and closure achieved. the patient was restless and thrashing around resulting in the sutures popping, causing bleeding which required manual pressure to be held. the patient was given protamine for heparin reversal and femostop was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
27646 | 1220648-2024-10258 | 19160279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19160279&pc=OZD | 04/20/2022 | 04/01/2025 |  | 04/22/2024 |  |  | 04/20/2022 | 04/22/2024 |  |  | 01/31/2024 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188048 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, patient had vascular damage. when the impella pump was explanted, the patient experienced severe pain in right groin and an angiogram revealed dissection of the right femoral artery (rfa). a stent was successfully placed in rfa via rght radial artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27647 | 1220648-2024-10259 | 19160342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19160342&pc=OZD | 04/21/2022 | 04/22/2024 |  | 04/22/2024 |  |  | 04/21/2022 | 04/22/2024 |  |  | 07/31/2023 |  | 04/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121812 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 53-year-old patient on impella cp support that developed a thrombus in the left ventricle and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27648 | 1220648-2024-10105 | 19142266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142266&pc=OZD | 03/21/2024 | 05/07/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024435795 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 48-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support a reperfusion sheath was placed for better flow down the leg. that was not successful, so a fasciotomy was performed. the impella was removed and replaced. after the impella cp explant the patient had an above knee amputation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
27649 | 1220648-2024-10106 | 19142295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142295&pc=OZD | 03/26/2024 | 05/20/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451025 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 03/26/2024 |  |  |  |  | The user facility reported a 54 year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient developed access site bleeding while on support. the bleeding was resolved within ten minutes by use of a balloon tamponade and a perclose suture. the patient was transferred to the intensive care unit and remains stable. | Additional information was received. the user facility reported that while on impella cp support, the sheath had kinked upon implant. the pump was inserted and started without issue despite the kink and position was optimized. furthermore, the patient had nearly full loss of pulsatility with every balloon inflation with impella supported mean arterial pressure. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation has been completed for introducer damage - mechanical. the pump and introducer were not returned for investigation. the cause of the introducer damage issue was most likely an introducer sheath kink due to patient vessel size and condition. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.1 product problem was inadvertently left off the previously submitted report and was added. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site name and address section as previously entered incorrectly. h.6 code 4755 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | Section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.6 codes 1942, 2588 and 3038 were inadvertently omitted from the previously submitted report. code 2993 was reported incorrectly on the original submitted report. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation into the reported issues has been completed since the original report was submitted. the pump and introducer were not returned for investigation. the cause of the vascular damage issue was patient condition related with sheath kinking due to patient vessel size and artery being deep. the cause of the introducer damage issue was most likely an introducer sheath kink due to patient vessel size and condition.
27650 | 1220648-2024-10107 | 19142356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142356&pc=OZD | 03/28/2024 | 04/26/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED INC. |  |  | US |  | 1579132 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a patient of unknown demographics in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following the impella explant staff tried to pull out purge disc and it was discovered the blue disc holder was broke. the patient expired after the explant. there was no association between impella use and outcome. | [5 empty] | The automated impella controller was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. d4-updated model number. in accordance with updated procedures, section d2 has been revised from the initial submission.
27651 | 1220648-2024-10108 | 19142378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142378&pc=OZD | 03/28/2024 | 05/23/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024421988 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support the patient suffered an embolic stroke and expired. there was no association between impella use and outcome. patient's pre-procedural condition was critical. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore the investigation for the reported stroke has been completed., the root cause of the stroke could not be determined. the instructions for use referenced in the initial report is from ifu impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states), which now includes "(including ventricular perforation).¿ in accordance with updated procedures, section d6 has been revised from the initial submission.
27652 | 1220648-2024-10109 | 19142475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142475&pc=OZD | 05/16/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient's feet were pink and cool. the patient was taken to cath lab to reestablish fem-fem bypass. attempts were unsuccessful attempt to regain antegrade access. the physician felt that the leg had enough perfusion. the device was exchanged to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27653 | 1220648-2024-10110 | 19142522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142522&pc=OZD | 05/24/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 02/29/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194993 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2022 |  |  |  |  | The user facility reported a 74-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, distal pulses were absent and the patient was taken to the cardiac catheterization laboratory for intervention to open iliac artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27654 | 1220648-2024-10111 | 19142523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142523&pc=OZD | 05/04/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190424 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. the medical staff observed a large hematoma of the left groin extending into the scrotum. manual pressure was held. during the evening, the systemic heparin was discontinued due to access site bleeding. the patient was cannulated for extracorporeal membrane oxygenation with significant difficulty. the patient was started on venoarterial extracorporeal membrane oxygenation (va ecmo) and remained on va ecmo and impella support. cardia thoracic surgery (cts) added an additional suture to the impella sheath. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27655 | 1220648-2024-10112 | 19142624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142624&pc=OZD | 07/10/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 07/31/2022 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1499663 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced supraventricular tachycardia that was resolved by shock. the patient also experienced limb ischemia that was resolved by removal of the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the cause of the limb ischemia and vf/vt/af/at was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27656 | 1220648-2024-10113 | 19142665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142665&pc=OZD | 05/26/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 03/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191502 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced bleeding at the impella access site. due to the bleeding, the decision was made to prematurely explant impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27657 | 1220648-2024-10114 | 19142781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142781&pc=OZD | 07/26/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184713 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 40-year-old male undergoing cardiac surgery was attempted to be implanted with an impella device for mechanical circulatory support. it was reported that the device was unable to be placed due to the patient¿s anatomy. the physician implanted an intra-aortic balloon pump (iabp) instead. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27658 | 1220648-2024-10116 | 19142967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142967&pc=OZD | 06/05/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 02/29/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197407 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support the patient was treated for bleeding issues and a hematoma. it was reported the hematoma was present prior to impella insertion. due to the bleeding, heparin was discontinued and manual pressure was applied to achieve hemostasis. the patient was transfused with 1 unit of fresh frozen plasma, and 2 units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27659 | 1220648-2024-10117 | 19142991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142991&pc=OZD | 08/31/2022 | 05/23/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 05/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Limb Pain | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced left leg pain. the patient was given 80 mg of lasix before being weaned due to the patient¿s end-diastolic pressure. the impella was weaned down without issue and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27660 | 1220648-2024-10118 | 19143016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143016&pc=OZD | 04/23/2022 | 05/21/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165809 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/23/2022 |  |  |  |  | The user facility reported a 53-year-old patient needing mechanical circulatory support. it was reported while on support, low placement signal alarms occurred all night with p level set to p2. doctor explanted bedside and manual pressure was applied to access site for 35 minutes. distal pulses were not present after removal and ultrasound of right lower extremity (rle) revealed clot in the common femoral artery. surgical cutdown was performed, clot was removed, and pulses were present in rle. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27661 | 1220648-2024-10119 | 19143035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143035&pc=OZD | 09/28/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 06/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240360 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma during support. during explant it was reported that a single perclose was deployed with good hemostasis after 5 min of manual pressure. due to the hematoma, the physician employed a belt pressure femstop for precautionary measures. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27662 | 1220648-2024-10121 | 19143138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143138&pc=OZD | 09/28/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 06/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238368 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 78-year-old female on impella cp for high-risk percutaneous coronary intervention (hrpci) support. after several attempts to insert wire past occlusion, the physician was unsuccessful. every time the pump was advanced, the wire went into the coronaries. it would enter the subintimal space. the patient started to experience chest pain. the case was aborted, and patient was sent to the icu on impella cp support. the operating room (or) team readied for an urgent coronary artery bypass graft (cabg). | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27663 | 1220648-2024-10120 | 19143141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143141&pc=OZD | 05/01/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187512 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, there was access site oozing (bleeding) during implant noted by the groin, which was treated by re-suturing the site. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27664 | 1220648-2024-10122 | 19143175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143175&pc=OZD | 04/26/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Insufficient Flow or Under Infusion | Patient-Device Incompatibility | Positioning Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/26/2022 |  |  |  |  | The user facility reported a 77-year-old patient needing mechanical circulatory support. it was reported the impella was positioned adequately for the left ventricle; however, the patient was unable to tolerate the impella cp support due to the small left ventricle cavity and shape. the staff also noted the patient had intermittent suction events that caused loss of pulsatility. in addition, there was bleeding at the groin site, one unit of packed red blood cells was given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27665 | 1220648-2024-10123 | 19143296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143296&pc=OZD | 10/27/2023 | 05/01/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 09/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024394647 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | A notification was received long-term outcome and quality indicator impella registry that the patient endured ventricular tachycardia (vt) storm with a possible relationship to the impella 5.5 device. subject admission course was complicated by onset of ventricular tachycardia storm. the patient had torsades, no chest compressions, shocked, then had ventricular fibrillation. a procedure of stellate ganglion block was performed, ongoing oral amiodarone and mexiltine as needed throughout course. the patient was ultimately bridged to left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ minor and the vf/vt/af/at issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding cannot be determined since the product was not determined and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
27666 | 1220648-2024-10124 | 19143308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143308&pc=OZD | 04/26/2022 | 05/20/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 10/31/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158668 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/26/2022 |  |  |  |  | The user facility reported a 69-year-old patient was placed onto impella cp support due to cardiogenic shock. while on support, the pump advanced deep into the left ventricle and prolapsed. attempts were made to reposition, but they were unsuccessful. a new impella had to be placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27667 | 1220648-2024-10125 | 19143326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143326&pc=OZD | 05/12/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced lost pulsatility and the doctor placed a distal perfusion catheter for the fem-fem bypass. the patient received two units of red blood cells. towards the end of support, the patient had distal limb ischemia in the left lower extremity and the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27668 | 1220648-2024-10126 | 19143343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143343&pc=OZD | 05/01/2022 | 05/13/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 12/31/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170718 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced hemolysis. due to the hemolysis the physician weaned and explanted the impella after a successful mitral valve replacement was done. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27669 | 1220648-2024-10127 | 19143372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143372&pc=OZD | 04/27/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a 62-year-old patient needing mechanical circulatory support. it was the reported that the patient had limb ischemia noted by the complete distal occlusion due to the patient's poor femoral anatomy. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27670 | 1220648-2024-10128 | 19143523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143523&pc=OZD | 05/13/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 03/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191949 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was noted that the impella kinked during insertion due to anatomy in the patient's aortic arch, so the procedure was aborted and a second pump was inserted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27671 | 1220648-2024-10129 | 19143695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143695&pc=OZD | 05/03/2022 | 05/20/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161063 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient was taken to the cardiac catheterization laboratory for impella cp placement. the impella cp was placed in the right femoral artery and the pump got caught under the papillary muscle. the flows continued to decrease despite good positioning, an echocardiogram revealed pericardial effusion. the physician performed a pericardiocentesis procedure, the flows returned to 3.51/m and the patients pressure stabilized. the impella was weaned and explanted to reverse heparinization and close the perforation, the sheath was pulled, and manual pressure held for closure with no bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27672 | 1220648-2024-10130 | 19143851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19143851&pc=OZD | 03/16/2024 | 05/07/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 06/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023243424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient's pulses were unable to be found. blood-flow lower extremity (ble) doppler was ordered; however, the patient decompensated and the family withdrew care. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. this report is being filed as part of a retrospective review of historical records. corrected information provided in d6a and d6b.
27673 | 1220648-2024-10132 | 19144040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144040&pc=OZD | 04/27/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 06/30/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107039 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a 75-year-old patient on impella cp had delivery issues. the impella was having issues with the pump was unable to pass beyond left iliac aortic junction and was noted to be caused by previous iliac aortic stent. md tried to advance impella beyond left iliac twice to no avail, unable to pass through vasculature. the decision was made to remove pump and intra-aortic balloon pump (iabp) support was done. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27674 | 1220648-2024-10131 | 19144044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144044&pc=OZD | 06/11/2022 | 04/03/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181261 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding and systemic heparin was reduced. it was also reported the patient experienced hemolysis while on support through device explant. the impella position was verified and optimized by echocardiography, but hemolysis did not resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27675 | 1220648-2024-10133 | 19144136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144136&pc=OZD | 04/27/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188375 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The complainant reported a patient was taken to the cardiac catheterization laboratory (ccl) for impella placement. it was reported the physician made multiple attempts were made to cross the aorta, to place the impella, due to patient¿s anatomy it was unable to pass through the vasculature. implantation was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27676 | 1220648-2024-10134 | 19144162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144162&pc=OZD | 04/27/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 12/31/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174883 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was report while on support, the patient experienced right groin oozing at impella site. the medical team was advised to withheld systemic heparin. additionally, heparin was continued when appropriate. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27677 | 1220648-2024-10135 | 19144191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144191&pc=OZD | 04/09/2024 | 05/21/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467744 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced slight oozing and hematoma at the right groin. femostop was placed over the site and two units of packed red blood cells were administered. support with the implanted pump continued following the intervention. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27678 | 1220648-2024-10136 | 19144323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144323&pc=OZD | 05/03/2022 | 04/07/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 07/31/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118199 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, some hematuria was noted and a femoral thrombectomy was performed post-explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27679 | 1220648-2024-10137 | 19144362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144362&pc=OZD | 04/09/2024 | 05/28/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455133 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed right groin oozing. as a result, heparin was withheld. the bleeding was stopped with manual pressure and femostop. the patient remained on impella cp support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleeding issue was not determined since the product was not returned and not much information was provided in the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has been completed since the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27680 | 1220648-2024-10139 | 19144463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144463&pc=OZD | 04/10/2024 | 06/06/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434299 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 64-year-old male in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after one day of support, the impella cp came out of appropriate left ventricle positioning. the malposition caused the team to explant and replace the impella cp with a new pump. the second impella cp was placed without harm or injury and the patient is stable and remains on support to date. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the product was not returned for investigation. as per the provided clinical details, the pump migrated past the valve into the left ventricle and it was pulled back out. it could not be repositioned at the bedside. the cause of the positioning issue was most likely use related, based on given clinical details. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 medical device problem code was revised in accordance with updated procedures. new code has been added to medical device problem codes. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has been completed since the previously submitted report.
27681 | 1220648-2024-10140 | 19144466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144466&pc=OZD | 05/04/2022 | 04/07/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 12/31/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181011 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. it was noted that the patient presented with a tortuous iliac anatomy and despite multiple attempts to advance the impella, the long sheath would kink. the motor housing would not advance, so the pump was removed and a manta closure was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27682 | 1220648-2024-10141 | 19144514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144514&pc=OZD | 07/06/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the impella was determined to be too shallow, sitting just 1.5 cm across the aortic valve and the impella was repositioned successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27683 | 1220648-2024-10143 | 19144607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144607&pc=OZD | 10/27/2022 | 05/08/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 07/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254349 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support for urgent revascularization. it was reported that the impella cp was inserted and started but explanted shortly after due to growing hematoma concerns. during the (pci) sheath and impella sheath exchange hematoma was revealed, patient became hypovolemic and unstable. hematoma expression, contralateral 10x80 percutaneous transluminal (coronary) angioplasty balloon inflated to control bleed. percutaneous stent graft placed but still required transfer to vascular operating room for right common femoral artery repair. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vascular damage has been completed. the root cause of the vascular damage issue was not determined since product was not returned and not much information is provided in the clinical description. this report is being filed as part of a retrospective review of historical records. corrected information provided in d6a and d6b.
27684 | 1220648-2024-10144 | 19144687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144687&pc=OZD | 07/23/2023 | 05/13/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339439 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that a patient who was being transitioned from an impella cp to an impella 5.5 developed hemolysis. the escalation procedure was successful. | [5 empty] | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. however, data logs were reviewed finding several pump in aorta alarms that resolved quickly. therefore, the root cause of the access site bleed and hemolysis could not be determined. d4-updated model number added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
27685 | 1220648-2024-10145 | 19144696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144696&pc=OZD | 09/10/2022 | 04/07/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 06/25/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236205 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 27-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the wire was dissected through the left anterior descending artery (lad) and the lad was found to be full of thrombus. the patient heart stopped and was shocked nine times. access was quickly gained to the right femoral artery (rfa) and upsized to an impella 14 fr sheath. the impella insertion was ultimately aborted and the intra-aortic balloon pump (iabp) was placed due to the large thrombus noted in the left ventricle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27686 | 1220648-2024-10146 | 19144720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144720&pc=OZD | 05/06/2022 | 10/20/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 12/31/2023 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165936 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma at the impella cp site. the patient remained on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27687 | 1220648-2024-10142 | 19144823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144823&pc=OZD | 04/29/2022 | 03/11/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was advanced, and flows would not increase above 1.4l/min. a kink was noticed in the impella catheter under the motor housing near outlet cage. the impella was then explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27688 | 1220648-2024-10147 | 19144906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144906&pc=OZD | 02/26/2024 | 05/17/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 10/31/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024419235 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. a notification was received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was reported that the patient on impella 5.5 support endured ventricular tachycardia (vt) with a "possible" relationship to the impella device used. vt requiring cardioversion. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at issue has been completed since the original report was submitted. the device was not returned for investigation. it was reported that the patient endured ventricular tachycardia (vt) which required a cardioversion. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. impella rp system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27689 | 1220648-2024-10148 | 19144947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19144947&pc=OZD | 07/06/2022 | 10/29/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1637920 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 74-year-old patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp had been initiated in auto, the placement signal did not appear, so aic was replaced. the placement signal appeared and no other complications observed. procedural rn stated that the white connector cable was probably not pushed hard enough into the first aic because the second aic felt a ¿snap¿ when connected. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27690 | 1220648-2024-10149 | 19145044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145044&pc=OZD | 04/28/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191695 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/28/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, staff noted the pump looked a little deep on the fluoroscopy and was pulled back. the impella sprung back and buried the pigtail into the papillary muscle. the doctor attempted to push the pump forward and the pump buckled. the pump was removed, rewired, and reinserted into optimal position without difficulty. additionally, it was reported that there was a placement signal low alarm. an echocardiogram was called for to check the position and the impella was pulled back 3 cm to the 80 cm marking. the patient was given diuretics and one unit of blood. they were getting suction alarms above p7. lastly, the staff noted the patient lost all pulsatility twice during a stent procedure. the first event, the patient quickly recovered after. the second event, the patient's map dropped 5 points. the impella sheath was removed and replaced. there were limited pulses available post sheath exchange with a faint pulse noted in popliteal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27691 | 1220648-2024-10150 | 19145048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145048&pc=OZD | 05/06/2022 | 04/04/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 01/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182448 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the pump was wedged under papillary muscle. they were unable to obtain proper positioning, so the impella was removed and successfully replaced with proper position. additionally, the patient experienced atrial fibrillation and supraventricular tachycardia, which were resolved with cardioversion and antiarrhythmics. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27692 | 1220648-2024-10151 | 19145054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145054&pc=OZD | 08/12/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 05/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023222820 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Myocarditis | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The complainant reported a patient with de-novo cardiomyopathy and non-st elevated myocardial infarction was was implanted with an impella cp for mechanical circulatory support. during support, the patient showed signs of sepsis, and a biopsy showed signs of infection (myocarditis). the patient remained on impella cp support and was scheduled to be escalated to an impella 5.5 the following week with no information on treatment or intervention provided to patient for sepsis. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the device passed all post-sterile inspection checks. the cause of the sepsis/ infection was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27693 | 1220648-2024-10152 | 19145192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145192&pc=OZD | 03/30/2024 | 05/16/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 09/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446667 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 26-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp delivery failed and kinked and was removed due to patient¿s anatomy. the impella cp was removed, and the patient went to surgery where extracorporeal membrane oxygenation was placed with plans for direct aortic placement of the impella 5.5 pump. the impella 5.5 was not delivered as the patient passed in the operating room. per the medical safety officer review, there is no association between impella use and outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since the patient had small vessels and small aortic arch. in accordance with updated procedures, section d6 has been revised from the initial submission.
27694 | 1220648-2024-10153 | 19145206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145206&pc=OZD | 11/19/2022 | 05/16/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 05/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219531 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was being placed on an impella cp for mechanical circulatory support. it was reported that the medical staff was unable to deliver the impella pump to the left ventricle due to the patient's vascular anatomy and the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27695 | 1220648-2024-10154 | 19145219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145219&pc=OZD | 10/04/2022 | 05/07/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218095 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing at the impella cp insertion site (groin). the nurse turned off the heparin drip until resolved. heparin was later started in the purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was determined to be anticoagulation management as hemostasis was achieved after heparin was stopped as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
27696 | 1220648-2024-10155 | 19145226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145226&pc=OZD | 03/22/2024 | 05/09/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 05/31/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed¿s long-term outcome and quality indicator impella registry reported a 39-year-old male patient that endured acute renal insufficiency with a "possible" relationship to the impella cp device used. the patient with a baseline creatinine around 1.5. creatinine started to rise the day after the impella procedure and continued to rise throughout hospitalization. peak creatinine of 9.12 on support. the patient was started on dialysis. tunneled dialysis catheter was placed on 11-oct-2023 and removed on 30-oct-2023 when renal function normalized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported renal failure has been completed. the product was not returned for investigation. data logs were review and no abnormalities were observed. the cause of the renal failure issue was not determined since product was not returned and due to insufficient clinical information. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27697 | 1220648-2024-10156 | 19145249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145249&pc=OZD | 03/21/2024 | 05/08/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 05/31/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365564 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 54-year-old male patient developed hemolysis during impella cp support. the hemolysis was identified via lab values and subsequently led to escalation to impella 5.5 support. the hemolysis resolved following 5.5 implant and no further intervention was required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
27698 | 1220648-2024-10157 | 19145384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145384&pc=OZD | 03/26/2024 | 08/14/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1163688 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller (aic) had a system self-check error leading to another device being used. there was no reported patient harm. during product maintenance, the failure was unable to be reproduced. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation of aic - system self-check error has been completed. the data logs confirmed that multiple unexpected controller shutdowns had occurred. the system self-check error could not be reproduced. the console booted up correctly. the unexpected shutdown/reboot issue could not be reproduced. the console was physically inspected, and no physical defects were identified. the cause of the issue was most likely a defective component. f6/f8-should have been left blank on manufacturer device report 1220648-2024-10157.
27699 | 1220648-2024-10158 | 19145542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145542&pc=OZD | 08/04/2022 | 04/19/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223938 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/21/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient possibly had a stroke. washout and explant noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27700 | 1220648-2024-10159 | 19145631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145631&pc=OZD | 11/09/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 06/03/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233892 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was being placed on impella cp for mechanical circulatory support. the medical staff attempted to insert the impella, but caused vascular damage with the guidewire and aborted. the vascular damage was managed with a balloon to achieve hemostasis. the medical team considered impella insertion again, but the blood vessel diameter was assessed to be about 5 millimeters and was determined to be too small to attempt another insertion. the impella implant procedure was aborted and an intra-aortic balloon pump was implanted instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27701 | 1220648-2024-10160 | 19145942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145942&pc=OZD | 04/06/2024 | 05/13/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 08/31/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443760 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 37-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding from the access site. a femstop was placed and two units of blood products were provided to the patient. the clinical representative did not know why the patient received the blood units. they stated that the patient was ill and it could have been for multiple reasons. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27702 | 1220648-2024-10161 | 19145961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19145961&pc=OZD | 10/07/2022 | 03/31/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 07/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support the pulses were absent in the patient's lower extremities. the medical team made the decision to escalate the patient to an impella 5.5 for increased support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows : ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27703 | 1220648-2024-10162 | 19146574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19146574&pc=OZD | 08/21/2022 | 03/10/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023218089 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, low purge flow and high purge pressure occurred when the patient and the catheter was moved, which caused the in aorta alert. the physician repositioned the catheter with the echo and pushed the catheter across the valve and into the left ventricle only 2 cm to correct the catheter position. however, a decrease in the purge flow still occurred even with a change of the cassette and purge tubing. it was decided to start tissue plasminogen activator for purge solution replacement which would run until empty. eventually the impella cp was explanted and the impella 5.5 was to follow post coronary artery bypass graft. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27704 | 1220648-2024-10163 | 19146909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19146909&pc=OZD | 11/14/2022 | 05/16/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207284 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was bein implanted with an impella cp device for mechanical circulatory support and encountered resistance during removal of the 0.018 guide wire and it became difficult to remove the peel-away sheet without forcing it. as a result, the introducer was removed surgically by the vascular surgeon and no bleeding was found after the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27705 | 1220648-2024-10164 | 19147327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19147327&pc=OZD | 11/22/2022 | 05/16/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 08/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253905 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. and the procedure had to be aborted due to the patient¿s vascular anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27706 | 1220648-2024-10165 | 19147549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19147549&pc=PYX | 04/11/2024 | 05/30/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 07/31/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024374057 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Frayed | Patient-Device Incompatibility | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported an 83-year-old female with a right heart failure was to be implanted with an impella rp device for mechanical circulatory support post cardiac surgery. after the sheath and dilator were inserted, the dilator put pressure on the sheath causing it to tear prematurely inside the body. the sheath was not fully inserted. there was no harm to the patient reported. the decision was made to remove the sheath and replace. once the dilator and sheath exchange was performed, placement of the impella rp was then successful. | [5 empty] | The investigation for the reported introducer damage has been completed. the 23 fr introducer sheath was returned with dilator. the returned introducer was observed to be split at the tip per visual inspection. no other abnormalities were found. the cause of the introducer damage issue was a sheath tip split due to patient tortuous vessels. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27707 | 1220648-2024-10166 | 19149595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19149595&pc=OZD | 06/15/2023 | 07/26/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 04/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024339383 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Sepsis | Respiratory Failure | [9 empty] | Required Intervention | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/21/2024 |  |  |  |  | A notification was received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that the patient on impella 5.5 support endured gl bleed, sepsis, and a suspicion of hemolysis with a "possible" relationship to the impella device used. as treatment, 1 unit rbc given and the patient was treated with antibiotics throughout his stay. despite ldh of 630 with bilirubin trend increasing, the impella's speed was adjusted and the medical team "decided patient does not have hemolysis." | Additional information received that patient experiment respiratory failure and was hospitalized to address the issue. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Additional information received from the clinical study coordinator indicated that the event is no longer related to the impella device. no further reports will be forthcoming. | Instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27708 | 1220648-2024-09114 | 19149599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19149599&pc=OZD | 07/29/2022 | 03/10/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 03/31/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023214379 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a 44-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device that the device was previously removed for mechanical circulatory support. it was decided to re-implant the device, due to the left ventricular end-diastolic pressure. the introducer was not removed, due to the failed preclose and the failed antegrade stick that was just below the sheath which required femstop to be applied around the groin site. a review of an x-ray showed the sheath kinked and the physician was advised on the risk of breakage, due to the peel away nature of the sheath. therefore, the physician decided to leave the sheath in place. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding and positioning issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
27709 | 1220648-2024-09841 | 19149600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19149600&pc=OZD | 04/09/2024 | 06/07/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2024 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024294897 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 72-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during the procedure of impella cp the sheath split while trying to remove sheath out of the artery. the patient developed bleeding and had hematoma. reposition unit forward tension sutured and angled match dressing applied. the physician held manual pressure and expressed some of the hematoma. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the access site bleed and introducer damage has been completed. the impella nor the introducer was returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and the clinical information provided was inconclusive. per the clinical information provided, the root cause of the introducer damage - mechanical issue was an introducer sheath split along the scoreline. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. added code 925 as it was inadvertently left off the previously submitted manufacture device report. code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
27710 | 1220648-2024-09872 | 19149601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19149601&pc=OZD | 04/10/2024 | 03/06/2025 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  |  |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1668177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported the automated impella controller (aic) was stuck in the purge menu. the biomedical engineer could not replicate the issue and noticed sticky residue on the screen. the residue was cleaned off. no patient involvement has been reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the console logs from the reported day of event are consistent with complaint and show that after the last purge cassette change, there was no purge pressure build-up, and console registered purge cassette being removed (or dislodged) followed by ¿piston blocked¿ error, and subsequent pud reset which resulted in momentary loss of pud communication. after pump and purge cassette were transferred to a backup console, the purge issue did not reproduce and support continued. during the device analysis the console was and the reported issue was not reproduced. console¿s purge system operated within specification. the root cause of the issue was most likely use related. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report in accordance with updated procedures. e2 revised on manufacturer device report in accordance with updated procedures. f6/f8-should have been left blank on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was revised as it was not included on initial manufacturer device report.
27711 | 1220648-2024-09839 | 19149602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19149602&pc=OZD | 04/09/2024 | 05/15/2024 |  | 04/19/2024 |  |  |  | 04/19/2024 |  |  | 11/30/2025 |  | 04/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC.ABIOMED, INC.ABIOMED, INC.ABIOMED, IN |  |  | US |  | 2024430144 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 55 year old male patient with coronary artery bypass graft was electively implanted with an impella device for mechanical circulatory support. the patient developed access site oozing and hematoma. the patient had an axillary washout performed. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-updated model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27712 | 1220648-2024-10003 | 19130228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130228&pc=OZD | 05/23/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/23/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179192 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. insertion of 14fr sheath was very smooth without complication at left common femoral artery (lcfa). however, aortic valve was found to be severely calcified with stenosis. after several trials of cannulation and ultrasound imaging, no wire successfully crossed the valve and reached the lv (left ventricle). physician changed support to intra-aortic balloon pump (iabp). | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27713 | 1220648-2024-10004 | 19130243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130243&pc=OZD | 05/26/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/26/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The complainant reported a patient with high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient experienced multiple runs of ventricular tachycardia (vt) during and after the case requiring defibrillation at 200 joules. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | To make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
27714 | 1220648-2024-10006 | 19130245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130245&pc=OZD | 05/07/2022 | 04/17/2024 |  | 04/18/2024 |  |  | 05/07/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194876 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/07/2022 |  |  |  |  | The user facility reported a 63-year-old patient was on impella cp support and some oozing (bleeding) was noted at the patient¿s groin. perclose was tightened with sutures, angioseal applied, and the groin dressing was reported as clean dry and intact. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27715 | 1220648-2024-10007 | 19130270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130270&pc=OZD | 06/04/2022 | 04/03/2025 |  | 04/18/2024 |  |  | 06/04/2022 | 04/18/2024 |  |  | 03/25/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197476 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, there were suction alarms on the device. impella flow was reduced over a few hours from p6 to p2. position was verified and optimized with transthoracic echo (tte) but there was no change. no additional information was provided for the reported event. | [5 empty] | Additional information added to section d6a/d6b implant and explant dates. the investigation for the reported low pump flow has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27716 | 1220648-2024-10005 | 19130273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130273&pc=OZD | 05/26/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/26/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195188 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient developed infection/sepsis. no additional information related to the resolution of the infection/sepsis was provided. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
27717 | 1220648-2024-10010 | 19130316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130316&pc=OZD | 05/07/2022 | 04/17/2024 |  | 04/18/2024 |  |  | 05/07/2022 | 04/18/2024 |  |  | 10/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151361 |  | 0048-0024-HP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/07/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and that the patient encountered hemolysis during support. the treatment was unknown. the pump was eventually explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27718 | 1220648-2024-10008 | 19130334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130334&pc=OZD | 06/07/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 06/07/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191988 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hematoma. no additional information related to the resolution of hematoma was provided. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27719 | 1220648-2024-10009 | 19130378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130378&pc=OZD | 06/12/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 06/12/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192734 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella pump had delivery issues due to a kink in the aortic arch and bleeding at the groin access site was noted. no additional information was received for the reported issues. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27720 | 1220648-2024-10012 | 19130484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19130484&pc=OZD | 05/03/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/03/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184926 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, hematoma was noted. it was expressed out and the pump was explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27721 | 1220648-2024-10014 | 19131126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19131126&pc=OZD | 10/01/2022 | 05/20/2025 |  | 04/18/2024 |  |  | 10/01/2022 | 04/18/2024 |  |  | 07/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251487 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2022 |  |  |  |  | The complainant reported a 54 year old male patient on impella cp support. during the morning, the patient was active in their hospital bed and the device jumped back into the aorta. the impella cp was removed at the patient's bedside by medical staff and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27722 | 1220648-2024-10011 | 19131419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19131419&pc=OZD | 06/13/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 06/13/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181344 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/13/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the peel away sheath was left inside the patient after the procedure due to vascular injury. the sheath was being used as a vascular bleeding prevention. additionally, it was reported that the patient developed distal limb ischemia on the insertion site due to peripheral arterial vascular disease (pavd). the limb was noted as being cold. impella cp was weaned and explanted. no further information was provided. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ this report is being filed as part of a retrospective review of historical records.
27723 | 1220648-2024-10013 | 19131441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19131441&pc=OZD | 05/04/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 02/23/2022 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176507 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was slightly oozing overnight from insertion site but has slowed now. oscor sheath removed and pump secured with forward tension sutures/angle matched, preclosed x2 sutures protected and sterile. femoral swan in place. plan to keep impella in to rest the heart and to be slowly weaned and explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27724 | 1220648-2024-10015 | 19132238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132238&pc=OZD | 05/04/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176372 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, they encountered delivery issues due to patient anatomy. the decision was made to explant the pump. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27725 | 1220648-2024-10016 | 19132293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132293&pc=OZD | 05/05/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/05/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178374 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was noted there was a scant ooze at the right femoral access site and the patient became bradycardic. due to continued dropping heart rate, no pulsatility, and flows slightly lower on the impella, a low dose epinephrine drip was started. the patient presented with bilateral lower limb ischemia, rewarming patient to assess neuro status. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27726 | 1220648-2024-10017 | 19132316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132316&pc=OZD | 05/05/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/05/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183697 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during post closure of the impella 2 attempted, however due to calcified iliac, post closure was unsuccessful. they were unable to control bleeding. covered stent was placed and hemostasis was achieved. the patient remained on impella support and was later successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27727 | 1220648-2024-10018 | 19132353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132353&pc=OZD | 05/05/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/05/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient was found to be heparin induced thrombocytopenia positive. it was also noted that labs hemolyzed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27728 | 1220648-2024-10019 | 19132372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132372&pc=OZD | 05/05/2022 | 05/22/2025 |  | 04/18/2024 |  |  | 05/05/2022 | 04/18/2024 |  |  | 11/30/2022 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021020157 |  | 0048-0032 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was reported that after about 20 minutes of support, the impella abruptly stopped due to clot ingestion. reportedly several attempts were made to restart without success and the impella was explanted. the patient was transferred to another hospital for further support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27729 | 1220648-2024-10020 | 19132483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132483&pc=OZD | 10/20/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 10/20/2022 | 04/18/2024 |  |  | 07/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023249667 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella 5.5 as a bridge to heart transplant. it was reported that while the patient was on impella cp support, the patient experienced bleeding and oozing at the impella access site at the groin and heparin and cangrelor were withheld to resolve the bleeding. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records.
27730 | 1220648-2024-10021 | 19132498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132498&pc=OZD | 05/06/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193559 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Material Twisted/Bent | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support complication arose due to kink in the cannula. sheath was removed with pump exchanged for 14fr cook sheath. it was also noted that patient required increased medical management since explant and was possibly septic. | [5 empty] | When diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27731 | 1220648-2024-10022 | 19132549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132549&pc=OZD | 11/08/2022 | 05/16/2025 |  | 04/18/2024 |  |  | 11/08/2022 | 04/18/2024 |  |  | 04/30/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212355 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced a kinked cannula that required a new impella to be used. additionally, the patient experienced access site bleeding that resolved with deploying a long peelaway sleeve and the patient was transfused with blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27732 | 1220648-2024-10023 | 19132643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132643&pc=OZD | 10/03/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 10/03/2022 | 04/18/2024 |  |  | 10/31/2022 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021003968 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The user facility reported a 72-year-old male on impella cp support due to st-elevated myocardial infarction / cardiogenic shock. while on support, the patient experienced access site bleeding that was addressed by stat-locking down the blue suture hub and redressing, which decreased the bleeding but it was still present. no further resolution was reported. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
27733 | 1220648-2024-10024 | 19132668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132668&pc=OZD | 05/06/2023 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2023 | 04/18/2024 |  |  | 08/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022117272 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2023 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction issues and was reported to moved out of position. when checked under fluoroscopy, it was found that the pigtail was just under the aortic valve. staff attempted to reinsert it many times, but were unsuccessful. the pump was then replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27734 | 1220648-2024-10025 | 19132908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132908&pc=OZD | 05/06/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, when getting ready to transfer patient, bleeding was noted from site. manual pressure applied for 20 minutes but site continued to steadily ooze. the decision was made to wean and explant impella since the patient was stabilized. the impella was successfully weaned, and site closed using proglide x2 and dry closure from contralateral access with good angiographic result. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27735 | 1220648-2024-10026 | 19132936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132936&pc=OZD | 04/13/2024 | 06/05/2024 |  | 04/18/2024 |  |  | 04/13/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453921 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that upon attempting insertion of the impella 5.5, the catheter was very difficult to advance. the physician was not able to torque once inside the left ventricle. the decision was to remove and attempt to rewire and try again. upon removal, a significant kink was observed behind the motor housing. a second impella 5.5 was implanted successfully. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to use issue since the physician was not able to advance and torque once the pump was in the lv. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings states ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added- d9 device return date in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27736 | 1220648-2024-10027 | 19132992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19132992&pc=OZD | 05/01/2022 | 05/19/2025 |  | 04/18/2024 |  |  | 05/01/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193985 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  |  | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The user facility reported a 32-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support the impella device was accidently pulled back past the valve and was unable to be advanced back into the ventricle. a new impella cp was placed without issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27737 | 1220648-2024-10028 | 19133000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133000&pc=OZD | 05/06/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 05/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022086945 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, there was severe hemolysis noted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27738 | 1220648-2024-10029 | 19133126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133126&pc=OZD | 05/08/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was stable but labile with any changes, had diminished pulses in the right foot. per vascular evaluation, it was noted that no duct-dependent pulmonary circulation was needed, however was unable to be above p4 without suction, discussed central venous pressure 7-8 and tachycardic. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27739 | 1220648-2024-10030 | 19133304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133304&pc=OZD | 05/02/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/02/2022 | 04/18/2024 |  |  | 10/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149039 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient was moved to an intensive critical unit (icu) bed where access site bleeding was noticed. manual pressure was applied, and fem stop was placed. the physician was informed that the sheath was not properly sutured down and the sheath was able to move. the physician decided to wean and explant the impella due to the bleeding and improperly sutured sheath. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27740 | 1220648-2024-10031 | 19133314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133314&pc=OZD | 05/08/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184868 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a patient had thrombus and was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed and distal left arterial descending clotted with balloon inflation and required thrombectomy. the patient still had significant lesion to the left circumflex and distal left anterior descending artery which will be addressed. impella successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27741 | 1220648-2024-10032 | 19133371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133371&pc=OZD | 06/26/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 06/26/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support the patient developed hemolysis. it was noted that the patient had hematuria and hemolyzed labs. repositioning was attempted to resolve the hemolysis but was not successful. additionally, the p level was unable to be lowered without the patient becoming unstable. it is unknown how the hemolysis was resolved; however, the pump was discontinued after it was noted that the patient had a brain bleed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27742 | 1220648-2024-10033 | 19133383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133383&pc=OZD | 05/06/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while attempting to place the impella cp, due to groin complication while accessing sas the wire was unable to track up the artery and kept buckling this caused a vascular perforation, vascular was notified, and the patient was taken to or for repair of the common femoral artery. the case was ended at that time and the impella was removed, perclose x2 was achieved and manual pressure was applied. patient was stable during this complication and remained stable. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
27743 | 1220648-2024-10034 | 19133436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133436&pc=OZD | 05/02/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/02/2022 | 04/18/2024 |  |  | 01/01/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192772 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, patient developed oozing at the impella access site. the preclose suture was tightened and forward tension dressing applied to resolve the bleeding. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27744 | 1220648-2024-10035 | 19133483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133483&pc=OZD | 05/06/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186259 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. insertion was aborted due to peripheral vascular complications with blood products being administered. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
27745 | 1220648-2024-10036 | 19133489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133489&pc=OZD | 05/08/2022 | 05/20/2025 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 11/30/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Renal Failure | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support due to dilated cardiomyopathy. while on support, the patient was heparin-induced thrombocytopenia positive and was on bivalirudin and no other anticoagulation being administered. the patient also became anuric overnight and required dialysis with -25ml/hr. it was noted that the pump either dives in deep or very shallow even with clockwise and counterclockwise rotation. the doctor attempted to reposition the impella cp. the impella cp was explanted and escalated to impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27746 | 1220648-2024-10037 | 19133514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133514&pc=OZD | 06/27/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 06/27/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199335 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported the impella cp was placed to help with unloading. the patient went into cardiogenic shock post myocardial infarction. extra corporeal life support was discontinued. the device suddenly stopped at high motor current during care and was unable to restart. the device was explanted, and a sizable thrombus was found at the inlet area. no further information was provided. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. h1: type of reportable event. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27747 | 1220648-2024-10038 | 19133567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133567&pc=OZD | 05/06/2022 | 04/01/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170828 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, staff noted the patient was oozing at right groin site. the patient received packed red blood cells, cryo, platelets and fresh frozen plasma. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27748 | 1220648-2024-10041 | 19133844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133844&pc=OZD | 06/20/2022 | 08/05/2025 |  | 04/18/2024 |  |  | 07/11/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191886 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced access site bleeding with no information reported regarding the resolution. additionally, the pump was positioned too deep; no attempt was made to reposition due to planned exchange for impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10041. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-10041. d6a and d6b revised from the initial manufacturer device report number 1220648-2024-10041 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10041 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10041 in accordance with updated procedures. h6 health effect - impact code 4648 and medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2024-10041. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10041. h6 component code revised from the on the initial manufacturer device report number 1220648-2024-10041 in accordance with updated procedures.
27749 | 1220648-2024-10042 | 19133964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19133964&pc=OZD | 05/09/2022 | 04/17/2025 |  | 04/18/2024 |  |  | 05/09/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. implantation of the impella cp access through the left femoral artery was attempted, and difficulty was encountered when advancing the guidewire. in the interest of time and patient's stability, the decision was made to remove the pump, the impella cp was rapidly weaned. the physician inserted the extra-corporeal membrane oxygenation cannula instead through the same pump's arterial access site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27750 | 1220648-2024-10043 | 19134015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134015&pc=OZD | 04/13/2024 | 07/10/2024 |  | 04/18/2024 |  |  | 04/15/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471889 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. | The user facility reported a 68-year-old female post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient developed hemolysis. it was noted that pump flows at the time of hemolysis were insufficient and a potential clot was suspected at the inflow cage. the pump was replaced as a result of the noted issues and a white clot was confirmed inside the cannula of the pump. despite resolution of the hemolysis following pump replacement, it was noted that the patient's renal function deteriorated as a result of the prior condition. continuous hemodiafiltration and haptoglobin were administered to treat the renal deterioration and the patient was expected to recover in the coming weeks. |  |  |  |  | B.5 additional information was received and abiomed's medical safety officer review was completed. the investigation for the hemolysis, low pump flow, and renal failure has been completed. neither the impella device nor data logs were returned for evaluation. as per clinical observation, there was a clear white blood clot inside the cannula of the impella 5.5 pump when it was removed. after replacement, the hemolysis subsided, and the flow was maintained. the root cause of low flow was most likely biomaterial based on clinical information. the root cause of hemolysis and renal failure was unable to be determined as limited clinical details were provided and no product was returned for review. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a implant date, d.6b explant date, g.1 reporting contact fax number and h.10 additional manufacturer narrative instructions for use were revised in accordance with updated procedures since the initial manufacturer device report was submitted. f.6 and f.8 dates were reported on manufacturer device report number and should not have been. h.6 code 2954 was reported incorrectly on manufacturer device report . | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
27751 | 1220648-2024-10044 | 19134097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134097&pc=OZD | 07/17/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 07/17/2022 | 04/18/2024 |  |  | 03/13/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213472 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 50-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient was placed on impella cp support and the patient developed significant oozing around the sheath at the impella groin access site. dressing created tenting, so an in-service was performed to redress the site. after approximately 1 hour of being on site, the groin site oozing started to slow down. bicarb was in the purge solution but heparin was not. | [5 empty] | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27752 | 1220648-2024-10045 | 19134147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134147&pc=OZD | 05/09/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/09/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196745 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and that a hematoma at the impella cp access site was observed, which was managed with manual pressure and withholding heparin. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27753 | 1220648-2024-10046 | 19134186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134186&pc=OZD | 05/09/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/09/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191058 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient had hemolysis. the patient was found to have hemolysis from the insertion to removal the of the pump. the pump was explanted. no further information was reported. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27754 | 1220648-2024-10047 | 19134210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134210&pc=OZD | 04/14/2024 | 06/12/2024 |  | 04/18/2024 |  |  | 04/15/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025476195 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The user facility reported a 71 year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. a left ventricle perforation occurred, and a heart patch closure surgery was performed. the patient survived the event. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the ventricle perforation was not determined as insufficient clinical information was provided. d6a, d6b. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
27755 | 1220648-2024-10048 | 19134247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134247&pc=OZD | 09/14/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 09/14/2022 | 04/18/2024 |  |  | 06/30/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236871 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 74-year-old female on impella cp for mechanical circulatory support. it was reported the patient experienced retroperitoneal (rp) bleed. the patient received 6 units packed red blood cells (prbc) due to left femoral rp bleed. hemoglobin (hgb) remained in low 7s. purge fluid was switched to bicarbonate. | [5 empty] | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27756 | 1220648-2024-10049 | 19134252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134252&pc=OZD | 05/02/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/02/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding from groin was noted. repo sheath and catheter were re-sutured. additionally, the impella cp oriented towards mitral valve and pigtail appeared to be under papillary muscle. patient went experienced episodes of atrial fibrillation with occasional non-sustained ventricular tachycardia (vt). the patient received an amiodarone infusion to stabilize the arrhythmias. | [5 empty] | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. .
27757 | 1220648-2024-10050 | 19134406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134406&pc=OZD | 03/21/2024 | 05/06/2024 |  | 04/18/2024 |  |  | 03/21/2024 | 04/18/2024 |  |  | 02/28/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325243 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a patient of unknown age and gender with a high risk of percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. the impella device could not be implanted due to the patient¿s anatomy. resistance from the patient¿s vasculature was noted. the implant procedure was aborted and no patient harm has been reported. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of delivery issue is determined to be patient anatomy because the pump was unable to pass through the vasculature and case was aborted. potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". "to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis." "to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels." physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27758 | 1220648-2024-10051 | 19134458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134458&pc=OZD | 05/06/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174655 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, continuous bleeding was noted circumferentially from puncture site. manual pressure initially placed over groin above repositioning sheath. hemostasis was difficult to maintain without significant manual pressure. ultimately, the decision was made to remove impella. a dry closure balloon was placed with fem stop. patient's vitals remained stable throughout. | [5 empty] | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27759 | 1220648-2024-10053 | 19134526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134526&pc=OZD | 05/09/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/09/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023193192 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced groin site oozing that was managed overnight with manual pressure, redressing to match the angle stick, and discontinuing the integrellin drip to help control the bleeding. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27760 | 1220648-2024-10052 | 19134528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134528&pc=OZD | 03/30/2024 | 05/10/2024 |  | 04/18/2024 |  |  | 03/30/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441938 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 50-year-old male of unknown race with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient experienced some oozing from the access site. one unit of packed red blood cells was subsequently transfused to treat the condition. support with the implanted pump was maintained following the intervention. | [5 empty] | The investigation into the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
27761 | 1220648-2024-10054 | 19134664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134664&pc=OZD | 04/05/2024 | 05/15/2024 |  | 04/18/2024 |  |  | 04/05/2024 | 04/18/2024 |  |  | 03/31/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329649 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing coronary artery bypass grafting (cabg) was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient developed a small hematoma. pressure was held at site. it was further reported that after cabg procedure, the hematoma was evacuated. the patient is stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
27762 | 1220648-2024-10055 | 19134774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134774&pc=OZD | 05/06/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195576 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, staff noted a significant hematoma and oozing. the physician decided to remove the device and place in the right femoral artery. | [5 empty] | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27763 | 1220648-2024-10056 | 19134777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134777&pc=OZD | 03/19/2024 | 05/13/2024 |  | 04/18/2024 |  |  | 03/19/2024 | 04/18/2024 |  |  | 08/31/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384789 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant notified via abiomed¿s long-term outcome and quality indicator impella registry that a patient endured ventricular tachycardia (vt) with a "possible" relationship to the impella 5.5 pump. after the impella implant, the patient had episodes of coughing. patient had one episode of vt lasting around 40 seconds, which broke with further coughing. bedside imaging showed the impella in place with right ventricle akinesis. a few minutes later, patient went into 4 minutes of vt, which broke with cardioversion (was given 1 mg midazolam before). given 150 mg iv amiodarone x 2. dob was increased to 5 mcg/kg/min (previously 4). | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27764 | 1220648-2024-10057 | 19134795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134795&pc=OZD | 05/09/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/09/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184705 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and sinus tachycardia-elevation accompanied by 3-vessel disease. while on support, the patient experienced access site bleeding that was resolved via surgical closure. the patient remained stable and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27765 | 1220648-2024-10059 | 19134838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134838&pc=OZD | 05/03/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/02/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191713 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support a large clot was noted around the access site, the groin site was redressed. keliox and 4x4 bandages were used to prop up the angle distal to site, and the patient received 1 unit of packed red blood cells. | [5 empty] | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27766 | 1220648-2024-10060 | 19134910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134910&pc=OZD | 05/10/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/10/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187386 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and left ventricular end diastolic pressure. the peel away was exchanged for a repositioning sheath for the patient to remain on support for 24 hours for unloading the left ventricle. the patient was taken to the cardiac catheterization lab for removal of the impella. the impella and sheath were removed, there was a brisk bleed back noted. previously deployed perclose sutures were tightened, and hemostasis obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27767 | 1220648-2024-10058 | 19134959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19134959&pc=OZD | 04/05/2024 | 05/22/2024 |  | 04/18/2024 |  |  | 04/05/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438016 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 71-year-old female patient with low cardiac output syndrome and post-cardiotomy cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella device could not be implanted due to the patient¿s anatomy. the implant procedure was aborted, and an impella cp device was successfully implanted. the patient remains stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since the pump could not pass due to the patient anatomy. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings states ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ it also states ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added- d6a/d6b.
27768 | 1220648-2024-10063 | 19135080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135080&pc=OZD | 05/06/2022 | 03/07/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 11/30/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169031 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support thrombus formed in the left femoral artery (lfa) occluding vessel. an additional angiogram was performed, and the patient was given additional heparin. a penumbra catheter was used to remove clot. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Analysis summary the failure mode of "thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
27769 | 1220648-2024-10061 | 19135108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135108&pc=OZD | 05/06/2022 | 03/07/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 05/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090414 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the pump was hung up on the papillary muscles, which was resolved by removing the pump and re-accessing the left ventricle. additionally, the patient experienced access site bleeding that was resolved by cut down and surgical closure. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning and access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
27770 | 1220648-2024-10062 | 19135160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135160&pc=OZD | 04/06/2024 | 05/23/2024 |  | 04/18/2024 |  |  | 04/06/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449982 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. during support, there were purge alarms. the team replaced the purge cassette and the automated impella controller (aic) was swapped with no improvement. the patient was not on anti-coagulation medication despite recommendations. the purge pressure rose and the team instead determined they would use tissue plasminogen activator in the purge. the surgeon decided to replace the impella 5.5. the patient was stable with the new impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue and the priming issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the high purge pressure was most likely due to a gradual buildup of biomaterial over a period of 24 minutes due to improper anticoagulation of the patient. the root cause of the priming issue was undetermined. impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27771 | 1220648-2024-10064 | 19135235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135235&pc=OZD | 05/06/2022 | 03/07/2025 |  | 04/18/2024 |  |  | 05/06/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179432 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the user reported issues with suction. volume status optimization was ongoing; insertion angle of the repo sheath were optimized to help resolve the reported issue. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27772 | 1220648-2024-10065 | 19135254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135254&pc=PYX | 04/09/2024 | 06/07/2024 |  | 04/18/2024 |  |  | 04/09/2024 | 04/18/2024 |  |  | 03/31/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024394937 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 57-year-old male of an unknown race in non-st segment elevation myocardial infarction was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the impella was opened and prepped, but the placement signal was very high. the impella rp flex sensor was reading in the 90's. the impella rp flex was replaced. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical issue has been completed. the impella rp flex pump was not returned. the data logs were reviewed and no placement signal unreliable alarm was observed. the optical 5s parameters for the last few seconds of the six minutes case run on rt logs were seen to be varying. no definitive failure pattern was observed. the cause of the optical signal issue was not determined since the product was not returned and due to inconclusive evidence per log review. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has since been completed.
27773 | 1220648-2024-10066 | 19135325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135325&pc=OZD | 05/10/2022 | 05/19/2025 |  | 04/18/2024 |  |  | 05/10/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191627 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported a patient, noted as high risk percutaneous coronary intervention was scheduled to be implanted with an impella cp device for mechanical circulatory support. during implantation, the sheath kinked after removal of the dilator, a 14f 25 cm sheath used in the second step, however it was impossible to advance the impella over the sheath. the insertion was aborted due to the patient¿s vascular anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27774 | 1220648-2024-10067 | 19135393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135393&pc=OZD | 03/11/2024 | 05/31/2024 |  | 04/18/2024 |  |  | 03/19/2024 | 04/18/2024 |  |  | 11/30/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024433527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support the cp pump support was also enrolled in the long-term outcome and quality indicator (loqi) protocol. the pump support was ongoing when there was hemolysis with a "possible" relationship to the impella use per the loqi database. the patient outcome was unknown. | There is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation. no data logs available to review. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27775 | 1220648-2024-10068 | 19135454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135454&pc=OZD | 05/07/2022 | 03/07/2025 |  | 04/18/2024 |  |  | 05/07/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191791 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, there was a hematoma at the impella site that grew rapidly. the impella was explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
27776 | 1220648-2024-10069 | 19135482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135482&pc=OZD | 05/02/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/02/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195383 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella support, the patient had poor pedal pulses and their bilateral lower extremities were mottled. the pulses improved on bilateral lower extremities the next day with no information provided if there was any intervention performed. the patient¿s laboratory samples were also hemolyzed. the patient's plasma free hemoglobin was elevated and the patient was started on continuous renal replacement therapy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27777 | 1220648-2024-10070 | 19135492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135492&pc=OZD | 05/07/2022 | 03/07/2025 |  | 04/18/2024 |  |  | 05/07/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195007 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support a possible hit (thrombocytopenia) was noted. argatraban systemic and bicarb purge was administered. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27778 | 1220648-2024-10071 | 19135511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135511&pc=OZD | 04/10/2024 | 06/08/2024 |  | 04/18/2024 |  |  | 04/10/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440644 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. a distal perfusion catheter and fasciotomy was performed to resolve the limb ischemia. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added codes 4625, 4114 and 3038 as they were inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27779 | 1220648-2024-10072 | 19135530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135530&pc=OZD | 06/18/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 06/18/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200344 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and renal failure. the condition continued from insertion to device removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27780 | 1220648-2024-10073 | 19135559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135559&pc=OZD | 05/13/2022 | 05/14/2025 |  | 04/18/2024 |  |  | 05/13/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196958 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Difficult to Insert | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/13/2022 |  |  |  |  | The user facility reported a 75-year-old patient needing mechanical circulatory support. it was reported that the patient was presented with a left ventricular ejection fraction under fifteen percent. as the patient¿s condition declined, the impella was implanted for left ventricular unloading. the sheath from the percutaneous coronary intervention was still in the right femoral artery from the prior unsuccessful attempts to place the thirteen-centimeter sheath, which was replaced with a twenty-five-centimeter peel away and the impella cp was advanced without issue. also, the patient¿s lab results identified hemolysis. transthoracic echocardiogram showed the inlet in the left ventricular outflow tract. the inlet was repositioned. per a right atrium thrombus was observed in the vicinity of the forearm. it is stated this condition will be monitored closely. | Additional info received that patient expired and medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-10073 was provided in sections b2, b5, d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27781 | 1220648-2024-10074 | 19135671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135671&pc=OZD | 04/10/2024 | 05/30/2024 |  | 04/18/2024 |  |  | 04/10/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450433 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the next day after placement, there were aortic placement alarms. repositioning was unsuccessful, and the impella pump was removed at bedside. there was no report of patient harm. | [5 empty] | The investigation into the positioning issues and the access site bleeding ¿ minor issue has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27782 | 1220648-2024-10075 | 19135842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135842&pc=OZD | 05/16/2022 | 05/23/2025 |  | 04/18/2024 |  |  | 05/16/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the pump experienced flows lower than expected with no specified resolution. additionally, the patient experienced ventricular tachycardia that resolved on its own. upon removal, thrombus was observed on the impella in association with potential left ventricular thrombus observed on images. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27783 | 1220648-2024-10076 | 19135914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135914&pc=OZD | 05/03/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/03/2022 | 04/18/2024 |  |  | 10/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155156 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient required an external femoral-femoral bypass on the right leg for perfusion due to limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27784 | 1220648-2024-10077 | 19135922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19135922&pc=OZD | 04/11/2024 | 06/06/2024 |  | 04/18/2024 |  |  | 04/11/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456072 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Improper Flow or Infusion | Infusion or Flow Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced increased pump flows/motor current and subsequently purge fluid dropped and there was an increase in purge pressure. the patient's urine become pink-tinged. the pump was replaced to resolve the issue. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported issues has been completed since the original report was submitted. the pump was not returned for investigation. the data logs showed a four minute rise in purge pressure from 400 mmhg to 650 mmhg without triggering any purge-related alarms. the purge pressure rise was within the normal specification range, and its trend was decreasing gradually. the pump flow was noted to be slightly higher than the normal limit during the initial spike in purge pressure. no abnormalities were reported with motor current spikes, placement signals, or positioning issues. clinically, pink-tinged urine was reported during the case run. the cause of the hemolysis issue was not determined since the product was not returned and sufficient clinical information was not provided. the cause of the saturated pump flow issue was most likely biomaterial, which could lead to a rise in purge pressure. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 medical device problem code 2954 was revised to 2964 in accordance with updated procedures. h.6 added code 41115 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27785 | 1220648-2024-10078 | 19136042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19136042&pc=OZD | 05/03/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/03/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194795 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced access site bleeding issues and the purge was switched to bicarbonate with heparin discontinued to resolve the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27786 | 1220648-2024-10079 | 19136048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19136048&pc=OZD | 05/17/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/17/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191886 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped two times, but remained stable for the rest of the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27787 | 1220648-2024-10080 | 19136234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19136234&pc=OZD | 05/20/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/20/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180948 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The complainant reported a patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, patient developed persistent fever. as a result, the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27788 | 1220648-2024-10081 | 19136565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19136565&pc=OZD | 04/12/2024 | 05/28/2024 |  | 04/18/2024 |  |  | 04/12/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448721 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 71 year-old male with cardiomyopathy and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient is under consideration for a ventricular assist device and/or transplant. the patient developed signs of hemolysis and access site bleeding. the access site was stitched and had manual pressure applied. the patient received two units of red blood cells. the impella cp device was explanted and replaced with an impella 5.5 device three days later. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported access site bleed and hemolysis has been completed. product was not returned for review. data logs showed pump ran without issue during support. there were some positioning alarms at the end of the case but resolved quickly. there were also suction alarms triggered but not impact to the flow of impella. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use referenced in the initial report is from ifu for the impella cp with smart assist system, sections: general patient care considerations, entering cardiac output, potential adverse events (united states), which now includes statement "(including ventricular perforation).¿, use of echocardiography for positioning of the impella catheter, hemolysis, and suction . added-d6a/d6b
27789 | 1220648-2024-10082 | 19136666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19136666&pc=OZD | 05/08/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189093 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a patient was placed on impella cp due to ventricular tachycardia. while attempting to place the impella, the pump was unable to be delivered via radiofrequency catheter ablation, so it was placed via lateral circumflex femoral artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27790 | 1220648-2024-10083 | 19136755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19136755&pc=OZD | 05/03/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/03/2022 | 04/18/2024 |  |  | 09/30/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after being successfully weaned and the impella explanted the patient experienced a persistent bleed. reportedly 1 of 2 initial perclose sutures were successful, but the bleed did not stop and a third perclose was deployed with little change in the bleed size. an 8 french dilator was inserted to measure the remaining bleed and an 8 french angio seal was deployed to cover the remaining hole. manual pressure was held on the small remaining ooze for about five minutes and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27791 | 1220648-2024-10085 | 19137254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19137254&pc=OZD | 05/03/2022 | 04/07/2025 |  | 04/18/2024 |  |  | 05/03/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172844 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2022 |  |  |  |  | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support it was reported that the patient had a lack of perfusion in the impella leg. the medical team decided to explant the impella due to the limb ischemia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27792 | 1220648-2024-10084 | 19137260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19137260&pc=OZD | 04/12/2024 | 05/15/2024 |  | 04/18/2024 |  |  | 04/12/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 74-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding from the impella insertion site in the left femoral artery. once the peel away sheath was removed and repositioning sheath was advanced, forward tension sutures were placed. sheath angle matched appropriately. the insertion site required one hour of manual pressure was held to control bleeding. the patient was then administered one unit each of red blood cells and platelets. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the act was 358 which is above the range limit. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27793 | 1220648-2024-10086 | 19137633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19137633&pc=OZD | 05/08/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 10/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157020 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support due to heart disfunction. the patient started bleeding from both ecmo and impella site. femostop in place with venous pressure was applied and received 2 units of packed red blood cells administered. later, vascular surgery came to bedside and added a stitch to the impella site to control some oozing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27794 | 1220648-2024-10087 | 19137707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19137707&pc=PYX | 04/11/2024 | 07/30/2024 |  | 04/18/2024 |  |  | 04/13/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Power Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 39-year-old male that underwent cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that when leaving the procedure area after impella rp insertion, the automated impella controller (aic) alarmed and displayed impella pump stopped and aic failure. restarting did not solve the issue and the aic was exchanged. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) pump stop and red alarm issue has been completed. the aic was returned for evaluation. upon functional testing, a firmware check revealed that the epc was unable to communicate with the pmd. testing of the impellatronic pcb found that the rs232 communication lines between the epc and the pmd confirmed that no communication was occurring between the epc and the pmd. the voltage between test points was measured as 0v (5v expected), confirming that no voltage was being delivered to the pmd dsp circuitry. data logs were reviewed which showed a controller failure alarm had been triggered due to a loss in communication between the epc and the pmd. the controller failure resulted in a pump stop. following the controller failure alarm, a separate controller error alarm was triggered due to the voltage delivered by the pbm being identified as being out-of-spec. this confirmed the reported complaint. the specific voltages identified as out-of-spec were the voltages delivered from the pbm to the pmd. the cause of the aic issue was a defective pbm board. d2-corrected common device name f6/f8 should have been left blank in the initial report.
27795 | 1220648-2024-10088 | 19137883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19137883&pc=OZD | 05/04/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced a positioning issue and hemolysis. it was noted that there was a ¿placement signal low¿ alarm when the patient arrived at the intensive care unit and that the deep positioning issue was resolved by pulling the inlet back to an alternate depth. hemolyzed labs noted that the patient¿s urinary output was red. hemolysis was noted to have cleared. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27796 | 1220648-2024-10089 | 19138098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19138098&pc=OZD | 04/13/2024 | 06/04/2024 |  | 04/18/2024 |  |  | 04/13/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025447352 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male post-cardiotomy cardiogenic shock and low cardiac output syndrome was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was placed post mitral valve replacement. the physician experienced issues and made multiple attempts to get the impella pump to the left ventricle. however, post placement, it was noted that the pump was sitting in a "good" position away from the mitral valve in the left ventricle. a transesophageal echocardiography was performed and showed a loose chord in the left ventricle with a possible clot formed on the chord. the physician expressed uncertainty whether the impella pump or mitral valve replacement was the cause of the loose chord. the patient continued with impella pump support. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported heart valve damage has been completed. no product was returned for evaluation. clinical details noted that mitral valve (mv) repair was performed pre-impella insertion and it was unclear if the loose cord was caused by the mvr or impella insertion. data logs were reviewed and no impella positioning alarms noted. the root cause of the heart valve damage cannot be determined with limited clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (unites states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ in accordance with updated procedures, sections d6, g3, and h10 have been revised from the initial submission.
27797 | 1220648-2024-10090 | 19138322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19138322&pc=OZD | 05/04/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185195 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the angiography showed no distal flow, a 14 f peel away sheath was removed, peeled away and flow was restored. the impella was weaned without compromise and explanted. after the device was explanted, 2 perclose singed down however there was bleeding over the access site. the staff used a 1 8 f angio seal as an additional closure measure. the angio seal did not resolve the issue, so pressure was held for 30 minutes, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27798 | 1220648-2024-10091 | 19138602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19138602&pc=OZD | 04/13/2024 | 06/05/2024 |  | 04/18/2024 |  |  | 04/13/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446682 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male with cardiomyopathy planned for an impella cp device for mechanical circulatory support. while the pump was being prepped and while it was on the wire, the impella was started outside of the body. due to concerns for damage to the pump, that pump was discarded and a new impella pump was open and successfully placed without issue. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue is determined to be use issue because the pump was turned on outside the body and discarded because of any damage issues to the pump. the instructions for use for the impella cp with smartassist system in the following section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27799 | 1220648-2024-10092 | 19139026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139026&pc=OZD | 04/13/2024 | 06/05/2024 |  | 04/18/2024 |  |  | 04/13/2024 | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481333 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 69-year-old female of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during percutaneous coronary intervention (pci), the patient received a shock for ventricular fibrillation and woke up and tried to sit up. later during the procedure, it was discovered the patient had pericardial effusion. a pericardiocentesis was performed and after two liters were pulled off, there were concerns for continued bleeding source. a right coronary artery angiogram was done to rule out coronary perforation. an echocardiogram was performed with definity injection and found to have contrast in the left ventricle (lv) and pericardium. lv perforation of the apical aspect of the lv was confirmed in the operating room and repaired with a suture. the patient then was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation for the reported ventricle perforation has been completed. no product was returned for evaluation. clinical details noted that patient was given a shock for vf arrest and patient awoke and moved violently and tried to sit up while on impella support. a pericardial effusion was later noted and echo was performed and found a lv perforation which was repaired in the or. data logs do not show any positioning issues with the pump. the root cause of the ventricle perforation is most likely due to patient movement as the patient moved violently and attempted to sit up with impella cp implanted after being shocked. the instructions for use referenced in the initial report for the impella cp with smartassist for use during cardiogenic shock and high risk pci in section: potential adverse events (united states) now includes statement "cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27800 | 1220648-2024-10093 | 19139461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139461&pc=OZD | 04/13/2024 | 06/03/2024 |  | 04/18/2024 |  |  | 04/13/2024 | 04/18/2024 |  |  | 08/31/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380545 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 56 year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. a thrombus was discovered on the impella cannula within the left ventricle (lv). anticoagulant therapy was increased in response, but it was reported that the patient suffered a stroke overnight. the thrombus was no longer present in the lv following the stroke. support continued with the implanted impella device, and the patient showed signs of recovery. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issues has been completed since the original report was filed. pump was returned and investigated. the catheter was stretched from multiple areas and the kink protector was loose from the red handle area. data logs show pump stop 115 alarm which correlates to the stretched catheter. the pump failed electrical resistance test with 0ohms for all the 3 phases further confirming electrical wire open. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27801 | 1220648-2024-10094 | 19139514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139514&pc=OZD | 05/04/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195380 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a 87-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella cp fourteen (14) french (fr) peel away sheath was placed over the device stiff wire. the impella was inserted over the wire. a twenty-three-centimeter peel away sheath was inserted, but it was noted that the sheath had difficulty passing through the tortuous anatomy. while removing the dilator, the peel-away was identified to have a small kink and an abrupt turn at the bifurcation of the common iliacs. the impella was inserted over a 0.018 wire via the left ventricle. however, the device did not pass through the valve on the 14 french sheath. the sheath was upsized to a sixteen (16) french, thirty-three (33) centimeter dry seal gore sheath and the device past through the vasculature with ease. the right femoral artery was accessed for percutaneous coronary intervention (pci). however, when the guidewire was inserted, it was observed that the medical professional was unable to cannulate due to the presence of the 33-centimeter 16 french sheath. the right radial approach (rra) was accessed for pci and the technique was performed. hypotensive events throughout pci, but the patient recovered well on impella support. the 16 fr gore sheath was removed, the device was explanted, re-close technique was utilized, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27802 | 1220648-2024-10095 | 19139585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139585&pc=OZD | 05/08/2022 | 09/30/2025 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 08/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. trans-esophogeal echocardiogram revealed the impella to be shallow in position. the physician advanced the impella 1cm to 87cm and re-secured by touhy valve/lock. the impella still appeared shallow. the physician continued to monitor the patient. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-10095 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10095 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10095 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10095.
27803 | 1220648-2024-10096 | 19139668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139668&pc=OZD | 05/04/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 11/30/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165990 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion the patient experienced bleeding at the access site, a perclose was used and failed to resolve the issue. additionally, two angio seals were attempted and also failed. manual pressure was held while the patient decompensated, neo and levo medications were administered. contralateral access enabled the use of an occlusive balloon to control the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27804 | 1220648-2024-10097 | 19139738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139738&pc=OZD | 05/08/2022 | 04/08/2025 |  | 04/18/2024 |  |  | 05/08/2022 | 04/18/2024 |  |  | 03/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191983 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding that was treated by changing the dressing, applying sutures and securing the angle. additional bleeding was encountered later and treated by repositioning of the pump. bleeding and hematoma persisted; eventually the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27805 | 1220648-2024-10098 | 19139764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139764&pc=OZD | 05/13/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/13/2022 | 04/18/2024 |  |  | 12/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding with no resolution specified. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. section h6 : impact code has been updated to correct code, which was inadvertently reported incorrectly in the initial report.
27806 | 1220648-2024-10099 | 19139801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139801&pc=OZD | 05/16/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/16/2022 | 04/18/2024 |  |  | 01/31/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023189565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed leg ischemia. the leg was improperly perfused, and the leg was cold. resolution is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27807 | 1220648-2024-10100 | 19139815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139815&pc=OZD | 05/16/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/16/2022 | 04/18/2024 |  |  | 08/31/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, some oozing was noted around peel away site. purse string was placed, proglides were tightened, and pressure dressing was placed with good hemostasis. the following day, the patient was bleeding from all access sites. a pressure bag was placed on the groin site and this reduced impella access bleeding. the patient received units of blood. follow up confirmed that the oozing slowed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27808 | 1220648-2024-10101 | 19139818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139818&pc=OZD | 05/04/2022 | 04/18/2024 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 02/29/2024 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194887 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support it was reported that an echocardiogram was completed and there was bleeding from the impella access site. manual pressure was held for 40 minutes, however there was severe bleeding from the sheath site and hematoma. the decision for a compression device was added, a c-clamp, and partial thromboplastin time was 180. the heparin purge solution was held due to bleeding complications. the bleeding at the impella site has subsided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27809 | 1220648-2024-10102 | 19139865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139865&pc=OZD | 05/16/2022 | 03/31/2025 |  | 04/18/2024 |  |  | 05/16/2022 | 04/18/2024 |  |  | 11/30/2023 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166529 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the left groin access site while in the intensive care unit. heparin was discontinued. the staff held manual pressure and a femstop was placed to contain the hematoma. over an hour later, the medical professional received information that the hematoma had increased in size; therefore, the patient was sent for explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27810 | 1220648-2024-10103 | 19139928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139928&pc=PYX | 12/12/2023 | 09/26/2024 |  | 04/18/2024 |  |  | 03/20/2024 | 04/18/2024 |  |  | 10/31/2025 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2024418625 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Renal Failure | Cardiogenic Shock | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 03/20/2024 |  |  |  |  | The user facility reported a 29-year-old male patient was implanted with an impella rp device for mechanical circulatory support. a notification was received long-term outcome and quality indicator impella registry that the patient endured worsening cardiogenic shock with a possible relationship to the impella rp device. low impella flow was also noted with a probable relationship to the impella rp device used on a patient. care was withdrawn and the patient expired. outcome associated with the use of the device. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Additional information received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that worsening renal failure with a possible relationship to the impella rp device used on this patient. patient with progressively worsening state and malfunctioning right internal jugular impella requiring impella exchange emergently. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the low pump flow and the renal failure issues has been completed since the original report was submitted. the device was not returned for investigation. it was reported that there were suction alarms at the start of the case. the root cause of the low pump flow and the renal failure issues was not determined since product was not returned and there is a lack of clinical details. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5 and h1. | B5 and h6 was updated as per additional information received.
27811 | 1220648-2024-10104 | 19139938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19139938&pc=OZD | 05/04/2022 | 05/14/2024 |  | 04/18/2024 |  |  | 05/04/2022 | 04/18/2024 |  |  | 11/30/2022 |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022023245 |  | 0048-0004 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and that the impella was repositioned under echo for hemolysis and hematuria. | [5 empty] | The initial report, 1220648-2024-10104, was submitted in error. the reported event does not meet medical device reporting requirements. corrected information in h10. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27812 | MW5153979 | 19142220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19142220&pc=OZD | 02/29/2024 | 04/16/2024 |  | 04/18/2024 |  |  |  | 04/18/2024 |  |  |  |  | 04/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Unknown | [6 empty] | 79 YR |  | Male | 112 KG | Non Hispanic | White | Death | Yes | Yes | [7 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Failure of Implant | Renal Failure | Low Oxygen Saturation | Vascular Dissection | Thrombosis/Thrombus | [6 empty] | Life Threatening | Death | Disability | Hospitalization | Required Intervention | [24 empty] | Initial submission | [9 empty] | On (b)(6) 2024 impella 5.5 placed to bridge to lvad. it was a successful placement, however, hypotension, decrease perfusion, hemolysis, renal failure occurred due to device failure. family was told there were concerns for shearing of the left ventricle resulting in blood clot formation. emergency bedside removal was performed and balloon pump had to be placed. this resulted in death due to the complications following the impella placement. the patient did really well until the complications. the impella was sent for lab testing.
27813 | 1220648-2024-09964 | 19123949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19123949&pc=OZD | 03/21/2024 | 05/13/2024 |  | 04/17/2024 |  |  | 03/21/2024 | 04/17/2024 |  |  | 10/31/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418773 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the physician encountered delivery issues during implant. the graft was sewn too low on the aorta and the patient had anatomical issue/calcifications. the implant was aborted. the physician opted to centrally cannulate the patient for venoarterial extracorporeal membrane oxygenation. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related as the patient had calcified vessels which made placement of the graft difficult and led to additional resistance. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27814 | 1220648-2024-09966 | 19124009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124009&pc=OZD | 03/22/2024 | 05/14/2024 |  | 04/17/2024 |  |  | 03/22/2024 | 04/17/2024 |  |  | 10/31/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | Yes | Difficult to Insert | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 66 year-old male patient with a high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. while inserting the impella device, a hematoma formed on the left femoral site following the micro puncture and sheath placement. the medical staff decided to abort the procedure due to the patient¿s anatomy, the hematoma, and continued bleeding. the patient remains stable. | [5 empty] | The investigation for the reported access site bleed and delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine root cause. the root cause of the delivery issue was most likely due to patient condition since the case was aborted due to patient's anatomy. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. h.6 code 2993 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27815 | 1220648-2024-09968 | 19124146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124146&pc=OZD | 03/31/2024 | 06/19/2024 |  | 04/17/2024 |  |  | 03/31/2024 | 04/17/2024 |  |  | 06/30/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367933 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella 5.5 support for 62-year-old male patient, the clear side arm of the purge system attached to the red impella plug was broken. alarm alerted of the damage. there was no reported patient harm. it was also noted that the patient had a thrombus of unknown origin. | [5 empty] | The investigation for the reported thrombus and purge leak has been completed. clinical details state that a purge system alarm started and they discovered that the clear sidearm attached to red pump plug was broken. the device was not removed due to failure mode. thrombus noted upon explant. no product or data logs were returned. the root cause of the purge leak - pump, was not determined as no product was returned and limited clinical information provided. as the necessary information was not provided, the root cause of the thrombus was not determined instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ b.2 revised as previously entered incorrectly. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site name and address section as previously entered incorrectly. h.6 code 481 was reported incorrectly on the previously submitted report. a new code has been added to component codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation) "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
27816 | 1220648-2024-09969 | 19124196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124196&pc=OZD | 04/01/2024 | 05/23/2024 |  | 04/17/2024 |  |  | 04/01/2024 | 04/17/2024 |  |  | 08/31/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443552 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. same day, the impella device was removed due to occlusion of distal flow to the left leg. support transitioned to veno-arterial extracorporeal membrane oxygenation. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27817 | 1220648-2024-09970 | 19124274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124274&pc=OZD | 04/06/2024 | 05/22/2024 |  | 04/17/2024 |  |  | 04/07/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460413 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old male with cardiomyopathy had an impella cp device implanted and placed on veno-arterial extracorporeal membrane oxygenation for mechanical circulatory support. the next day, the catheter migrated into the aorta and impella pump was replaced. it was noted the catheterization lab personnel broke the t-lock, and thereafter was unable to secure the catheter to sheath. there was no report of patient harm. | Additional information was provided that states that due to being unable to secure the catheter to sheath, the catheter migrated into the aorta. a decision was made to replace the catheter. |  |  |  |  | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was use related due to improper technique as the t-lock broken by the cath lab personnel. added-b5 additional narrative, h6 med dev prob code 1069 added- d6a/d6b.
27818 | 1220648-2024-09971 | 19124299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124299&pc=OZD | 04/02/2024 | 05/23/2024 |  | 04/17/2024 |  |  | 04/02/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448777 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. both the introducer and pump were found to be occlusive during impella cp support. it was initially noted following following a planned procedure, that the in place 14 fr introducer was occlusive. the sheath was swapped for the repositioning sheath, but it too was noted to be occlusive. pt had significant dampening during the inflation. after pulsatility recovered a single stent was placed and post dilated. the lm was then stented and pt had no pulsatility during inflation. pt recovered and ivus was used to verify stent sizes. as a result of the condition, the decision was made to wean and explant the pump. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27819 | 1220648-2024-09972 | 19124381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124381&pc=OZD | 04/05/2024 | 05/14/2024 |  | 04/17/2024 |  |  | 04/05/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461219 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation of care from an impella cp device. the patient was also escalated to extracorporeal membrane oxygenation. during support, blood was observed at the red plug. the blood source was not known, but was monitored. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The investigation for the reported product damage has been completed. at the conclusion of the case, blood was noted via visual inspection at location of red handle. no alarms or impact to pump function at time of reporting. no product or data logs were returned. the root cause of product damage (hole in catheter) was not determined as no product was returned and insufficient clinical information was provided. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27820 | 1220648-2024-09973 | 19124412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124412&pc=OZD | 04/05/2024 | 09/19/2024 |  | 04/17/2024 |  |  | 04/05/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a "controller error" alarm upon powering up. the power was cycled in an attempt to troubleshoot the issue, but the failure remained. the aic was swapped as result of the noted issue and the device was removed from service. there was no patient involvement. | [5 empty] | The investigation for the console (aic) controller alarm yellow has been completed. data logs were reviewed. the reported complaint was confirmed. two controller errors were triggered due to the console identifying the optical sensor lamp as defective (event set #1039). preceding each controller error, the console attempted to reset the optical bench to resolve the issue, and this process also revealed that the bad lamp caused the 5s parameters to be out of spec. the console was returned for evaluation. the 5s parameters were tested and confirmed to be out of spec, and matched the pattern of a defective lamp. the console was opened up, and the lamp was confirmed to not be emitting any light while the console was powered on. the lamp was confirmed to be defective. the reported complaint was reproduced. the reported complaint was determined to have been caused by a defective lamp within the optical bench. d2-corrected common device name. d4-corrected model number, catalog number. e2 changed to no. f6/f8 should have been left blank in the initial report. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
27821 | 1220648-2024-09974 | 19124427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124427&pc=OZD | 05/17/2022 | 09/30/2025 |  | 04/17/2024 |  |  | 05/17/2022 | 04/17/2024 |  |  | 02/29/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the pump was repositioned in the catheterization laboratory. the flow was reported to be in range with no alarms. echocardiography revealed the impella cp was in papillary muscle. the medical team decided to escalate patient to impella 5.5 and did not readjust the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-09974 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-09974 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-09974 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-09974. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-09974.
27822 | 1220648-2024-09975 | 19124563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124563&pc=OZD | 04/07/2024 | 05/21/2024 |  | 04/17/2024 |  |  | 04/07/2024 | 04/17/2024 |  |  | 12/31/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293566 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2024 |  |  |  |  | The user facility reported a 14-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that three days following impella 5.5 implant, the patient was diagnosed with a large brain bleed. heparin was introduced following implant per protocol and was subsequently stopped upon discovery of the bleed. a drain was placed and impella support continued without the use of heparin. the patient was reported to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." | The investigation into the stroke/cva issue has been completed since the original report was submitted. the pump was not returned for investigation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27823 | 1220648-2024-09976 | 19124681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124681&pc=OZD | 04/07/2024 | 05/29/2024 |  | 04/17/2024 |  |  | 04/07/2024 | 04/17/2024 |  |  | 10/31/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411959 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient was implanted with an impella cp device and turned on with low pump flows. fluoroscopy showed the cannula had kinked just distal of the motor. the impella pump was removed and replaced with a new one. there was no report of patient harm. | [5 empty] | The investigation for the product damage has been completed. returned complaint cp pump was visually inspected. cannula kink was observed near outlet flow cage. no broken blades found. no biomaterial, no other abnormality observed. the cause of the low pump flow was cannula kink due to improper technique. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27824 | 1220648-2024-09977 | 19124844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124844&pc=OZD | 04/08/2024 | 05/20/2024 |  | 04/17/2024 |  |  | 04/08/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454562 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. the physician was unable to pass the pump through the axillary artery once through the anastomosis. the procedure was aborted. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since the pump could not pass through the axillary artery once through the anastomosis. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27825 | 1220648-2024-09978 | 19124950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19124950&pc=OZD | 04/09/2024 | 05/29/2024 |  | 04/17/2024 |  |  | 04/09/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451661 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 63 year-old female patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support following a defibrillator shock for atrial fibrillation. while turning the patient, the pump in aorta alarm sounded. the medical staff attempted to reposition the pump unsuccessfully and removed the impella device. following the removal, the medical staff discovered that the touhy lock was not completely secure. the patient survived the event. | [5 empty] | The investigation into the positioning issues and the vf/vt/af/at (afib with rvr) issues has been completed since the original report was submitted. the pump was not returned for investigation. based on clinical description & data analysis, the root cause of positioning issue was most likely use related. as the necessary information was not provided, the root cause of the vt/vf was not determined. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27826 | 1220648-2024-09979 | 19125086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19125086&pc=PYX | 04/10/2024 | 07/17/2024 |  | 04/17/2024 |  |  | 04/10/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025461560 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male was implanted with an impella rp flex. upon removal of the impella rp flex device, the pump fractured at the motor housing. it was noted that insertion and removal were per protocol with no other issues. there was no report of patient harm. | [5 empty] | The investigation for the removal issues has been completed. the returned pump was imaged and showed a relatively straight breakage across the mating surfaces across the 2 peek bearing caps at the motor housing. the optical sensor line was in tact yet kinked, and holding the pump together. the two halves of the pump were clean with no blood on the interior. sem showed the fracture point to be near the sensor groove on the motor house, starting on the exterior and propagating inwards. there were also no signs of fatigue fracture which shows that the break was likely a fast fracture. data logs indicate that the pump ran successfully for the duration of the case with no indications of the breakage occurring while on patient support. the last rt log did not show any hardware failure alarms, which indicates that the sensor cables did not break during support. the rt log also does not show a drop in purge pressure indicating a crack in the motor housing did not occur on support. the clinical states that removal was smooth with no resistance felt. results of this investigation indicate that: the break occurred outside of the patient¿s body. there is no evidence of deviation or non-conformance from the manufacturing process. there is no evidence that the design does not meet specifications. the break occurred outside of the operation of the device, with force exceeding. the root cause of the motor house breakage was not determined. added- d6a/d6b. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
27827 | 1220648-2024-09980 | 19125226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19125226&pc=OZD | 11/19/2023 | 05/07/2024 |  | 04/17/2024 |  |  | 03/18/2024 | 04/17/2024 |  |  | 06/30/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356976 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | A notification was received regarding a 73-year-old male who was escalated to impella 5.5 device for mechanical circulatory support. the notification was received from long-term outcome and quality indicator impella registry, which noted that worsening kidney failure and worsening thrombocytopenia occurred with a probable relationship to the impella 5.5 device used on a patient. | [5 empty] | The investigation into the thrombocytopenia and the renal failure issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as per the clinical details, doctor reported possible relation to the impella pump use during renal failure and thrombocytopenia. the cause of the thrombocytopenia issue was not determined due to insufficient clinical details. the cause of the renal failure issue was not determined due to insufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27828 | 1220648-2024-09981 | 19125383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19125383&pc=OZD | 04/11/2024 | 06/06/2024 |  | 04/17/2024 |  |  | 04/11/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443204 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported due to a 54-year-old male patient's co-morbidities and low ejection fraction, a decision was made to place an impella 5.5 device to support the patient for planned bariatric surgery. shortly after placement of the impella 5.5 device, high purge pressure alarms were experienced. tissue plasminogen activator was added to the purge in response to the noted issue. support with the implanted pump continued, and there was no report of patient harm. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was submitted. the pump was not returned for investigation. the data log showed a five minute rise in purge pressure, which persisted for two hours before gradually reducing to normal levels. no motor current spikes were reported during the high purge pressure event. as per the clinical details, high purge pressure alarms went off and tissue plasminogen activator (tpa) was added to the purge. the cause of the high purge pressure issue was most likely biomaterial, based on data log review since tpa in the purge resolved the issue. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 4114 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
27829 | 1220648-2024-09982 | 19125487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19125487&pc=OZD | 04/11/2024 | 06/20/2024 |  | 04/17/2024 |  |  | 04/11/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441890 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 45-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during support, fluid was seen leaking from the pressure reservoir. one day following the discovery, it was noted that the pressure reservoir was approximately half filled with purge fluid, and the liquid was dripping onto the floor. as no extreme changes were seen in the purge pressure or flow, the decision was made to maintain support while observing the purge data to see if it worsens. there was no report of patient harm. | [5 empty] | The investigation for the reported purge leak has been completed. no product was returned and no data logs were returned. clinical details state that a leak from purge reservoir was noted on day 3 and 4 of support. decision made to continue support as leak caused no extreme changes. the root cause of the purge leak - pump was most likely due to a damaged sidearm, but the cause of damage is undetermined due to lack of sufficient clinical information and no product was returned for analysis. the instructions for use referenced in the initial report is for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock in section: cleaning. added- d6a. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
27830 | 1220648-2024-09989 | 19126679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19126679&pc=OZD | 04/12/2024 | 05/29/2024 |  | 04/17/2024 |  |  | 04/12/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025452419 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 74 year-old male with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. during the implant, the devices catheter could not advance due to the patient¿s probable complex vascular anatomy. the patient was instead implanted with an impella cp device. no patient harm has been reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue is determined to be patient condition because of the complex vascular anatomy of the patient. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions states "warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27831 | 1220648-2024-09990 | 19126780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19126780&pc=OZD | 04/12/2024 | 06/07/2024 |  | 04/17/2024 |  |  | 04/12/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466860 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 63-year-old male of unknown ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding. the decision was made to explant the pump and transfuse two units of blood products as treatment. the outcome of the procedure was unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. added code 925 as it was inadvertently left off the previously submitted manufacture device report. code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury".
27832 | 1220648-2024-09991 | 19126945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19126945&pc=OZD | 05/29/2022 | 04/01/2025 |  | 04/17/2024 |  |  | 05/29/2022 | 04/17/2024 |  |  | 12/31/2023 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174045 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported a male on impella cp for mechanical circulatory support. the patient was noted to have limb ischemia and a left arterial thrombus that had not been initially identified which was later able to resolve itself. additionally, the pump occluded blood flow to the patient's limb noted by an absent pedal pulse and the patient's dusky colored right foot. this was treated by explanting the pump. | [5 empty] | The complainant reported that the patient had a thrombus and limb ischemia. the patient was noted to have a left arterial thrombus that had not been initially identified which was later able to resolve itself. additionally, the pump occluded blood flow to the patient's limb noted by an absent pedal pulse and the patient's dusky colored right foot. this was treated by explanting the pump. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27833 | 1220648-2024-09992 | 19127009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127009&pc=OZD | 04/12/2024 | 06/03/2024 |  | 04/17/2024 |  |  | 04/12/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461206 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 73 year-old female patient with a coronary artery bypass graft was implanted with an impella 5.5 device for mechanical circulatory support. the medical staff found the impella 5.5 device against the septum. the doctor retrieved the impella and determined it had been too shallow. once advancing the impella pump again, the doctor noticed pooled blood in the patient. the left ventricle was perforated, and the automated impella controller was alarming for flow saturation. the impella ingested tissue or a clot. the impella was removed, and the perforation was repaired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. data logs showed the pump flow saturation and purge pressure rise trend and several impella position alarms in the aorta/ventricle were reported. based on given clinical details and the data log review, the cause of the ventricle perforation issue was improper positioning of the device, and the cause of the saturated pump flow issue was most likely biomaterial.
27834 | 1220648-2024-09993 | 19127270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127270&pc=OZD | 05/07/2022 | 04/17/2024 |  | 04/17/2024 |  |  | 05/07/2022 | 04/17/2024 |  |  | 01/31/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187710 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 05/07/2022 |  |  |  |  | The user facility reported a patient was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella was unable to pass through the long peel away sheath. wire access was maintained, and the long sheath was removed and replaced with short sheath. the impella was delivered successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27835 | 1220648-2024-09994 | 19127301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127301&pc=PYX | 04/12/2024 | 06/07/2024 |  | 04/17/2024 |  |  | 04/12/2024 | 04/17/2024 |  |  | 10/31/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413663 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp flex device for mechanical circulatory support. it was reported that impella rp flex was inserted into the right internal jugular over 0.27 wire and into the right ventricle but was unable to pass into the pulmonary artery. multiple attempts were made but due to the right ventricle structure, the impella rp flex kept pushing towards the apex and not making the turn to the pulmonary artery. additionally, the guidewire kept bending and would not let the impella advance. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella was returned for evaluation. there were no abnormalities on the returned guidewire, no kinks or damages were observed. there were no damages or abnormalities on the pump. no catheter or cannula kinks or other visual damages. multiple different wires were used without success to deliver the pump and due to the right ventricle structure, the pump kept pushing towards the apex and not making the turn to the pulmonary artery. the root cause of the delivery issues was most likely patient condition related. instructions for use for the related event are as follows: impella rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the product has been returned since then and the investigation has been completed. h.10 additional manufacturer narrative instructions for use were added in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27836 | 1220648-2024-09995 | 19127312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127312&pc=OZD | 09/28/2022 | 04/01/2025 |  | 04/17/2024 |  |  | 09/28/2022 | 04/17/2024 |  |  | 03/31/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209832 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a hematoma at the groin and femstop was applied and pressure was held to resolve hematoma and the patient was transfused with two units of blood. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27837 | 1220648-2024-09996 | 19127607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127607&pc=OZD | 04/13/2024 | 06/05/2024 |  | 04/17/2024 |  |  | 04/13/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479508 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 39-year-old male undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the impella rp was explanted due to the patient developing hemolysis. it was further reported a ¿whirring/wine¿ sound occurred upon auscultation. the patient had a durable left ventricle assist device and continuous renal replacement therapy (crrt) in place. there were concerns that crrt may be interfering with the impella rp flex inflow due to proximity. the impella rp flex was still functioning as intended with good flows/waveforms. the patient is stable. | [5 empty] | The investigation for the reported hemolysis has been completed. no data logs were returned for the second console that was used during the reported hemolysis. the impella was returned for evaluation. upon visual inspection, biomaterial was found inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. pump was run with biomaterial to observe if it was noisy. the whirring sound while pump running was confirmed. the biomaterial was removed and pump was run again. at this time, the pump ran without whirring sound indicated that the noisy sound was due to biomaterial wrapped around the impeller result in the pump got to run on high gear and caused noisy sound. the root cause of the hemolysis was biomaterial ingestion. the root cause of the noisy sound was due to biomaterial wrapped around the impeller result in the pump got to run on high gear and caused that sound. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller in section: potential adverse events (united states) states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ added- d9 device return date added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27838 | 1220648-2024-09997 | 19127742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127742&pc=OZD | 04/13/2024 | 06/03/2024 |  | 04/17/2024 |  |  | 04/13/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 71-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, there was a pump stop due to high motor current. the pump appeared to have saturated flows. another pump stop occurred, and the team was unable to restart the pump. the team suspected a clot was ingested. the impella was removed, and the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. clinical details note potential saturated flows and suspected mitral valve (mv) thrombus in the patient. per the clinical information provided and the data logs, the root cause of the saturated flow and the pump stop is most likely due to biomaterial. the root cause of the thrombus was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27839 | 1220648-2024-09998 | 19127779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127779&pc=OZD | 09/28/2022 | 04/01/2025 |  | 04/17/2024 |  |  | 09/28/2022 | 04/17/2024 |  |  | 05/31/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 48-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp was successfully implanted, multiple attempts were made to engage the right coronary artery and to wire the lesion but all were unsuccessful with multiple wire/guide combinations. the impella was turned down and eventually pulled back into the descending aorta. the ostium was engaged and the lesion was wired once, but the guide slipped out and lost the wire position. the physician could not get the guide to seat in ostium. the percutaneous coronary insertion was aborted and the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27840 | 1220648-2024-09999 | 19127867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19127867&pc=OZD | 04/13/2024 | 06/03/2024 |  | 04/17/2024 |  |  | 04/13/2024 | 04/17/2024 |  |  |  |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447450 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 57 year-old male with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the implant, the pump would not start. the automated impella controller was replaced, which did not solve the issue. the medical staff removed the pump and implanted an intra-aortic balloon pump. the medical staff then noted that the pump was never started on auto mode. the pump setting was at zero and the staff did not know to manually adjust the level. the patient was then sent to the intensive care unit. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump did not start issue has been completed since the original report was submitted. the device was not returned for investigation. per the data logs, optical signal issue was observed as well. the root cause of the pump not starting was due to use as clinical team did not select p-auto to start pump. the root cause of the optical signal issue could not be determined as the product was not returned for evaluation. in accordance with updated procedures, section d6 has been revised from the initial submission.
27841 | 1220648-2024-10000 | 19128045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19128045&pc=OZD | 04/12/2024 | 07/02/2024 |  | 04/17/2024 |  |  | 04/13/2024 | 04/17/2024 |  |  | 06/30/2025 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366225 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to exclude use of the device from the outcome. | The user facility reported an 82 year-old male was implanted with an impella cp device for mechanical circulatory support and for pre-operation placement. vascular dissection of the artery occurred when the impella cp device was removed. the vessel was repaired with a stent. the patient remains stable. |  |  |  |  | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist for use during cardiogenic shock and high risk pci in section: warnings & cautions: cautions added- b5 mso review, h6 component code 925, b7 relevant history added e1 international phone number added d6a/d6b f6/f8-should have been left blank in the initial report. h6 component code 925 added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta."
27842 | 1220648-2024-10001 | 19128046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19128046&pc=OZD | 10/04/2022 | 04/01/2025 |  | 04/17/2024 |  |  | 10/04/2022 | 04/17/2024 |  |  | 04/30/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225133 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced a hematoma and bleeding around the repositioning sheath. sutures were placed to inferior two suture holes and then removed by the physician to place superior suture holes. the physician placed two superior suture holes while staff held sheath at angle of insertion. the hematoma was stable per staff after 30 minutes of manual pressure was applied. a femstop was placed above the insertion site and the insertion site bleed was resolved. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27843 | 1220648-2024-10002 | 19128984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19128984&pc=OZD | 10/17/2022 | 04/01/2025 |  | 04/17/2024 |  |  | 10/17/2022 | 04/17/2024 |  |  | 01/31/2024 |  | 04/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189065 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella insertion, the physician perforated the gastric wall vessel and caused bleeding. the vascular surgeon cut down and clipped the epigastric vessel to stop the bleeding. the patient required blood transfusions and fluids. | [5 empty] | According to the clinical, shortly after beginning impella cp implantation the physician perforated the patient's gastric wall vessel. a vascular surgeon then arrived to resolve the issue by cutting down and clipping the vessel. after the procedure was completed the patient crashed a number of times and required multiple pints of blood to stabilize. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the peripheral vessel vascular damage was use related. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27844 | 1220648-2024-09891 | 19114669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19114669&pc=OZD | 05/16/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/16/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a 47-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a cool right leg. the medical team decided to discontinue impella cp support and remove the device early. upon removal of the impella cp, a fasciotomy was required to resolve a dissection found at the access site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27845 | 1220648-2024-09892 | 19114830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19114830&pc=OZD | 05/16/2022 | 03/28/2025 |  | 04/16/2024 |  |  | 05/16/2022 | 04/16/2024 |  |  | 03/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192330 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced jaundice and signs of hemolysis. the patient experienced a drop in hemoglobin without evidence of bleeding. the impella cp position was verified with transesophageal echocardiography. information regarding intervention for hemolysis was not provided, however the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27846 | 1220648-2024-09893 | 19114931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19114931&pc=OZD | 05/16/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/16/2022 | 04/16/2024 |  |  | 01/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190653 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient developed a hematoma at the impella insertion site. the physician opted to not start systemic heparin due to the small hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27847 | 1220648-2024-09894 | 19115073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115073&pc=OZD | 06/09/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 06/09/2022 | 04/16/2024 |  |  | 03/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205202 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/09/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, oozing was noted at the impella site. the decision was made to replace impella due to bleeding at site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records
27848 | 1220648-2024-09895 | 19115168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115168&pc=OZD | 05/14/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/16/2022 | 04/16/2024 |  |  | 12/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported during support there were suction events, echocardiography was performed to verify impella position but the medical team was unsure of measurements of the impella in the ventricle. the medical team was advised to confirm placement of impella with a additional echocardiography. prior to repeating the echocardiography, the impella alarmed for impella in aorta. the medical team moved patient to operating room and explanted the impella due to positioning issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27849 | 1220648-2024-09896 | 19115197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115197&pc=OZD | 05/21/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 05/21/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192847 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and had some bleeding and hematoma. the blue suture hub was buried to the point that only one suture ring on the hub was showing, angle matched and dressing was applied to resolve the bleeding but was not resolved even with the perclose that was tightened. the physician decided to explant to control the bleeding. the impella was weaned and explanted, the previously deployed percloses were tightened down and manual pressure was held for 30 minutes until hemostasis was obtained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27850 | 1220648-2024-09897 | 19115491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115491&pc=OZD | 05/21/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/21/2022 | 04/16/2024 |  |  | 10/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022156727 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and there was a concern of sepsis. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27851 | 1220648-2024-09898 | 19115512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115512&pc=OZD | 09/21/2023 | 04/16/2024 |  | 04/16/2024 |  |  | 03/29/2024 | 04/16/2024 |  |  | 02/28/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311541 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Blocked Connection | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, there was a temporary purge blocked alarm following a cord getting caught in a chair. the problem was not resolved by the de-air process. the purge cassette was changed and the issue was resolved. there was no patient harm. | [5 empty] | The investigation has been completed. the device was returned for evaluation and data logs were available. data logs confirmed the failure mode. the logs showed purge pressure (pp) went up to 1100 mmhg from 500 mmhg and remained in this level for about 45 minutes, then went back to 500 mmhg and in this range for the rest of the case. the pump was visually inspected and no issues were found. the pump was tested by running for > 22 hours to reproduce the failure mode. the high pp occurred in the first 30 minutes of running due to dry blood and glucose residing in the purge gap. after the dry blood and glucose residue was expelled, the pp was in specification for the rest of the test. the root cause of high purge pressure was use related as the pump tubing was get caught in the patient's chair during support.
27852 | 1220648-2024-09899 | 19115513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115513&pc=OZD | 05/17/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/17/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The user facility reported a 53-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a loss of pulses in the leg. no information was provided on any intervention provided to resolve the limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27853 | 1220648-2024-09900 | 19115681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115681&pc=OZD | 05/17/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/17/2022 | 04/16/2024 |  |  | 09/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had oozing and developed a small hematoma at the insertion site. in response, manual pressure, mattress suture, and femostop were applied to resolve the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27854 | 1220648-2024-09901 | 19115722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115722&pc=OZD | 05/26/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/26/2022 | 04/16/2024 |  |  | 12/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, 2 thrombosis were noted. thrombectomy procedure was performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27855 | 1220648-2024-09902 | 19115870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115870&pc=OZD | 03/31/2024 | 05/13/2024 |  | 04/16/2024 |  |  | 03/31/2024 | 04/16/2024 |  |  | 08/31/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443541 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was placed but the limb accessed had signs of limb ischemia. the team observed the leg to have lost pulses. there was knowledge that the patient had a history of peripheral vascular disease. the team discussed possible explant and care escalation to the 5.5 pump but, instead the family declined any further interventions and the patient expired while on support. critical limb ischemia in a patient with cardiogenic shock. association of effect on patient outcome could not be determined. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27856 | 1220648-2024-09903 | 19115990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19115990&pc=OZD | 04/01/2024 | 05/17/2024 |  | 04/16/2024 |  |  | 04/01/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454362 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Malfunction | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 77-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. after the insertion, the left ventricle wave forms were well above aortic waveforms and the flows demonstrated flow saturation. the impella was also alarming in the aorta. the doctor removed the impella and observed biomaterial in the outflow. the impella device was replaced. | [5 empty] | The investigation into the saturated flow and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. data logs were reviewed and there were motor current spikes seen followed by saturated flow along with steady rise in purge pressure. per the clinical information and data log review, the root cause of the saturated flow and the thrombus issues was most likely biomaterial. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ d6a, d6b. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27857 | 1220648-2024-09904 | 19116020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116020&pc=OZD | 05/17/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/17/2022 | 04/16/2024 |  |  | 01/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187787 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock and occluded left anterior descending was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced hemolysis. the patient received fluids to add volume and the impella was repositioned, but the hemolysis did not resolve. the medical team decided to escalate the patient to an impella 5.5 due to increase support and resolve the hematuria. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27858 | 1220648-2024-09905 | 19116050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116050&pc=OZD | 06/17/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 06/17/2022 | 04/16/2024 |  |  | 08/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022125426 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2022 |  |  |  |  | The user facility reported a 62-year-old male on impella 5.5 for mechanical circulatory support and that there was continued groin oozing (bleeding). no other information provided. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27859 | 1220648-2024-09906 | 19116077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116077&pc=PYX | 04/02/2024 | 05/17/2024 |  | 04/16/2024 |  |  | 04/02/2024 | 04/16/2024 |  |  | 04/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337540 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported an 81-year-old female with pulmonary embolism was implanted with an impella rp device for mechanical circulatory support. it was reported that physician advanced the impella rp into the pulmonary artery, but flows were lost upon advancement. the pump was then removed for suspected thrombus. the patient was then assessed for left ventricle failure and the patient's blood pressure was lost. advanced cardiac life support was started but after multiple rounds of cardiopulmonary resuscitation, return of spontaneous circulation was not obtained and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was returned for evaluation. upon visual inspection, a cannula kink was observed. no biomaterial in the impeller. no other abnormalities found. data logs were reviewed which showed that about 16 minutes into the case there is a sudden drop in flows. there were no motor current spikes prior to the loss of flows. there were no alarms. the root cause of the positioning issues of the pump being pulled back into the rv was most likely patient movement as the patient was agitated and was moving on the table. the pump was then confirmed to be in improper position and had to be repositioned. the loss of flows was reported immediately after repositioning. the patient's condition was also deteriorating. upon investigation of the product a cannula kink was observed. the root cause of the low pump flow issue was most likely a kinked cannula. this most likely occurred during repositioning. added- d9 device return date, h6 impact code 4628 added- d6a/d6b.
27860 | 1220648-2024-09907 | 19116287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116287&pc=OZD | 05/17/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/17/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the patient developed bleeding at the impella access site despite applying manual pressure and femostop, hemostasis was unable to be obtained. on 5/18/2022 it was reported the impella site continued to bleed and four units of packed red blood cells (prbc¿s) were administered to the patient. the team was unable to obtain hemostasis, leading to the pump removal. the impella cp pump speed level was decreased to p2 for 30 minutes to determine if the patient¿s anatomy could tolerate removal. it was determined that the patient was hemodynamically stable, the impella was explanted at bedside and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27861 | 1220648-2024-09908 | 19116456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116456&pc=OZD | 04/05/2024 | 05/21/2024 |  | 04/16/2024 |  |  | 04/16/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025459272 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Disability | Required Intervention | Life Threatening | Death | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to having coronary ischemia, pulmonary hypertension and valvular disease. while on support, the patient coded, and cardiopulmonary resuscitation was performed. there was subsequent bleeding from the axillary access site. the patient had to have vascular surgery and more than ten units of blood products were provided. the patient had irreversible loss of all functions of the brain after the event. care was withdrawn and the patient expired. the relationship between the cause of the death and the impella is unknown. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Impact code death was inadvertently left off the previously submitted manufacture device report. | The investigation for the reported vessel damage has been completed. the pump was not returned for investigation. bleeding was reported from the graft anastomosis after cpr and defibrillation were administered. the patient was sent to the or for vascular repair. the cause of the vascular damage issue was most likely use related, based on given clinical details. the instructions for use referenced in the initial report is from ifu for impella 5.5 with smart assist for use during cardiogenic shock sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27862 | 1220648-2024-09909 | 19116669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116669&pc=OZD | 05/29/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/29/2022 | 04/16/2024 |  |  | 11/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166272 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported a 62-year-old female on impella 5.5 for mechanical circulatory support. while on support, the patient developed atrial fibrillation. the physician repositioned the unit sutured to skin and the pocket was sutured close. an echocardiogram was done and placement was confirmed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
27863 | 1220648-2024-09910 | 19116907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19116907&pc=OZD | 04/08/2024 | 09/26/2024 |  | 04/16/2024 |  |  | 04/08/2024 | 04/16/2024 |  |  | 10/31/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427120 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Disability | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 65-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced multiple complications. four days into support, the patient developed hemolysis. it was determined that there was insufficient volume in the left ventricle requiring routine volume to be added. due to the patient's overall condition, a do not attempt resuscitation (dnar) status was applied, and volume stopped being given. the hemolysis persisted. the same day, the patient developed a cerebral hemorrhage (and the relationship between impella pump and cerebral hemorrhage is unknown). it was noted because active treatment was discontinued and volume loading was not performed, venoarterial extra corporeal membrane oxygenation was no longer available. the following day, the patient died due to lack of circulation. the relationship between cerebral hemorrhage and death is unknown. the cause of death is unknown as well. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the stroke/cva issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of hemolysis was most likely due to patient condition. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections h1 and h6.
27864 | 1220648-2024-09911 | 19117156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117156&pc=OZD | 04/08/2024 | 05/28/2024 |  | 04/16/2024 |  |  | 04/08/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025452638 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella cp insertion, the impella access site was oozing. no hematoma was present. pressure was held and heparin was on hold until the bleeding was mitigated. the patient is stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was most likely anticoagulation management as heparin was stopped and the bleed was controlled. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27865 | 1220648-2024-09912 | 19117279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117279&pc=OZD | 04/09/2024 | 05/22/2024 |  | 04/16/2024 |  |  | 04/09/2024 | 04/16/2024 |  |  | 11/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434532 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 could not deliver despite attempts due to anatomy and so an impella cp was instead placed. the impella cp was delivered successfully via the axillary. the patient was reported as stable. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since the pump could not pass due to the patient anatomy. the instructions for use for impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings states ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ it also states ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27866 | 1220648-2024-09913 | 19117399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117399&pc=OZD | 04/10/2024 | 05/09/2024 |  | 04/16/2024 |  |  | 04/10/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684985 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had an overheating alarm, battery failure (red alarm). the aic was replaced. there was no report of patient involvement. | [5 empty] | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the reported battery overheating alarm issue was unable to be reproduced in the lab. this issue was most likely due to a rare, momentary communication glitch caused by the pbm. | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27867 | 1220648-2024-09914 | 19117498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117498&pc=OZD | 04/10/2024 | 05/17/2024 |  | 04/16/2024 |  |  | 04/10/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1126797 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported during servicing, the automated impella controller displayed a system self-check failure and was removed from service. there was no report of patient involvement. | [5 empty] | The investigation into the aic - system self-check error has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure mode was reproduced. upon bootup, the console rebooted automatically and showed the ¿system self check failure¿ screen. the root cause of system self check failure was pbm contamination. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27868 | 1220648-2024-09915 | 19117522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117522&pc=OZD | 06/04/2022 | 05/21/2025 |  | 04/16/2024 |  |  | 06/04/2022 | 04/16/2024 |  |  | 12/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179556 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2022 |  |  |  |  | The facility reported the 69 year old female patient was on impella cp support. while placing the initial sheath, the artery was dissected. the impella was switched to contralateral groin. while on support, the staff completed a bilateral ultrasound of lower extremities. there were no present pulses in either of the feet. impella cp was explanted and escalated to 5.5 faint doppler pulses were later present. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27869 | 1220648-2024-09916 | 19117686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117686&pc=OZD | 04/10/2024 | 04/29/2024 |  | 04/16/2024 |  |  | 04/10/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205051 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported during servicing, the automated impella controller was not detecting the purge cassette. there is no report of patient involvement. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. the purge disc retainer was found to be broken. the cause of the issue was confirmed to be a broken purge retainer. | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27870 | 1220648-2024-09917 | 19117892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19117892&pc=OZD | 04/10/2024 | 06/18/2024 |  | 04/16/2024 |  |  | 04/10/2024 | 04/16/2024 |  |  | 11/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431331 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella cp was implanted, a hematoma was observed at the right groin. the nurse pushed out the hematoma and held pressure. |  |  |  |  | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details was provided and no product was returned for analysis. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations .¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has been completed. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27871 | 1220648-2024-09918 | 19118009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118009&pc=OZD | 06/13/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 06/13/2022 | 04/16/2024 |  |  | 03/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199211 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 35 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a 35-year-old female on impella cp for mechanical circulatory support. it was reported that the patient on impella cp had cerebral vascular accident. the patient¿s neurological status is unclear. an evaluation was on going, sedation was stopped, but no adequate response. the patient is stable on support. the patient opens eyes but no response. the stroke is confirmed, still ventilated, cardiac computed tomography planned for patient and left ventricular assist device scheduled. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. timing of the stroke in relation to impella support (during or post-implant). detailed description of the symptoms experienced by the patient. neurological examination findings and any focal deficits observed. diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). laboratory test results, including coagulation profiles and cardiac markers. any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records. g2 foreign: from no to yes.
27872 | 1220648-2024-09919 | 19118133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118133&pc=OZD | 04/11/2024 | 05/07/2024 |  | 04/16/2024 |  |  | 04/11/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145436 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) had a broken purge disc and was removed from use. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. the root cause of the issue was a broken purge retainer. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
27873 | 1220648-2024-09921 | 19118277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118277&pc=OZD | 04/11/2024 | 05/30/2024 |  | 04/16/2024 |  |  | 04/11/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during implantation of the impella cp, the transition from dilator to sheath kinked back and would not advance. the sheath was removed, and implantation was successful with a replacement sheath. | [5 empty] | The investigation for the reported introducer damage has been completed. data log review not applicable. the product was not returned for review. based upon images returned, if the needle stick in the 14fr introducer is very far out towards the edge of the 14fr valve, the user is most likely to run into the sharp edge of the inner geometry of the peel away hub when doing single access. the cause of the introducer damage was most likely user related since the needle stick was likely too lateral. added- h6 med dev prob code 2889 in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27874 | 1220648-2024-09920 | 19118282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118282&pc=OZD | 06/14/2022 | 05/14/2025 |  | 04/16/2024 |  |  | 06/14/2022 | 04/16/2024 |  |  | 11/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160693 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2022 |  |  |  |  | The complainant reported that the 41-year-old male patient was placed on impella cp support. while on support, the patient experienced bleeding including hematoma requiring transfusion, and hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27875 | 1220648-2024-09922 | 19118663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118663&pc=OZD | 05/19/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/19/2022 | 04/16/2024 |  |  | 03/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a(b)(6)-year-old female was being placed on impella cp for mechanical circulatory support. it was reported that the medical staff was unable to deliver the impella pump due to the patient's vascular anatomy. after multiple attempts, the impella implant procedure was aborted | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27876 | 1220648-2024-09923 | 19118765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118765&pc=OZD | 04/11/2024 | 05/17/2024 |  | 04/16/2024 |  |  | 04/11/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614416 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility's biomedical engineer reported purge fluid not detected alarm occurred on the automated impella controller (aic). the aic was removed from use. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. purge disc retainer was found to be broken (broken at the post where the purge disc retainer was supposed to be fastened to the purge assembly. the logs were reviewed which confirmed the inability to detect the purge disc. the root cause of this issue was a broken purge disc retainer post. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27877 | 1220648-2024-09924 | 19118859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118859&pc=OZD | 06/17/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 06/17/2022 | 04/16/2024 |  |  | 03/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207366 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that a hematoma and significant amount of bleeding were observed at the access site post implant of the impella cp. one unit of blood was administered to the patient and systemic heparin and cangrelor were temporarily discontinued to manage the complications. the device was explanted on the following day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27878 | 1220648-2024-09925 | 19118990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19118990&pc=OZD | 04/11/2024 | 05/28/2024 |  | 04/16/2024 |  |  | 04/11/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456429 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after nine days of support, the patient was weaned and explanted. upon explant, the pump was observed to have a thrombus. the thrombus was adhered to the pigtail of the pump and there was known anticoagulation issues during the support in which the patient was on and off of anticoagulation medical support. the patient was reported to be stable. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cp in section: direct aortic insertion states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27879 | 1220648-2024-09926 | 19119018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119018&pc=OZD | 05/18/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/18/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183279 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient showed signs of sepsis while on impella support. it was also reported that during impella support hemolysis was suspected with high lactate dehydrogenase (ldh) results and a low hemoglobin. two units of packed red blood cells (prbcs) were transfused due to low hemoglobin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27880 | 1220648-2024-09927 | 19119140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119140&pc=OZD | 04/11/2024 | 05/29/2024 |  | 04/16/2024 |  |  | 04/11/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025447206 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the patient was on impella 5.5 support due to having non-radiating chest pressure with associated dyspnea. while on support, there were positioning issues. the impella was attempted to be repositioned at bedside. the doctors were unable to pull the catheter back. the physician popped the patients suture and pulled the graft and the impella back. the impella was too shallow and the pump metrics were off. the staff noted low flows and increased pressor support. the doctor decided to remove the pump. during the removal, the catheter sheared off and the impella motor and pump were still inside the aorta. the patient was emergently taken to the cardiovascular operating room for removal of the rest of the impella pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was filed. the pump was returned for investigation. the pump catheter was found stretched and separated into two pieces. biomaterial was present on the impeller. the cause of the positioning issue was use related, based on given clinical details. improper positioning could lead to a low pump flow issue. the cause of the product damage during removal was use related since pump catheter was sheared off due excessive force was applied. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27881 | 1220648-2024-09928 | 19119271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119271&pc=OZD | 06/18/2022 | 10/17/2025 |  | 04/16/2024 |  |  | 06/18/2022 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was then reported that following a purge cassette exchange during impella 5.0 support, an aic failed to recognize the newly installed cassette. a second cassette was placed in an attempt to troubleshoot the issue but the failure remained. the aic was subsequently swapped to resolve the failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | D2 revised common device name since the initial manufacturer device report number was submitted in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h5 was erroneously selected on the initial manufacturer device report number. h6 medical device problem code 2900 and component code 925 was erroneously entered on the initial manufacturer device report number. h6 investigation conclusions code 11 was erroneously selected on the initial manufacturer device report number. h8 was erroneously selected on the initial manufacturer device report number.
27882 | 1220648-2024-09929 | 19119280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119280&pc=OZD | 05/19/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/19/2022 | 04/16/2024 |  |  | 06/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104238 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported the 86 year old male patient was brought to the cardiac catheterization laboratory for high-risk percutaneous coronary intervention and was being placed on impella cp support. while attempting to place the patient on the impella cp support the physician was attempting to advance the guidewire and the physician felt the guide was too tight and the guide kinked. the physician switched out for 7x45 terumo destination sheath and was able to successfully implant the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27883 | 1220648-2024-09930 | 19119308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119308&pc=OZD | 04/11/2024 | 09/30/2025 |  | 04/16/2024 |  |  | 04/11/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165516 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 04/11/2024 |  |  |  |  | The complainant reported that the device exhibited a placement signal not reliable alarm. there was no reported patient harm. | The user facility reported a 68-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that automated impella controller (aic) had a controller error alarm and screen flickering while in use. the aic was exchanged. there was no patient harm. |  |  |  |  | The investigation for console (aic) controller alarm-yellow alarm has been completed. data logs were reviewed which confirmed the reported failure mode given there multiple controller error alarms (opm comm crc error ¿ event set # 1045) as well as multiple placement signal not reliable (opm invalid signal, optical pump connected - event set#1036) triggered during the clinical procedure. real time (rt) logs confirmed that signals received from optical bench ( snr, contrast, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. there was no observable screen flickering while console was in the lab for investigation. the console's front optical connector was inspected before cleaning and after cleaning. the consoles fringe pattern revealed no observable abnormality. there were numerous pbm communications errors observed in the imc logs. pbm was visually inspected for possible pbm contamination and pbm corrosion. there was no evidence for it. the controller error and loss of placement signal is due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. in addition, placement signal was lost immediately after the shockwave device used. the cause of the optical signal issues was damaged sensor due to the shockwave device interaction, based on given clinical details and data log review. the root cause for the screen flickering could not be determined since the failure could not be reproduced. added- d9 device return date corrections to the initial mfr report: d2-common device name. f6/f8 should have been left blank. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data analysis: imc logs confirmed the reported failure mode given there multiple controller error alarms (opm comm crc error ¿ event set # 1045) as well as multiple placement signal not reliable (opm invalid signal, optical pump connected - event set#1036) triggered during the clinical procedure. device analysis: console ((b)(6)) was sent back fs for further investigation. there was no observable screen flickering while console was in the lab for investigation there were numerous pbm communications errors observed in the imc logs. pbm was visually inspected for possible pbm contamination and pbm corrosion. there was no evidence for it. conclusion: the controller error and loss of placement signal is due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. in addition, placement signal was lost immediately after the shockwave device used. the cause of the optical signal issues was damaged sensor due to the shockwave device interaction, based on given clinical details and data log review. however, the root cause could not be determined. the root cause for the screen flickering could not be determined since the failure could not be reproduced. repair: before returning the console to the customer, field service instructed to; ensure all optical connectors are clean. perform a full functional check (0042-7316). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name updated. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27884 | 1220648-2024-09931 | 19119764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119764&pc=OZD | 06/20/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 06/20/2022 | 04/16/2024 |  |  | 03/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192548 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 60-year-old male on impella cp for mechanical circulatory support had leaking from the groin and thrombocytopenia. the surgeon added some sutures, and discussed the alternative anticoagulation regime, heparin via impella cp and iv. the patient's hemodyanmics are stable. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27885 | 1220648-2024-09932 | 19119910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19119910&pc=OZD | 04/11/2024 | 07/29/2024 |  | 04/16/2024 |  |  | 04/12/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1359283 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | A user facility in japan reported that when the automated impella controller (aic) was turned on the failure alarm sounded, and operation of this console was not possible. there was no patient harm reported. | [5 empty] | The investigation for the console (aic) red alarm has been completed. the console was returned for evaluation. the console was tested but the reported failure could not be reproduced. however, data logs were reviewed which confirmed that the console recorded ¿purge: sensor defective (voltage out of range),¿ which resulted in the ¿controller failure (purge pressure sensor defective) ¿ alarm, event #418. then, the console was shut down and restarted six times, but event #418 was persistent. this confirms the reported failure mode. the root cause of the intermittent controller failure alarm was most likely due to the purge pressure sensor malfunction. f6/f8 should have been left blank in the initial report. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
27886 | 1220648-2024-09933 | 19120354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19120354&pc=OZD | 07/26/2022 | 09/30/2025 |  | 04/16/2024 |  |  | 07/26/2022 | 04/16/2024 |  |  | 11/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 70-year-old male on impella cp for mechanical circulatory support. while on support, the doctor attempted to remove the peel away sheath. during the removal, the pump was inadvertently pulled back into aorta and removed entirely from body at p-8. a 14fr gore sheath was placed in the right femoral artery and patient was transported to icu. the patient is in stable condition. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-09933 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-09933 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-09933 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-09933. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-09933.
27887 | 1220648-2024-09934 | 19120434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19120434&pc=OZD | 04/12/2024 | 06/07/2024 |  | 04/16/2024 |  |  | 04/12/2024 | 04/16/2024 |  |  |  |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453878 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the support was ongoing when on day two of the support, the pump had a high purge pressure (pp) alarm which was troubleshooted by replacing the purge cassette. when this did not remedy the issue, the tissue plasminogen activator in purge was utilized for the high pp. | [5 empty] | The investigation for the reported high purge pressure has been completed. the product was not returned for investigation. the data logs showed a three-hour rise in purge pressure until the high purge pressure alarm appeared. this trend gradually decreased after some time. according to the clinical details, tissue plasminogen activator was added to the purge after seeing the purge pressure alarm. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and provided clinical details. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.6a revised from the initial submission in accordance with updated procedures. h.6 health effect - impact code was revised as previously entered incorrectly. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has since been completed. h.10 additional manufacturer narrative instructions for use were added in accordance with updated procedures. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27888 | 1220648-2024-09935 | 19120913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19120913&pc=OZD | 08/16/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 08/16/2022 | 04/16/2024 |  |  | 11/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157875 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and that a hematoma formed at the access site. blood products were administered and the use of a catheter ball to apply pressure were utilized to resolve the issue. the patient was successfully weaned from the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27889 | 1220648-2024-09936 | 19121040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19121040&pc=PYX | 08/30/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 08/30/2022 | 04/16/2024 |  |  | 04/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022083415 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [5 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the impella rp was added for support due to the patient¿s overall condition. the patient became septic and needed oxygenation through extra-corporeal membrane oxygenation (ecmo). the impella rp was explanted and the patient continued on ecmo support. patient pending a donor heart. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27890 | 1220648-2024-09937 | 19121937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19121937&pc=OZD | 05/12/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/12/2022 | 04/16/2024 |  |  | 12/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023177145 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the patient received volume for central venous pressure and a decrease in urine output. a urinalysis was ordered for pink urine and the heparin induced thrombocytopenia panel was ordered for thrombocytopenia, which was trending down. there was heparin in the purge only and it was decided to decrease the purge concentration. eventually, the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27891 | 1220648-2024-09938 | 19122009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122009&pc=OZD | 05/12/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 05/12/2022 | 04/16/2024 |  |  | 04/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022078963 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 56-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced continuous oozing at the groin site, regardless of the site having proper redressing techniques applied. oozing occurred overnight from the site, which was controlled with femostop placement. the device was explanted and hemostasis was achieved with manual pressure and compression. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27892 | 1220648-2024-09939 | 19122037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122037&pc=OZD | 05/12/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 05/12/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194664 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that there was access site bleeding observed. this was treated by redressing and matching the angle. additionally, the double anticoagulation therapy was addressed concerning aggrastat and heparin, where the aggrastat was discontinued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27893 | 1220648-2024-09940 | 19122109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122109&pc=OZD | 05/12/2022 | 05/13/2025 |  | 04/16/2024 |  |  | 05/12/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194916 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support it was noted that the pump was turned down due to hypertension and potential for worsening hemolysis. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27894 | 1220648-2024-09941 | 19122132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122132&pc=OZD | 05/12/2022 | 05/19/2025 |  | 04/16/2024 |  |  | 05/12/2022 | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196076 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support positioning issues occurred; the positioning alarms were on and it was confirmed by echocardiography (echo) and disabled. the impella was repositioned at the patient's bedside and pulled back under echo guidance. there were good flows and waveforms. pressure was increased to p8. the tuohy-borst was tightened. patient tolerated wean well and hemodynamics remained stable. no bleeding or hematoma noted to site post closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27895 | 1220648-2024-09942 | 19122185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122185&pc=OZD | 05/13/2022 | 04/16/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 02/29/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196781 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. pulsatile oozing and access site bleeding from the peel away valve was noted as the single access site was too close to the center of the valve causing loss of hemostasis. the second attempt at single access farther from the center was successful. hubbing the peel away sheath resolved the access site bleeding. it was reported that, due to the length of the case and the issues with bleeding, the procedure was halted for address at a later date. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27896 | 1220648-2024-09943 | 19122197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122197&pc=OZD | 05/13/2022 | 04/16/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 12/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177279 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. some oozing was noted at the site. this was resolved with angle matching. bleeding started again around the reposition sheath after the patient was turned several times. the bleeding was controlled with fem stop. heparin was withheld from purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27897 | 1220648-2024-09944 | 19122235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122235&pc=OZD | 05/13/2022 | 04/16/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 01/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188214 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a 47-year-old patient on impella cp support and the patient experienced a hematoma and bruising. three units of red blood cells (rbc) were given, and manual pressure was held. there were intermittent suction alarms at p4. the alarm was not sounding appropriately, so the aic monitor was exchanged. the air in line, change purge line, and change bag alarms started. the purge cassette was changed, alarms resolved, and the wave forms had better pulsatility. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27898 | 1220648-2024-09945 | 19122251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122251&pc=OZD | 05/14/2022 | 05/21/2025 |  | 04/16/2024 |  |  | 05/14/2022 | 04/16/2024 |  |  | 08/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124010 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 05/14/2022 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the pump stopped. the impella cp was explanted then the impella 5.5 was implanted. it was reported that the patient was successfully escalated to the new device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27899 | 1220648-2024-09947 | 19122480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122480&pc=OZD | 06/29/2023 | 10/07/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 05/31/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351878 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | Unspecified Respiratory Problem | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 36-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that worsening respiratory failure occurred with a probable relationship to the impella 5.5 device used on a patient. | [5 empty] | The investigation for the respiratory failure has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of respiratory failure could not be determined as the device was not returned nor sufficient clinical details. added-h6 component code 925 corrections to the initial mfr report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27900 | 1220648-2024-09948 | 19122481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122481&pc=OZD | 08/07/2023 | 05/06/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 06/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359958 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Respiratory Failure | [11 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that worsening respiratory failure and hemolysis occurred with a probable relationship to the impella 5.5 device used on a patient. | [5 empty] | The investigation for the reported respiratory failure and hemolysis has been completed. the device was not returned for evaluation. however, data logs were reviewed which showed pump ran without issue during support. there were intermittent alarms impella position in aorta triggered but resolved quickly. therefore, the root cause of the respiratory failure and hemolysis could not be determined. added-h6 impact code 4628 adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
27901 | 1220648-2024-09949 | 19122483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122483&pc=OZD | 01/08/2024 | 05/03/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 11/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430010 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that access site bleeding occurred with a probable relationship to the impella 5.5 device used on a patient. the bleeding was resolved with additional sutures. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since the bleeding was resolved by adding additional sutures. added- d6a/d6b, d4 updated model number, h1 corrected mfr site name and address | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
27902 | 1220648-2024-09950 | 19122484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122484&pc=OZD | 02/12/2024 | 09/12/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 11/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024426613 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | It was reported due to the patient's worsening right ventricular function after being implanted with the impella 5.5 device, the patient was taken to the cardiovascular operating room and implanted with an impella rp flex device. it was noted that the patient expired while on impella support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that worsening right ventricular failure occurred with a probable relationship to the impella 5.5 device used on a patient. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-09950 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27903 | 1220648-2024-09951 | 19122491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122491&pc=OZD | 02/16/2024 | 05/29/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 09/30/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395896 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 54 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The patient expired on support. upon abiomed's medical safety officer review, there is not enough information to associate the use of the device with the patient's outcome. | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that vt (tachycardia) occurred with a probable relationship to the impella 5.5 device used on a patient. |  |  |  |  | The investigation into the vf/vt/af/at (ventricular tachycardia) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. a.5 ethnicity was previously entered incorrectly. the patient's ethnicity is unknown. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 first name and middle name were previously entered incorrectly. first name was revised and middle name is unknown. g.2 foreign was entered incorrectly on the previous submitted report. h.5 revised as previously entered incorrectly. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation code. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation code has been updated to reflect discarded. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27904 | 1220648-2024-09952 | 19122492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122492&pc=OZD | 02/29/2024 | 05/07/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 07/31/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443420 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that limb ischemia occurred with a probable relationship to the impella 5.5 device used on a patient. the impella was removed to resolve this issue. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27905 | 1220648-2024-09954 | 19122496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122496&pc=OZD | 10/23/2023 | 05/08/2024 |  | 04/16/2024 |  |  |  | 04/16/2024 |  |  | 08/31/2025 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385989 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 56-year-old white undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that a hematoma occurred with a probable relationship to the impella cp device used on a patient. the hematoma resolved without treatment. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27906 | 1220648-2024-09955 | 19122523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122523&pc=OZD | 10/11/2022 | 03/17/2025 |  | 04/16/2024 |  |  | 10/11/2022 | 04/16/2024 |  |  | 05/31/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023231229 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 75-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had a small vessel size and a 14 x 13cm sheath would be adequate for use. however, attempts to advance and gently torque the catheter did not help with placement of the device. therefore, the 14 x 13cm sheath was exchanged for the impella 14 x 25 cm sheath and the second pump's sheath had an internal kink. a new pump catheter was used and was advanced past the kink area and placed successfully in the left ventricle pre-percutaneous coronary intervention. eventually, the device was weaned and and it was decided to remove the pump and place a dilator then remove the pump's sheath. both pump and sheath was removed without complications and preclose x2 closure was used for hemostasis. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
27907 | 1220648-2024-09956 | 19122543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122543&pc=OZD | 10/23/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 10/23/2022 | 04/16/2024 |  |  | 09/30/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146422 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2022 |  |  |  |  | The user facility reported a 71-year-old male on impella cp for mechanical circulatory support and stabilization. the patent was bleeding profusely from the impella groin site. the doctor attempted placing suture. the bleeding was unresolved. team reported possible crack in repositioning sheath. holding manual pressure at site, pausing systemic heparin, and immobilizing the patient. decision made to have vascular surgery to remove the impella cp and repair artery. the device was removed without issue. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27908 | 1220648-2024-09957 | 19122552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122552&pc=OZD | 11/28/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 11/28/2022 | 04/16/2024 |  |  | 10/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152760 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The user facility reported a 57-year-old male on impella cp for mechanical circulatory support. it was reported while the patient was on support, the angiogram of the left femoral artery via peel away revealed there was no flow. due to a lot of calcifications distal to the femoral insertion, the doctor chose not to perform an external bypass. additionally, mean arterial pressure dropped to 60, however the doctor decided to remove the impella cp. the patient remained stable. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27909 | 1220648-2024-09958 | 19122624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122624&pc=OZD | 08/18/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 08/18/2022 | 04/16/2024 |  |  | 03/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 2022061351 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/18/2022 |  |  |  |  | The complainant reported that there was access site bleeding noted at the groin on the 72 year old male patient. this was treated by giving the patient three units of blood, repositioning of the sheath angle and tightening of preclose sutures. however, access site bleeding persisted and the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27910 | 1220648-2024-09959 | 19122638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122638&pc=OZD | 08/19/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 08/19/2022 | 04/16/2024 |  |  | 04/30/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224326 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/19/2022 |  |  |  |  | The complainant reported a 65 year old male patient was implanted the impella device via single access and the saphenous vein graft (svg) was opened and stented. when the patient went for computed tomography (ct) scan a few days later, bleeding in the ventricles was noted. an external ventricular drain (evd) was placed to drain the fluid and help relieve the increased intracranial pressure (icp). the pump speed level was increased to p8 due to the patient requiring pressor support. impella wean that was initially scheduled was put on hold as the pressors were soft and the patient required two (2) units of blood. the heart rate was elevated with more ectopy and amiodarone was administered. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27911 | 1220648-2024-09960 | 19122641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122641&pc=OZD | 09/26/2022 | 04/01/2025 |  | 04/16/2024 |  |  | 09/26/2022 | 04/16/2024 |  |  | 12/31/2023 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166705 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 51 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The complainant reported a patient with heart failure/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the impella implant, the doctor tried to use the pulmonary artery sheath to no avail. the doctor noticed that the dilator lock (transparent semicircle) was attached to the pulmonary artery sheath. the staff broke the transparent disc with pliers and removed it once it was fixed in the indwelling sheath. the impella was left in the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d4: model number g2: foreign: yes.
27912 | 1220648-2024-09961 | 19122663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122663&pc=OZD | 09/29/2022 | 05/20/2025 |  | 04/16/2024 |  |  | 09/29/2022 | 04/16/2024 |  |  | 12/31/2022 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022025796 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The complainant reported that a 51 year old female patient was implanted with an impella cp for mechanical circulatory support. during support, the echocardiogram indicated the pump into pap muscle, the blue suture hubs were not secured as peel away was still in. the impella was independently explanted, weaned and patient no longer needed support. physician was happy with support; however, a small clot was noted in outflow after removal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27913 | 1220648-2024-09962 | 19122674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122674&pc=OZD | 09/30/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 09/30/2022 | 04/16/2024 |  |  | 04/30/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221016 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/30/2022 |  |  |  |  | The complainant reported an 86 year old patient brought to cath lab for hrpci (high risk percutaneous coronary intervention) after surgical turndown. patient placed on impella cp for hemodynamic support. while on impella support, oozing was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
27914 | 1220648-2024-09963 | 19122685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19122685&pc=OZD | 10/01/2022 | 04/16/2024 |  | 04/16/2024 |  |  | 10/01/2022 | 04/16/2024 |  |  | 06/30/2024 |  | 04/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236017 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2022 |  |  |  |  | The complainant reported that a 63 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed oozing at the impella cp access site. the site dressing was changed and the patient was switched to bicarb in purge and heparin was withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27915 | 1220648-2024-09831 | 19104481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104481&pc=OZD | 03/26/2024 | 04/23/2024 |  | 04/15/2024 |  |  | 03/26/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while the patient was on support, the automated impella controller (aic) failed at the purge disc. the team replaced the aic and impella cp support proceeded without harm or injury from the interruption. impella support was initiated after the error, patient was transferred in stable condition. no association to patient outcome can be determined. | [5 empty] | The aic was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. the issue with the purge disc not consistently being detected was reproduced. the root cause of the issue was determined to be that the purge disc retainer was loose due to a broken internal mounting post.
27916 | 1220648-2024-09832 | 19104519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104519&pc=OZD | 03/27/2024 | 06/03/2024 |  | 04/15/2024 |  |  | 03/27/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5. |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2025460702 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-male off-pump coronary artery bypass surgery was implanted with an impella device for support. while on support the patient experienced cardiac arrhythmia, limitation of ventricular arrhythmia. the pump placement was verified with echocardiogram. the impella pump remained and was successfully explanted after three days of support. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (cardiac arrhythmia, limitation of ventricular arrhythmia) issues has been completed since the original report was submitted. as the necessary information was not provided, the root cause of the vt/vf was not determined.
27917 | 1220648-2024-09833 | 19104553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104553&pc=OZD | 03/27/2024 | 04/18/2024 |  | 04/15/2024 |  |  | 04/02/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1549903 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported by the user facility during preparation, the automated impella controller (aic) did not recognize the purge cassette and disc. the disc was noted to be broken and consequently, the console was exchanged. there was no report of patient involvement. | [5 empty] | The automated impella controller (aic) has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. the console logs are consistent with complaint and show multiple purge change wizard error messages. inspection of the console revealed broken purge disk retainer, which resulted in console¿s inability to measure purge pressure and detect purge disk. the cause of the issue was a defective component.
27918 | 1220648-2024-09834 | 19104601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104601&pc=PYX | 04/04/2024 | 05/14/2024 |  | 04/15/2024 |  |  | 04/02/2024 | 04/15/2024 |  |  | 06/30/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024367566 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Exsanguination | Hemorrhage/Blood Loss/Bleeding | Asystole | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 49 year-old female patient with in post-operation for cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella device was placed via the right femoral vein without issue. the medical staff used a figure eight suture along with a suture retention device. during preparation to transfer the patient off the catheter lab table, oozing at the access site was noticed. manual pressure was held, and a femoral stop was placed. due to the patient¿s size, every time the right leg moved, the access site would ooze again. medical staff decided to reposition the femoral stop and remove the figure eight suture along with the suture retention device. the bleeding then began to increase. the femoral stop was removed, and manual pressure was again applied. the femoral stop was then replaced again. the patient was transferred to the intensive care unit, and upon arrival, had their heart rate drop to the 20s, with asystole. cardiopulmonary resuscitation was then administered. after the patient had coded for approximately one hour, the medical staff made the decision to call target organ damage, and the patient expired. | [5 empty] | The investigation has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since the doctor removed the mattress suture and then opted to remove the figure 8 suture and flowstasis along with activated clotting time (act) being 206 which is above the limit range. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ b.2 date of death and b.3 date of event revised as previously entered incorrectly. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 name prefix/title revised as previously entered incorrectly. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. h.10 h.6 codes 4755 was reported incorrectly on the previously submitted report. a new code was added to health effect - impact codes as it was inadvertently left off the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
27919 | 1220648-2024-09835 | 19104648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104648&pc=OZD | 04/06/2024 | 05/22/2024 |  | 04/15/2024 |  |  | 04/06/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449541 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced hemolysis. the physician was concerned the impella pump was affecting the patient's kidneys; and therefore, a decision was made to remove the impella pump to see if the patient's kidney function would recover. there was note of possible consideration for placing an impella 5.5 in the future. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. complaint device not returned for investigation. data logs reviewed finding no motor current spikes, no suction alarms and no positioning alarms observed. no indications to the hemolysis issue observed. the cause of the hemolysis issue cannot be determined since insufficient data provided. the instructions for use referenced in the initial report is from ifu for impella cp with smartassist system, sections: use of echocardiography for positioning of the impella catheter, hemolysis, suction, potential adverse events (united states) which now includes cardiac or vascular injury (including ventricular perforation). d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27920 | 1220648-2024-09836 | 19104678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104678&pc=OZD | 04/07/2024 | 05/24/2024 |  | 04/15/2024 |  |  | 04/07/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464055 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced oozing around the drain and was taken to the operating room to achieve hemostasis. support continued following the intervention for access site complication. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and not much information was provided in the clinical description. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27921 | 1220648-2024-09837 | 19104714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104714&pc=OZD | 04/06/2024 | 05/30/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453866 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 63 year old female patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis, identified by increased lactate dehydrogenase with coinciding hematuria. one unit of blood products was given, and impella device placement was optimized. the patient stabilized. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. clinical details state that hemolysis occurred due to hypovolemia & improper position of the pump. hemolysis was resolved with addition of volume and optimization of placement. data logs were reviewed which showed multiple suction alarms were triggered and resolved corresponding with clinical narrative of hemolysis. there was no spiked motor current (mc) for entire case. the impella was returned for evaluation. upon visual inspection, biomaterial was found inside the pump housing. no other issues were found on the rest of the pump. most likely, the biomaterial was ingested into the pump housing during removal of the pump. the root cause of hemolysis was most likely due to pump was not in the proper position & the patient's hypovolemia. the instructions for use referenced in the initial report is from ifu for impella 5.5 with smartassist for use during cardiogenic shock in sections: : potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿; use of echocardiography for positioning of the impella catheter; hemolysis; and suction. added- d9 device return date d4-updated model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27922 | 1220648-2024-09838 | 19104748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104748&pc=OZD | 04/08/2024 | 05/22/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  | 11/30/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430192 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 36 year old female patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the patient developed right leg ischemia, losing pulses. the patient was scheduled to be transitioned to another impella device to resolve the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is from the ifu for impella cp with smartassist system section: potential adverse events (united states) which now also includes cardiac or vascular injury (including ventricular perforation).¿ d4-updated model number added- d6a/d6b.
27923 | 1220648-2024-09843 | 19104899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19104899&pc=OZD | 04/09/2024 | 06/03/2024 |  | 04/15/2024 |  |  | 04/09/2024 | 04/15/2024 |  |  | 09/30/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024404638 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was leaking blood from the blue suture hub. the leaking was observed at the blue suture hub and the catheter distal to the reaccess y port. a considerate amount of blood was leaking. staff used dermabond around the blue suture hub at the connection and the bleeding stopped. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned for investigation and the clinical information provided is inconclusive.
27924 | 1220648-2024-09844 | 19105078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105078&pc=OZD | 04/10/2024 | 09/05/2024 |  | 04/15/2024 |  |  | 04/10/2024 | 04/15/2024 |  |  | 11/30/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426662 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp smart assist support and they experienced hemolysis while on support. the patient's urine was red and the impella was repositioned. the hemolysis improved, however a blood transfusion was performed as well. | Additional information received, that stated the blood transfusion was not for the hemolysis, it was for the patient's low lactate values. the patient's hemolysis was resolved with repositioning the pump. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | B5 updated with additional information. the hemolysis resolved with repositioning, so no further investigation will be reported. h6 component code added. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ the following have been reported incorrectly or missing from manufacturer device report. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. according to new information received h6 health effect - impact code 4643 blood transfusion is no longer applicable.
27925 | 1220648-2024-09845 | 19105116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105116&pc=OZD | 04/10/2024 | 06/12/2024 |  | 04/15/2024 |  |  | 04/10/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1251266 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 60 year old male patient with acute myocardial infarction and cardiogenic shock was implanted with an impella device for mechanical circulatory support. the automated impella controller (aic) had two purge cassettes leak. the aic was replaced successfully. the patient remains on support. | [5 empty] | The automated impella controller was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the aic - display issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure mode was unable to be reproduced. the root cause of the leakage was not determined due to the inability to reproduce. in accordance with updated procedures, sections d6 and d9 have been revised from the initial submission.
27926 | 1220648-2024-09846 | 19105154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105154&pc=OZD | 04/10/2024 | 12/11/2025 |  | 04/15/2024 |  |  | 04/10/2024 | 04/15/2024 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1493604 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 58-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the automated impella controller (aic) abruptly dropped 100 percent to 50 percent two times during transport. both occurences the aic would quickly return back to 100 percent. the aic was switched to a backup and there were no more issues. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - battery charging/other issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the reported battery level low (50%) issues were unable to be reproduced. the root cause of this issue was a defective pbm. random power cycling of powermod1 in pbm resulting in abrupt loss of communication to the batteries.
27927 | 1220648-2024-09847 | 19105279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105279&pc=OZD | 04/07/2024 | 06/05/2024 |  | 04/15/2024 |  |  | 04/15/2024 | 04/15/2024 |  |  | 10/31/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407799 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the impella insertion site during transfer from the community to tertiary center. upon arrival at the tertiary center, the team observed the repo sheath was sticking up almost at a 90 degree angle, and the hematoma was large. the patient was taken to the operating room for a planned impella 5.5 insertion. after impella 5.5 insertion, vascular surgery was performed to due an arterial cutdown to the right femoral artery to remove the impella cp and evacuate the hematoma. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was most likely patient movement. upon arrival to the hospital, the repositioning sheath was sticking up almost at a 90-degree angle and a large hematoma was noted as per the clinical description provided. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | H.6 health effect - impact code 4629 was inadvertently left off on the previously submitted manufacture device report.
27928 | 1220648-2024-09848 | 19105291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105291&pc=OZD | 04/10/2024 | 05/20/2024 |  | 04/15/2024 |  |  | 04/10/2024 | 04/15/2024 |  |  | 10/31/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024410003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 73-year-old in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during the impella cp setup and calibration, cp was placed in patient while pump was calibrating. when the impella was turned on the left ventricle (lv) waveforms were above ao. flows were high at 4.4l/min mean. the impella was exchanged. no thrombus was noted in the lv or on the impella. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow issue has been completed since the original report was submitted. the device was not returned for investigation. per the data logs, at the time of insertion, when the pump was started, motor current spikes were seen, and saturated flow was observed at p9. the root cause of the saturated flow issue was most likely biomaterial. in accordance with updated procedures, section d6 has been revised from the initial submission.
27929 | 1220648-2024-09850 | 19105327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105327&pc=OZD | 04/07/2024 | 05/23/2024 |  | 04/15/2024 |  |  | 04/07/2024 | 04/15/2024 |  |  | 11/30/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428662 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2024 |  |  |  |  | The user facility reported a 55-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was that the impella 5.5 abruptly stopped and then briefly restarted but stopped again. the impella was explanted. the patient was hemodynamically stable and no harm was reported. | [5 empty] | The investigation for the reported pump stop has been completed. pump was returned for evaluation. upon visual inspection, blood was found in the red handle. electrical testing was performed and electrical resistance between motor lines 1&2 and 1&3 were both found to be reading out of spec approximately between 1.2 megaohms and 1.6 megaohms. further inspection of the catheter was performed and a hole in the catheter was found at around the 31cm mark. data logs were reviewed and rt logs at the end of the case showed that an alarm #119 pump stop occurred. pump was able to be restarted for 3 minutes and again for 1 minute before alarm #119 occurred and pump stopped. no issues with motor current or purge observed. the root cause of the pump stop was due to an electrical short due to blood ingress into the red handle. the instructions for use referenced in the initial report is from ifu for impella 5.5 with smartassist for use during cardiogenic shock in sections: direct aortic insertion, using the automated impella controller with the impella catheter, and impella stopped. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
27930 | 1220648-2024-09851 | 19105448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105448&pc=OZD | 04/06/2024 | 05/23/2024 |  | 04/15/2024 |  |  | 04/07/2024 | 04/15/2024 |  |  | 07/31/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442731 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2024 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that following explant of an impella cp pump, the patient developed bilateral groin hematomas. an unspecified quantity of blood and fresh frozen plasma was transfused in response to the condition. support with a subsequently implanted pump remained ongoing following the intervention. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was determined to be anticoagulation management because of the high act (302) of the patient. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury."
27931 | 1220648-2024-09852 | 19105642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105642&pc=PYX | 04/08/2024 | 05/23/2025 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  | 10/31/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413686 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 74-year-old female patient with acute myocardial infarction, cardiogenic shock, and angina was implanted with an impella rp flex device for mechanical circulatory support. while the impella device was being prepped for the implant, the placement signal had an unsure reading, and the medical staff decided to implant another impella rp flex device instead. the second implant was successful. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical issue has been completed. no product was returned for evaluation. data logs were reviewed and the right log shows no issues with the 5s parameters on the pump. no placement signal spikes or "placement signal not reliable" alarms. clinical detail noted that the pump was not inside the body or running when the clinical team noted the suspected issue. no problem was found with the pump upon investigation. in accordance with updated procedures, section d6 has been revised from the initial submission. | The investigation of optical signal issue has been completed. data logs were reviewed and the real time (rt) log shows no issues with the 5s parameters on the pump. no placement signal spikes or "placement signal not reliable (psnr)" alarms. clinical detail noted that the pump was not inside the body or running when the clinical team noted the suspected issue. the returned product was visually inspected and there were no damages to the dp sensor. no psnr alarms produced, the 5s parameters were within specification, and it passed electrical testing. there was no problem was found with the pump upon investigation. d9 device returned to manufacturer date added f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-09852. g1 reporting contact email revised on manufacturer device report 1220648-2024-09852 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-09852. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-09852.
27932 | 1220648-2024-09853 | 19105765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105765&pc=OZD | 04/08/2024 | 06/05/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450972 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 37-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had blood tinged urine and elevated lactate dehydrogenase. the patient is stable. | [5 empty] | The investigation has been completed. the data logs showed no issue for the pump during support. the product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has since been completed. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
27933 | 1220648-2024-09854 | 19105938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19105938&pc=OZD | 04/08/2024 | 05/29/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  | 08/31/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443686 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, there were aorta alarms. the pump was repositioned, and the left ventricle signal significantly increased and was higher than the placement signal. impella flows have been sporadic throughout the case. the max and min flows were very narrow with average flows of 4.1liter per minute at p8. th staff believed this could be displaying saturated flows and there could be an obstruction to the cannula. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow issue has been completed since the original report was filed. the pump was returned for investigation. upon visual inspection, no biomaterial or blood ingress was seen on the pump; its parameters were found to be in range. however, data logs showed that pump experiencing saturated flows, and "impella in aorta" alarms occurred after saturated flows. the root cause of the saturated flow cannot be determined as the failure was unable to be reproduced. in accordance with updated procedures, section d6 has been revised from the initial submission.
27934 | 1220648-2024-09855 | 19106103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106103&pc=OZD | 06/20/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 06/20/2022 | 04/15/2024 |  |  | 12/31/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023167738 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the patient had bleeding at the access site including a hematoma, requiring a blood transfusion. the bleeding lasted for the entirety of the implant period, and the device was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g2 foreign: from no to yes.
27935 | 1220648-2024-09856 | 19106157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106157&pc=OZD | 07/14/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 07/14/2022 | 04/15/2024 |  |  | 09/03/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022135445 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding occurred at the intra-aortic balloon pump (iabp) insertion site. additionally, it was difficult to stop the bleeding. percutaneous coronary intervention was completed with impella cp support, and circulation was stable. the impella was removed to control bleeding from the contralateral side. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g2 foreign: from no to yes.
27936 | 1220648-2024-09859 | 19106334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106334&pc=OZD | 04/08/2024 | 05/28/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 36-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. the team attempted to deliver the 5.5 pump, however, there was anatomic limitations and the impella 5.5 failed to deliver. instead the team allowed for an impella cp and extracorporeal membrane oxygenation to be placed. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the patient had smalls vessels and small stature as per the clinical description. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27937 | 1220648-2024-09858 | 19106456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106456&pc=OZD | 07/21/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 07/21/2022 | 04/15/2024 |  |  | 03/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023200347 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | NA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was later noted to have significant lower limb ischemia. an anterograde sheath was placed, and the impella device remained in the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records. g2 foreign: from no to yes.
27938 | 1220648-2024-09860 | 19106474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106474&pc=OZD | 08/05/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 08/05/2022 | 04/15/2024 |  |  | 03/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206249 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a 65-year-old female on impella cp for mechanical circulatory support. it was reported that the patient on impella cp had arrhythmia. the patient was conscious with breathing issues and impella cp running at p4 due to arrhythmia. plans to explant and escalate to impella 5.5. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the root cause of the vf/vt/af/at could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
27939 | 1220648-2024-09862 | 19106633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106633&pc=OZD | 04/06/2024 | 06/07/2024 |  | 04/15/2024 |  |  | 04/09/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454482 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the catheter sleeve disconnected from the proximal portion near the insertion site leaving the catheter exposed and unsterile. the catheter sleeve was secured by tape following the event and support continued with the impella 5.5 pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) issue has been completed since the original report was filed. the pump was returned for investigation. the sterile sleeve was torn. as per the clinical information provided and product investigation, the root cause of the product damage (sterile sleeve) issue was use related to improper technique while repositioning the pump. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the product has been returned since then and the investigation has been completed.
27940 | 1220648-2024-09863 | 19106925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106925&pc=OZD | 04/09/2024 | 05/28/2024 |  | 04/15/2024 |  |  | 04/09/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458543 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 81-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. the user facility reported that an impella 5.5 could not be implanted due to excessive bleeding caused by the patient's anatomy. two units of blood were transfused to treat the condition and the impella 5.5 implant was aborted. an impella cp was successfully implanted in the patient. | [5 empty] | The investigation into the delivery issues and the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was most likely due to patient condition since excessive bleeding was caused by the patient's anatomy. the root cause of the delivery issue was most likely due to patient condition since due to patient anatomy the pump was unable to successfully be inserted. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27941 | 1220648-2024-09864 | 19106943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106943&pc=OZD | 08/06/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 08/06/2022 | 04/15/2024 |  |  | 04/30/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216474 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 74-year-old male on impella cp for mechanical circulatory support. it was reported that the patient on impella cp had thrombus. the impella cp was successfully implanted and switched to i-cor 2 hours later due to left ventricular thrombus. the device was explanted within 3 hours of admission to unit. no other information is available. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the thrombus was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27942 | 1220648-2024-09865 | 19106976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19106976&pc=OZD | 04/09/2024 | 05/28/2024 |  | 04/15/2024 |  |  | 04/09/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450086 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 50-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the during support there was high purge pressures. tissue plasminogen activator protocol was initiated, and the pressure alarms resolved. no patient harm was reported, and the patient remained on support. | [5 empty] | The investigation for the reported high purge pressure has been completed. pump was not returned for investigation. data logs were returned and reviewed. there was a gradual rise in purge pressure over a few hours before the purge pressure related alarms. no motor current spikes were seen and the pump was at p9. tissue plasminogen activator (tpa) was added to the purge which helped resolved the alarms. per the clinical information and data log review, the root cause of the high purge pressure was most likely biomaterial. the instructions for use for the impella 5.5 with smartassist in section "using the automated impella controller with the impella catheter" states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 impact code 4644 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27943 | 1220648-2024-09866 | 19107359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19107359&pc=OZD | 04/09/2024 | 08/14/2024 |  | 04/15/2024 |  |  | 04/09/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1731799 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 75-year-old female of unknown race and ethnicity in non-st segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during the priming process prior to patient implant, the nurse could hear the sound of the cassette trying to prime but the automated impella controller (aic) displayed a purge fluid not detected message. the setup was verified to be correct. despite two more attempts, the failure persisted. as a result, both the aic and purge cassette were replaced to resolve the issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of priming issue has been completed. the logs confirmed the reported complaint. there were multiple instances of a purge system failure alarm due to a blocked piston. visual inspection of the purge assembly revealed excessive purge fluid residue, indicating leakage of purge fluid during clinical use. purge fluid residue was also found within the purge motor assembly, preventing the purge motor from spinning both by hand and via telnet commands to the console. the purge system failure was due to buildup of glucose on the purge stepper motor gasket. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09866. d2 common device name, f6/f8-should have been left blank, g1 reporting contact fax number missing.
27944 | 1220648-2024-09867 | 19107516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19107516&pc=OZD | 03/15/2024 | 04/15/2024 |  | 04/15/2024 |  |  | 03/15/2024 | 04/15/2024 |  |  | 03/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203671 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 70-year-old male patient experienced a loss of pulses (ischemia) while on the impella cp for mechanical circulatory. as a result of the noted issue, peel away sheath was removed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed.
27945 | 1220648-2024-09868 | 19107762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19107762&pc=OZD | 04/06/2024 | 05/17/2024 |  | 04/15/2024 |  |  | 04/09/2024 | 04/15/2024 |  |  |  |  | 04/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1376689 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the automated impella controller (aic) alarmed for a temperature rise on the battery. this red alarm resolved. the pump remained on for support and to date the aic has not been replaced. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for reported console (aic) red alarm has been completed. the aic was returned for evaluation. the controller was tested finding ¿battery temperature high alarm¿ was a false alarm, and battery temperature never rose above 29.4 degc. this resulted from a single erroneous data sample reported by battery pack bms to pbm. data logs were reviewed which are consistent with complaint and show ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. . therefore, the cause of the false "battery temperature high" alarm was most likely a defective pbm board. added- h6 impact code 4629 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27946 | 1220648-2024-09873 | 19110131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110131&pc=PYX | 04/10/2024 | 05/30/2024 |  | 04/15/2024 |  |  | 04/10/2024 | 04/15/2024 |  |  | 08/31/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024407118 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pressure Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 59-year-old male post operative left ventricular assist device (lvad) was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the patient developed hemolysis and the pump displayed high purge pressure alarms. the position of the pump was checked, and tissue plasminogen activator was infused. a thrombus was suspected and therefore, the impella rp flex was removed. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was filed. the device was returned for investigation. per the product investigation, data logs and clinical information provided, the root cause of the hemolysis issue was biomaterial found around the impeller blades. the root cause of the high purge pressure issue was biomaterial. found on the impeller and in the inflow of the pump. the root cause of the low pump flow issue was biomaterial found in the inflow cage and around impeller. section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27947 | 1220648-2024-09874 | 19110288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110288&pc=OZD | 05/15/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 05/15/2022 | 04/15/2024 |  |  | 02/29/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191480 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed bleeding, including hematoma, and hemolysis. the patient was transfused with blood products as a result. the bleeding and hemolysis was reported to have occurred from insertion through explanation of the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27948 | 1220648-2024-09875 | 19110455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110455&pc=OZD | 05/15/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 05/15/2022 | 04/15/2024 |  |  | 02/28/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055715 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing at the impella access site. heparin was discontinued and the patient was transfused with packed red blood cells and platelets. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27949 | 1220648-2024-09876 | 19110575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110575&pc=OZD | 05/19/2022 | 05/19/2025 |  | 04/15/2024 |  |  | 05/19/2022 | 04/15/2024 |  |  | 12/31/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176078 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient female patient was on impella cp support. while on support, the patient had hemolysis that cleared up overnight. the patient had a cool right foot, which is indicative of limb ischemia. the patient also experience ischemic stroke. the pump was repositioned overnight and found that the locking system was not locked, causing the pump to migrate when the patient was turned. flows were significantly reduced. the pump was repositioned and locked properly. flows were returned to the previous support level of p7. the patient also experienced a groin site bleed and received blood for hemoglobin drop from the groin site bleed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27950 | 1220648-2024-09879 | 19110594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110594&pc=OZD | 05/19/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 05/19/2022 | 04/15/2024 |  |  | 12/31/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient with high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. access site bleeding was noted later and the sheath needed to be removed as the initial sheath stick was too close to the impella catheter, and the sheath was restricted. a guide was placed and repositioning of the sheath was required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27951 | 1220648-2024-09878 | 19110645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110645&pc=OZD | 05/09/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 05/09/2022 | 04/15/2024 |  |  | 01/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189529 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had hemolyzed labs along with ongoing hematuria. furthermore, the patient developed a hematoma in the groin area. levophed was initiated, and under echocardiogram the impella was repositioned. it was noted that the flows improved, and urine began to clear. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27952 | 1220648-2024-09881 | 19110719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110719&pc=OZD | 05/09/2022 | 04/02/2025 |  | 04/15/2024 |  |  | 05/09/2022 | 04/15/2024 |  |  | 08/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128357 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp device for mechanical circulatory support. during impella support, the physician attempted to reposition cp and pulled it out into the aorta. the physician could not get it back in place, so they removed it and left the still-in-place peel away in. due to tamponade of the heart and lungs combined with shallow impella placement, flow issues arose. the repositioning helped the impella to have optimal flows. additionally, a clot was discovered after opening the abdomen. the patient was given packed red blood cells, sodium bicarbonate and systemic heparin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27953 | 1220648-2024-09883 | 19110779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110779&pc=OZD | 05/09/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/09/2022 | 04/15/2024 |  |  | 01/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported that a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the pump encountered resistance halfway through the aorta, leading to limb ischemia in the left femoral artery. surgical cutdown and removal were performed to address the limb ischemia. the impella cp implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27954 | 1220648-2024-09884 | 19110839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110839&pc=OZD | 05/10/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/10/2022 | 04/15/2024 |  |  | 01/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190996 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp device for mechanical circulatory support. it was noted on esophageal echo that a significant amount of air was detected in the artery resulting in the diagnosis of air embolism shock. the cause of the air embolism is unknown. at the initiation of extracorporeal membrane oxygenation, a large amount of air was aspirated. the impella device was effectively discontinued and replaced with an intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27955 | 1220648-2024-09885 | 19110852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110852&pc=OZD | 05/11/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/11/2022 | 04/15/2024 |  |  | 02/29/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump suction remained above p1 despite repositioning. the impella cp was turned off and explanted and a new impella cp was successfully placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27956 | 1220648-2024-09887 | 19110882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110882&pc=OZD | 05/11/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/11/2022 | 04/15/2024 |  |  | 02/29/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023196573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The user facility reported a 84-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the peel-away sheath was removed and the patient experienced bleeding at the access site after repositioning the sheath. therefore, the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events : acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27957 | 1220648-2024-09888 | 19110988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19110988&pc=OZD | 05/20/2022 | 05/21/2025 |  | 04/15/2024 |  |  | 05/20/2022 | 04/15/2024 |  |  | 02/29/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and that there were aortic position alarms overnight. the false alarms were due to motor current threshold had no real pulsatility. a new echo confirmed that the catheter had flipped towards mitral and elevated plasma free. it was further reported that the patient had an ischemic lower extremity and the impella cp was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27958 | 1220648-2024-09889 | 19111017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19111017&pc=OZD | 05/11/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/11/2022 | 04/15/2024 |  |  | 01/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023191172 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The user facility reported that a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. due to intermittent reduced flow/suction alarms, the positioning of the pump was adjusted, causing the device to be removed and the catheter to be reinserted. the overall position was then optimal; however, the patient developed leg ischemia and it was decided to wean and remove the device. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. h1 type of reportable event: from malfunction to serious injury.
27959 | 1220648-2024-09890 | 19111090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19111090&pc=OZD | 05/12/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/12/2022 | 04/15/2024 |  |  | 12/31/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023171263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported that a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the insertion site which good angle matching and forward tension sutures were applied for treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma." ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27960 | 1220648-2024-09857 | 19111277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19111277&pc=OZD | 04/08/2024 | 05/28/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  | 12/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309633 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that following removal of the peel away sheath and insertion of the repositioning sheath, the patient¿s right lower extremity (rle) was mottled. doppler pulse was not found in the leg. the patient continued to complain of increasing pain to rle, numbness, and unable to move toes. the medical doctor decided to place a distal perfusion catheter, followed by tissue plasminogen activator infusion through the catheter. the patient was transferred to another facility for higher level of care and surgical evaluation. vascular surgery was performed. the patient became unstable and was taken emergently to the operating room for impella removal, open femoral repair, and rp hematoma evacuation. four units of red blood cells were given. computed tomography surgery in operating room was performed. hemodynamics were evaluated and transesophageal echocardiography stated that instability was more than likely from the hemorrhagic shock, and therefore, elected to not place another impella at this time. the site was closed. post repair angiogram was performed and showed adequate perfusion bilaterally. | [5 empty] | The investigation for the reported limb ischemia and vascular damage has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and vascular damage could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is from ifu for impella cp with smart assist system in section: potential adverse events (united states), which now includes a statement "cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿
27961 | 1220648-2024-09861 | 19111470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19111470&pc=OZD | 04/08/2024 | 05/28/2024 |  | 04/15/2024 |  |  | 04/08/2024 | 04/15/2024 |  |  | 06/30/2025 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 71 year-old-male patient with cardiogenic shock and acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed access site bleeding. packed red blood cells were transfused and the patient survived the event. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was determined to be anticoagulation management because the act values were high that resulted in the cause of the issue and heparin was also stopped to control the bleeding. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
27962 | 1220648-2024-09849 | 19111907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19111907&pc=OZD | 04/10/2024 | 06/07/2024 |  | 04/15/2024 |  |  |  | 04/15/2024 |  |  |  |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450354 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being delivered to support a patient admitted in a stemi (non-st segment elevation myocardial infarction) with ami/cgs (acute myocardial infarction / cardiogenic shock) and history of mvd (mitral valve disease) cad (coronary artery disease) with cabg (coronary artery bypass grafting) . the team had delivery issues and observed the pump-wire to be in need of replacement due to wire damage within the pump. | [5 empty] | The investigation for the delivery and removal issues has been completed. only the pump was returned for investigation. no kink was present on cannula. no physical damage was noted on the returned pump. additionally, the pump ran as expected during the test. according to the given clinical report, the medical doctor felt resistance while removing the guidewire after the pump placement. the medical doctor removed the pump along with the guidewire. the same resistance was felt the second time. fluoroscopy revealed a bent guidewire. additionally, replacing the device did not resolve the issue. the cause of the delivery issue was most likely patient condition related since during rewire and replacement device had the same issue with delivering the pump into the aorta. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a.1 revised as this information was inadvertently omitted from the previously submitted report. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.1 revised as previously entered incorrectly. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27963 | 1220648-2024-09886 | 19112259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19112259&pc=OZD | 05/19/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/19/2022 | 04/15/2024 |  |  | 02/29/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. two suctions events were observed on the arrhythmia induced cardiomyopathies and on arrival, the placement signal and left ventricle waveform indicated that the decoupled pump was likely not in the proper position and was hemolyzing. the physician decided to explant the impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27964 | 1220648-2024-09877 | 19112901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19112901&pc=OZD | 05/09/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/09/2022 | 04/15/2024 |  |  | 08/31/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had limb ischemia with cool extremities. to address this, an angiogram was conducted down to the right lower extremity (rle) to restore blood flow to the right leg and foot, thereby reinstating pulsatility. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of vf/vt/af/at was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27965 | 1220648-2024-09880 | 19113051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19113051&pc=OZD | 05/09/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/09/2022 | 04/15/2024 |  |  | 01/31/2024 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp device for mechanical circulatory support. during support, percutaneous coronary intervention was performed, but post-closure was unsuccessful. removal of the long sheath resulted in hematoma formation and a shorter 13cm impella sheath was then inserted. during angiography, slow flow was observed down the iliac, with positive pulses detected via doppler. however, upon retracting the sheath, another hematoma formed, accompanied by limb ischemia. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27966 | 1220648-2024-09882 | 19113115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19113115&pc=OZD | 05/19/2022 | 04/01/2025 |  | 04/15/2024 |  |  | 05/19/2022 | 04/15/2024 |  |  | 07/31/2023 |  | 04/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022134978 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient with heart dysfunction was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced some oozing at the site. no additional information about the resolution of the bleeding was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27967 | 1220648-2024-09829 | 19102602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19102602&pc=OZD | 06/29/2023 | 05/02/2024 |  | 04/13/2024 |  |  | 03/14/2024 | 04/13/2024 |  |  | 05/31/2025 |  | 04/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344634 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 56 year old male patient presenting with cardiomyopathy. one day after insertion of the device, the patient endured "worsening ventricular tachycardia" the prompted administration of "amiodarone adn 1l ivf(weak, diaphoretic)" and a cardioversion. approximately 4 days later, the patient endured atrial fibrillation that prompted a second cardioversion. impella support continued until the pump was explanted and the patient was bridged to transplant. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-updated model number. added- d6a/d6b.
27968 | 1220648-2024-09830 | 19102612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19102612&pc=OZD | 02/03/2024 | 03/06/2025 |  | 04/13/2024 |  |  | 03/14/2024 | 04/13/2024 |  |  | 10/31/2025 |  | 04/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | Cardiogenic Shock | Pericardial Effusion | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/14/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 71 year old female patient presenting with acute myocardial infarction and cardiogenic shock for impella support. after support, approximately 2 minutes post removal of the device, the patient became hypotensive and tachycardic. pericardial effusion/cardiac tamponade was noted. the allegation comes with a "possible" relationship to the impella device. intervention was administered via a pericardiocentesis and cpr, however, the patient expired. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured cardiogenic shock- worsening with a "possible" relationship to the impella device used: ¿the patient was diagnosed with cardiogenic shock, scai stage c. despite mechanical circulatory support via the impella cp and other life-saving measures, the patient failed to recover. on (b)(6) 2024, the patient had the impella cp removed and became symptomatic for cardiac tamponade (reported separately). during the pericardiocentesis, the patient developed ventricular fibrillation. brief cpr was administered as well as unsuccessful direct current cardioversions. time of death was called at 1646.¿. the patient outcome fatal. |  |  |  |  | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states) states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ d4-updated model number. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B5: revised with additional information received from the clinical site. h6: revised to reflect the additional information received d6a and d6b revised from the initial submission in accordance with updated procedures on manufacturer device report 1220648-2024-09830. f6/f8: should have been left blank on manufacturer device report 1220648-2024-09830. g1: reporting contact email revised on manufacturer device report 1220648-2024-09830 in accordance with updated procedures. g1: reporting contact fax number was missing on manufacturer device report 1220648-2024-09830. h6: component code revised from the initial submission in accordance with updated procedures on manufacturer device report 1220648-2024-09830.
27969 | 1220648-2024-09779 | 19095084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095084&pc=OZD | 05/20/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/20/2022 | 04/12/2024 |  |  | 10/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138517 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support the patient had bleeding including a hematoma from insertion to removal that required a blood transfusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27970 | 1220648-2024-09780 | 19095196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095196&pc=OZD | 05/16/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/16/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The user facility reported patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at the access site and patient was febrile and looked more septic. heparin was discontinued, the site was angle matched, and was redressed. hemostasis was unable to be achieved, and an arterial sheath was placed to control bleeding and patient returned to cardiac catheterization laboratory for surgical closure. no information was provided on interventions used for alleged sepsis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27971 | 1220648-2024-09781 | 19095261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095261&pc=OZD | 05/20/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/20/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198450 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s access site was observed to be bleeding (slight oozing). the patient's heparin was discontinued in the purge solution to resolve the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27972 | 1220648-2024-09782 | 19095469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095469&pc=OZD | 05/21/2022 | 04/17/2025 |  | 04/12/2024 |  |  | 05/21/2022 | 04/12/2024 |  |  | 12/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174292 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella placement, the guidewire was unable to cross the aortic valve and the intervention was aborted. additionally, it was reported that the patient had hemolyzed labs. to address the issue, the p level was dropped to p6, and an echocardiogram was performed, and the pump was repositioned. the hemolysis did not resolve and the impella pump was explanted to treat the patient's septic shock that was unrelated to the impella cp device. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27973 | 1220648-2024-09783 | 19095613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095613&pc=OZD | 05/14/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/14/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPLLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190651 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, slight oozing was noted from the swan site. angle of insertion was maintained, and the device was locked. the following day, manual pressure was held until the dressing was taken down. additional oozing was identified overnight. a femostop was used and the patient received 2 units of packed red blood cells. three days later, the access site was oozing again. the dressing was removed and re-dressed with angle matching. the oozing subsided after a new dressing was applied. there were no further reports of oozing after the new dressing was applied. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27974 | 1220648-2024-09784 | 19095630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095630&pc=OZD | 05/22/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/22/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188088 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2022 |  |  |  |  | The user facility reported patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced low purge pressure, which was resolved by switching the purge fluid to d20. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27975 | 1220648-2024-09785 | 19095677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095677&pc=OZD | 05/17/2022 | 03/10/2025 |  | 04/12/2024 |  |  | 05/17/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023193536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was repositioned, due to the depths of the original position; however, the device was considered to be shallow and was removed. there were no reported issues regarding the pump and the patient was still on 2 pressor support and needed escalation. eventually, the device was successfully weaned. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
27976 | 1220648-2024-09786 | 19095802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095802&pc=OZD | 05/23/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/23/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023189147 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device there was significant bleeding at the pump site, therefore, sutures were added around the sheath and manual pressure was held for 60 minutes, which controlled the bleeding. it was also reported that during impella support the patient's pedal pulses could not be found and the limb was cool. the device was explanted to resolve limb ischemia as patient was stable and weaned successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27977 | 1220648-2024-09787 | 19095806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095806&pc=OZD | 05/17/2022 | 05/21/2025 |  | 04/12/2024 |  |  | 05/17/2022 | 04/12/2024 |  |  | 12/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171206 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and while on support, the impella pulled back into the aorta, resulting in a large amount of blood in the chest cavity. a chest tube was placed, and the patient was brought to the operating room to remove the impella cp, repair the artery and surgically close the impella site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿. ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27978 | 1220648-2024-09788 | 19095912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19095912&pc=OZD | 05/18/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 12/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181140 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that no pulses were able to be found in either foot, but the legs were warm to the touch and the mean arterial pressure the increased to the 90s in mmhg under impella support | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27979 | 1220648-2024-09790 | 19096248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096248&pc=OZD | 03/14/2024 | 09/17/2025 |  | 04/12/2024 |  |  | 03/19/2024 | 04/12/2024 |  |  | 10/31/2025 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415441 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that placement signal not reliable (psnr) alarm occurred on impella 5.5. psnr alarm audio was turned off and placement monitoring was disabled. it was further reported that the patient developed a gi bleed and heparin was held. the patient is stable. placement signal (ps) issue is not reportable as there is no/limited potential for adverse impact to patient and ps does not affect impella performance. | [5 empty] | The investigation into the placement signal issue and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the optical signal issue was a damaged optical fiber line, but the exact details of the damage cannot be determined as product was not returned. the root cause of the vascular damage gi bleed and its relation to impella were not determined since product was not returned and there is a lack of clinical details. d6a, d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ . ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. h6 medical device problem code revised to include 2917 device sensing problem and component code 925 pump from the initial manufacturer device report number 1220648-2024-09790.
27980 | 1220648-2024-09789 | 19096263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096263&pc=OZD | 05/18/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 11/30/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150513 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the impella cp was explanted after the patient developed leg thrombosis and would have another pump implanted via the axillary approach. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the thrombus was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
27981 | 1220648-2024-09791 | 19096412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096412&pc=OZD | 03/21/2024 | 05/21/2024 |  | 04/12/2024 |  |  | 03/21/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025450954 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Infusion or Flow Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An 80-year-old patient received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. it was reported that after inserting and starting the impella cp, the pump flow increased to 3 l/min and immediately decreased to a pump flow <1 l/min. attempts at repositioning did not solve the problem and the correct position of the impella cp in the left ventricle was confirmed fluoroscopically. before inserting the impella cp, the user checked that there were no thrombi in the left ventricle. the user removed the impella cp, cleaned it and reinserted it. the impella cp then worked as intended. a day after, the patient was noted to have expired. the official cause of death is not available, however, there is currently not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the low pump flow issue has been completed since the original report was submitted. the pump was not returned for investigation. the cause of the low pump flow issue cannot be determined since no sufficient clinical data provided and product was not returned. in accordance with updated procedures, section d6 has been revised from the initial submission.
27982 | 1220648-2024-09792 | 19096416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096416&pc=OZD | 03/22/2024 | 04/23/2024 |  | 04/12/2024 |  |  | 03/22/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299754 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Run on Battery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient had the impella device and intra-aortic balloon pump concurrently running for support. when in the intensive care unit, the automated impella controller (aic) alarmed for battery failure. the aic was removed and replaced. there was no harm or injury from interruption. | [5 empty] | The automated impella controller (aic) has been received from the customer, however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) battery failure alarm has been completed. the data logs from the reported day of event are consistent with complaint and show battery failure (#360) and battery level low (#23) alarms about 4 hours after pump was started. battery data embedded in ltlog files revealed cell imbalance of battery 2. the console was returned for evaluation and testing of the battery confirmed battery lock-out due to permanent failure condition. the cause of the aic issue was a defective battery. added- h6 impact code 4629, h6 component code 765 d4-updated model number.
27983 | 1220648-2024-09793 | 19096509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096509&pc=OZD | 03/21/2024 | 05/23/2024 |  | 04/12/2024 |  |  | 03/25/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED INC |  |  | US |  | 2025438424 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported from the study that an 80-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during the support the patient had anemia. the patient experienced blood loss through arterial accesses and hemolysis. the urine was noted to be bloody. the event was determined to be moderate in severity and possibly related to the impella device and its procedure. no secondary intervention has been taken at the time of this mdr writing. the patient is being monitored. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿.
27984 | 1220648-2024-09794 | 19096664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096664&pc=OZD | 05/18/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188736 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was found to be positioned very deep but still in the loading patient. the night prior, ¿low outflow¿ alarms occurred and the plan was to take the patient to reposition the impella. the patient was taken to the lab to reposition, and the pump came out. the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
27985 | 1220648-2024-09795 | 19096832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096832&pc=OZD | 05/29/2022 | 04/08/2025 |  | 04/12/2024 |  |  | 05/29/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202213 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient developed some oozing at the impella cp insertion site at the groin. heparin was discontinued and the patient was transfused with 2 units of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27986 | 1220648-2024-09796 | 19096848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096848&pc=OZD | 03/29/2024 | 06/04/2024 |  | 04/12/2024 |  |  | 03/29/2024 | 04/12/2024 |  |  | 04/30/2021 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1422055 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was noted from publication, a 72-year-old male diagnosed with acute fulminant myocarditis was admitted with cardiogenic shock was started treatment with impella device and extracorporeal membrane oxygenation (ecmo) the support was for over two weeks when the team determined the patient's condition was not improving. the ecmo was explanted with sepsis concerns, and the impella pump remained until explant. the team later published on the event and the piece was read/reviewed in april 2024. the patient had suffered thrombosis at the location of the pump in the aorta and expired from septic shock. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation into the thrombus and the unrelated death has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. the root cause of the death cannot be determined due to insufficient clinical details. d6a, d6b.
27987 | 1220648-2024-09797 | 19096886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096886&pc=OZD | 05/18/2022 | 05/14/2025 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 11/30/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Cardiogenic Shock | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced a delivery issue as the physician crossed the valve with the j wire and pigtail. the impella inlet got stuck in the valve. the impella was removed and placed in a heparin saline bowl. a v18 was used on the second attempt but was unsuccessful again. the valve ballooned again with the truflow balloon and reattempted to place the impella pump. the impella crossed with difficulty but was started immediately. patient remained on support and was sent to unit. care was withdrawn and the patient expired. no allegations were made towards the impella cp for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial 1220648-2024-09797 was provided in sections b2, b5, d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27988 | 1220648-2024-09798 | 19096895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096895&pc=OZD | 03/29/2024 | 05/28/2024 |  | 04/12/2024 |  |  | 03/29/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 85 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The complainant in japan reported that a literature observational study on long term impella use reported the patient had impella support due to having st elevated myocardial infarction. the patient had support for 14 days and when the pump was explanted, the catheter had blood clots attached to both the inlet and outlet cage parts, and a blood clot was also found inside the cannula. the patient expired after the pump was removed. there is no relationship between the impella and the cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. ¿ | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, section h10 has been revised from the initial submission.
27989 | 1220648-2024-09799 | 19097108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097108&pc=OZD | 05/18/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200006 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was initially in the papillary muscle and the doctor attempted to reposition the pump. when the pig was released from the papillary muscle, the pump popped into the aorta. the pump was then removed, and any clots were removed by running the pump in sterile water on the table. the impella was recrossed with better position in the mid-left ventricle. moreover, the patient had ablated ventricular tachycardia. the impella was weaned removed in the lab. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27990 | 1220648-2024-09800 | 19097307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097307&pc=OZD | 05/18/2022 | 05/13/2025 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188358 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support and that the patient had signs of hemolysis. therefore, the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27991 | 1220648-2024-09801 | 19097322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097322&pc=OZD | 04/02/2024 | 05/13/2024 |  | 04/12/2024 |  |  | 04/02/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1441953 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) console reports "purge fluid not detected." there was no reported harm to an unknown age patient. | The console was exchanged for a backup. this complaint was related to a clinical procedure, but there were no adverse events related to this event. |  |  |  |  | The automated impella controller (aic) has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) purge disc/flag issues has been completed. console logs from the reported day of event are mostly consistent with complaint and show ¿purge disk not detected¿ alarm (#402) which was resolved in 15s, however the case start wizard was unable to proceed due to purge system malfunction, resulting in ¿purge fluid not detected¿ error message. logs show multiple ¿piston blocked¿ messages as well as repeating pud communication issues, not resolved after pud resets. after few failed attempts to prime/de-air, the console was exchanged for backup, where issues were resolved. the aic was returned for evaluation which revealed immobilized purge motor shaft (by dried-up glucose), which was causing ¿piston blocked¿ messages as well as pud communication issues, due to increased current draw from the motor. after cleaning the shaft and its gasket, the purge was operating within specification. the cause of the aic issue was contamination. added- b5-addl event narrative, h6 impact code 4629 d4-updated model number, corrected udi. in accordance with updated procedures, sections d2 has been revised from the initial submission.
27992 | 1220648-2024-09802 | 19097426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097426&pc=OZD | 04/02/2024 | 05/21/2024 |  | 04/12/2024 |  |  | 04/02/2024 | 04/12/2024 |  |  | 11/30/2025 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Difficult to Insert | Activation, Positioning or Separation Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella device and extracorporeal membrane oxygenation (ecmo) in place prior to a transcatheter aortic valve replacement (tavr) procedure. upon evaluation patient had limited access available. the physician decided to removed the impella pump and proceed with the tavr procedure, which was successfully completed. the same impella pump was replaced for continued support with the ecmo. however, the artery was noted to immediately have bleeding issues and hemostasis could not be obtained. the physician utilized the reaccessed port to maintain vascular access and removed impella pump and placed a dry seal sheath. then valve recrossed utilizing dry seal sheath and impella pump was placed through dry seal sheath. it was noted the impella placement difficult but was ultimately placed at the time successfully. at conclusion of case, the impella pump started to have continual suction alarms. it was suggested that the impella pump possibly may not be optimally placed any longer. fluoroscopy performed to evaluate the impella pump position, and it was noted the impella pump had migrated too shallow, towards mitral. the physician attempted to advance, but ultimately was untentionally pulled back into aorta. the impella pump had to be removed with the decision to prep the patient again to place new impella pump. the patient's condition was noted as stable. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported access site bleed, positioning issue, and low pump flow has been completed. pump and data logs were not provided foe investigation. as per provided clinical details, the edwards sheath used for the tavr was exchanged for a 14 french peel-away sheath to facilitate the reintroduction of the impella. however, immediately upon reintroduction of the impella, extreme arterial bleeding occurred. initial attempts to control the bleeding through the peel-away sheath were unsuccessful. to manage this complication and maintain vascular access, the team utilized the reaccess port, removed the impella, and inserted a dry seal sheath. despite the difficulty, the impella was eventually reintroduced successfully through the dry seal sheath. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. also, fluoroscopy revealing that pump had migrated too shallow and was too close to the mitral valve. the doctor attempted to reposition the device by advancing it, but it was accidentally pulled back into the aorta. consequently, the impella had to be removed. the cause of the low pump flow issue was improper positioning of the device, based on given clinical details. the cause of the positioning issue was use related since pump was pulled back into aorta during repositioning. the instructions for use referenced in the initial report is from ifu for impella cp with smart assist system sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation.¿
27993 | 1220648-2024-09803 | 19097487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097487&pc=OZD | 06/03/2022 | 04/07/2025 |  | 04/12/2024 |  |  | 06/03/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195034 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there were suction alarms while the patient was on impella cp support. the patient experienced ventricular fibrillation arrest while on impella cp support, and they were treated by being placed on a lidocaine drip. the device placement signal showed that the device was in the aorta and improper impella placement was confirmed with echocardiography. the impella was repositioned and advanced into the left ventricle. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27994 | 1220648-2024-09804 | 19097548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097548&pc=OZD | 05/19/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/19/2022 | 04/12/2024 |  |  | 07/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022105337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient¿s access site was observed to be bleeding (oozing). sutures were placed and heparin was discontinued to resolve the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27995 | 1220648-2024-09805 | 19097725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097725&pc=OZD | 06/17/2022 | 04/07/2025 |  | 04/12/2024 |  |  | 06/17/2022 | 04/12/2024 |  |  | 03/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following insertion, the pump was pulled out accidentally by staff while removing drapes. a replacement pump was placed without complication. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27996 | 1220648-2024-09806 | 19097791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097791&pc=OZD | 04/02/2024 | 05/06/2024 |  | 04/12/2024 |  |  | 03/31/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437566 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following a scheduled transcatheter aortic valve replacement procedure, the patient experienced uncontrollable bleeding from the femoral access site. the pump was explanted as a result of the condition and a vascular cutdown to surgically closed the site was performed. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
27997 | 1220648-2024-09807 | 19097873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097873&pc=OZD | 04/03/2024 | 05/14/2024 |  | 04/12/2024 |  |  | 04/03/2024 | 04/12/2024 |  |  | 10/31/2025 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421306 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an attempt to place an impella cp device in an unknown age male with acute myocardial infarction/cardiogenic shock was unsuccessful due to the patient's anatomy. the physician was unable to advance past the iliac in the right femoral artery, and the left femoral artery was too calcified to attempt. the procedure was aborted for the impella device placement. there was no report of patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since it was unable to advance past the iliac in the right femoral artery and the left femoral artery was too calcified to attempt. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27998 | 1220648-2024-09808 | 19097890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19097890&pc=OZD | 05/19/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/19/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the patient experienced bleeding around the impella site. a stitch was placed around the site and femostop was applied, which resolved the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27999 | 1220648-2024-09809 | 19098006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098006&pc=OZD | 04/03/2024 | 05/16/2024 |  | 04/12/2024 |  |  | 04/03/2024 | 04/12/2024 |  |  | 05/31/2025 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352551 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an attempt to implant an impella device in a 76-year-old male in acute myocardial infarction/cardiogenic shock was unsuccessful. resistance was encountered during the attempted impella device implant. it was noted that after backloading the pump over the .018 guidewire, the pump was unable to advance through the sheath. the anatomy was tortuous, and manipulation of the sheath did not resolve the noted problem. as a result, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the difficulty inserting/removing pump through introducer and the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that patient had tortuous anatomy and pump was unable to pass through the peel away sheath. attempts were made to buddy wire and manipulate the sheath to advance the cannula without success. the root cause of the difficulty inserting the pump through the introducer cannot be determined as no product was returned for evaluation. in accordance with updated procedures, section d6 has been revised from the initial submission.
28000 | 1220648-2024-09810 | 19098061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098061&pc=OZD | 05/12/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/12/2022 | 04/12/2024 |  |  | 01/31/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188807 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 56-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced significant venous ooze. the venous ooze was resolved with sutures and pressure dressing. additionally, the patient experienced a thrombus for which a thrombectomy was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28001 | 1220648-2024-09811 | 19098106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098106&pc=OZD | 05/13/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/13/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199397 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during the time of wire removal, the physician was unable to pull the wire out of impella cp catheter. the device was removed. the wire was noted to have kink on the distal tip. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
28002 | 1220648-2024-09812 | 19098220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098220&pc=OZD | 05/13/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/13/2022 | 04/12/2024 |  |  | 08/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132358 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a 84-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. due to bleeding at the sheath, single access was aborted and additional access to left groin was obtained with ultrasound. the guidewire and a diagnostic catheter were utilized, then the impella was backloaded onto wire. the device was inserted and advanced to the left ventricle and turned on with no issue. it was reported that the impella was explanted and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28003 | 1220648-2024-09813 | 19098336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098336&pc=OZD | 05/13/2022 | 04/07/2025 |  | 04/12/2024 |  |  | 05/13/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023199806 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the access site. femstop was applied at belt pressure and the bleeding controlled. the heparin was held due to partial thromboplastin time (ptt). patient received 2 units packed red blood cells (prbc). follow up was performed and the bleeding was reportedly resolved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28004 | 1220648-2024-09814 | 19098344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098344&pc=OZD | 04/03/2024 | 05/20/2024 |  | 04/12/2024 |  |  | 04/05/2024 | 04/12/2024 |  |  | 06/30/2025 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369993 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had ectopy when the p-level was higher than p-6. the pump position was confirmed. after the pump was removed, distal pulses were not present. angiogram revealed a clot in the tibioperoneal trunk. a thrombectomy was performed and flow was restored. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported arrhythmia, limb ischemia, and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia, limb ischemia, and thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is from ifu for impella cp with smartassist system section: potential adverse events (united states), which now includes cardiac or vascular injury (including ventricular perforation).¿ added- h6 clinical code 1721 added- d6a/d6b.
28005 | 1220648-2024-09815 | 19098418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098418&pc=OZD | 05/13/2022 | 10/06/2025 |  | 04/12/2024 |  |  | 05/13/2022 | 04/12/2024 |  |  | 12/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced lower extremity ischemia from insertion to removal of the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication. vascular size or variations thereof detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. g1: updated manufacturer email.
28006 | 1220648-2024-09816 | 19098496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098496&pc=OZD | 04/06/2024 | 05/15/2024 |  | 04/12/2024 |  |  | 04/06/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463781 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 46-year old male patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the access site at the axillary developed a hematoma and a clotted surgical drain. medical staff returned the patient to the operating room for evacuation and surgical stripping. the patient had a hemoglobin drop and received one unit of packed red blood cells. the patient remains on support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use state: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28007 | 1220648-2024-09817 | 19098518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098518&pc=OZD | 05/28/2022 | 04/07/2025 |  | 04/12/2024 |  |  | 05/28/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 05/28/2022 |  |  |  |  | The user facility reported a 49-year-old patient was presented with st elevation myocardial infarction (stemi) and in cardiogenic shock (csg). a diagnosis catheter revealed a ninety-nine (99) % occlusion of the left main (lm) and severe disease in the mid/distal left anterior descending artery (lad) and circumflex arteries. it was reported that after the introducer and repositioning sheaths were removed, a hematoma was noted to be present at the insertion site with significant oozing. manual pressure was held for twenty (20) minutes, however the site bleed continued. the patient was receiving intracoronary integrilin and the physician ordered a fem stop to be placed. bleeding was successfully managed with fem stop placement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28008 | 1220648-2024-09818 | 19098576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098576&pc=OZD | 05/14/2022 | 10/20/2025 |  | 04/12/2024 |  |  | 05/14/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient exhibited atrial fibrillation, absence of distal pulses (attributed to non-dopplerable pre-procedure pulses), gastrointestinal (gi) bleeding, bleeding at access site, and a slight fever, suggesting the presence of severe sepsis. systemic heparin was withdrawn, and the patient received packed red blood cells (prbc). after treatment the physician elected to wean and explant the impella device. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
28009 | 1220648-2024-09819 | 19098632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098632&pc=OZD | 05/15/2022 | 04/02/2025 |  | 04/12/2024 |  |  | 05/15/2022 | 04/12/2024 |  |  | 08/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134223 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced oozing around the insertion site. the site was redressed to maintain angle of entry. the dressing was reinforced, and forward suturing of blue butterfly was effective in the reduction of oozing. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28010 | 1220648-2024-09820 | 19098671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19098671&pc=OZD | 04/06/2024 | 05/23/2024 |  | 04/12/2024 |  |  | 04/06/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that when placing the impella cp, the impella cp kept getting hooked in the papillary muscle. the physician had to remove the impella, re-wired, and placed it again. once the patient was in the room, the impella cp came across the aortic valve and the physician was not able to put it back. the patient was taken to the catheterization lab and the impella cp was replaced. the patient is stable. the new impella device was successfully placed. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the pump was not returned for investigation. the cause of the positioning issue most likely was use related due to improper technique as the pump shape got hooked in papillary and then pulled out from the aortic valve with pump cannula kinking during repositioning. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28011 | 1220648-2024-09821 | 19099932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19099932&pc=OZD | 04/06/2024 | 07/12/2024 |  | 04/12/2024 |  |  | 04/06/2024 | 04/12/2024 |  |  |  |  | 04/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145121 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Date/Time-Related Software Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 52-year-old male of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the clinician was unable to change the date and time or see the purge infusion screen on the automated impella controller (aic). there were no active alarms. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console display issues has been completed. data logs from the date of occurrence show numerous interactions with display menu and the infusion screen. there is no observable evidence from the logs to indicate denial of input when interacting with the console screens. the console was returned for evaluation. both the date time screen and the infusion history screens were navigated to and interacted with. there were found to be operational to console specification. the root cause for device interaction issues could not be determined as the failure could not be reproduced. added- d9 device return date d2-corrected common device name f6/f8 should have been left blank in the initial report.
28012 | 1220648-2024-09822 | 19100139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19100139&pc=OZD | 03/13/2024 | 04/24/2024 |  | 04/12/2024 |  |  | 04/10/2024 | 04/12/2024 |  |  | 04/30/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213919 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported that a 57-year-old male was implanted with an impella cp pump for mechanical circulatory support. following impella cp implant via femoral access, a swan was placed through the right side of the neck for right heart catheterization. a planned percutaneous coronary intervention was performed to the left anterior descending artery with impella support. shortly after, an emergent computed tomography scan was performed as a large hematoma developed at the swan access site and extended into the upper chest and superior mediastinum. exploration of the right neck was performed and the carotid artery was found to be punctured. a surgical repair was completed, but active bleeding caused the patient to go into cardiogenic shock. the patient was given two units of packed red blood cells to help stabilize the patient. support with the implanted pump was maintained for two more days before planned explant. two days following explant, the patient was reported to have a stage three pressure injury. initial reports believed that the deep pressure injury was caused by the patient's illness and their immobility due to pump implant. triad and foam dressings were placed to help treat the condition. each of these events were documented to be potentially related to the impella procedure. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage is most likely due to use as the artery was punctured during access. the active bleeding caused the patient to go into cardiac shock. the root cause of the vt/vf is most likely due to use. a few days later it was reported that the patient was experiencing a pressure injury that was caused by the complications experienced. the root cause of the nerve compression is most likely due to use as the patient experienced vascular damage and vt/vf due to the punctured carotid artery. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use (ifu) for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ as noted in the impella ifu, cardiogenic shock, bleeding, and vascular injuries are known potential adverse event associated with impella support.
28013 | 1220648-2024-09824 | 19100456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19100456&pc=OZD | 04/06/2024 | 05/20/2024 |  | 04/12/2024 |  |  | 04/06/2024 | 04/12/2024 |  |  | 11/30/2025 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434115 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 would not pass through an anatomically challenging portion of the subclavian into the innominate artery. a decision was made to abort the 5.5 axillary insertion and attempt a right axillary cp insertion. the cp was placed successfully. the patient is stable, and no harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue is most likely due to patient condition as the patient had diseased vessels. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings states ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ it also states ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28014 | 1220648-2024-09825 | 19100570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19100570&pc=OZD | 05/15/2022 | 09/30/2025 |  | 04/12/2024 |  |  | 05/15/2022 | 04/12/2024 |  |  | 07/31/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114156 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed hemolysis and the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number. h6 component code revised from the initial submission in accordance with updated procedures.
28015 | 1220648-2024-09826 | 19100806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19100806&pc=OZD | 05/16/2022 | 04/02/2025 |  | 04/12/2024 |  |  | 05/16/2022 | 04/12/2024 |  |  | 07/31/2022 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1495146 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, it was observed that the patient had a thrombotic proximal stent occlusion. additionally, the patient went into a pulseless electrical activity (ventricular tachycardia/ventricular fibrillation) rhythm. advanced cardiac life support protocol was initiated. distal limb perfusion to impella cp insertion site was also observed. as a result, device was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia, thrombus, vf/vt/af/at was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
28016 | 1220648-2024-09823 | 19101167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19101167&pc=OZD | 05/15/2022 | 04/02/2025 |  | 04/12/2024 |  |  | 05/15/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the patient had some bleeding at the access site. they went over angle matching and they decreased systemic heparin. in addition, suction alarms were reported when running over p7. the patient received 1 liter of fluid. the physician was at bedside for placement echo which was measuring shallow and inlet close to anterior leaflet of mitral valve. the physician advanced the impella. the patient was successfully weaned and impella explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28017 | 1220648-2024-09827 | 19101632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19101632&pc=OZD | 05/18/2022 | 03/10/2025 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 02/29/2024 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023196084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a 80-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the patient experienced bleeding a the insertion site which was possibly due to a deep incision prior to the pump sheath's insertion. however, the bleeding was controlled and required continuous manual pressure. to control the bleeding better, the device was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28018 | 1220648-2024-09828 | 19101980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19101980&pc=OZD | 05/18/2022 | 04/12/2024 |  | 04/12/2024 |  |  | 05/18/2022 | 04/12/2024 |  |  | 09/30/2023 |  | 04/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144026 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported an (b)(6) patient needing mechanical circulatory support. it was reported that the patient presented with chest pain and experienced a myocardial infarction. the patient was taken to the cath lab and was found to be hypotensive with a reduced ejection fraction and left anterior descending artery (lad) down. the impella cp was placed, and percutaneous coronary intervention (pci) of lad was done without complication. echo was done and revealed a large clot in the left ventricle. it was reported that the hemodynamics were good. patient was weaned and the impella removed. manual pressure was held and a femstop placed after manual pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
28019 | 1220648-2024-09740 | 19086496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19086496&pc=OZD | 05/13/2022 | 10/01/2025 |  | 04/11/2024 |  |  | 05/13/2022 | 04/11/2024 |  |  | 01/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient experienced oozing from the access site and had signs of hemolysis. additionally, the patient was prepped for femoral artery repair by cardiac surgeon. the medical team performed a cut-down procedure and then repaired the femoral artery using vessel loops and cauterization. after treatment, the physician elected to wean and explant the impella device. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28020 | 1220648-2024-09741 | 19087007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087007&pc=OZD | 03/19/2024 | 05/21/2024 |  | 04/11/2024 |  |  | 03/19/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 58 YR |  | Female |  |  |  | Malfunction | No | Yes | Electromagnetic Interference | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was noted from a publication, a 58-year-old female with acute coronary syndrome and cardiogenic shock had an impella device placed before performance of percutaneous coronary intervention to stabilize hemodynamics during the procedure. when the impella device was set at p-4, it was noted the device caused magnetic interference with the carto 3 system, resulting in an inability to acquire positional information and perform mapping. however, when the team dropped the p level to p-2, issue resolved, and the ablation procedure was completed. there was no report of adverse effects to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. sections d4 and h4: information is unavailable. no product-specific information other than the brand name and manufacturing location was noted in the publication. | The investigation for the reported ventricular fibrillation (vf) and positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf and positioning issue could not be determined. added-h6 med dev prob code 3009. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28021 | 1220648-2024-09742 | 19087068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087068&pc=OZD | 03/21/2024 | 05/06/2024 |  | 04/11/2024 |  |  | 03/26/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1352449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) had a purge pressure transmitter installation failure/purge line click-on not detected alarm. staff checked the connection on the purge cassette and there were no issues. the aic was replaced. there was no patient harm. | [5 empty] | The automated impella controller device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge issue ¿ other has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the issue with detecting the purge disc was confirmed, and the purge flag was discovered to be broken. the root cause for the purge not detected alarm was a broken purge flag.
28022 | 1220648-2024-09743 | 19087113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087113&pc=OZD | 03/26/2024 | 05/01/2024 |  | 04/11/2024 |  |  | 03/26/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025457976 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support due to patient being in shock for four days. while on support, the patient emergently got up and accidentally pulled the impella and the access site started to bleed. the patient bled out almost three liters out of the jackson pratt drain. heparin was withheld and the site stabilized. the patient remained on impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to patient movement since the patient accidentally pulled the impella when he got up and site started bleeding. added- d6a/d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28023 | 1220648-2024-09744 | 19087173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087173&pc=OZD | 03/28/2024 | 03/10/2025 |  | 04/11/2024 |  |  | 03/28/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1315813 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported during teaching, a broken automated impella controller (aic) purge disc channel was noted. there was no report of patient involvement. | [5 empty] | The automated impella controller (aic) was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) purge disc/flag issues has been completed. console logs did not contain any data relevant to this complaint. the console was returned for evaluation. the purge disc retainer was found to be broken (retainer post broken). the root cause of the issue was a broken purge disk retainer. added d9 device return date corrections to the initial mfr report: d2 updated common device name d4-updated model number e4 should have been no g1 mdr reporting contact email.
28024 | 1220648-2024-09745 | 19087197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087197&pc=OZD | 04/03/2024 | 05/28/2024 |  | 04/11/2024 |  |  | 04/03/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450259 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 41-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was successfully placed but it was observed that the limb was ischemic. the ischemia diagnosed by absent distal pulses was treated by a bypass. the pump remained on for support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the limb ischemia ¿ reversible has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. positioning issue was also reported. per the clinical information, the root cause of the positioning issue was most likely use related as during repositioning the pump was pulled ack across the valve.
28025 | 1220648-2024-09747 | 19087276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087276&pc=OZD | 03/29/2024 | 06/04/2024 |  | 04/11/2024 |  |  | 03/29/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Heart Block | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | Information obtained by japan from literature reported that seven days following impella cp implant, the patient suffered a cardiac arrest due to a complete atrioventricular block. hypoxemia and metabolic acidosis progressed gradually in response to the 25 minute long arrest. a transvenous pacemaker was placed temporarily and conduction was restored. the patient survived the intervention. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported arrhythmia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia could not be determined as the device was not returned nor sufficient clinical details.
28026 | 1220648-2024-09749 | 19087304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087304&pc=OZD | 03/30/2024 | 06/06/2024 |  | 04/11/2024 |  |  | 03/30/2024 | 04/11/2024 |  |  | 08/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402887 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a patient of unknown age and gender was implanted with an impella device for mechanical circulatory support. the impella device showed a sensor failure alarm when connected with the automated impella controller (aic). the implanted impella device was exchanged, which resolved the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported optical issue has been completed. the impella was returned for evaluation. this was an out-of-the-box failure. when plugged into lab console, impella controller error was observed. the pump failed the light continuity test and there was a light leak observed in the red pump plug. there was no bond issue or sign of excessive force, but it was observed that the cables were wound around each other. the root cause of the optical signaling issue was determined to be a damaged fiber from a manufacturing process (internal) issue. added- d9 device return date d4-updated model number. added- d6a/d6b.
28027 | 1220648-2024-09748 | 19087330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087330&pc=OZD | 04/07/2024 | 06/10/2024 |  | 04/11/2024 |  |  | 04/07/2024 | 04/11/2024 |  |  | 05/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346630 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2024 |  |  |  |  | The user facility reported a male patient of an unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp pump experienced complications with the placement signal. it was noted that the signal would disappear temporarily during support. as a result of the noted issue, the decision was made to replace the component. there was no patient harm related to the issue. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical signal has been completed. the data logs were reviewed which showed that the controller was the cause of the optical signal, and not the pump. the device was returned for evaluation. upon evaluation, no problem was found. there was no issue found during the case. the device functioned/operated to specification. added- d9 device return date added- d6a/d6b.
28028 | 1220648-2024-09750 | 19087449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087449&pc=OZD | 03/30/2024 | 08/20/2025 |  | 04/11/2024 |  |  | 03/30/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116498 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed a high purge pressure alarm while bicarb purge solution was running. the aic was replaced and the issue was resolved. there was no patient harm. | [5 empty] | The automated impella controller was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1: mdr reporting contact email. d2: device name has been updated. d4: model number updated. f6: an f8 should have been left blank.
28029 | 1220648-2024-09752 | 19087481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087481&pc=OZD | 03/30/2024 | 05/08/2024 |  | 04/11/2024 |  |  | 03/30/2024 | 04/11/2024 |  |  | 08/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 49-year old patient was implanted with an impella device for mechanical circulatory support. during escalation to another impella device¿s support, blood began to pool from under the patient's access site. a cutdown was performed to control the bleeding, but the patient became hypotensive and coded. an unspecified quantity of packed red blood cells and fresh frozen plasma were administered. cardiopulmonary resuscitation (cpr) was also initiated. the patient stabilized and pacing wires were placed to assist with bradycardia, but the patient coded a second time. after thirty additional minutes of cpr, medical staff decided to halt support and called the time of death. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28030 | 1220648-2024-09753 | 19087506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087506&pc=OZD | 03/30/2024 | 05/16/2024 |  | 04/11/2024 |  |  | 03/30/2024 | 04/11/2024 |  |  | 10/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420669 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 55-year old female patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. after seven days of support, the medical staff observed a hematoma that resulted in an operating room visit for evacuation. the patient remained on support for 14 days before the impella device was explanted. the patient remains in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement, patient moved their arm outside of the approved range of motion and caused trauma to the arm and created a hematoma. the instructions for use in the initial report was from the ifu for impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states), section: general patient care considerations, and section: entering cardiac output. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission.
28031 | 1220648-2024-09746 | 19087579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087579&pc=OZD | 03/29/2024 | 05/31/2024 |  | 04/11/2024 |  |  | 03/29/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2025455140 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during insertion of the guidewire for impella implant the tip of the guidewire was stretched/frayed and could not be used. there was no harm to the patient reported. the guidewire was replaced. the defective guidewire was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - peripheral vessels) and the stroke/cva issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the product damage (guidewire damage - peripheral vessel) was not determined as no product or images were returned. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28032 | 1220648-2024-09754 | 19087632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087632&pc=OZD | 04/01/2024 | 10/20/2025 |  | 04/11/2024 |  |  | 04/11/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314435 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/04/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller (aic) placement signal was not accurate upon initiation of impella support. all troubleshooting was performed without resolution. the aic was replaced. | The user facility reported that the automated impella controller (aic) placement signal was not accurate upon initiation of impella support. all troubleshooting was performed without resolution. the aic was replaced to resolve the issue. there was no progression with patient status. care team and family decided to withdraw care and the patient expired. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. h.3 device not returned to manufacturer for evaluation was inadvertently checked on the original manufacture device report (mdr). the device was returned for evaluation. h.10 additional manufacturer narrative in the previously mdr mentioned that the device was reported. this field as been updated to say that the device has been received and an evaluation is underway. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data analysis: imc logs were reviewed, and no evidence was found of any placement-related alarms. imc logs did not reveal information relevant to the reported complaint (inaccurate placement signal). rt logs were reviewed and analyzed for evidence of an inaccurate placement signal. there were no clear/obvious signs of an inaccurate placement signal evident in the rt logs. potential anomalies in other 5s parameters were identified. there were anomalous spikes in the raw optical sensor data that corresponded to anomalous spikes in the signal not reliable (snr) and contrast data. one instance of the optical data anomalies corresponded to a placement signal low alarm. see attachments for more details. the root cause of the anomalies was not able to be determined from the data analysis. device analysis: console (b)(6) was sent back to field service for further investigation. the reported complaint could not be reproduced during testing. the console¿s optical system was tested with a fixed-length optical test cavity, and the 5s parameters were within specification. the console¿s optical system was physically inspected and no defects were found. the root cause of the reported complaint was not determined, and no issues were found with the console. before returning the console to the customer, field service was instructed to replace the optical bench as a precaution. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added death "1802" for the automated impella controller (aic) "765" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | D10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated.
28033 | 1220648-2024-09755 | 19087651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087651&pc=OZD | 03/18/2024 | 05/20/2024 |  | 04/11/2024 |  |  | 04/01/2024 | 04/11/2024 |  |  | 02/28/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311142 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced access site bleeding. bleeding was noted on the graph site and inside the sterile repositioning sleeve. the doctor applied another 0-0 silk suture to the graft around the blue suture hub and the bleeding subsided. the patient remained on impella support for four additional days and was successfully explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, doctor applied another 0-0 silk suture to the graft around the blue suture hub and the bleeding subsided. the instructions for use referenced in the initial report is from the ifu for impella 5.5 with smartassist for use during cardiogenic shock, section: general patient care considerations, section: entering cardiac output , and section: potential adverse events (united states) which also now includes cardiac or vascular injury (including ventricular perforation).¿ added-d6a/d6b.
28034 | 1220648-2024-09751 | 19087656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087656&pc=OZD | 03/30/2024 | 05/15/2024 |  | 04/11/2024 |  |  | 03/30/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was opened for implantation and the automated impella controller (aic) console displayed a "pump defective" error when plugging the impella in. staff tried to unplug and plugging the impella back in to no avail. staff made the choice to use a second impella and it was successful. | [5 empty] | The investigation into the pump not detected issue has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that they made multiple attempts at un-plugging and re-plugging in pump into the console. the root cause of the pump not detected could not be definitively determined without returned pump to attempt to recreate failure. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28035 | 1220648-2024-09756 | 19087709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087709&pc=OZD | 04/01/2024 | 05/17/2024 |  | 04/11/2024 |  |  | 04/01/2024 | 04/11/2024 |  |  | 10/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411523 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 68-year old male patient with acute on chronic decompensated congestive heart failure and coronary artery bypass graft was implanted with an impella device for mechanical circulatory support. post-operation, the patient was found to have a failure and an epicardial bleed. the impella device supported the patient until day 13 when it was explanted. the explant was complicated from either the graft having a clot or the pulmonary artery pinching the impella device¿s cannula. the impella¿s components became separated during the explant ¿ the cannula became detached from the motor. medical staff were able to complete the explant without issue. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: avoid manual compression of the inlet and outlet areas of the cannula assembly. handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time. | The investigation into the cannula/pigtail detachment - peripheral vessels issue has been completed since the original report was filed. the pump was returned for investigation. the cannula was discovered detached from the motor end, and a noticeable clamping force kink was observed on the cannula and in the metal outlet cage. the cause of the cannula detachment issue was use related, based on returned product review and provided clinical details.
28036 | 1220648-2024-09757 | 19087843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19087843&pc=OZD | 04/04/2024 | 05/17/2024 |  | 04/11/2024 |  |  | 04/04/2024 | 04/11/2024 |  |  | 06/30/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390774 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp was unable to be placed due to patient¿s vessel condition. the case was aborted and planned to be rescheduled with alternative impella support. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since the pump could not pass due to the tortuosity in the aortic arch. impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28037 | 1220648-2024-09758 | 19088090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088090&pc=OZD | 04/04/2024 | 05/28/2024 |  | 04/11/2024 |  |  | 04/04/2024 | 04/11/2024 |  |  | 10/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411240 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a 46 year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while the patient was on support, the impella device had a controller error and/or optical sensor error on two consecutive consoles. the impella device was replaced. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was submitted. the pump was not returned for investigation. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. d6a, d6b
28038 | 1220648-2024-09759 | 19088143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088143&pc=OZD | 04/02/2024 | 05/28/2024 |  | 04/11/2024 |  |  | 04/02/2024 | 04/11/2024 |  |  | 08/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397722 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 60-year old female patient with a high risk of percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during the implantation, there was an issue with the sheath. the doctor had struggled with the sheath and opted for a smaller sheath. the doctor had some friction advancing but was successful in placing the sheath alongside the impella catheter. after the implant, there was excessive oozing. a third femoral suture and a collagen plug were successfully applied. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed and the delivery issues has been completed. product was not returned for investigation. as per the clinical, md initially struggled with advancing the recommended terumo destination 7fr x 45cm sheath and opted for a shorter 7fr x 25cm sheath, ultimately succeeding in placing it alongside the impella catheter. 14fr x 13cm abiomed peel away was used during procedure. the cause of the single access issue was most likely use related, based on given clinical details. bleeding at the access site was reported after explanation, and it was resolved with perclose and manual pressure. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use referenced in the initial report is from ifu for impella cp with smart assist system sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation.¿ d4-updated model number added- d6a/d6b.
28039 | 1220648-2024-09760 | 19088204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088204&pc=OZD | 04/05/2024 | 05/15/2024 |  | 04/11/2024 |  |  | 04/05/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1352449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The complainant reported that an aic (automated impella controller) displayed a controller failure alarm during start up. the power was cycled to troubleshoot the issue but the failure remained. it was noted that while flipping the on/off switch for the battery, there normal "click" noise was not present. the aic was swapped as a result of the persistent failure. there was no patient involvement:. | [5 empty] | The automated impella controller device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) red alarm has been completed. the console was returned for evaluation. upon functional testing, there was no click sound when pressing the battery on/off, likely due to a battery switch malfunction. data logs were reviewed, which confirmed a controller failure alarm due to epm software corruption. the root cause for controller failure was a defective impellatronic board, and the cause for epm software corruption was not determined. the root cause for controller failure was a defective impellatronic pca, and the cause for epm software corruption was not determined. the root cause for the inaudible click sound on the battery switch was most likely due to the defective battery switch assembly. added-d9 device return date. d4-updated model number in accordance with updated procedures, section d2 has been revised from the initial submission.
28040 | 1220648-2024-09761 | 19088242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088242&pc=OZD | 04/05/2024 | 05/23/2024 |  | 04/11/2024 |  |  | 04/05/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433993 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that soon after impella cp was placed in the cath lab as the patient was waking up, the patient moved abruptly, and hematoma developed. the physician held pressure, replaced forward tension sutures, and redressed access site. one unit of blood was given. | [5 empty] | Additional, changed, and/or corrected data: b5. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28041 | 1220648-2024-09762 | 19088343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088343&pc=OZD | 04/05/2024 | 05/28/2024 |  | 04/11/2024 |  |  | 04/05/2024 | 04/11/2024 |  |  | 09/30/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401647 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 47-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported a broken purge reservoir and damaged sidearm while using an impella cp. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported purge leak has been completed. no product and no data logs were returned. clinical details states that a broken purge reservoir was observed and a leak was reported. the root cause of the purge leak - pump was most likely a damaged sidearm but the location of damage was not determined as no product was returned. the instructions for use referenced in the initial report is from ifu for impella cp with smartassist for use during cardiogenic shock and high risk cpi, section: cleaning. d4-updated model number. added- d6a/d6b. added h6 component code 4307.
28042 | 1220648-2024-09763 | 19088425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088425&pc=OZD | 04/05/2024 | 05/17/2024 |  | 04/11/2024 |  |  | 04/05/2024 | 04/11/2024 |  |  | 07/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025442286 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported difficulty implanting an impella cp due to the patient's anatomy. interventional radiology was called to assist in successfully implanting a second impella cp. no patient harm was reported. | [5 empty] | The device has not been returned from the customer. upon investigation completion, a supplemental mdr will be submitted. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because of calcification in the vessels that resisted in advancement of pump through right iliac artery. impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28043 | 1220648-2024-09764 | 19088960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19088960&pc=OZD | 04/05/2024 | 05/20/2024 |  | 04/11/2024 |  |  | 04/05/2024 | 04/11/2024 |  |  | 03/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329150 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | IMPELLA 5.5 | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a 40-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was in the catheterization lab, distal limb ischemia was noted in the right leg. no doppler pulses were present. an antegrade sheath was placed and a fem-fem bypass was performed. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is from the ifu for impella cp with smartassist system section: potential adverse events (united states) which now also includes cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28044 | 1220648-2024-09765 | 19089369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19089369&pc=OZD | 04/05/2024 | 05/20/2024 |  | 04/11/2024 |  |  | 04/05/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469069 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported an 84-year-old male noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that due to the patient¿s anatomy, the impella implant was aborted. the physician attempted to use a buddy wire without success. due to concern in damaging the pump, a new pump was used. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely use related since non-abiomed products like 14fr cook sheath were used during the delivery. impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. iimpella cp with smartassist for use during cardiogenic shock and high risk pci. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella. rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28045 | 1220648-2024-09770 | 19090958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19090958&pc=OZD | 10/31/2023 | 08/20/2024 |  | 04/11/2024 |  |  | 10/31/2023 | 04/11/2024 |  |  | 08/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393632 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male was implanted with an impella device for mechanical support. high purge pressure was encountered during support. heparin was added to the purge system to counteract the condition and tpa protocols were discussed with the staff. the high purge pressure stabilized overnight without the need of the protocol, and there was no patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. product was not returned for investigation. data logs were investigated, and it was observed that after 6 days there was a gradual rise in purge pressure over the span of 2 hours. the "purge pressure high" was reducing after several hours of heparin addition to the purge solution. the purge pressure gradually resolved. later at the end of case the purge pressure seems to rise gradually. no motor current spikes were observed corresponding to the purge rise. the pump flow dipped at the start of the pressure rise and quickly rose again. therefore, the root cause of the high purge pressure issue was most likely biomaterial. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 impact code 4644, h6 med dev prob code 1491 d4-corrected model number added-d6a/d6b
28046 | 1220648-2024-09771 | 19091031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19091031&pc=OZD | 10/31/2023 | 04/16/2024 |  | 04/11/2024 |  |  | 11/04/2023 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1133364 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male was implanted with an impella device for mechanical support. during a purge system change, the blue purge disc holder was noted to be damaged, and the automated impella controller was swapped as a result of the noted damage. there was no patient harm. | [5 empty] | The automated impella controller was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. it was reported that field service inspected the console and discovered a broken purge retainer. field service replaced the broken retainer and the flag. the root cause of this issue was a broken purge retainer.
28047 | 1220648-2024-09772 | 19091056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19091056&pc=OZD | 07/18/2022 | 05/02/2025 |  | 04/11/2024 |  |  | 07/18/2022 | 04/11/2024 |  |  | 04/30/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The complainant reported a patient presenting for ablation was implanted with an impella cp device for mechanical circulatory support. impella cp was implanted per standard fashion without complications. after sheath was put in place, an angiogram was performed which showed minimal distal perfusion and a faint doppler pulse. after stabilization, impella support was weaned, pump was explanted, and manual pressure was held for a reported 45 minutes to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28048 | 1220648-2024-09773 | 19091083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19091083&pc=OZD | 02/14/2024 | 04/11/2024 |  | 04/11/2024 |  |  | 02/14/2024 | 04/11/2024 |  |  |  |  | 04/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1119116 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Dislodged or Dislocated | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella cp was prepped for implant and once the white cord was plugged in the alarm error ¿device defective, change pump¿ came on. pump was then changed out to a new one. the pump was not implanted in the body. it only happened during set up. a new impella cp was implanted successfully. | [5 empty] | The impella device was received from the customer on 02-apr-2024. data analysis: the cart broke and (b)(6) was dropped on the way to a case but had not been used yet. the console was not on at this time so there are no logs to review. device analysis: the gray knob was missing on return and the console had a whining sound that went away when the pbm was swapped out. otherwise the console was completely functional. root cause: the failures were caused by the cart breaking. repair: please replace the gray knob, the pbm ((b)(4)) and then perform a functional check on the console ((b)(4)).
28049 | 1220648-2024-09774 | 19091492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19091492&pc=OZD | 05/13/2022 | 05/23/2025 |  | 04/11/2024 |  |  | 05/13/2022 | 04/11/2024 |  |  | 01/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189131 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella support, a bedside ultrasound was performed to confirm positioning, and inlet was found to be at the aortic valve. the pump was advanced, bend of catheter flipped, and inlet then appeared to be in mitral valve. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28050 | 1220648-2024-09775 | 19093249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093249&pc=OZD | 01/19/2024 | 10/08/2024 |  | 04/11/2024 |  |  | 03/12/2024 | 04/11/2024 |  |  | 11/30/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426157 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Cardiogenic Shock | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured sepsis and worsening cardiogenic shock due to nicm with a "possible" relationship to the impella device used. antibiotics were used as treatment for sepsis, however, the patient was ultimately transitioned to "comfort measures only" status, the device was turned off, and the patient ultimately expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular tachycardia (vt) and sepsis has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vt and sepsis could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added- h6 component code 925 corrections to the initial mfr report: d4-corrected model number. added- d6a/d6b. f6/f8 should have been left blank.
28051 | 1220648-2024-09776 | 19093258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093258&pc=OZD | 08/30/2023 | 04/24/2024 |  | 04/11/2024 |  |  | 03/13/2024 | 04/11/2024 |  |  | 05/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355456 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured left gastric arterial rupture with a "possible" relationship to the impella device used. intervention was required and the patient received a blood transfusion. the patient survived after intervention. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. d4-updated model number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28052 | 1220648-2024-09777 | 19093267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093267&pc=OZD | 01/08/2024 | 05/02/2024 |  | 04/11/2024 |  |  | 03/13/2024 | 04/11/2024 |  |  | 10/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024415866 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | Respiratory Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hematoma at r axillary impella site, respiratory failure, and "accelerated junctional rhythm" with a "possible" relationship to the impella device used. hematoma was treated with surgical intervention, respiratory failure was treated via mechanical vent, and accelerated junctional rhythm was treated via inotropes. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the vf/vt/af/at issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28053 | 1220648-2024-09778 | 19093825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093825&pc=OZD | 05/14/2022 | 04/11/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 02/29/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 05/14/2022 |  |  |  |  | The user facility reported a 75-yesr-old patient placed on impella cp support and the patient developed bleeding from several sites and large hematoma at the groin. manual pressure and femstop were both applied to access site. angle matched dressing remained unchanged and the groin became stable. no systemic heparin due to the hematoma and the patient was provided 1 units of packed red blood cells to improve hemodynamics. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28054 | 1220648-2024-09304 | 19093878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093878&pc=OZD | 03/19/2024 | 05/17/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025454440 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Peripheral Nervous Injury | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 60-year-old patient received hemodynamic support from an impella 5.5 smartassist heart pump after coronary artery bypass graft surgery (cabg). it was reported that the patient was transfused with two units of platelet concentrate to treat thrombocytopenia. it was reported that an infection occurred at the access site and the patient suffered from paresthesia in the arm after removal of the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombocytopenia and the peripheral nerve injury issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the thrombocytopenia and the peripheral nerve injury was not determined due to insufficient clinical details. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28055 | 1220648-2024-09222 | 19093879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093879&pc=OZD | 03/25/2024 | 05/01/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445476 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the pump was placed to bridge the patient to transplant. on day three of the support, the team found an access site bleed. the bleed was treated by intervention in the form of applying an absorbable hemostat and changing the dressing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for impella cp with smart assist system, section: general patient care considerations. states ¿assess access site for bleeding and hematoma.¿ added-d6a. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. explant date: unknown.
28056 | 1220648-2024-09224 | 19093880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093880&pc=OZD | 10/17/2022 | 03/12/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 05/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023222790 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 86-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a complainant reported that the patient was placed on impella cp smart assist support and developed bleeding at the insertion site and a hematoma. the patient required blood transfusions. the bleeding was resolved by compression hemostasis. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28057 | 1220648-2024-09303 | 19093881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093881&pc=OZD | 06/05/2022 | 05/02/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 10/31/2023 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157855 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon insertion, the impella was slightly deep and caused ventricular tachycardia, which was treated by delivering three shocks and pulling back the pump's catheter. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. cardiac arrythmia. analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
28058 | 1220648-2024-09225 | 19093882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093882&pc=OZD | 06/06/2022 | 04/02/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 02/29/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202111 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site bleeding. despite attempts to tighten the perclose sutures and hold pressure, there was active bleeding at the site. calcium was also noted to be present at the site. as a result, the decision was made to explant the impella cp device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28059 | 1220648-2024-09221 | 19093883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093883&pc=OZD | 10/15/2022 | 03/11/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 07/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249865 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2022 |  |  |  |  | The user facility reported a 62-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and experienced limb ischemia on the impella cp side (left). the physician explanted the impella. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28060 | 1220648-2024-09305 | 19093884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093884&pc=OZD | 03/20/2024 | 05/17/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025443889 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the purge pressure had as sudden increase. the physician decided to explant the pump and place a new one. the new pump was started without issue and continued to support the patient. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure issue has been completed since the original report was filed. only the pump was returned for investigation. a significant amount of biomaterial was observed in the purge gap. as per the clinical details, high purge pressure alarm was reported during start of the case. the cause of the high purge pressure issue was biomaterial, based on returned product. in accordance with updated procedures, sections d6 has been revised from the initial submission.
28061 | 1220648-2024-09223 | 19093885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093885&pc=OZD | 07/21/2022 | 04/02/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 03/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200195 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a (b)(6) female was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella site and developed a hematoma. the patient was administered units of packed red blood cells (prbcs). | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28062 | 1220648-2024-09306 | 19093887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093887&pc=OZD | 06/12/2022 | 04/17/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 03/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient had a small thrombus on the end of the impella catheter. the impella flow was decreased to p7. it was reported that overnight the patient became much more dependent on the ventilator and was septic with blood pressure swings. the patient was weaned from medications and the impella cp was successfully removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28063 | 1220648-2024-09307 | 19093888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093888&pc=OZD | 03/24/2024 | 05/09/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 11/26/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435016 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. later, on day five of the support, the team observed loss of pulses in the limb and placed perfusion catheter as well as performed a fasciotomy. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28064 | 1220648-2024-09310 | 19093889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093889&pc=OZD | 03/25/2024 | 05/16/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1327996 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Prime | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) would not move past the priming screen during a routine purge cassette exchange. the aic was subsequently swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the priming issue has been completed since the original report was filed. the pump was returned for analysis. while the failure mode was not reproduced, there were signs of glucose around the purge and the purge assembly. glucose drying onto the piston could cause the piston to be unable to move. the cause of the piston block was a glucose ingress. d6a, d6b, d9.
28065 | 1220648-2024-09308 | 19093891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093891&pc=OZD | 07/11/2022 | 04/07/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 08/31/2023 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129765 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 61-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported during impella support, the patient experienced limb ischemia in right leg. the physician delivered angiomax and heparin and was continued to monitor. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28066 | 1220648-2024-09388 | 19093892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093892&pc=OZD | 06/20/2022 | 05/02/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 11/30/2023 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160697 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 79-year-old patient with heart failure and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had aortic regurgitation throughout pump support. this was treated by explanting the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28067 | 1220648-2024-09490 | 19093893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093893&pc=OZD | 05/31/2022 | 04/17/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 03/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204753 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, after the peel away was removed the patient was transferred to a room, however they were unable to find a doppler pulse, and a hematoma was noted. the patient's urine was yellow and clear. the echo showed the impella was shallow, so it was advanced. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28068 | 1220648-2024-09387 | 19093894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093894&pc=OZD | 03/25/2024 | 05/15/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMPTEMPORARY NON-ROLLER TYPE |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025451755 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 72 year-old female with cardiogenic shock and acute myocardial infarction was implanted with an impella device for mechanical circulatory support. on the patient's third day of impella device support, the device was removed due to positioning issues. multiple attempts at repositioning the impella device under echo were unsuccessful. the patient remains stable. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. data logs were reviewed and several suction alarms with low flow were observed since the start for few days. positioning issue alarms in aorta and ventricle were observed at the end of the case. the cause of the positioning issue was most likely use related per clinical. the cause of the low pump flow issue was not determined since product was not returned and due to insufficient evidence per log review.
28069 | 1220648-2024-09311 | 19093895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093895&pc=OZD | 03/25/2024 | 05/08/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455818 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that while on support of an impella cp the patient developed hemolysis. the impella cp p level was dropped to p5 without resolution. therefore, the impella cp was removed and the patient was implanted with an impella 5.5. there was no harm to the patient reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. according to the clinical, shortly after beginning impella cp support the patient's urine became a pinkish/red color and the team developed hemolysis concerns. imaging confirmed the pump was in the incorrect position and in contact with the mitral apparatus. no attempts were made to reposition the pump as the decision was made to explant the device. after the pump was removed the hemolysis resolved. no product was returned for a visual inspection. data log analysis did not confirm the presence of any abnormalities. the most likely cause of the hemolysis was improper positioning. added-h6 impact code 4628. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. d4 updated model number.
28070 | 1220648-2024-09680 | 19093897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093897&pc=OZD | 03/27/2024 | 05/10/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441825 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unsuccessful attempt of an impella 5.5 device placement in a 68-year-old female in acute myocardial infarction/cardiogenic shock for mechanical circulatory support. the delivery of the impella pump was unsuccessful due to the inability to pass through the patient's vasculature. consequently, the procedure was aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to the patient's anatomy as per clinical description. in accordance with updated procedures, section d6 has been revised from the initial submission.
28071 | 1220648-2024-09769 | 19093898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093898&pc=OZD | 04/06/2024 | 05/15/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445207 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 74-year-old male was implanted with an impella cp device for mechanical circulatory support. it was that the patient was transferred to another facility and upon arrival, access site bleeding was noted. the decision was made to explant the pump due the bleeding issues. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement, on arrival to another facility bleeding was noted at the impella access site as per the clinical description. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28072 | 1220648-2024-09681 | 19093899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093899&pc=OZD | 05/23/2022 | 04/17/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 02/29/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200609 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella insertion, the ¿j¿ wire would not advance, a guidewire was used to advance the wire with much difficulty. a hand injection was performed of the iliac which showed a lesion in the iliac, and the decision was made for a long sheath and eventually the guidewire advanced into the descending aorta. it was reported while on impella cp support, the patient¿s right foot was noted to be discolored in comparison to the left foot. the decision was made to place fem-fem bypass as no doppler signals could be found. the antegrade sheath in the superficial femoral artery was found to be kinked, and the physician exchanged the sheath for a stiffer sheath which resulted in much better doppler signals in the right foot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28073 | 1220648-2024-09767 | 19093900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093900&pc=OZD | 09/20/2023 | 04/15/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 05/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024366178 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 58 year old male patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the patient developed left leg ischemia, and the impella device was removed. after the explant, the patient¿s leg ischemia recovered. the patient continued support via extra-corporeal membrane oxygenation. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
28074 | 1220648-2024-09309 | 19093901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093901&pc=OZD | 03/25/2024 | 05/06/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 10/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 46-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following implant, an impella cp pump immediately began to trigger suction alarms. standard troubleshooting was unable to resolve the noted issue and the decision was subsequently made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed finding motor current spikes observed. suction alarms occurred. the cause of the low pump flow issue most likely was biomaterial ingestion. d4 updated model number. in accordance with updated procedures, section d6 has been revised from the initial submission.
28075 | 1220648-2024-09389 | 19093903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093903&pc=OZD | 03/24/2024 | 05/20/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 11/30/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434206 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 62 year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support and hemodynamic support. during support, the patient had a hematoma. after the nurse milked the drain, the access site started bleeding. heparin was then withheld, and the bleeding started to resolve. the patient also went into ventricular tachycardia while on support. the staff confirmed that the imeplla device inlet was not touching any heart walls, and the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major and the ventricular tachycardia (vt) issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding is determined to be anticoagulation management because it was noted that hematoma was resolving upon hold due to heparin. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt was not determined. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 was inadvertently omitted from the previously submitted report.
28076 | 1220648-2024-09528 | 19093904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093904&pc=OZD | 09/27/2022 | 04/08/2025 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 03/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203693 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient continued to utilize the leave the 14fr impella sheath in place and recover upstairs until the act dropped to 170 then the patient to be brought back down to the ccl for sheath removal. during explant the patient experienced access site oozing that required preclose w perclose proglide x2 sutures, in addition to the femstop and manual compression. imoella was explanted in procedure area and sent to icu. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d4: model number.
28077 | 1220648-2024-09768 | 19093905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093905&pc=OZD | 10/04/2023 | 04/15/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 07/31/2024 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250345 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2023 |  |  |  |  | The user facility reported a 76 year old male patient with acute myocardial infarction and cardiogenic shock was implanted with an impella device for mechanical circulatory support. impella device support lasted for three hours and then was explanted after a successful percutaneous coronary intervention and wean off the pump. the access site was bleeding and was remedied by manual hold and another sheath. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the pump was not returned for investigation. according to clinical information, bleeding following the repositioning sheath insertion persisted despite 2 hours of manual pressure, leading to the explanation of the pump due to unresolved issues. the patient had a calcified vessel condition. the cause of the access site bleeding was most likely lack of vessel recoil onto the sheath, based on provided clinical information.
28078 | 1220648-2024-09390 | 19093908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093908&pc=OZD | 03/26/2024 | 04/19/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075753 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a leak was observed and noted the automated impella controller had a broken purge disc. there was no report of patient involvement. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | The automated impella controller device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
28079 | 1220648-2024-09766 | 19093909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093909&pc=PYX | 04/06/2024 | 05/23/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 10/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024431794 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Insufficient Flow or Under Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The user facility reported a 72-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp flex for mechanical circulatory support. it was reported that an implanted impella rp flex pump began to experience high purge pressure and low pump flows during patient support. a tissue plasminogen activator protocol was initiated each time high purge pressure was encountered, but the pump flow remained below one liter per minute. as a result of the persistent issue, the decision was made to wean and explant the device. there was no patient harm. | [5 empty] | The investigation into the high purge pressure and the low pump flow issues has been completed since the original report was submitted. the pump was not returned for investigation. per the clinical data provided, there was a gradual increase in high purge pressure and low purge flow which resolved quickly after tpa addition. the root cause of high purge pressure was most likely biomaterial ingestion as issue was resolved after tpa addition. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28080 | 1220648-2024-09392 | 19093910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093910&pc=OZD | 03/27/2024 | 05/07/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 07/31/2025 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025442556 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma around their impella and venous extracorporeal membrane oxygenation (ecmo) cannula sites during impella cp support. manual pressure was held on the sites with a sandbag and heparin infusions were paused to manage the condition. hemostasis was subsequently achieved after pressure and surgicel were applied. support continued as planned with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." | The investigation into the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was determined to be anti-coagulation management since heparin infusion was paused to reduce the ptt values to the goal range of 60-80 and hemostasis achieved after manual pressure, surgicel and sandbag placement.
28081 | 1220648-2024-09391 | 19093913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093913&pc=OZD | 03/26/2024 | 05/14/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported an 81 year-old male with a high risk of percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the implanted impella device experienced high purge pressure, a spike in motor current, and flow saturation roughly fifteen minutes after implant. the pump was then replaced, and no patient harm has been reported. | [5 empty] | The investigation for the reported high purge pressure and saturated flow issues has been completed. no product was returned for investigation. data logs show that upon starting the pump, purge pressure begins to ruse and purge flow decreases for 7 minutes without triggering any purge alarms. it then stabilizes around 700 mmhg for 7 minutes and then begins to rise again for 2 minutes and triggering hpp and purge system blocked alarms. there are no motor current (mc) spikes. flow is saturated at 4.3 l while on p-9 during this time. the pump then has a high motor current pump stop with no mc spike prior. the root cause of the pump stop was most likely due to high purge pressure. the root cause of the high purge pressure was unknown because no pump was returned for investigation. the root cause of the saturated flow issue was not determined since product was not returned. the instructions for use states for impella cp with smartassist for use during cardiogenic shock and high risk cpi in section: using the automated impella controller with the impella rp with smartassist system catheter that ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ it also states in section: impella stopped that ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
28082 | 1220648-2024-09572 | 19093915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19093915&pc=OZD | 05/31/2022 | 04/08/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  | 12/31/2023 |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166703 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported that the patient on impella cp support had pump stop. no other information provided. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
28083 | MW5153781 | 19096481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19096481&pc=PYX | 02/28/2024 | 04/09/2024 |  | 04/11/2024 |  |  |  | 04/11/2024 |  |  |  |  | 04/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PYX | [5 empty] | ABIOMED, INC. | [8 empty] | No | [9 empty] | Voluntary report | Not Available | Patient Family Member or Friend | [6 empty] | 75 YR |  | Female | 70 KG | Non Hispanic | White | Death | Yes | Yes | [7 empty] | Ischemia | Reduced Blood Flow | [11 empty] | Death | [5 empty] | ALENDRONATE | ATORVASTATIN | CARVEDILOL | DOTERRA META POWER MIXES | EPLERENONE | FLUTICASONE | FUROSEMIDE | GABAPENTIN | JOINT/BONE VITALITY PO (DOTERRA) | OCCASIONAL ACETAMINOPHEN | PACEMAKER | POTASSIUM CHLORIDE | WARFARIN | [10 empty] | Initial submission | [9 empty] | On (b)(6) 2024, my mother was admitted to the er because the mitral valve was failing. her heart was only pumping out about 2 l/min when i guess it should be pumping about 5 l/min. she was extremely unstable and the doctor told us that he wanted to put in the impella pump, through the cath lab and it would be going in through her right groin area. they said the pump would help keep her alive until they could transport her to (b)(6). she was in really bad shape all night long and it was difficult to keep her alive, even with the pump. the next morning, they got her transported to (b)(6). once up there, they ran a whole bunch of tests and a few hours later, the doctor came and told us that because of the placement of the impella, her right leg was not getting enough blood flow and that if they didn't move it, she would lose her right leg. keep in mind that her right leg was her good leg as she had a stroke 8 years prior and her left leg was still affected by that. i asked if, since at this point the impella been in her and likely restricting blood flow to her legs (since it was placed about 16-18 hours prior to that) if it is possible that her leg already sustained enough damage that it would not be possible for her to walk on it, even if they moved it to the other side? the doctor said that yes, that was possible. then, i asked if it was likely that she would have the same issue with her other leg and he said that it was very possible. we were in an impossible position. a little bit later, they came and got my siblings to talk to us. they were still needing to run more tests and doubts if she would have been strong enough for a mitral valve replacement or putting in a clamp of some kind (i'm not really sure about that piece). they said that even just moving the impella from her right leg, to her left leg, would be a tricky procedure and that they were concerned she wouldn't survive the surgery. for us, a lot of this came down to the fact that she one good leg that was just compromised because of the impella device and that she would be miserable living in that state. it would be excruciating pain to lose a leg. she been through so much because my father was killed by a drunk driver 14 months prior and she was already living in that pain. she been intubated for almost an entire day at this point, which was awful for her and all the options in front of us, resulted in her dying. if this was a known issue with the impella, i really wish they used an alternative or just transferred her to the other hospital sooner, so they could have just tried to do a mitral valve replacement. i am sure that the impella and the circumstances after the installation, contributed to her death. from her medical records this is just one of the test results. feel free to reach out for more information. i don't know how much of this is helpful. in the right lower extremity, the right superficial femoral artery has antegrade nonpulsatile continuous flow with velocities ranging between 7 and 14 cm/s. no color flow or spectral signal is detected in the distal posterior tibial or dorsalis pedis arteries. in the left lower extremity, pulsatile low resistant flow is present in the left superficial femoral artery with peak systolic velocity ranging between 21 and 34 cm/s. pulsatile low resistant flow is seen in the distal left posterior tibial artery with peak systolic velocity of 20 cm/s. similar pulsatile flow is sampled from the left dorsalis pedis artery with a peak systolic velocity of 7cm/s. impression: ischemic right lower extremity with nonpulsatile flow in the thigh and no detectable arterial flow below the knee; pulsatile low resistance flow is present in the left lower extremity.
28084 | 1220648-2024-09673 | 19077596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077596&pc=OZD | 04/01/2024 | 05/16/2024 |  | 04/10/2024 |  |  | 04/01/2024 | 04/10/2024 |  |  | 09/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445459 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the doctor attempted to place the impella cp pump from the left side and was unable due to the patient's anatomy. the doctor feared that they may have damaged the pump during the implant and requested a new pump be opened for subsequent placement through the opposite femoral artery. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because of patient anatomy and pump was unable to pass through vasculature. in accordance with updated procedures, sections d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28085 | 1220648-2024-09674 | 19077624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077624&pc=OZD | 04/01/2024 | 05/16/2024 |  | 04/10/2024 |  |  | 04/01/2024 | 04/10/2024 |  |  | 09/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403596 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported an 86-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a cold foot due to ischemia. both extremities showed severe peripheral arterial disease. the decision was made to place an impella 5.5 axillary and remove the impella cp. a thrombectomy and stent were done in the left leg after explanting. the patient was reported to be stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial is from the ifu impella cp with smart assist system, section: potential adverse events (united states). section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury".
28086 | 1220648-2024-09675 | 19077655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077655&pc=OZD | 04/01/2024 | 05/06/2024 |  | 04/10/2024 |  |  | 04/01/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458112 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 63-year-old female patient was implanted with an impella 5.5 as a bridge to transplant. it was reported the that a hematoma formed at the cutdown site during impella 5.5 support. the resulting pressure in the pocket caused severe pain for the patient, and they were subsequently taken for a washout. a jackson pratt (jp) drain was placed, and heparin was put on hold to assist with managing the condition. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for impella 5.5 with smartassist® for use during cardiogenic shock instructions for use in section: warnings and cautions states to ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury".
28087 | 1220648-2024-09676 | 19077728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077728&pc=OZD | 04/01/2024 | 05/20/2024 |  | 04/10/2024 |  |  | 04/01/2024 | 04/10/2024 |  |  | 10/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422762 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported an 86-year-old female patient was to be escalated from the impella cp to the impella 5.5 for more support. it was reported that the impella 5.5 was unable to pass through the anastomosis and after several attempts a new impella cp was placed instead via the axillary artery. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through as there was a lot of resistance at anastomosis as per the clinical description. the instructions for use for impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings states ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28088 | 1220648-2024-09677 | 19077772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077772&pc=OZD | 03/21/2024 | 10/01/2025 |  | 04/10/2024 |  |  | 03/21/2024 | 04/10/2024 |  |  | 10/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024412188 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 68-year-old male, who is a heart transplant rejection, was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed persistent ischemia in the extremities during impella support. right lower extremity was mottled as well along with upper extremities noted approximately two days after impella device placement. the issue persisted, and the patient was administered milrinone in an attempt to decrease the patient's systemic vascular resistance. support was maintained with the impella pump despite the noted issues. | The patient outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of device with outcome. | The complainant also reported that the placement signal alarm triggered during support and was subsequently disabled temporarily. support was maintained with the impella 5.5 pump despite the noted issues. care was eventually withdrawn, and the patient expired. |  |  |  | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report impella 5.5 with smartassist for use during cardiogenic shock in section: potential adverse events (united states), which now includes statement "cardiac or vascular injury (including ventricular perforation.¿ b5-added mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump was not returned for investigation. data logs were returned and reviewed. placement signal low alarms were seen for a span of 12 days. the 5s parameters were within the specification. no other abnormalities were seen in the logs. the root cause of the optical signal issue was not determined since product was not returned and due to inconclusive evidence per log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b2 updated to death. b5 updated to include death and placement signal issue description. h1 updated to death. h6 updated to include death and placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09677. b7 other relevant history, including preexisting medical conditions updated d10 concomitant medical products and therapy dates updated. g1 mdr reporting contact email and reporting contact fax number updated.
28089 | 1220648-2024-09678 | 19077811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077811&pc=PYX | 03/21/2024 | 06/05/2024 |  | 04/10/2024 |  |  | 03/21/2024 | 04/10/2024 |  |  | 11/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024425638 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pressure Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | There is not enough information to associate use of device with outcome. | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that multiple complications were encountered during impella rp flex support. throughout the course of support, the patient developed persistent ischemia in their extremities. their right lower extremity became mottled and milrinone was administered in an attempt to decrease the patient's systemic vascular resistance. it was further noted that the implanted impella rp flex pump triggered high purge pressure/low purge flow alarms fourteen days into support. tissue plasminogen activator was subsequently ran through the purge to resolve the noted issue. support remained ongoing. |  |  |  |  | The investigation into the high purge pressure and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as per clinical findings, the cause of the high purge pressure was most likely a kinked purge line because there was a spike in the data logs and the issue resolved when the purge system was temporarily opened. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b5, d6a, d6b | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
28090 | 1220648-2024-09679 | 19077833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19077833&pc=OZD | 03/26/2024 | 05/13/2024 |  | 04/10/2024 |  |  | 03/26/2024 | 04/10/2024 |  |  | 11/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 92-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had an attempted impella implant and the 14 french sheath kinked upon delivery. the doctor used a wire to straighten the femoral artery for the 14 french sheath and the impella was inserted through the 14 french x 25 centimeter sheath. the cannula of the impella kinked after exiting the sheath. the device was removed from the body. however the impella was turned on while it was outside the body and the motor stopped eventually. a new impella and sheath were implanted successfully. | [5 empty] | The investigation for the reported product damage has been completed. the impella was not returned for evaluation. clinical details states the patient had a bent femoral artery. therefore, the root cause of the introducer damage - mechanical issue was a introducer sheath kink due to patients anatomy. the root cause of the product damage (cannula kink) issue was use related to improper technique of inserting impella in a kinked sheath. added- d4 model number, catalog number, serial number, lot number, exp date, and udi., h4 device mfr date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28091 | 1220648-2024-09682 | 19078095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078095&pc=PYX | 03/27/2024 | 05/15/2024 |  | 04/10/2024 |  |  | 03/27/2024 | 04/10/2024 |  |  | 10/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024428103 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported that a 50-year-old female on impella rp flex support developed hemolysis and was started on continuous renal replacement therapy a couple of days later. the impella rp flex was removed. a temporary right ventricular assist device was placed. | [5 empty] | The investigation for the hemolysis has been completed. the impella device was not returned for evaluation. the data logs were returned for evaluation and confirmed two suction events occurred before hemolysis began and were resolved by lowering the p-level. according to the clinical, on the first day of impella rp support the patient developed signs of hemolysis with very dark colored urine. correct pump position was confirmed via x-ray imaging. the following day the patient was no longer able to produce urine and they were started on crrt. plasma free hgb was drawn to rule out hemolysis with an initial test result of 230 and a following result of 130. a couple days later the decision was made to explant the pump. no biomaterial was observed upon explant and plasma free hgb returned to normal shortly after the device was removed. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. the instructions for use for impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28092 | 1220648-2024-09683 | 19078180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078180&pc=OZD | 05/23/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/23/2022 | 04/10/2024 |  |  | 01/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into ventricular fibrillation that required the patient to be defibrillated once to resolve. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28093 | 1220648-2024-09684 | 19078278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078278&pc=OZD | 03/26/2024 | 05/13/2024 |  | 04/10/2024 |  |  | 03/26/2024 | 04/10/2024 |  |  | 07/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438173 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Separation | Positioning Problem | [5 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that an implanted impella 5.5 pump experienced persistent positioning issues during patient support. it was noted that the pumps positioning was flipped from the preferred orientation for which repositioning was unsuccessful in resolving. in addition, roughly four days into support, the sterile sleeve broke apart overnight. the site was cleaned and the area was wrapped in transparent film dressings. the staff provided an extra fixation at the site to reduce movement. three days later, the patient had runs of ventricular tachycardia (vt) for which defibrillation was administered. concerns of the pumps positioning were renewed as a result of the vt and the patient's left ventricle was noted to be decompressed compared to the previous week. due to the torn sleeve and unsuccessful attempts at repositioning, the decision was made by the staff to maintain support with the current positioning until a final decision could be made on the patient's plan of care. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. the root cause of the positioning issue was most likely use related to improper fixation technique. the root cause of the product damage (sterile sleeve) issue was not determined since product was not returned and the clinical information is inconclusive. potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28094 | 1220648-2024-09685 | 19078377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078377&pc=OZD | 05/25/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/25/2022 | 04/10/2024 |  |  | 12/31/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179484 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced thrombocytopenia, which was resolved by discontinuing systemic heparin. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28095 | 1220648-2024-09686 | 19078383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078383&pc=OZD | 03/29/2024 | 07/02/2024 |  | 04/10/2024 |  |  | 03/29/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161393 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 03/29/2024 |  |  |  |  | The user facility reported the automated impella controller (aic) displayed 'optical sensor' error during preparation. the console was exchanged. no patient was involved. | [5 empty] | H.6 type of investigation codes were added as they were inadvertently omitted from the previously submitted report. f6 and f8 should have been blank in the initial mdr g1. corrected site of manufacture address, city g1. corrected reporting contact e-mail address. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure modes were not naturally reproduced, and the console functioned as expected. the logs from the field could be reproduced if the pump was not fully pushed in during case start and there was an air gap present. the cause of the optical signal failure was most likely an air gap from between the automated impella controller and the patient cable plug since it was able to resolve on the third connection. a.3 gender, d.2 common device name and g.1 manufacturing site fax number were revised from the initial submission in accordance with updated procedures. h.6 medical device problem code was revised in accordance with updated procedures. new code has been added to medical device problem codes. code 3233 results pending completion of investigation was inadvertently omitted from the previously submitted report and has been added to this report. the investigation has been completed since the previously submitted report.
28096 | 1220648-2024-09687 | 19078424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078424&pc=OZD | 05/19/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/19/2022 | 04/10/2024 |  |  | 01/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023190640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma around the access site. dextrose was added to the purge and heparin was withheld. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28097 | 1220648-2024-09688 | 19078476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078476&pc=OZD | 05/21/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/21/2022 | 04/10/2024 |  |  | 01/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023183630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported a 48-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, echo showed the pump was placed too deep. it was repositioned under image guidance. once completed, echo imaging was performed a second time showing the new placement was fine. however, the patient experienced renal failure which developed into a gross hematuria. the complication was managed with reposition of the pump and continuous renal replacement therapy. eventually, the patient was transferred for ventricular assist device workup. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28098 | 1220648-2024-09689 | 19078534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078534&pc=OZD | 06/17/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 06/17/2022 | 04/10/2024 |  |  | 03/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204402 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during delivery the physician was not able to get the inlet in the free space of the left ventricle. additionally, the access site was oozing blood. the issue was resolved by administering 250ml of albumin, switching the purge to 1000ml d5 with 50meq bicarb, and a femstop was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28099 | 1220648-2024-09690 | 19078552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078552&pc=OZD | 05/21/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/21/2022 | 04/10/2024 |  |  | 02/29/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023202120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported a 53-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing (bleeding) and a hematoma at the pump site. the device was removed and perclose was used to establish successful hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28100 | 1220648-2024-09691 | 19078583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078583&pc=OZD | 10/03/2022 | 05/12/2025 |  | 04/10/2024 |  |  | 10/03/2022 | 04/10/2024 |  |  | 04/30/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The user facility reported a 69-year-old patient was on impella cp support due to shortness of breath and coronary artery disease. it was reported while on support, the staff noticed the patient had red urine indicative of hemolysis. to correct the hemolysis, the impella was repositioned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28101 | 1220648-2024-09692 | 19078649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078649&pc=OZD | 10/19/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 10/19/2022 | 04/10/2024 |  |  | 06/30/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246427 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported an 83-year-old patient placed on impella cp support after being admitted through the emergency room with non-st elevated myocardial infarction. while on impella support patient received multiple units of packed red blood cells. a trans-thoracic echo was done and clot compressing right ventricle was found, patient returned to surgery for removal of clot. impella was weaned and removed. manual pressure was successfully utilized over a 40min hold. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28102 | 1220648-2024-09693 | 19078685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078685&pc=OZD | 03/29/2024 | 05/16/2024 |  | 04/10/2024 |  |  | 03/29/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1728184 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that during a reported suction alarm, no sound was emitting from the console. the staff went into the audio settings and disable the audio and reenabled again. support was maintained with no noted patient harm resulting from the failure. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - buzzer/ alarm malfunction has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure mode was reproduced, and the alarm did not sound when a suction condition was created. the console was analyzed, and the speaker cable was found to be unplugged which explains why no alarm sounds were produced. the cause of the inaudible alarm was a loose connector that was unplugged.
28103 | 1220648-2024-09694 | 19078686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078686&pc=OZD | 10/26/2022 | 03/11/2025 |  | 04/10/2024 |  |  | 10/26/2022 | 04/10/2024 |  |  | 07/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced right lower extremity mottled distal to ventricle and arterial cannula site, and a hematoma at the extra-corporeal membrane oxygenation (ecmo) access site. reportedly multiple units of blood products were transfused. planned to decannulate ecmo and maintain hemodynamics with impella cp. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28104 | 1220648-2024-09695 | 19078751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078751&pc=OZD | 08/12/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 08/12/2022 | 04/10/2024 |  |  | 10/31/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157572 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, an assessment of pedal pulses was conducted, and no pulse was obtained from the patient's right foot, the patient had developed ischemia. also, it was noted that the patient developed hemolysis. the physician explanted the impella cp and replaced with the impella 5.5 to resolve the reported issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28105 | 1220648-2024-09696 | 19078785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078785&pc=OZD | 12/06/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 12/06/2022 | 04/10/2024 |  |  | 04/30/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221549 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient was placed on impella cp support after transfer with incessant ventricular tachycardia. patient continued to intermittently have episodes of ventricular tachycardia; implantable cardioverter defibrillator successfully paces out of ventricular tachycardia. left femoral sheath from intervention left in place and side arm connected to right femoral antegrade sheath for external fem-fem bypass. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Vf/vt/af/at: to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. limb ischemia: when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28106 | 1220648-2024-09697 | 19078907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19078907&pc=OZD | 09/27/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 09/27/2022 | 04/10/2024 |  |  | 02/28/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022052558 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a patient was placed on impella cp support for a percutaneous coronary intervention to left anterior descending artery. while on support, the patient developed a hematoma at the impella cp insertion site. the hematoma was able to be reduced with manual pressure. systemic heparin was on hold due to the hematoma and a drop in the hemoglobin and hematocrit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28107 | 1220648-2024-09698 | 19079061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079061&pc=OZD | 04/03/2024 | 07/17/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1719481 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 65 year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the automated impella controller (aic) had a purge system failed alarm. the aic was replaced, and the procedure was completed. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g.,infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir. clean the connector cable with 70% isopropyl alcohol. maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency. | The investigation for the reported purge leak has been completed. data logs were reviewed which showed that the user was changing the purge cassette before the purge system failure alarm occurred. after the new cassette was initially connected, it was briefly (1 second) disconnected and reconnected during deairing. then, the purge system failure alarm triggered because the piston had only 35% range remaining (<50% range). right data logs confirm that the new cassette was disconnected briefly during deairing which caused system to believe there was a piston blocked event (<50% range). the console was returned for evaluation. upon visual inspection, there was no build up dried glucose, and the purge stepper motor spindle was moving freely with the console powered off. the console performed normally with a known-good pump and purge. no issues detecting the purge cassette. by disconnecting the purge cassette briefly during purge cassette change deairing (recreating the situation in the clinical setting), the purge system failure alarm was reproduced. the root cause of the purge issue was due to brief disconnection of the purge cassette during deairing. added- d9 device return date d3-corrected common device name. d4-corrected model number. f6/f8 should have been left blank in the initial report. h10 instructions for use for this event were for a impella cp device and not an automated impella controller. the instructions should be removed.
28108 | 1220648-2024-09699 | 19079106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079106&pc=OZD | 03/29/2024 | 05/10/2024 |  | 04/10/2024 |  |  | 03/29/2024 | 04/10/2024 |  |  | 07/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441402 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump was unable to cross the aortic valve during attempted delivery. the implant was aborted as a result of the noted resistance. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. in accordance with updated procedures, section d6 has been revised from the initial submission.
28109 | 1220648-2024-09700 | 19079189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079189&pc=OZD | 03/25/2024 | 05/13/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025439537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 55 year-old male with a high risk of percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient developed an infection at the access site due to improper cleansing of the insertion site. the patient started antibiotics and the infection improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: the infusion filter prevents bacterial contamination and air from entering the purge lumen. | The investigation into the infection/sepsis issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information and the relevant evidence were not provided, the root cause of the infection/sepsis could not be determined. section: table 3.2 impella catheter components: "the infusion filter prevents bacterial contamination and air from entering the purge lumen." section: warnings and cautions: warnings: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
28110 | 1220648-2024-09701 | 19079234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079234&pc=OZD | 04/03/2024 | 05/14/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  | 11/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024428608 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported an 80-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 was attempted to be inserted into the left ventricle multiple times, but patient¿s anatomy would not allow the pump to be delivered. a decision was made to remove 5.5 and stay on cp support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related; the patient anatomy would not allow the pump to be delivered as per the clinical description provided. in accordance with updated procedures, section d6 has been revised from the initial submission.
28111 | 1220648-2024-09702 | 19079267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079267&pc=OZD | 04/03/2024 | 05/14/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025437355 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported an 86-year-old female noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the medical team was unable to deliver the cp - the pump was unable to be delivered due to heavily calcified aorta. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. in accordance with updated procedures, section d6 has been revised from the initial submission.
28112 | 1220648-2024-09703 | 19079344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079344&pc=OZD | 04/04/2024 | 06/21/2024 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC |  |  | US |  | 2025462450 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | A japan user facility reported a 74-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that after 7 days of support, the pump stopped. the pump stop prompted pump explant. upon explant thrombosis was seen at the pump. the pump was not replaced as the weaning was ongoing at the time of the pump stop. the patient was reported to as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop and thrombus has been completed. the pump was returned for evaluation. the pump stop occurred on day seven of use. the data logs showed multiple positioning alarms and confirmed pump stop during weaning. the product was inspected and dextrose build up was observed beneath the impeller blades which cleared during soaking. the pump was run in the lab and the pump stop did not reproduce. there was no problem found as the root cause of the pump stop since the issue was unable to be reproduced. as the necessary information was not provided, the root cause of the thrombus was not determined instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.1 added adverse event as it was inadvertently left off the previously submitted report. b.2 selection was inadvertently left off the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. h.1 serious injury was inadvertently left off the previously submitted report. h.6 added code 4440 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
28113 | 1220648-2024-09704 | 19079358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079358&pc=OZD | 03/30/2024 | 06/08/2024 |  | 04/10/2024 |  |  | 03/30/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453437 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | Increase in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support for 56-year-old male, the purge tube ruptured when the patient was turned over in computed tomography (from the purge system, not from the pump). as a result, the purge system was completely changed, and the purge pressure was noted to be too high. the purge flow was less than 1 milliliter per hour. the purge cassette was replaced but did not resolve the issue. a pump exchange was recommended. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the reported issues has been completed since the original report was submitted. it was reported that the purge tube ruptured when the patient was turned over in the computed tomography from the purge system, not from the pump. the purge system was completely changed. high purge pressure was reported after the purge system change with pump flows less than 1 milliliter per hour. the pump was mostly weaned, but the physician reported that the pump stopped shortly before explanting. no product or data logs were returned. the root cause of the purge leak - cassette, was not determined as no product was returned and insufficient clinical information was provided. the root cause of the pump stop was most likely high purge pressure; however, the cause of the high purge pressure was unable to be determined as no product or logs were returned. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 added code 2503 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
28114 | 1220648-2024-09705 | 19079423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079423&pc=PYX | 04/04/2024 | 05/02/2025 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  | 11/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024434895 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a 51 year-old male with a left ventricular assist device, and a history of ischemic cardiomyopathy was implanted with an impella device for mechanical circulatory support. while the patient was on support, the impella had placement signal alarms and the motor current was unchanged. additionally, the patient had renal failure while on support, and continuous renal replacement therapy was started. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of renal failure has been completed. the impella device was not returned for evaluation. the clinical details noted that patient's pfhgb was noted to be 31 on day 3 of rp flex support. patient was placed on crrt support and patient was in renal failure on day 6 of support and pfhgb was trending down from previous days on support. the cause of the renal failure was not determined due to insufficient clinical details. d9 date device returned to manufacturer added. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-09705. d6a and d6b revised from the initial submission 1220648-2024-09705 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-09705. g1 reporting contact email revised on manufacturer device report 1220648-2024-09705 in accordance with updated procedures. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2024-09705.
28115 | 1220648-2024-09706 | 19079479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079479&pc=OZD | 04/04/2024 | 05/16/2024 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025455746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Pressure Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following implant, an impella cp pump began to experience high purge pressures coinciding with low purge flows. a tissue plasminogen activator protocol was initiated in the the purge line and the issue resolved. two days later, the patient's leg became cold, mottled, and pulseless. an external fem-fem bypass was placed and pulses returned to the extremity. support was maintained with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the limb ischemia ¿ reversible and the high purge pressure issues has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. per the clinical information provided, the root cause of the high purge pressure issue was most likely biomaterial. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28116 | 1220648-2024-09707 | 19079547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079547&pc=OZD | 05/25/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 05/25/2022 | 04/10/2024 |  |  | 01/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a 57-year-old patient was implanted with an impella cp device for mechanical circulatory support following an st elevation myocardial infarction. it was reported that while the patient was on impella cp support, hemolysis was confirmed in the patient. the impella was explanted prior to resolving the hemolysis. no further information was provided on any other intervention provided for the treatment of hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28117 | 1220648-2024-09708 | 19079568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079568&pc=OZD | 04/04/2024 | 04/22/2024 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1379919 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a patient of unknown age and gender with acute myocardial infarction and cardiogenic shock was implanted with an impella device for mechanical circulatory support. while the patient was on support, the automated impella controller (aic) had a purge disc not detected alarm that could not be cleared during support ¿ purge flag breakage was suspect. the aic was replaced. the patient remained on support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, supplemental report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. d4-updated model number.
28118 | 1220648-2024-09709 | 19079663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079663&pc=OZD | 04/04/2024 | 06/24/2024 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1679798 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant initially reported that an aic (automated impella controller) displayed a controller failure alarm during impella 5.5 support. after switching to a backup console, a review on impella connect noted that a controller error had occurred. there was no patient harm. | [5 empty] | The impella device was received from the customer on 09-apr-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure mode as reported could not be reproduced upon boot up in the lab. the root cause of the defective optical sensor lamp controller error could not be determined since the issue could not be reproduced. d2.
28119 | 1220648-2024-09710 | 19079688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079688&pc=OZD | 03/30/2024 | 04/10/2024 |  | 04/10/2024 |  |  | 03/30/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1395260 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller (aic) did not recognize the purge disc. the console was exchanged to the backup aic to resolve the issue. there was no reported patient harm. | [5 empty] | The investigation has been completed. the product was returned for investigation. console logs from the reported day were consistent with the complaint and showed multiple instances of intermittent purge disk detection followed by "purge disk not detected" alarms. there was no attempt to connect to the pump, considering the purge system was not operational. the console was inspected and revealed a broken/missing purge disk detect flag- which was responsible for being unable to detect purge disk. the cause of the issue was a defective component.
28120 | 1220648-2024-09711 | 19079692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079692&pc=OZD | 04/04/2024 | 04/29/2024 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1613100 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported an automated impella controller (aic) had a battery failure. there was no patient involvement, and the aic was removed from service. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported console (aic) red alarm has been completed. the console was returned for evaluation. battery failure alarm issue was able to be reproduced. data logs were reviewed which confirmed the battery failure issue. there were multiple instances of battery failure alarms (transport connection blown/missing) and a single instance of battery failure alarm (low voltage) [v < 12v] found from the reported occurrence date. results of running batttest on telnet revealed that ¿fu¿ was missing on battery 1 and that vcell2 in battery 1 had a voltage that was significantly less than the rest of the cells in that battery. issue was determined to be caused by internal battery cell imbalance (found vcell2 of battery 1). added- d9 device return date d4-updated model number.
28121 | 1220648-2024-09712 | 19079719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079719&pc=OZD | 04/04/2024 | 06/03/2024 |  | 04/10/2024 |  |  | 04/04/2024 | 04/10/2024 |  |  | 10/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024416326 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the patient had limb ischemia. the patient's foot was cold. the patient's acts between was noted to be between 160/180. the impella cp was therefore explanted and the issue resolved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28122 | 1220648-2024-09713 | 19079744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079744&pc=OZD | 05/24/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/24/2022 | 04/10/2024 |  |  | 02/29/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194902 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 05/24/2022 |  |  |  |  | The user facility reported a 63-year-old patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. the angle of insertion was changed with pressure held to alleviate the bleeding and systemic heparin was reduced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28123 | 1220648-2024-09714 | 19079768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079768&pc=OZD | 04/03/2024 | 06/21/2024 |  | 04/10/2024 |  |  | 04/05/2024 | 04/10/2024 |  |  | 10/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024414206 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a patient of unknown age and gender and with acute myocardial infarction and cardiogenic shock had an impella device implant procedure for mechanical circulatory support. while the impella device was being implanted, a kink was found at the bottom of the motor. the intended procedure was completed without the impella device, and the patient remains on support. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation has been completed. the device was not returned for investigation. based on clinical description, the root cause of kinked cannula was most likely patient condition. a.2 revised as inadvertently left off the previously submitted report and was added. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 revised reporting contact facility name and reporting contact fax number from the initial submission in accordance with updated procedures. h6 codes 486 and 4114 were reported incorrectly on the previously submitted report. medical device problem code was revised in accordance with updated procedures.
28124 | 1220648-2024-09715 | 19079773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079773&pc=OZD | 04/01/2024 | 05/15/2024 |  | 04/10/2024 |  |  | 04/01/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457212 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 64-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had a chest x-ray and it revealed the impella cp had moved back past the valve to the ascending aorta. echocardiogram confirmed the position. fluoroscopy was used to reposition but was unsuccessful. the impella was surgically removed without any problems. no patient harm was reported, and the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. the instructions for use for impella cp with smartassist states in use of echocardiography for positioning of the impella catheter section the following: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it" added- h6 impact code 4628 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28125 | 1220648-2024-09716 | 19079934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19079934&pc=OZD | 04/02/2024 | 05/17/2024 |  | 04/10/2024 |  |  | 04/02/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451967 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 83-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had no pedal pulses on the insertion site and the staff was unable to doppler pulses. the doctor made the decision to place a reperfusion catheter before transferring to the intensive care unit. pluses were present after the placement. additionally, the patient had ventricular fibrillation upon insertion of the pigtail to obtain left ventricle access. the patient was defibrillated once. the patient also had bleeding from the groin. two units of blood products were provided. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf and the limb ischemia was not determined. the root cause of the access site bleeding was determined to be patient condition because bleeding has occurred from multiple sites. in accordance with updated procedures, section d6 has been revised from the initial submission.
28126 | 1220648-2024-09717 | 19080793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19080793&pc=OZD | 04/06/2024 | 05/15/2024 |  | 04/10/2024 |  |  | 04/06/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025438109 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support, and while on support, the pump generated high purge pressure alarms. tissue plasminogen activator was added to the purge to resolve the issue. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the high purge pressure was most likely due to gradual buildup of biomaterial in the purge gap as it was resolved with tpa. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28127 | 1220648-2024-09718 | 19080995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19080995&pc=OZD | 06/29/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 06/29/2022 | 04/10/2024 |  |  | 11/30/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168484 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 53-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support the patient's distal pulses were absent on right leg with impella insertion. no additional information was provided on any interventions or treatments performed for limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28128 | 1220648-2024-09719 | 19081345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19081345&pc=OZD | 10/01/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 10/01/2022 | 04/10/2024 |  |  | 07/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249902 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experience oozing at access site, therefore the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28129 | 1220648-2024-09720 | 19081462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19081462&pc=OZD | 04/02/2024 | 05/13/2024 |  | 04/10/2024 |  |  | 04/02/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446029 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 73 year-old female with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there had been a concern about the patient developing limb ischemia therefore, the impella device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28130 | 1220648-2024-09721 | 19081519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19081519&pc=OZD | 05/18/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 05/18/2022 | 04/10/2024 |  |  | 10/31/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158361 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was noted the patient's urine was darkening and labs hemolyzed. the impella was reduced, and the right pressures were decreased, bolus of fluids was administered. there was a soft, small hematoma that had begun to form medial to insertion site. the physician placed a low-pressure femstop medial and cranial to insertion site. hematuria was not resolving, and the physician decided to explant the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28131 | 1220648-2024-09722 | 19081612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19081612&pc=OZD | 07/04/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 07/04/2022 | 04/10/2024 |  |  | 10/31/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142210 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The complainant reported a patient with myocarditis was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced an infection at the impella access site. the site infection was resolved but no information was provided on the treatment given to the patient to resolve the infection | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records.
28132 | 1220648-2024-09723 | 19081619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19081619&pc=OZD | 05/21/2022 | 04/02/2025 |  | 04/10/2024 |  |  | 05/21/2022 | 04/10/2024 |  |  | 02/29/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188947 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding and hematoma at the groin. as a result, impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28133 | 1220648-2024-09724 | 19081939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19081939&pc=OZD | 04/02/2024 | 05/20/2024 |  | 04/10/2024 |  |  | 04/02/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438712 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 63-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was no flow past repo sheath and that tissue plasminogen activator was used for low purge flows and high purge pressures. a decision was made to explant the impella cp. embolectomy was done before vessel closure. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the high purge pressure and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. per the data log review and clinical information, the root cause of the high purge pressure issue was most likely biomaterial. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. d6a, d6b.
28134 | 1220648-2024-09725 | 19082009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19082009&pc=OZD | 04/03/2024 | 05/16/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1388031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had an alert of system fault error. no patient was involved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) self check error has been completed. the console was returned for evaluation. upon functional testing, the reported failure mode was reproduced. upon bootup, the console rebooted automatically and showed the ¿system self check failure¿ screen. visual inspection confirmed the pbm contamination. data logs also confirmed the reported issue during initial console bootup, there was a communication issue with the pbm. the root cause of system self check failure was pbm contamination. added- d9 device return date d4-updated model number. in accordance with updated procedures, section d2 has been revised from the initial submission.
28135 | 1220648-2024-09726 | 19082097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19082097&pc=OZD | 06/03/2022 | 04/10/2024 |  | 04/10/2024 |  |  | 06/03/2022 | 04/10/2024 |  |  | 01/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced hemolysis. the staff lowered the power setting and repositioned the device. the patient was started on continuous renal replacement therapy and taken off levosimendan. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28136 | 1220648-2024-09727 | 19082153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19082153&pc=OZD | 04/03/2024 | 05/28/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  |  |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464706 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported while on support and after the implant, there was placement signal alarms. the pump was started, but the alarm did not resolve. no appreciable difficulty was observed with the insertion. the physician made the decision to replace the impella. a new impella was implanted successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was filed. the pump and data logs were returned for investigation. the failure was not reproduced during testing. however, per the data log review, the root cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug or within the middle of the red handle ferrule since the failure was not reproduced.
28137 | 1220648-2024-09728 | 19082265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19082265&pc=OZD | 06/11/2022 | 03/28/2025 |  | 04/10/2024 |  |  | 06/11/2022 | 04/10/2024 |  |  | 01/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191228 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced weak pulses. doppler was performed showing adequate flow. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28138 | 1220648-2024-09729 | 19082372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19082372&pc=OZD | 07/01/2022 | 04/11/2025 |  | 04/10/2024 |  |  | 07/01/2022 | 04/10/2024 |  |  | 03/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192017 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was being placed on an impella cp for mechanical circulatory support. the complainant reported that an impella cp pump implant attempt had to be aborted due to the patient's difficult anatomy, as it was kinked and tortuous. the pump could be advanced into the aortic arch, but could not pass it. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28139 | 1220648-2024-09730 | 19082549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19082549&pc=PYX | 04/02/2024 | 05/17/2024 |  | 04/10/2024 |  |  | 04/03/2024 | 04/10/2024 |  |  | 07/31/2024 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023248935 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. it was reported that the patient developed worsening hemolysis during impella rp support. as a result of the condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the hemolysis issue has been completed since the original report was filed. the pump was returned for investigation. the product was inspected, and biomaterial was found throughout the cannula and around the impeller. biomaterial was cleared and the pump passed a hemolysis test. the root cause of hemolysis was determined to be biomaterial. in accordance with updated procedures, section d6 has been revised from the initial submission.
28140 | 1220648-2024-09731 | 19085217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085217&pc=OZD | 10/04/2022 | 05/02/2025 |  | 04/10/2024 |  |  |  | 04/10/2024 |  |  | 09/30/2023 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022137565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding. heparin was withheld, the patient went for washout and was provided blood products. it was decided to escalate the patient to the impella 5.0 since additional support was required. the patient was stable and the imeplla cp was explanted achieving hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28141 | 1220648-2024-09732 | 19085378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085378&pc=OZD | 08/13/2023 | 03/18/2025 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 04/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329739 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | Death | Hospitalization | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 03/11/2024 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry notification received with a "probable" relationship between thrombocytopenia and the impella device: while the issue had resolved, the patient's hospitalization was prolonged, as a result. | The expiration of the patient after being withdrawn from support has been reviewed by abiomed's medical safety officer. it was determined that there is not enough information to associate use of device with outcome. | It was reported that a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿the patient developed hemolysis. it was previously reported that the patient expired. | It was reported that a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿the patient developed hemolysis. the patient's outcome was previously reported as death. |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b5-added medical safety officer review. d4-updated model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | B.3 revised date of event as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-09732. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-09732. f.6 and f.8 dates were reported on the initial manufacturer device report number 1220648-2024-09732 and should not have been. g.1 revised reporting contact email on manufacturer device report 1220648-2024-09732 in accordance with updated procedures. g.1 added reporting facility name and fax number on manufacturer device report 1220648-2024-09732 in accordance with updated procedures. g.2 added selection that was inadvertently omitted from manufacturer device report 1220648-2024-09732. h.6 component code was added since the previous manufacturer device reports number 1220648-2024-09732 were submitted in accordance with updated procedures. h.6 added health impact code for death as it was inadvertently omitted on manufacturer device report number 1220648-2024-09732-2. hospitalization code was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-09732. h.10 thrombus was mentioned in error on manufacturer device report number 1220648-2024-09732-2 b.2 death outcome should have been selected on manufacturer device report number 1220648-2024-09732-3 b.5 duplicate of manufacturer device report number 1220648-2024-09732-2. h.1 should have been death on manufacturer device report number 1220648-2024-09732-3. h.6 entered in error manufacturer device report number 1220648-2024-09732-3. | Added-b5 new information was received for additional failure modes of hemolysis and thrombus. h6 clinical codes 1886. | Added: b5- additional failure mode of hemolysis was added. b6-added lab values. h6 clinical code 1886.
28142 | 1220648-2024-09733 | 19085448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085448&pc=OZD | 03/11/2024 | 09/26/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 05/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024355446 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/11/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured worsening ventricular tachycardia and worsening hypotension with a "possible" relationship to the impella device. the device was repositioned and the vt resolved, however, the patient became hypotensive 3 days later and required norepinphrine and vasopressin as intervention. approximately 13 days later, due to worsening condition, the patient and family decided to withdraw care and the patient expired. we are unable to rule out a relationship between these allegations and the patient's outcome. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections h1 and h6.
28143 | 1220648-2024-09734 | 19085477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085477&pc=OZD | 11/21/2023 | 04/30/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 07/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured hemolysis with a "probable" relationship to the impella device. hospitalization was prolonged. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hemolysis issues has been completed. the pump was not returned for investigation. data logs show an active suction alarm with the impella position in aorta alarm during the middle of the case run for about 2 hours. there were no noted trends in motor current spikes or placement signal trends. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smart assist system states in section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ it also states in section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28144 | 1220648-2024-09736 | 19085487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085487&pc=OZD | 10/21/2023 | 05/01/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 05/31/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024368179 |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | A notification via abiomed¿s long-term outcome and quality indicator impella registry that a 60 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient endured hemolysis and thrombocytopenia with a "definite" relationship to the impella device used. prolonged hospitalization had resulted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis and the thrombocytopenia issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical details, the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the thrombocytopenia issue was not determined since product was not returned.
28145 | 1220648-2024-09735 | 19085489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085489&pc=OZD | 10/15/2023 | 05/24/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 02/28/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337749 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | A notification via abiomed¿s long-term outcome and quality indicator impella registry that a 76-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient endured hemolysis and thrombocytopenia with a "definite" relationship to the impella device used. prolonged hospitalization had resulted. | Upon review by abiomed's medical safety officer, it was determined that there was not enough information to associate use of the device with patient outcome of death. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombocytopenia and hemolysis has been completed. the pump was not returned for investigation. data logs do not report any trends related to motor current spikes or placement signals. according to the clinical details, the patient experienced thrombocytopenia and hemolysis during pump use. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the thrombocytopenia issue was not determined since product was not returned. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4- updated model number b5-added patient relationship from review by medical safety officer in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28146 | 1220648-2024-09737 | 19085511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085511&pc=OZD | 09/27/2023 | 04/30/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 02/28/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024337772 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used. the patient's hospitalization was prolonged as a result. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. data logs does not report any trend related to motor current spikes or placement signals, during case run. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ to conform with updated procedures, sections d6 and h10 have been revised with updated information.
28147 | 1220648-2024-09738 | 19085649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085649&pc=OZD | 03/11/2024 | 05/01/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 09/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024400420 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The complainant reported that a 58-year-old female patient was implanted with an impella cp pump for mechanical circulatory support. during the course of support, the patient suffered from a thromboembolic stroke. systemic heparin was halted per neuro recommendation to prevent conversion. plans were made to wean the patient off the pump as tolerated following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the stroke/cva issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary clinical information was not provided, the root cause of the stroke/cva could not be determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
28148 | 1220648-2024-09739 | 19085650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19085650&pc=PYX | 03/11/2024 | 09/18/2024 |  | 04/10/2024 |  |  | 03/11/2024 | 04/10/2024 |  |  | 11/30/2025 |  | 04/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2024435349 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The complainant reported that a 58-year-old female patient was implanted with an impella rp flex pump for mechanical circulatory support. during the course of support, the patient suffered from a thromboembolic stroke. systemic heparin was halted per neuro recommendation to prevent conversion. plans were made to wean the patient off the pump as tolerated following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This complaint has already been captured and reported under mdr report #1220648-2024-09738. therefore, the initial mdr 1220648-2024-09739 is considered a duplicate of mdr report #1220648-2024-09738. no further supplemental reports will be sent under 1220648-2024-09739.
28149 | 1220648-2024-09602 | 19069489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069489&pc=OZD | 05/23/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/23/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023196946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient had a large hematoma at the groin and pressure was held for 20 minutes and reevaluated. however, the site continued to bleed and a femostop was placed and heparin was withheld to resolve the bleeding. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28150 | 1220648-2024-09603 | 19069543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069543&pc=OZD | 05/23/2022 | 03/07/2025 |  | 04/09/2024 |  |  | 05/23/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023173027 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access sit. sutures from the preclose device were present and angle matching was confirmed. the medical team was advised to tighten the sutures. no information was provided on intervention used to resolve bleeding. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28151 | 1220648-2024-09604 | 19069570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069570&pc=OZD | 03/24/2024 | 06/03/2024 |  | 04/09/2024 |  |  | 03/24/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455587 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the pump stopped. several attempts were made to re-start but were unsuccessful. additionally, the pump was pulled back into the aorta. the pump was ultimately explanted. there was no reported patient harm. | [5 empty] | The investigation for the pump stop has been completed. on day 4.5 of support, data logs show a high motor current pump stop alarm 13 after a motor current spike. the pump was unable to be restarted. the impella was returned for evaluation. an enlarged purge gap was observed. the root cause of the pump stop was determined to be bearing wear. the instructions for use referenced in the initial report was from ifu for the impella cp with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
28152 | 1220648-2024-09605 | 19069633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069633&pc=OZD | 03/24/2024 | 06/21/2024 |  | 04/09/2024 |  |  | 03/24/2024 | 04/09/2024 |  |  | 09/30/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399707 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, purge fluid was noted to be leaking out of the side arm retainer. there were no alarms. there was a slow increase of purge fluid of around 5 milliliters per hour in the last three days. there was no reported patient harm. | [5 empty] | The investigation for the reported purge leak has been completed. the product was not returned. data logs were reviewed. on the 4th day of support, the purge sidearm leak was reported. purge flow is slightly trending upwards with purge pressure trending downwards. the root cause of the purge leak pump was most likely a leak from damaged sidearm but the exact location and cause of the damage was unknown because product was not returned. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ added-d6a/d6b | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." "clean the connector cable with 70% isopropyl alcohol.¿
28153 | 1220648-2024-09606 | 19069697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069697&pc=OZD | 05/26/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/26/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182409 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced hemothorax and progression of bleeding tendencies. the device was explanted to control the bleeding. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28154 | 1220648-2024-09607 | 19069716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069716&pc=OZD | 03/25/2024 | 06/04/2024 |  | 04/09/2024 |  |  | 03/25/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456671 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 90-year-old male patient was placed on impella cp support for semi-elective left anterior descending percutaneous coronary intervention. the patient had initial oozing and swelling on the insertion area. while the patient was in the intensive care unit, lactate was increased. nasal, oral, and gastric bleeding was noted. redd blood cells, tranexamic acid, and volume to be given to the patient. the impella position was verified to be correct. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was not determined since the product was not returned and limited information was provided in the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. revised manufacturing site fax number from the initial submission in accordance with updated procedures. added code 925 as it was inadvertently left off the previously submitted manufacture device report. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28155 | 1220648-2024-09608 | 19069753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069753&pc=OZD | 05/27/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/27/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201802 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a large hematoma at the impella site due to improper suturing. the impella was advanced until the blue suture hub was at the insertion site and forward tension suturing was maintained. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28156 | 1220648-2024-09609 | 19069867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069867&pc=OZD | 05/27/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/27/2022 | 04/09/2024 |  |  | 09/30/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. removal was complicated due to pseudoaneurysm of the iliac artery. additionally, the patient experienced access site bleeding that was resolved by proglide sutures, balloon tamponade, manual pressure, and femstop. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ removal issues: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ impella rp: ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
28157 | 1220648-2024-09610 | 19069896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069896&pc=OZD | 03/25/2024 | 05/01/2024 |  | 04/09/2024 |  |  | 03/25/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1096323 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported a controller error occurred. there was no patient involvement. | [5 empty] | The investigation into the aic - controller error (yellow alarm) issues has been completed since the initial report was submitted. the console was returned and tested, but the failure was unreproducible. however, the reported failure mode was confirmed in the data log. the root cause for controller error was not determined due to the inability to be reproduced. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28158 | 1220648-2024-09611 | 19069911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19069911&pc=OZD | 05/28/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/28/2022 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203990 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a thrombus was noted to be occlusive in right coronary artery and left anterior descending (coronary artery branch). moreover, bleeding was noted at the site; therefore, impella was explanted. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28159 | 1220648-2024-09612 | 19070025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070025&pc=OZD | 05/29/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/29/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198559 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient had oozing from groin site. heparin was not started and fem stop was placed to assist with bleeding. impella was explanted and patient was hemodynamically stable. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d3 manufacturer name | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28160 | 1220648-2024-09613 | 19070031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070031&pc=OZD | 05/21/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 05/21/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177101 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced atrial fibrillation and the patient was placed on diltiazem drip, epinephrine and propofol. lactate and creatinine levels showed an upward trend, accompanied by minimal urine output. therapeutic heparin was administrated. additionally, the impella was out of position and explanted successfully. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28161 | 1220648-2024-09614 | 19070041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070041&pc=OZD | 05/25/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/25/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a 69-year-old patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced episodes of ventricular tachycardia and suction alarms on the impella device. the p level of the impella was decreased to break suction, and a 500 ml fluid bolus was administered to the patient and complications were resolved. it was also reported the patient experienced bleeding issues, and the medical team decided to explant the impella as patient was already in the process of being weaned from the impella device. no information was provided on any treatment or intervention for bleeding issues. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28162 | 1220648-2024-09615 | 19070088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070088&pc=OZD | 05/29/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/29/2022 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hematuria and an elevated partial thromboplastin time (ptt). the impella cp was unsuccessfully repositioned under echo. the device was successfully weaned and explanted. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. d4 (serial number, lot number, expiration date, unique identifier) and h4 (manufacturing date) are unknown. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28163 | 1220648-2024-09616 | 19070179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070179&pc=OZD | 05/22/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 05/22/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175947 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/22/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and experienced slight oozing at the access site. the physician added a stitch to help resolve the bleeding. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28164 | 1220648-2024-09617 | 19070310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070310&pc=OZD | 05/22/2022 | 03/07/2025 |  | 04/09/2024 |  |  | 05/22/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023182324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the pump had slightly moved during transfer. the device was found to be located under the pap muscle having a stable flow. utilization of the echo in repositioning of the pump into a better spot, was unable to be accomplished, therefore, the device was successfully weaned and explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28165 | 1220648-2024-09618 | 19070314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070314&pc=OZD | 05/31/2022 | 03/11/2025 |  | 04/09/2024 |  |  | 05/31/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported while on support, oozing was identified at the access site, mattress suture was placed, and pressure dressing was applied. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28166 | 1220648-2024-09620 | 19070374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070374&pc=OZD | 03/23/2024 | 05/02/2024 |  | 04/09/2024 |  |  | 03/25/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1284271 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support for a 73-year-old male patient, the automated impella controller (aic) did not detect the purge disc while starting up. the purge disc and cassette were detached and then reattached again, and the issue was resolved. there was no reported patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected; however, the failure was not reproducible. the pump detection issue was most likely due to the rarely occurring epm crystal issue on the impellatronic board which did not reproduce during investigation. to conform to updated procedures, section d2 was revised with updated information. | The automated impella controller (aic) was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28167 | 1220648-2024-09621 | 19070462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070462&pc=OZD | 05/22/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/22/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197303 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2022 |  |  |  |  | The user facility reported a 55-year-old male on impella 5.5 support due to cardiomyopathy. while on support, a purge bypass had to be done on the impella 5.5 due to leaking and cracking of the yellow connection and purge filter. the complication was managed by performing a purge bypass using a 20g angiocath needle for first try. it was then replaced this with art line pressure tubing for the bypass. no patient harm was reported. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
28168 | 1220648-2024-09622 | 19070503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070503&pc=OZD | 05/30/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/30/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200804 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had oozing from the access site and a large hematoma developed as a result of insertion. manual pressure was applied along with fem stop, angioseal and perclose. the impella was removed due to the complications. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28169 | 1220648-2024-09623 | 19070534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070534&pc=OZD | 03/22/2024 | 06/04/2024 |  | 04/09/2024 |  |  | 03/26/2024 | 04/09/2024 |  |  | 11/30/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430852 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella 5.5 was unable to start during implant procedure. the purge cassette and the automated impella controller (aic) were both exchanged but the issue did not resolve. after four unsuccessful attempts, a new impella 5.5 was used and was successfully implanted without issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump did not start issue has been completed since the original report was filed. the device was returned for investigation. pump and both purge cassettes returned. the failure was not reproducible. however, data logs show the purge fluid not detected alarm followed by the purge system failure alarm. the cause of the priming issue was most likely use related, based on returned product review. d6a, d9.
28170 | 1220648-2024-09624 | 19070572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070572&pc=OZD | 05/23/2022 | 03/07/2025 |  | 04/09/2024 |  |  | 05/23/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195917 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s distal peripheral pulses were unable to be found. a brief suction event resolved the issue without intervention. the device was successfully weaned and explanted with the repair of the femoral artery. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28171 | 1220648-2024-09625 | 19070585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070585&pc=OZD | 05/30/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/30/2022 | 04/09/2024 |  |  | 10/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159772 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction alarms for which resolution was unsuccessfully attempted with several boluses. additionally, the pump was mistakenly pulled back 20cm in the icu, which was resolved by replacing the pump. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28172 | 1220648-2024-09626 | 19070668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070668&pc=OZD | 03/26/2024 | 05/17/2024 |  | 04/09/2024 |  |  | 03/26/2024 | 04/09/2024 |  |  | 03/31/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024318553 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, the automated impella controller (aic) displayed impella not working alarm. the console was exchanged but the issue did not resolve. the pump was ultimately explanted. there was no reported patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. after 10 days of support with impella 5.5, clinical reported that the impella stopped working and there was reverse flow. the console was replaced and the message impella not working remained. no data logs were provided and no product was returned. the root cause of the pump stop was not determined as no product was returned and limited clinical information was provided. the instructions for use referenced in the initial report can be found in the ifu for impella 5.5 with smartassist for use during cardiogenic shock, section: direct aortic insertion, section: using the automated impella controller with the impella catheter, and section: impella stopped. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
28173 | 1220648-2024-09627 | 19070744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070744&pc=OZD | 05/31/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/31/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172192 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a loss of distal pulses in both lower extremities. treatment actions taken are unknown. | [5 empty] | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28174 | 1220648-2024-09628 | 19070778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070778&pc=OZD | 05/26/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/26/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient laboratory samples were reported as hemolyzed. the impella was repositioned and the patient began continuous renal replacement therapy, and hemolysis resolved. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records
28175 | 1220648-2024-09629 | 19070889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070889&pc=PYX | 03/20/2024 | 05/21/2024 |  | 04/09/2024 |  |  | 03/20/2024 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP EU |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023258119 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 75-year-old female patient on impella rp support had access site bleed that had to be managed with adrenalin dab. the bleeding resolved. following successful explant, the patient was noted to have expired as care was withdrawn. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the issue is resolved with adrenalin dab. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
28176 | 1220648-2024-09630 | 19070904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19070904&pc=OZD | 05/31/2022 | 11/25/2025 |  | 04/09/2024 |  |  | 05/31/2022 | 04/09/2024 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the automated impella controller (aic) had a, ¿switch to backup controller which supports optical pressure sensor¿ alarm. the aic was successfully replaced. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
28177 | 1220648-2024-09631 | 19071038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071038&pc=OZD | 03/05/2024 | 06/07/2024 |  | 04/09/2024 |  |  | 03/26/2024 | 04/09/2024 |  |  | 11/30/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426318 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed on impella cp to support crash and burn situation with ongoing chest compressions. the impella was placed and return of spontaneous circulation (rosc) was achieved. however, after 15-20 minutes of support, there was a red alarm for high purge pressure. there were no kinks noted and the system was deaired, but the issue remained. the purge fluid was 5% glutamic acid with bicarb. following percutaneous coronary intervention (pci), the purge cassette was exchanged but the purge pressure remained high. there was no left ventricular thrombus. the motor current increased and there were several pump stops leading to the pump being explanted and a new one being placed. the pump was running well. however, the patient's condition was noted to have deteriorated during pump exchange. | [5 empty] | D.6a and d.6b are unknown. the investigation into the pump stop issue has been completed since the original report was filed. the pump was returned for investigation. about 15 minutes after impella placement, there was a high purge pressure alarm which did not resolve after troubleshooting. motor current started to increase and there were several pump stops and pump was run at p2 till a pump swap. no data logs were returned. upon product evaluation, biomaterial was found in the purge gap. the pump was run in the lab and pump stop did not reproduce. the root cause of the temporary pump stop was most likely biomaterial. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.3 revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 added codes 4629, 4114 and 3233 as it was inadvertently left off the previously submitted manufacture device report. the product has been returned since then and the investigation has been completed. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was discarded. at the time of the previous report, the device was not returned but has since been returned. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
28178 | 1220648-2024-09632 | 19071048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071048&pc=OZD | 06/29/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 06/29/2022 | 04/09/2024 |  |  | 04/30/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214007 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and experienced a weak monophasic doppler popliteal in the impella leg, as well as access site bleeding. bicarbonate purge was utilized. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28179 | 1220648-2024-09633 | 19071081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071081&pc=OZD | 09/27/2022 | 05/21/2025 |  | 04/09/2024 |  |  | 09/27/2022 | 04/09/2024 |  |  | 06/30/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238090 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic |  | Malfunction | Yes | Yes | Decrease in Pressure | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported an 84-year-old female was on impella cp support, there were alarms which could not be resolved, and patient was becoming hypotensive. therefore, impella was removed. a new impella cp was inserted and started without issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28180 | 1220648-2024-09634 | 19071089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071089&pc=OZD | 03/20/2024 | 05/06/2024 |  | 04/09/2024 |  |  | 03/20/2024 | 04/09/2024 |  |  | 10/31/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025445616 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male postcardiotomy cardiogenic shock was implanted with an impella device for support. the patient was additionally supported by extracorporeal membrane oxygenation. on day after implant, the team had runs of ventricular tachycardia that prompted cardioversion and medical therapy. | [5 empty] | The investigation into the vf/vt/af/at issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. "arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia." | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28181 | 1220648-2024-09635 | 19071154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071154&pc=OZD | 03/20/2024 | 06/19/2024 |  | 04/09/2024 |  |  | 03/21/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441797 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of device with the patient's outcome. | The user facility reported an unknown age male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. a sudden increase in purge pressure during patient care was observed. the issue was not resolved by replacing the purge set, but rather, the pump was replaced to resolve the issue. there was no report of patient harm. |  |  |  |  | The investigation for the high purge pressure (hpp) has been completed. the data logs showed that about six days into support there was a sudden rise in purge pressure and drop in purge flow over five minutes. there are no motor current spikes prior to this. there are hpp and purge system blocked alarms. the purge flow tick stalls. the purge cassette was changed but did not resolve the issue. the pump was received for evaluation. hpp was not reproduced in a bucket of water when purged. hpp still did not reproduce when the pump was run. there was no biomaterial in the purge gap. there was no change in motor current at the time of the sudden spike in purge pressure. the root cause of the hpp was most likely a kinked purge line. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 address line 2 revised as previously entered incorrectly. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 medical device problem code was revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28182 | 1220648-2024-09636 | 19071320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071320&pc=OZD | 03/21/2024 | 05/04/2024 |  | 04/09/2024 |  |  | 03/21/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2025441809 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Difficult to Insert | [6 empty] | Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male postcardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella device for support. however, during delivery of the impella pump a left ventricular perforation occurred. it was noted no intervention was required for the small perforation due to thickened left ventricle, and support continued as planned with the pump. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. it was reported that there was difficulty implanting the impella. patient had a small left ventricle and during positioning, the placement was deep and there was a left ventricle perforation. the root cause of the ventricle perforation was most likely due to improper positioning of the impella. "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter." "to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.
28183 | 1220648-2024-09637 | 19071341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071341&pc=OZD | 07/29/2022 | 05/27/2025 |  | 04/09/2024 |  |  | 07/29/2022 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had hemolysis as the impella was positioned very shallow, with low flows. the catheter was adjusted to 3.5 cm from the valve and performance was adjusted to p6 where it will remain until more pressors can be removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28184 | 1220648-2024-09638 | 19071373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071373&pc=OZD | 03/22/2024 | 04/18/2024 |  | 04/09/2024 |  |  | 03/22/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1273762 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the automated impella controller (aic) did not recognize the purge disc. the aic was exchanged and the procedure was able to be completed without issue. there was no reported patient harm. | [5 empty] | The investigation for the console (aic) purge disc/flag issues has been completed. the impella device and data logs were returned for evaluation. inspection of the console performed revealed broken/missing purge disk detect flag. data logs from the reported day of event are consistent with complaint and show multiple attempts to install purge cassette and disk, each time with purge disk not detected alarm (#402). therefore, the cause of the issue was a defective component. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28185 | 1220648-2024-09639 | 19071419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071419&pc=OZD | 09/28/2022 | 05/12/2025 |  | 04/09/2024 |  |  | 09/28/2022 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214323 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The complainant reported that while the 63 year-old male patient was on impella cp support, there was no pulse on right foot. also, the patient had heparin intolerance. therefore, heparin was withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28186 | 1220648-2024-09640 | 19071430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071430&pc=OZD | 03/22/2024 | 05/08/2024 |  | 04/09/2024 |  |  | 03/22/2024 | 04/09/2024 |  |  | 04/30/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024349473 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. following the device placement, the patient experienced bleeding from the access site. a suture was placed, and pressure was held for one hour to achieve hemostasis. perclose sutures were subsequently tightened, and support proceeded with the impella device. | [5 empty] | The investigation into access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related. perclose sutures were subsequently tightened to reduce the access site bleeding as per the clinical description. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28187 | 1220648-2024-09641 | 19071571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071571&pc=OZD | 09/18/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 09/18/2022 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205958 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 66-year-old female patient in non-st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced hemolysis that was discovered when all laboratory blood draws were hemolyzed. the impella cp pump was repositioned and the patient was placed on continuous renal replacement therapy. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28188 | 1220648-2024-09642 | 19071581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071581&pc=OZD | 11/30/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 11/30/2022 | 04/09/2024 |  |  | 07/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256706 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The complainant reported that the 80 year-old female patient on impella cp had limb ischemia. impella was inserted in the right artery as arterial pressure decreased from 120's to 50¿s. patients right foot was numb and a bair hugger was placed with multiple rounds of bicarbonate and epinephrine given. patient was stabilized with maps of 80 and the pulse improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28189 | 1220648-2024-09643 | 19071594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071594&pc=OZD | 08/21/2022 | 05/22/2025 |  | 04/09/2024 |  |  | 08/21/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180979 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support due to heart disfunction. a surgical cutdown was performed, and arterial control was obtained. the impella cp was removed in usual fashion, and arteriotomy was surgically closed followed by surgical wound closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28190 | 1220648-2024-09644 | 19071808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19071808&pc=OZD | 08/26/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 08/26/2022 | 04/09/2024 |  |  | 05/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support for heart disfunction. while on support, the distal pulses were not found by doppler. additionally, the ultrasound revealed thrombus. vascular consult obtained and patient was taken for clot removal. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28191 | 1220648-2024-09645 | 19072245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072245&pc=OZD | 09/23/2022 | 10/29/2025 |  | 04/09/2024 |  |  | 09/23/2022 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251625 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a 80-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the purge disc of the automated impella controller (aic) was found to be broken. the aic was swapped out to resolve the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | D2 revised common device name since the initial manufacturer device report number was submitted in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 investigation conclusions revised from the initial submission in accordance with updated procedures.
28192 | 1220648-2024-09646 | 19072323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072323&pc=OZD | 03/27/2024 | 06/24/2024 |  | 04/09/2024 |  |  | 03/27/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported when starting a case, the deairing of the purge cassette and pump were not working. the purge wizard skips the deairing. no alarms occurred. the console was exchanged as a result. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the root cause of the pump stop was most likely contaminant built up from re-use which would explain why the initial run with pump (b)(4) would generate an error but subsequent runs would pass. the root cause of skipping step 2 of the case wizard was the purge disc already having been inserted. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
28193 | 1220648-2024-09647 | 19072446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072446&pc=OZD | 09/26/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 09/26/2022 | 04/09/2024 |  |  | 06/30/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247188 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 67-year-old female was being placed on an impella cp for mechanical circulatory support. the 8f and 6f dilators used for the impella cp insertion failed to advance through the patient¿s vasculature due tight stenosis in common iliac just below bifurcation and concentric calcification of the vessels. the implant was aborted and an intra-aortic balloon pump was placed. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28194 | 1220648-2024-09648 | 19072600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072600&pc=OZD | 03/27/2024 | 07/15/2024 |  | 04/09/2024 |  |  | 03/27/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1197225 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 51-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the automated impella controller (aic) had an aic failure alarm. the aic was replaced successfully. there was no adverse effect to the patient reported. |  |  |  |  | The investigation for the console (aic) failure alarm has been completed. the aic was returned for evaluation. upon visual inspection and functional testing, there were no issues found that could cause the reported controller failure. data logs were reviewed which confirmed the reported failure mode given there was a controller error alarm (opm comm stopped ¿ event set # 1043) triggered during the clinical procedure. the root cause of the controller error was due to an aic software issue. added- b5-mso review, d9 device return date d2-corrected common device name. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28195 | 1220648-2024-09649 | 19072769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072769&pc=OZD | 03/28/2024 | 01/02/2025 |  | 04/09/2024 |  |  | 03/28/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1163156 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during impella cp support for 50-year-old male patient, purge pressure was noted to be low. the nurse detected a broken part at the point where the purge disc must be arranged. the aic was replaced. the issue was resolved. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. there was no issue found during the case. the device functioned/operated to specification. d9 date device returned to manufacturer added. | The investigation for the console (aic) purge disc issues has been completed. the console nor the data logs were returned for evaluation. without additional information, the root cause of the console purge disc issues could not be determined. | The automated impella controller (aic) was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28196 | 1220648-2024-09650 | 19072804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072804&pc=OZD | 10/20/2022 | 03/13/2025 |  | 04/09/2024 |  |  | 10/20/2022 | 04/09/2024 |  |  | 05/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227651 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a patient experienced ventricular tachychardia/fibrillation after placement of the impella cp, an epinephrine bolus was given. the physicians discovered that the impella was unloading the left ventricle to the point that it was hyperdynamic, so the power was lowered to setting p6. during support, heparin was withheld due to oozing from impella site. femstop was placed above insertion site. then it was later reported that the patient received 1 unit of packed red blood cells for drop in hemoglobin. the patient was taken back to the operating room for washout/re-exploration and bleeding was noted from arterial and venous cardiopulmonary bypass sites. in addition, there was a clot noted in the mediastinum and washed out. the patient remained on impella support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. this report is being filed as part of a retrospective review of historical records.
28197 | 1220648-2024-09651 | 19072894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072894&pc=OZD | 10/07/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 10/07/2022 | 04/09/2024 |  |  | 06/30/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249468 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 56-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient developed an occluded superior femoral artery (sfa) from the reperfusion sheath. an angioplasty was completed to resolve the occlusion. | [5 empty] | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28198 | 1220648-2024-09652 | 19072999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19072999&pc=OZD | 10/24/2022 | 03/12/2025 |  | 04/09/2024 |  |  | 10/24/2022 | 04/09/2024 |  |  | 06/30/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023233902 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The complainant reported a patient with post-cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was report that a coronary artery bypass graft was performed while on impella support and the patient experienced bleeding from the ascending replacement anastomosis that was difficult to control. the impella was then explanted and replaced. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28199 | 1220648-2024-09653 | 19073138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19073138&pc=OZD | 10/11/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 10/11/2022 | 04/09/2024 |  |  | 07/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249941 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 62-year-old female female in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing at the impella access site. the access site was redressed, angle of entry corrected and re-sutured, and the bleeding was resolved. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28200 | 1220648-2024-09654 | 19073317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19073317&pc=OZD | 03/28/2024 | 05/10/2024 |  | 04/09/2024 |  |  | 03/28/2024 | 04/09/2024 |  |  | 09/30/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404450 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was found to be sitting too shallow in the ventricle. due to the noted positioning and the improved hemodynamics of the patient, the decision was made to remove the pump. there was no patient harm. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the positioning issue was due to patient condition. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28201 | 1220648-2024-09655 | 19074003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19074003&pc=OZD | 03/28/2024 | 05/21/2024 |  | 04/09/2024 |  |  | 03/28/2024 | 04/09/2024 |  |  | 11/30/2025 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467850 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 28-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after seven days of support, the patient developed an access site hematoma that was treated by pocket drainage and impella repositioning. the patient remained on impella for support and was reported as stable. | [5 empty] | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. d.4 revised serial number as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28202 | 1220648-2024-09656 | 19074246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19074246&pc=OZD | 03/25/2024 | 09/23/2024 |  | 04/09/2024 |  |  | 03/25/2024 | 04/09/2024 |  |  |  |  | 04/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1262851 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 46-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the pump was supporting when the patient was brought to the computed tomography (ct) imaging suite. the automated impella controller (aic) alarmed and stopped. the team replaced the aic and pump restarted. the ct was completed, and patient arrested again, and the family made the choice to withdraw care. the patient expired. per abiomed medical safety officer, the aic error resulted in loss of support on a pump dependent patient. association to death could not be excluded. | [5 empty] | The investigation for the console (aic) pump stop has been completed. data logs were reviewed which confirm the console software tried to reset the pmd due to lack of communication. the reset of the pmd was unsuccessful leading to the impella stopped controller failure alarm being issued. shutdown was then requested by the user. shortly after, another alarm controller error (pbm voltage out-of-spec) alarm was issued. eventually the console was rebooted but unable to reinitialize pmd leading to system self-check failure. the console was returned for evaluation. upon initiating the unit, there was a system self-check failure. performing the fw check with a test pc plugged into the console revealed that there was no information of the sau_motor_1 component of the test. visual inspection of the internal circuitry of the aic revealed that both leds of the impellatronic board were not illuminating as expected. there was voltage supply to the eep sub circuitry of the impellatronic board but the led was not illuminated. there was no voltage being supplied from the pbm to the pmd portion of the impellatronic board. 20v was expected. it was determined that the dc/dc converter on the pbm was not outputting voltage through isolative testing. when a known good pbm was swapped into the console under investigation. both leds of the impellatronic board would not illuminate indicating that the impellatronic was electrically compromised. the root cause for the impella stopped ¿ controller failure alarm and the system self-check failure was a defective mtm dc/dc converter on the pbm, unable to supply 20v. there is no indication if the component having an expiration/maximum usage. according to the dc/dc converter vendor, this was determined to not be a systemic, age-related failure. the instructions for use for a pump as it relates to an impella 5.5 being used with a controller states the following: section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ added- b2-date of death d2-corrected common device name. f6/f8 should have been left blank in the initial report. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
28203 | 1220648-2024-09657 | 19075065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075065&pc=OZD | 11/17/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 11/17/2022 | 04/09/2024 |  |  | 08/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella support 2 perclose were put in place at time of explant but cut by the tech when undoing the dressing. they decided to explanted pump and replaced with 14 fr sheath. sheath removed today and manual pressure held. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28204 | 1220648-2024-09658 | 19075482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075482&pc=OZD | 06/01/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 06/01/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174103 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/01/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced lower limb ischemia. no further information has been provided. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28205 | 1220648-2024-09660 | 19075726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075726&pc=OZD | 05/19/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 05/19/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185929 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction alarms that were resolved by lowering pump flow. additionally, the patient experienced bleeding at the access site that was resolved by withholding heparin. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28206 | 1220648-2024-09659 | 19075732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075732&pc=OZD | 06/02/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 06/02/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201763 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. after the pigtail catheter was inserted into the left ventricle, the patient developed ventricular tachycardia, prompting a single shock and subsequent restart of the impella cp. hemolysis was noted during the night. the patient was placed on cannulation of extracorporeal membrane oxygenation. the impella cp was successfully explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28207 | 1220648-2024-09661 | 19075759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075759&pc=OZD | 05/19/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/19/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199924 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced 4 to 5 beat run of ventricular tachycardia (vt) and patient was converted from normal sinus to a-flutter rhythm, but support remained solid at 3-3.3 l/min. the impella was rapidly weaned and removed without issue post procedure. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28208 | 1220648-2024-09662 | 19075818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075818&pc=OZD | 05/19/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/19/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient access site oozing was noted post removal of the 14fr peel away sheath, however multiple practices were utilized without success. the impella was weaned and removed from impella using perclose successfully providing hemostasis. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28209 | 1220648-2024-09663 | 19075832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075832&pc=OZD | 06/02/2022 | 04/09/2024 |  | 04/09/2024 |  |  | 06/02/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201437 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. oozing was noted from the groin site on arrival. the dressing was taken down, angle match was done, and new dressing was applied, achieving hemostasis. systemic heparin was on hold since due to bleeding and supra-therapeutic inr, activated clotting time, and partial thromboplastin time. right groin impella site dressing was found saturated in early morning and redressed/angle matched with scant oozing since and no hematoma noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28210 | 1220648-2024-09664 | 19075967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19075967&pc=OZD | 06/16/2022 | 03/07/2025 |  | 04/09/2024 |  |  | 06/16/2022 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023207371 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a clot had formed near the left main/ left anterior descending artery region, and a thrombectomy was performed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Analysis summary the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
28211 | 1220648-2024-09665 | 19076069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076069&pc=OZD | 06/17/2022 | 03/07/2025 |  | 04/09/2024 |  |  | 06/17/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023183501 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the micron filter on the purge sidearm was cracked and leaking purge solution, which caused the purge flow to increase. the physician used a 21g needle and the backup purge cassette to bypass the pre-attached sidearm. upon completion of the bypass, the purge flow was down, and the purge pressure was 454. the purge solution was change from d10w to d5w w/ 25 meq of sodium bicarb. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28212 | 1220648-2024-09666 | 19076143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076143&pc=OZD | 05/19/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/19/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199840 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, thee patient experienced oozing from the access site. the site was sutured, femstop was placed, manual pressure was applied, and dressings were changed. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28213 | 1220648-2024-09667 | 19076873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076873&pc=OZD | 05/31/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 05/31/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella cp was accidentally pulled out of the left ventricle when trying to reposition. the patient was doing well so the impella remained out. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28214 | 1220648-2024-09668 | 19076881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076881&pc=OZD | 06/01/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 06/01/2022 | 04/09/2024 |  |  | 12/31/2023 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170678 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2022 |  |  |  |  | The complainant reported that the patient¿s intra-aortic balloon pump (iabp) was scheduled to be escalated to the impella cp. the device had ¿position low¿ alarm resulting in suction alerts. the impella cp was pulled back from the atrioventricular annulus; however, the suction alarms persisted. one liter of normal saline was given to the patient to rule out hypovolemia as the cause of the suction alarms. the impella cp was explanted, and a new device was implanted. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28215 | 1220648-2024-09669 | 19076891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076891&pc=OZD | 06/17/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 06/17/2022 | 04/09/2024 |  |  | 03/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204854 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2022 |  |  |  |  | The complainant reported that a patient was on impella cp support due to st-segment elevation myocardial infarction (stemi). while on support, oozing was noted at the access site. no additional adverse patient effects were reported. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28216 | 1220648-2024-09670 | 19076900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076900&pc=OZD | 06/25/2022 | 04/02/2025 |  | 04/09/2024 |  |  | 06/25/2022 | 04/09/2024 |  |  | 01/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187493 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding noted at the site. as a result, heparin was stopped. moreover, there were no distal pulses. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28217 | 1220648-2024-09671 | 19076916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076916&pc=OZD | 05/28/2022 | 03/10/2025 |  | 04/09/2024 |  |  | 05/28/2022 | 04/09/2024 |  |  | 02/29/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023203046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed a hematoma at the insertion site, due to the steep angle of the pump. it was suggested for the angle to be decreased a little to attempt to decrease the growth of the hematoma. the hematoma boarders were marked and monitored while the patient received systemic heparin. eventually, the device was explanted independently with no issues. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28218 | 1220648-2024-09672 | 19076921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19076921&pc=OZD | 12/14/2022 | 05/23/2025 |  | 04/09/2024 |  |  | 12/14/2022 | 04/09/2024 |  |  | 07/31/2024 |  | 04/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252178 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported patient was implanted with an impella 5.5 device for mechanical circulatory support. while on impella support, the patient experienced heparin induced thrombocytopenia positive. heparin was withheld. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28219 | 1220648-2024-09559 | 19058052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058052&pc=OZD | 05/31/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 05/31/2022 | 04/08/2024 |  |  | 01/31/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022034521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced going in and out of atrial fibrillation with rapid ventricular response. amnio boluses were given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
28220 | 1220648-2024-09560 | 19058118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058118&pc=OZD | 05/31/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 05/31/2022 | 04/08/2024 |  |  | 01/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189114 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a 60-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to stress induced cardiomyopathy and left ventricular outflow tract obstruction. while on support, when the reperfusion catheter was removed a large clot was found at the superficial femoral artery and a penumbra was used successfully to remove it. another clot was found deep in the left femoral, but the physician decided to not remove it. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
28221 | 1220648-2024-09561 | 19058161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058161&pc=OZD | 05/31/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/31/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195194 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a 61-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient experienced oozing and hematoma at the access site. the patient received one unit of packed red blood cells and the site was redressed to resolve the bleeding. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28222 | 1220648-2024-09562 | 19058226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058226&pc=OZD | 05/31/2022 | 05/20/2025 |  | 04/08/2024 |  |  | 05/31/2022 | 04/08/2024 |  |  | 01/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support after an episode of ventricular tachycardia and complaints of chest pain. it was reported that the impella was pulled across the valve before the cross clamp was applied. multiple attempts were made to re-cross the valve following the coronary artery bypass graft (cabg), but was unsuccessful. the decision was made to remove the device. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28223 | 1220648-2024-09563 | 19058369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058369&pc=OZD | 03/26/2024 | 05/14/2024 |  | 04/08/2024 |  |  | 03/26/2024 | 04/08/2024 |  |  |  |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456785 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 41 year-old female was implanted with an impella device for mechanical circulatory support. the patient became agitated and pulled out the arterial line. ventricular fibrillation followed. the patient was resuscitated, and an echocardiogram found the impella device deep. the impella device was repositioned, and the patient was stable. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the impella cp with smart assist system states in section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
28224 | 1220648-2024-09564 | 19058410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058410&pc=OZD | 04/01/2024 | 05/20/2024 |  | 04/08/2024 |  |  | 04/01/2024 | 04/08/2024 |  |  | 03/31/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342302 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 45 year-old male with non-ischemic cardiomyopathy, non-st elevated myocardial infarction, and cardiogenic shock was implanted with an impella device for mechanical circulatory support. the device was removed due to the patient developing hemolysis. the impella device was replaced another impella device. the patient survived the procedure and continued support. | [5 empty] | The investigation for the reported hemolysis has been completed. pump was not returned for investigation. data logs show that there was placement signal low issues and positioning issue which was managed by pulling back the catheter. also, purge pressure high and blocked alarms were seen after positioning issue was resolved. after 4 hours of impella placement, red urine was noticed which corresponded to the positioning issue alarm. per the data log and clinical information provided, the root cause of the hemolysis issue was most likely improper positioning of the device. the instructions for use referenced in the initial report is from the ifu for impella cp with smartassist system. section: use of echocardiography for positioning of the impella catheter, impella catheter too far into the left ventricle, section: hemolysis, section: suction , and section: potential adverse events (united states) which now includes cardiac or vascular injury (including ventricular perforation).¿ d4-updated model number added- d6a/d6b. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine. hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure. patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.
28225 | 1220648-2024-09565 | 19058444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058444&pc=PYX | 04/02/2024 | 05/28/2024 |  | 04/08/2024 |  |  | 04/02/2024 | 04/08/2024 |  |  | 04/30/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024338003 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 60-year-old male patient with cardiomyopathy was implanted with an impella rp device for mechanical circulatory support. it was reported that a dilator kept popping off of the peel away sheath during attempted delivery. despite verifying a proper connection, the components still came apart when inserting. due to the noted issue, a second introducer was opened and utilized for pump implant. there was no patient harm. | [5 empty] | The investigation for the reported introducer damage has been completed. a 23fr peel away introducer sheath and 23fr dilator were returned for evaluation. 23fr dilator was able to be fully secured onto 23fr introducer sheath. when pressure was applied to the end of the dilator, dilator would pop off the introducer blue hub even when fully secured and screwed on. upon visual inspection of the clear locking cap on the 23fr dilator, no abnormalities were found. upon visual inspection of the blue hub on the 23fr, the securing tab was found to be damaged. measurements of the damaged blue hub were taken and one side of the blue introducer hub was found to not meet the 0052-3021 drawing specs. length of top of blue hub on the damaged side found to be out of spec at 9.08mm with drawing spec at 9.84+/-.02mm. height of the securing tab on the blue hub found to be out of spec at 0.772mm with drawing spec at 0.84+/-.02mm. length of the securing tab on the blue hub found to be out of spec at 0.611mm with drawing spec at 0.64+/-.02mm. clinical details noted that when inserting the 23fr peel away sheath with 23fr dilator, the dilator would pop off the introducer sheath when being inserted into the patient. physician noted that they attempted to screw on the dilator before inserting but would still pop off. the root cause of the issue was due one of the blue hub securing tabs dimensions not meeting drawing specs resulting in tab damage and the clear dilator cap popping off with applied force. added-d9 device return date, h6 impact code 4629 added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28226 | 1220648-2024-09566 | 19058480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058480&pc=OZD | 04/02/2024 | 05/28/2024 |  | 04/08/2024 |  |  | 04/02/2024 | 04/08/2024 |  |  |  |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025467665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 27-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the access site was persistently bleeding, and a hematoma developed. the decision was made to explant the impella cp and place intra-aortic balloon pump. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management since act was 261s after transfer to icu. the instructions for use referenced in the initial report is from the ifu impella cp with smart assist system in sections: general patient care considerations, entering cardiac output, and potential adverse events (united states) , which now includes "cardiac or vascular injury (including ventricular perforation¿. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ . ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ . ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
28227 | 1220648-2024-09567 | 19058571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058571&pc=OZD | 04/01/2024 | 05/13/2024 |  | 04/08/2024 |  |  | 04/02/2024 | 04/08/2024 |  |  |  |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025463254 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a female patient of an unknown age in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. as a result of the resistance caused by the patient's vascular anatomy, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition due to the patient's vascular anatomy and small artery size. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28228 | 1220648-2024-09568 | 19058582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058582&pc=OZD | 05/31/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 05/31/2022 | 04/08/2024 |  |  | 11/30/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161805 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Ischemia | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a 50-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that echocardiography was performed after impella insertion and revealed the impella cp pigtail was under papillary muscle and the medical team was unable to reposition towards mid left ventricle. the impella was successfully repositioned into the mid-ventricular space the next day. it was reported while on support, the patient experienced hematuria, plasma free hemoglobin results >500 mg/dl, and acute kidney failure. the patient was treated with, continuous renal replacement therapy, repositioning of the pump into the mid-ventricular space, and heparin was discontinued, but ultimately the pump was removed for the patient's hemolysis. it was reported the patient experienced a loss of pulses in the impella leg. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28229 | 1220648-2024-09569 | 19058642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058642&pc=OZD | 04/03/2024 | 06/12/2024 |  | 04/08/2024 |  |  | 04/03/2024 | 04/08/2024 |  |  |  |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella cp was prepped for implant and once the white cord was plugged in the alarm error ¿device defective, change pump¿ came on. pump was then changed out to a new one. the pump was not implanted in the body. it only happened during set up. a new impella cp was implanted successfully. | The patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that the there was no association between impella use and outcome. |  |  |  |  | The device has not been received from the customer. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump not detected issue has been completed. no pump was returned for investigation. data logs confirmed the impella defective alarm. the root cause of the pump not detected issue was not determined since product was not returned. added- b5 mso review, d4-corrected model number, g1-corrected mfr site name and address.
28230 | 1220648-2024-09570 | 19058735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058735&pc=OZD | 04/04/2024 | 10/10/2024 |  | 04/08/2024 |  |  | 04/04/2024 | 04/08/2024 |  |  |  |  | 04/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | No Display/Image | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | A japan user facility reported a 58-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was that that the automated impella controller (aic) had a controller error alarm, and the screen went dark. a reboot was required. the aic was replaced for patient¿s safety. there was no patient harm. | [5 empty] | The investigation for the reported console (aic) battery charging issues has been completed. data logs confirmed the reported complaint of an unexpected shutdown. however, there was no evidence of a preceding controller alarm. immediately preceding the unexpected shutdown, the following event occurred: ¿systemidle: process io-graphics 339995 died abnormally¿. previous investigations of related unexpected shutdown issues indicate that a crash of the io-graphics process within the console's underlying qnx operating system caused the unexpected shutdown of the console. the device was returned for evaluation. no apparent issues were identified. the reported complaint could not be reproduced. corrections to the initial mfr report:f6/f8 should have been left blank. | The automated impella controller was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
28231 | 1220648-2024-09571 | 19058769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058769&pc=OZD | 10/14/2023 | 09/24/2025 |  | 04/08/2024 |  |  | 10/14/2023 | 04/08/2024 |  |  | 08/31/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385965 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/14/2023 |  |  |  |  | The user facility reported a 43-year-old female patient implanted with an impella cp for mechanical circulatory support. heparin was withheld after bleeding was noted into the chest and patient was given blood products. | The complainant also reported that the impella had placement signal not reliable alarms. the alarm was disabled. |  |  |  |  | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the most likely cause of the peripheral vessel vascular damage was patient condition related. instructions for use as follows: impella 5.5 with smartassist for use during section: warnings and cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings and cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ d6a and d6b added. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09571 f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for a visual inspection. data log analysis confirmed a placement signal not reliable alarm occurred on 10/14. this was followed by an immediate drop in signal-to-noise ratio (snr), gain, and light, while the contrast began to barcode and lamp maxed out. additionally, motor current increased. the 5s parameters returned to normal after indicating that the fiber was most likely torqued. no other abnormalities were found. the most likely cause of the optical signal issue was a damaged fiber line. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. corrections that were made from the initial manufacturer device report number 1220648-2024-09571. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
28232 | 1220648-2024-09573 | 19058966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19058966&pc=OZD | 05/25/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 05/25/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. a hematoma was noted, and due to angle matching, platelets were given, heparin held and redressed site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28233 | 1220648-2024-09574 | 19059148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059148&pc=OZD | 10/17/2023 | 09/26/2025 |  | 04/08/2024 |  |  | 10/17/2023 | 04/08/2024 |  |  | 08/31/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382312 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, the patient had experienced oozing at the access site. pressure was applied and patient was given blood products. | Additional information indicated that, during patient support, the pump experienced placement signal not reliable alarm. the patient remained in p8-9 with normal flows, and the waveform was present. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the most likely cause of the major access site bleeding was patient condition related. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. no product was returned for an investigation. upon further review of the data logs, it is apparent that the console was the potential cause of the optical signal issue. no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-09574.
28234 | 1220648-2024-09575 | 19059213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059213&pc=OZD | 05/25/2022 | 05/20/2025 |  | 04/08/2024 |  |  | 05/25/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194978 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was being implanted with an impella cp for mechanical circulatory support. during implantation, the physician attempted to recross the valve with the impella and .018 wire, however this was unsuccessful. therefore, the impella and the wire were explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28235 | 1220648-2024-09576 | 19059260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059260&pc=OZD | 05/25/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 05/25/2022 | 04/08/2024 |  |  | 12/31/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the valve of the sheath. the bleeding stopped when the destination sheath was removed. it was also reported that delivery issues were experienced; the physician was unable to pass through vasculature. due to blood loss, one unit of packed red blood cells was administered. the patient's blood pressure stabilized and pulsatility returned. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28236 | 1220648-2024-09577 | 19059416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059416&pc=OZD | 05/25/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/25/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198443 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the patient experienced access site bleeding. the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28237 | 1220648-2024-09578 | 19059433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059433&pc=OZD | 05/27/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/27/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/27/2022 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. implanation was unsuccessful due to the physicians inability to get past the iliac even after shockwave attempt. the impella placement was aborted due to patient's anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28238 | 1220648-2024-09579 | 19059543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059543&pc=OZD | 05/30/2022 | 05/22/2025 |  | 04/08/2024 |  |  | 05/30/2022 | 04/08/2024 |  |  | 11/30/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164252 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Device Contaminated at the User Facility | [5 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 05/30/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the pump was removed and had to be replaced due to possible contamination/damage to the catheter. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28239 | 1220648-2024-09580 | 19059603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059603&pc=OZD | 05/25/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/25/2022 | 04/08/2024 |  |  | 01/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187498 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, once the wire entered the left anterior descending (coronary artery branch) it was suspected that a plaque shift/rupture which caused occlusion. the patient went into ventricular fibrillation and arrested, return of spontaneous circulation x2. the impella cp was immediately placed for support. ultimately the decision was made to wean and remove the impella as the left femoral artery was very small and the foot was not being perfused. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28240 | 1220648-2024-09581 | 19059638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059638&pc=OZD | 05/30/2022 | 05/28/2025 |  | 04/08/2024 |  |  | 05/30/2022 | 04/08/2024 |  |  | 01/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190992 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2022 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma and hemolysis during support. the patient was treated with a blood transfusion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28241 | 1220648-2024-09582 | 19059705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059705&pc=OZD | 05/31/2022 | 05/22/2025 |  | 04/08/2024 |  |  | 05/31/2022 | 04/08/2024 |  |  | 12/31/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient was weaned from bypass, but the pump speed level was not able to be run higher than p3-p4 without requiring suction and suboptimal flows. transesophageal echocardiography (tee) was performed to confirm the device position. bedside echocardiogram identified a kink in the catheter potentially due to the patient¿s anatomy. the decision was made to leave the device in place and remove the following day. the patient was weaned, the device was explanted bedside and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28242 | 1220648-2024-09583 | 19059755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059755&pc=OZD | 11/28/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 11/28/2022 | 04/08/2024 |  |  | 08/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268327 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician placed a distal perfusion catheter through the side port of the repo sheath. the physician was not satisfied with the distal flow and made the decision to place an external bypass. the patient remained on support and was transported to unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28243 | 1220648-2024-09584 | 19059772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059772&pc=OZD | 05/25/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/25/2022 | 04/08/2024 |  |  | 01/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188862 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced a left leg ischemia. the patient was administered one liter of normal saline, atropine x1, and started on dopamine. impella flows improved. monochorionic and pulmonary stenosis both pulsatile. the impella was weaned and removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28244 | 1220648-2024-09585 | 19059804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059804&pc=OZD | 06/01/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 06/01/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2022 |  |  |  |  | The complainant reported a patent in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, patient labs were intermittently hemolyzing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28245 | 1220648-2024-09586 | 19059827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059827&pc=OZD | 10/16/2023 | 10/20/2025 |  | 04/08/2024 |  |  | 10/16/2023 | 04/08/2024 |  |  | 07/31/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371248 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | Device Sensing Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the pump experienced high purge pressure and pump stop occurred after days of purge alarms. subsequently, the pump was removed from the left ventricle at the bedside and explanted in the operating room. an intra-aortic balloon pump (iabp) was utilized as a bridge to transplantation, and during explantation, a clot was observed at the outlet. | [5 empty] | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results.
28246 | 1220648-2024-09587 | 19059913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059913&pc=OZD | 06/14/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 06/14/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported a 22-year-old female was implanted with an impella cp device for mechanical circulatory support. device encountered low pump flows, limb ischemia and access site bleeding. low pump flows were encountered as suction alarms occurred and persisted when increasing the pump support level settings. this was treated by repositioning the pump. additionally, limb ischemia was noted as no doppler pulses were present in both lower extremities. this was treated by increasing pump flow to maintain a mean arterial pressure (map) above 60 and having the patient undergo aggressive diuresis. access site bleeding was noted by the small amount of oozing at the left groin insertion site. the site was redressed, angle matched and suture placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28247 | 1220648-2024-09381 | 19059962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19059962&pc=OZD | 03/21/2024 | 05/23/2024 |  | 04/08/2024 |  |  | 03/21/2024 | 04/08/2024 |  |  |  |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457126 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | Unintended Movement |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a male with cardiogenic shock and acute myocardial infarction was implanted with an impella device for mechanical circulatory support. during the impella device implantation, bleeding around the sheath was observed. another sheath was then used, but then the pump position was lost. the physician then decided explant the pump and not make a third implantation attempt. the patient remains on support. | [5 empty] | The investigation for the reported access site bleed and positioning issues has been completed. the cp pump and guidewire were returned for evaluation and visually inspected. guidewire kink damage observed, and small impact of guidewire end found on pigtail with slight bend or bulge outwards. no major damages or kink observed in the pump. no other damage along the guidewire and the tip observed. pump was functioning good with any issue. the cause of the positioning issue was use related due to improper technique. the cause of the access site bleeding was most likely patient condition related due to patient calcified vessel. the cause of the product damage (guidewire damage) issue was most likely use related per return product investigation. the instructions for use referenced in the initial report is from ifu for impella cp with smartassist system, sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes cardiac or vascular injury (including ventricular perforation).¿ added- d9 device return date d4-updated model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
28248 | 1220648-2024-09588 | 19060072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19060072&pc=OZD | 06/15/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 06/15/2022 | 04/08/2024 |  |  | 03/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203829 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 71-year-old male requiring impella cp for mechanical circulatory support. the device was placed with no issues. it is stated the patient was sitting, coughed and had a ventricular tachycardia (vt) episode while on support. the recommendation was that the position of the device be checked via echocardiogram. the correct position was confirmed. the patient was successfully weaned from the impella cp and survived explant. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
28249 | 1220648-2024-09589 | 19060185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19060185&pc=OZD | 06/17/2022 | 04/03/2025 |  | 04/08/2024 |  |  | 06/17/2022 | 04/08/2024 |  |  | 11/30/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165434 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, significant oozing was noted from the access site. as a result, thrombectomy procedure was performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
28250 | 1220648-2024-09590 | 19060402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19060402&pc=OZD | 08/14/2022 | 03/12/2025 |  | 04/08/2024 |  |  | 08/14/2022 | 04/08/2024 |  |  | 04/30/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella suddenly stopped after the pump was started. troubleshooting was carried out in the event of an emergency stop, but the pump did not restart and had to be removed and replaced. when the catheter was checked after removal, the guidewire guide tube remained in the cannula. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28251 | 1220648-2024-09591 | 19060589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19060589&pc=OZD | 02/26/2024 | 09/19/2025 |  | 04/08/2024 |  |  | 02/26/2024 | 04/08/2024 |  |  | 11/30/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC.. |  |  | US |  | 2024431202 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/26/2024 |  |  |  |  | Additional information states that the patient experienced access site bleeding that was resolved by withholding heparin, insertion of a jp drain, redressing, and blood products. the patient also went into ventricular tachycardia, which was resolved by shock and antiarrhythmic medication. | The user facility reported a 76-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on impella 5.5 support, the patient had low platelets, the patient received four units of blood and one unit of platelets. | The us complainant reported that an implanted impella 5.5 pump began to experience recurrent placement signal alarms during patient support. the alarm was disabled and the staff was advised that the pump can function normally without the placement signal. there was no patient harm associated with the placement signal issues and support was maintained. additionally, the patient experienced access site bleeding that was resolved by withholding heparin, jp drain, giving blood products, and redressing. the patient also went into ventricular tachycardia, which was resolved by shock and antiarrhythmic medication. the patient also experienced thrombocytopenia that was resolved by giving blood products. reportable due to patient harm requiring intervention for access site bleeding, ventricular tachycardia, and thrombocytopenia. mdr and pmda report required. |  |  |  | The investigation for the reported access site bleed, thrombocytopenia, and ventricular tachycardia (vt) has been completed. the impella pump was not returned for evaluation. data logs were reviewed which indicate a definitive aic failure pattern suspecting the console. no problem was found with the pump. without additional clinical information or the pump return, the cause of the access site bleed, thrombocytopenia, and ventricular tachycardia issue was not determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (unites states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ added-b5 additional information in the narrative, h6 clinical code 1888, 2095 added-d6a | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. upon review of the data logs, it indicates a definitive aic failure pattern suspecting the console. no problem was found with the pump, and it is apparent the console ic3638 was the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
28252 | 1220648-2024-09592 | 19060678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19060678&pc=OZD | 12/22/2023 | 09/30/2025 |  | 04/08/2024 |  |  | 12/22/2023 | 04/08/2024 |  |  | 10/31/2025 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414205 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/22/2023 |  |  |  |  | The user facility reported a 59-year-old male was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient experienced limb ischemia in both lower extremities. right lower extremity remained cool with no pulses at baseline. heparin purge was continued due to concern for clot on outflow. pulses were found to be good on both feet. upon explant of the impella, a clot was found. additionally, the patient had multiple episodes of ventricular tachycardia (vtach) overnight since p8 was dropped to p4 due to suction alarm. throughout impella cp support suction problems persisted. patient was monitored until impella cp device was successfully weaned and explanted. | The complainant also reported that there was a placement signal not reliable. |  |  |  |  | The investigation into the reported issues has been completed since the original report was filed. the pump was returned, evaluated, and visually inspected. there was biomaterial in the purge gap and wrapped around the impeller. based on the product returned and clinical details provided, the root cause of the low pump flow is most likely due to ingested biomaterial. the root cause of the optical signal issue cannot be determined as limited clinical information was provided and the data logs were not returned. as the relevant information regarding the vt/vf and the limb ischemia was not provided, the root cause could not be determined. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was returned and there was biomaterial in the purge gap and wrapped around the impeller. the pump's 5s parameters were taken and the laser continuity test was performed. the placement signal unreliable alarm did not reproduce. no issues were found with the pump besides the ingested biomaterial. the root cause of the optical signal issue cannot be determined as sufficient clinical information was not provided and the data logs were not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09592. f6 and f8 should not have had entries in the initial report. g1 reporting contact email and g1 manufacturing site name and address updated.
28253 | 1220648-2024-09593 | 19061196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19061196&pc=OZD | 09/23/2022 | 04/02/2025 |  | 04/08/2024 |  |  | 09/24/2022 | 04/08/2024 |  |  | 06/30/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247939 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support positioning issues, oozing, and thrombus was noted. repositioned multiple times, blood products were administered, and large amount of clots removed from chest. the patient remained on support and was sent to intensive care unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28254 | 1220648-2024-09594 | 19061333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19061333&pc=OZD | 09/27/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 09/27/2022 | 04/08/2024 |  |  | 07/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231479 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported patient was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp was implanted, the pump speed level could only be run at p4 without the use of suction. an echocardiogram was performed and observed that the impella placement was shallow. the physician repositioned the device to 3.5 cm below the aortic valve. pump speed flow was resolved as the reposition of the impella cp caused the flow to increase. additionally, the patient experienced large hematoma from the access site. the hematoma was expressed, and manual pressure was held. patient was monitored until impella cp device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28255 | 1220648-2024-09595 | 19061396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19061396&pc=OZD | 06/01/2022 | 04/02/2025 |  | 04/08/2024 |  |  | 06/01/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202865 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/01/2022 |  |  |  |  | The user facility reported a 77-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to chest discomfort felt at rest but exacerbated on exertion and associated with dyspnea. patient primary indication was high risk percutaneous coronary intervention. it was reported while on impella cp support, the impella required repositioning of the pig tail, and during repositioning the pig tail appeared to be hung up in papillary. the physician opted to leave the impella in a less than ideal position as the flows were adequate to support the patient with no adverse events. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. h6 medical device problem code: corrected from 2993 adverse event without identified device or use problem to 1158 positioning failure
28256 | 1220648-2024-09596 | 19061439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19061439&pc=OZD | 11/08/2022 | 04/17/2025 |  | 04/08/2024 |  |  | 11/08/2022 | 04/08/2024 |  |  | 05/31/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231134 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while inserting the device, there was introducer damage. the peel-away sheath fenestrations were slightly cracked and seemed to have appeared after forceful attempts were made with the destination sheath. the destination sheath was replaced with a standard sheath and the procedure was completed using the same introducer. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28257 | 1220648-2024-09597 | 19061534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19061534&pc=OZD | 11/09/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 11/09/2022 | 04/08/2024 |  |  | 09/03/2023 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142467 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/09/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was labile and hypotensive. pressure was being held above the impella site. the physician evaluated the patient and requested removal of the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28258 | 1220648-2024-09598 | 19061813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19061813&pc=OZD | 11/24/2022 | 04/02/2025 |  | 04/08/2024 |  |  | 11/24/2022 | 04/08/2024 |  |  | 06/30/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The complainant reported a patient with acute decompensated chronic cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, a unit nurse tripped over the impella pump power cord and the device was pulled out of position, into the aorta. the physicians' efforts to reposition the impella were not successful, and the device was explanted. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28259 | 1220648-2024-09599 | 19067487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19067487&pc=OZD | 05/20/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/20/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200008 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/20/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced slight oozing at the groin site, which was resolved by stopping heparin. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28260 | 1220648-2024-09600 | 19067615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19067615&pc=OZD | 05/20/2022 | 04/08/2024 |  | 04/08/2024 |  |  | 05/20/2022 | 04/08/2024 |  |  | 02/29/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193389 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/20/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. during the explant of the impella cp, two preclose were used. after the deployment of the perclose device, there was some oozing at the right groin site. the patient site was redressed, no new oozing was noted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28261 | 1220648-2024-09601 | 19067712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19067712&pc=OZD | 11/23/2022 | 04/02/2025 |  | 04/08/2024 |  |  | 11/23/2022 | 04/08/2024 |  |  | 06/30/2024 |  | 04/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236935 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient underwent surgical cutdown and repair of the femoral artery during the impella explant procedure. post-removal, the patient experienced limb ischemia requiring fogerty embolization. additionally, two 8cm clots were removed from the patient. pulses were reported to be audible via doppler after the procedure. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28262 | 1220648-2024-09557 | 19057082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19057082&pc=OZD | 09/15/2023 | 04/18/2024 |  | 04/07/2024 |  |  | 03/08/2024 | 04/07/2024 |  |  | 12/31/2024 |  | 04/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315918 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Atrial Fibrillation | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 76 year old female patient presenting with cardiomyopathy for impella support. during support, the patient endured acute renal failure, worsening of atrial fibrillation, and new onset anemia. for acute renal failure, 1 round of hemodialysis was administered. for worsening of atrial fibrillation, a cardioversion was performed. lastly, for new onset anemia, a transfusion was delivered. | [5 empty] | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the renal failure, the vf/vt/af/at (worsening of atrial fibrillation), and the vascular damage - peripheral vessel has been completed since the original report was submitted. the device was not returned for investigation. the root cause of renal failure and the vascular damage was not determined because no product was returned and limited clinical information was provided. as the necessary information was not provided, the root cause of the vt/vf was not determined.
28263 | 1220648-2024-09558 | 19057085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19057085&pc=OZD | 03/08/2024 | 04/18/2024 |  | 04/07/2024 |  |  | 03/08/2024 | 04/07/2024 |  |  |  |  | 04/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025438008 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The complainant reported a 68 year old male patient presenting with cardiomyopathy for impella support. during support, the patient developed an access site hematoma. manual pressure was held and the condition stabilized. it was noted that heparin was "restarted" after hematoma was controlled. | [5 empty] | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use states ¿assess access site for bleeding and hematoma.¿
28264 | 1220648-2024-09537 | 19056555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056555&pc=OZD | 03/10/2024 | 06/05/2024 |  | 04/06/2024 |  |  | 04/03/2024 | 04/06/2024 |  |  | 12/31/2024 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300895 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion of the guidewire for impella implant, the tip of the guidewire was slightly bent and could not be used. it was replaced with a new one. there was no harm to the patient reported. | [5 empty] | The investigation for the reported purge cassette leak and guidewire damage has been completed. the device was not returned for evaluation. however, clinical details were provided to assist in determining root cause. the tip of the 0.018 guidewire was bent bent when taken out of the box. the wire was replaced with a new one. about 24 hours after implant, there was a leak from the purge cassette. the purge cassette was replaced with a new one which resolve the issue. the root cause of the product damage (guidewire damage - peripheral vessel) was not determined as no product was returned and insufficient clinical information was provided. the root cause of the purge leak - cassette was not determined as no product was returned. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28265 | 1220648-2024-09538 | 19056557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056557&pc=OZD | 06/28/2023 | 10/20/2025 |  | 04/06/2024 |  |  | 08/23/2023 | 04/06/2024 |  |  | 02/28/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325127 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/23/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella was repositioned with echocardiogram due to ectopy. the patient is stable post issue. | The complainant also reported that there was a significant discrepancy (~45 mmhg) in aorta pressure (aop) from the arterial line and placement signal (ps) on impella. the patient was non-pulsatile. bedside monitor aop 102/100 and ps 55/52. proper placement was verified via transesophageal echocardiography. the patient was awake and alert in a chair, very talkative. the physician re-wired the arterial line to ensure its accuracy, and there were no additional issues reported. |  |  |  |  | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia could not be determined. the instructions for use for the impella cp with smart assist system in section: potential adverse events (united states) states the following: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ d4-corrected model number added- d6a/d6b f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. according to the clinical, on the 35th day of impella 5.5 support, frequent suction alarms began to occur that lasted until the end of the case. five days later, the pump was repositioned due to ectopy. on the 56th day of impella 5.5 support, a significant discrepancy (~45 mmhg) in aop from arterial line and ps on impella was recorded. the patient continued on support for another eight days until they were successfully transplanted. no product was returned for a visual inspection. data log analysis confirmed frequent suction alarms and that all 5s parameters were within normal bounds. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal issue description. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
28266 | 1220648-2024-09539 | 19056559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056559&pc=OZD | 10/07/2023 | 09/26/2025 |  | 04/06/2024 |  |  | 10/07/2023 | 04/06/2024 |  |  | 08/31/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392378 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/07/2023 |  |  |  |  | The complainant reported that the patient on impella support experienced a placement signal not reliable alarm occurring during rounds. alarm was disabled and echocardiography was called to confirm placement. p5 was confirmed with appropriate flows and waveforms. | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when the guidewire was placed into the left anterior descending artery, the patient had multiple rounds of ventricular fibrillation requiring the patient to be shocked eight times. the patient did not convert independently and cardiopulmonary resuscitation was performed. medications were administered. the patient was intubated. return of spontaneous circulation was achieved. |  |  |  |  | The investigation for the reported ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ added- h6 component code 925 d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. there was no issue found during the case. the device functioned/operated to specification. upon review of the data on the device it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. additional information: b1 product problem b5 b7 d10 concomitant medical products and therapy dates h3 h6 medical device problem code h6 type of investigation h6 investigation findings h6 investigation conclusions h10. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28267 | 1220648-2024-09540 | 19056561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056561&pc=OZD | 11/23/2023 | 04/29/2024 |  | 04/06/2024 |  |  | 11/23/2023 | 04/06/2024 |  |  | 06/30/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365362 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that all of the preparation screens were skipped following connection of the impella cable. while preparing for the implant, preparation screens including purge solution concentration input, impella priming, and optical sensor calibration failed to appear. as a result, a new impella pump was prepped and inserted for support. | [5 empty] | The investigation for the priming issue has been completed. the product was returned to the lab and it was able to go through priming and ran with no issues. the data logs were reviewed which revealed the pump skipping the priming step. however, the right data log shows the purge pressure was already high as the pump was already primed during multiple partial steps to start case start on a previous console, so when the pump was reconnected to the reported console, it recognized the pump as already fully primed. the pump was then switched out as it was thought to malfunction. the root cause of the priming issue is a defective component on the initial console causing the back pressure to be high and back up console to skip the priming steps at case start. d4-updated model number to conform to updated procedures, section d6 has revised with updated information. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28268 | 1220648-2024-09541 | 19056566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056566&pc=OZD | 11/14/2023 | 10/29/2024 |  | 04/06/2024 |  |  | 11/14/2023 | 04/06/2024 |  |  | 05/31/2024 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226927 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the staff noticed purge fluid leaking from the purge disc when they opened the door. the cassette was replaced. the physician then removed the impella from the body. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the priming issue has been completed. purge cassette was not returned for investigation. without the product back, it is not possible to investigate the complaint. therefore, the root cause of the purge-leak cassette issue was not determined since product was not returned. corrections to the initial mfr report: f6/f8 should have been left blank.
28269 | 1220648-2024-09542 | 19056577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056577&pc=PYX | 03/16/2024 | 05/03/2024 |  | 04/06/2024 |  |  | 03/16/2024 | 04/06/2024 |  |  | 06/30/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024369391 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The user facility reported that the wire was found to be kinked in the hoop/wire loop. the medical doctor noticed the wire was kinked before the wire was used in the patient or pump. the team used another wire and the original impella rp pump was successfully placed. | [5 empty] | The investigation for the product damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the bent guidewire - packaging issue was not determined since product was not returned for investigation. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28270 | 1220648-2024-09543 | 19056578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056578&pc=OZD | 03/20/2024 | 09/26/2025 |  | 04/06/2024 |  |  | 03/21/2024 | 04/06/2024 |  |  | 04/30/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339422 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Life Threatening | Death | Required Intervention |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support due to patient decompensating after an impella cp was removed. while on support, the patient had access site bleeding. heparin was withheld. multiple blood products were given. sandbags were placed and coagulation products were provided. the patient's white blood cell count was trending down and an antibiotic was given to the patient. it is unknown if the patient's infection was due to the impella. the patient remains on impella 5.5 support. | The complainant also reported that the optical sensor placement signal was not matching with the arterial line during the whole time patient was on support. care was eventually withdrawn, and the patient expired. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | D.4 catalog number, lot number, expiration date and udi number were incorrect on the previously submitted manufacture device report (mdr). these section has been updated. h.3 manufacture device date was incorrect on the previously mdr. it has been updated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue, infection, and access site bleeding. the pump was not returned for investigation. data logs show several low alarms for placement signal without noted abnormalities in 5s optical signal parameters. bleeding and discrepancies in the optical signal were reported during the case run. the cause of the access site bleeding issue was high patient act values, based on given clinical details. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. necessary evidence has not been provided, the root cause of the infection cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b2 updated to death. b5 updated to include death and placement signal issue description. h1 updated to death. h6 updated to include death and placement signal issue coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09543 d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
28271 | 1220648-2024-09544 | 19056584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056584&pc=OZD | 03/25/2024 | 09/23/2025 |  | 04/06/2024 |  |  | 03/25/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438142 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/25/2024 |  |  |  |  | The user facility reported an 60 year-old female with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient had recent viral and fungal infections and valvular disease as well. during support, suction alarms sounded. the medical staff tried to reposition but was unable to move the catheter. the patient remains on support. | The investigation uncovered placement signal issues. |  |  |  |  | The investigation for the reported low pump flow has been completed. no product was returned for a visual inspection. data log analysis confirmed frequent suction alarms and positioning issues. several impella position unknown alarms and some impella in ventricle alarms were observed. also, several placement signal low alarms found with placement signal trending down and no major impact to 5s parameters. the cause of the low pump flow issue was most likely improper positioning of the device per clinical and log review. the most likely cause of the positioning issues was use related. d4-updated model numberin accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the optical signal issue was not determined because no product was returned, and inconclusive evidence per log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09544. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank. g1 reporting contact email and reporting contact fax number updated.
28272 | 1220648-2024-09545 | 19056588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056588&pc=OZD | 03/28/2024 | 05/14/2024 |  | 04/06/2024 |  |  | 03/28/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464462 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion of the impella device, the 14 french peel away sheath was used with the dilator requiring enough force to insert that the sheath kinked on itself like an accordion. a second sheath was placed and went in with slightly more ease. however, the sheath stuck out about three centimeters and would not advance further due to the patient¿s vessel calcification. the pump was successfully placed. however, it was further reported that the impella pump was weaned and removed because the patient¿s leg was ischemic. | [5 empty] | The investigation for the reported introducer damage and limb ischemia has been completed. the introducer was not returned for evaluation. clinical details states that patient had calcified vessels and dilation was not done prior to pump insertion. upon pump insertion, the sheath buckled during insertion. the root cause of the introducer damage issue was patient condition related to small/ diseased vessels. the root cause of the limb ischemia could not be determined without additional clinical details. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28273 | 1220648-2024-09546 | 19056601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056601&pc=OZD | 03/28/2024 | 07/22/2024 |  | 04/06/2024 |  |  | 03/28/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1194348 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility's biomedical engineer reported that an automated impella controller experienced a controller error outside of service. there was no patient involved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller (aic) yellow alarm has been completed. data analysis: imc logs confirm that the (opm communication stopped) [pump not connected] - controller error alarm was issued on the day of the complaint. device analysis: the failure mode was reproducible upon initial boot up of the console during investigation. visual inspection of the internal circuitry of the console confirmed that ribbon cable between the embedded personal computer (epc) and the optical bench was loose/disconnected and securing it resolved the alarm. conclusion: the root cause for the (opm communication stopped) [pump not connected] - controller error alarm was a disconnected ribbon cable between the epc and the optical bench. preventive maintenance process. d9 added device return and date d2 corrected common device name. e2 corrected health professional from yes to no. f6 and f8 should be blank.
28274 | 1220648-2024-09547 | 19056610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056610&pc=OZD | 03/29/2024 | 04/12/2024 |  | 04/06/2024 |  |  | 03/29/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441848 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that heparin was on hold for a few days due to the patient having a subdural head bleed. the purge was switched to bicarb. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the stroke/cva, the access site bleeding ¿ minor, and the pain at insertion site issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided, the root cause of the stroke/cva could not be determined. the root cause of the access site bleeding and the pain at insertion site was unable to be determined as limited clinical details were provided and no product was returned for analysis.
28275 | 1220648-2024-09548 | 19056613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056613&pc=OZD | 03/29/2024 | 05/10/2024 |  | 04/06/2024 |  |  | 03/29/2024 | 04/06/2024 |  |  | 10/31/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMPTEMPORARY NON-ROLLER TYPE |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024422892 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 59 year-old male with cardiogenic shock and acute myocardial infarction was to be implanted with an impella 5.5 device for mechanical circulatory support. attempts to implant the device were unsuccessful due to the small size of the patient's vasculature. the implant was subsequently aborted, and a an impella device cp was placed. no patient harm has been reported. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because the patient vessels were small to place the impella 5.5 and they were successful in implanting impella cp. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28276 | 1220648-2024-09549 | 19056621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056621&pc=OZD | 03/29/2024 | 05/15/2024 |  | 04/06/2024 |  |  | 03/29/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. however, after the impella pump was placed, it was removed after it interfered with the right coronary artery catheter engagement. the team proceeded with a valvuloplasty. there was no report of an adverse effect to the patient. | [5 empty] | The investigation for the reported insertion/removal issues has been completed. pump was not returned for investigation. data logs were returned and it was observed exactly as mentioned in the clinical information that low flows of 1.5l/min were seen at p9. previous to the complaint pump, another pump was used which had a similar cannula kink observed leading to low flows. per the clinical information and data logs, the root cause of the low flow issue was most likely a kinked cannula due to tavr interaction. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28277 | 1220648-2024-09550 | 19056623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056623&pc=OZD | 03/26/2024 | 05/02/2024 |  | 04/06/2024 |  |  | 03/29/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453251 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old female with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the impella was repositioned under echocardiogram and since then is less ectopy. the purge was changed to heparin. two days later, bloody drainage was noted to the impella site. systemic heparin was placed on hold and bleeding has stabilized since. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed and ectopy has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to anticoagulation management since hemostasis was achieved by holding the heparin. without additional clinical details, the root cause of the ectopy could not be determined. to conform to updated procedures, section d6 was revised with updated information.
28278 | 1220648-2024-09551 | 19056631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056631&pc=OZD | 03/30/2024 | 05/16/2024 |  | 04/06/2024 |  |  | 03/30/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025449667 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 41 year-old male with cardiogenic shock and acute myocardial infarction was attempted to be implanted with an impella cp device for mechanical circulatory support. attempts to place an introducer were unsuccessful due to the small size of the patient's vasculature. the implant was subsequently aborted, and an intra-aortic balloon pump was implanted instead. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the difficulty inserting/removing introducer was most likely patient condition related; the introducer insertion was unsuccessful due to small size of the patient as per the clinical description. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28279 | 1220648-2024-09552 | 19056635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056635&pc=OZD | 03/30/2024 | 05/15/2024 |  | 04/06/2024 |  |  | 03/30/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457367 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 62-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that suction alarms occurred. the pump was repositioned and became subvalvular. while positioning back across the aortic valve, the proximal portion of the cannula became kinked and was removed. | [5 empty] | The investigation into the low pump flow and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue was most likely use related, as it occurred during repositioning, and the doctor was unable to cross the aortic valve. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28280 | 1220648-2024-09553 | 19056636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056636&pc=OZD | 03/31/2024 | 05/13/2024 |  | 04/06/2024 |  |  | 03/31/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025452348 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 65 year-old female with cardiogenic shock and non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient lost blood flow in their lower left extremity during impella cp support. the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28281 | 1220648-2024-09554 | 19056645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056645&pc=OZD | 04/01/2024 | 11/18/2025 |  | 04/06/2024 |  |  | 04/01/2024 | 04/06/2024 |  |  |  |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Misassembly by Users | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 47-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the automated impella controller failed to detect the purge disc on impella connect. it was confirmed by the clinical representative that the nurse did not push the disc in all the way until the click was heard. the reported issue was quickly resolved, and support was continued. no patient harm was reported. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email d2 device name has been updated f6 and f8 should have been left blank. | The investigation has been completed. analysis: 5.5 pump 435614 was started up on (b)(6) 2024 and ran for just under 23 days, disconnecting on 4/5. during this time, including the complaint date of (b)(6) 2024, there were no instances of purge disc not detected alarms. additionally, there were no purge disc alarms outside of step 2 of case start when the purge disc is inserted. root cause: there was no issue found during the case. the device functioned/operated to specification. g1 reporting contact fax number updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28282 | 1220648-2024-09555 | 19056666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056666&pc=OZD | 04/01/2024 | 06/13/2024 |  | 04/06/2024 |  |  | 04/01/2024 | 04/06/2024 |  |  | 07/31/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433897 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 55 year-old male was implanted with an impella cp device for mechanical circulatory support. the patient lost doppler pedal pulses on his right leg. a femoral-femoral external bypass was planned. | Additional information states that care was withdrawn and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of device with outcome |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system in section: potential adverse events (united states), which now includes that the statement "cardiac or vascular injury (including ventricular perforation.¿ added- b5-mso review d4-corrected model number. added- d6a/d6b.
28283 | 1220648-2024-09556 | 19056674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056674&pc=PYX | 04/01/2024 | 05/21/2024 |  | 04/06/2024 |  |  | 04/01/2024 | 04/06/2024 |  |  | 11/30/2025 |  | 04/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024434388 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 55 year-old male was implanted with an impella rp flex device for mechanical circulatory support. the impella device¿s introducer sheath was bleeding profusely through the hemostatic seal when the dilator was removed. the dilator was placed back in the sheath to stop the bleeding. bleeding continues as soon as the dilator was removed. there was no obvious visible damage to the sheath. a new introducer kit was opened to use a new sheath. no issues occurred with the replacement sheath. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported introducer damage has been completed. the 23fr small introducer and dilator were returned for investigation. no damage was noted on the dilator. upon investigation of the introducer, the clear hemostatic valve was found to be damaged on the blue ring side. as per the provided clinical details, bleeding was reported after removing the dilator from introducer sheath. the issue was resolved with sheath replacement. the cause of the introducer damage issue was a damaged hemostatic valve. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission.
28284 | 1220648-2024-09477 | 19049188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049188&pc=OZD | 05/31/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 05/31/2022 | 04/05/2024 |  |  | 11/30/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169701 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a 46-year-old patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the impella was alarming for suction when higher than p-1 in conjunction with extracorporeal membrane oxygenation (ecmo). the medical team continued to monitor, and continued impella support. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
28285 | 1220648-2024-09478 | 19049409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049409&pc=OZD | 06/01/2022 | 10/09/2025 |  | 04/05/2024 |  |  | 06/01/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195385 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Pressure Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/01/2022 |  |  |  |  | Us complainant reported the patient was on impella cp support due to nstemi (non-st-elevation myocardial infarction). while on support high purge flow, purge system leakage from between cassette and disc, vf (ventricular fibrillation), and uncontrollable oozing at access site was noted. patient received 1 shock with immediate return to nsr (normal sinus rhythm). it was also noted that leakage was due to improper cassette insertion. reportable due to high purge pressure, vf/vt/af/at, and access site bleeding. mdr and pmda report required. | The user facility reported a 59-year-old patient in non-st-elevation myocardial infarction was implanted with an impella cp for mechanical circulatory support. it was reported that after the impella cp was placed, the patient had an episode of ventricular fibrillation. the patient received one shock and returned to normal sinus rhythm. it was reported that during impella cp support, there were high purge pressure alarms, and a leak at the blue leur lock on the cassette. the purge cassette was changed and resolved the high purge pressure issues and it was noted the original cassette was improperly inserted. it was also reported the patient experienced uncontrolled oozing at the impella access site after transport. the cardiac team found the repo sheath short of the artery. the team decided to wean and explant the impella as the patient was hemodynamically stable and the right femoral artery was closed with a manta closure device. |  |  |  |  | Investigation of the reported issue has been completed. no device or data logs have been returned .due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
28286 | 1220648-2024-09479 | 19049469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049469&pc=OZD | 10/20/2023 | 09/26/2025 |  | 04/05/2024 |  |  | 10/20/2023 | 04/05/2024 |  |  | 08/31/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387190 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/14/2023 |  |  |  |  | The user facility reported a 41-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient developed a thrombus and was taken to the operating room for inferior vena cava clot removal. | Additional information indicated that, during patient support, the pump experienced placement signal not reliable alarm. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: d4-model number. added- d6a/d6b. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. no product was returned for an investigation. data log analysis confirmed multiple placement signal spikes from (b)(6). however, when looking at the pumps used before and after on the same console no pattern could be found. the root cause of the optical signal issue was not determined because no product was returned and not enough clinical information was given. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem b.5. updated with additional information. g.1. updated with correct mdr reporting contact email. h.6. medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number: 1220648-2024-09479.
28287 | 1220648-2024-09480 | 19049532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049532&pc=OZD | 03/07/2024 | 03/03/2025 |  | 04/05/2024 |  |  | 03/07/2024 | 04/05/2024 |  |  | 08/31/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397682 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 03/07/2024 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. while on impella support, the patient developed serious drainage from the access site and a small hematoma. manual pressure was applied, and the patient received one unit of red blood cells and platelets. the patient returned to the icu on approximately the same inopressors as prior to surgery. additionally noted, the patient's prognosis had been noted as poor. patient was on dialysis, and unable to wean extracorporeal membrane oxygenation (ecmo) the day prior. most likely route for the patient was noted as palliative care. approximately 17 days later, the patient expired. | It was further reported by abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient endured cardiac arrest asystole with a "unknown (undetermined)" relationship to the impella device used: "throughout the morning pt's hr was dropping at 12:14 pt was noted to be in a stole, and was pronounced dead at 12:15. no spontaneous respirations observed, no audible breath sounds or cardiac activity, and no pulse". the patient outcome is expired. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry the patient endured cardiac arrythmias- vtach with "possibly" related relationship to the impella. per loqi: overnight (b)(6), pt had several suction alarms as well as a 45 s run of vt which resolved spontaneously. impella was decreased to p7 with no further alarms. overnight on (b)(6) pt was noted to have vtach hr 150s ~2-3 min around 1:45am. converted back to sinus tachy to 100s on own. again at 2:45 am back in vtach again to 150s. dccv -> converted back to sinus. gave amio bolus and started amio drip. |  |  |  | The investigation for the reported access site bleed and thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and thrombocytopenia could not be determined. d4-updated model number. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added: b5 addl details about death failure mode. patient outcome of death did not change from initial report. corrections to the initial mfr report #1220648-2024-09480: b2-selected death and added date of death. f6/f should have been left blank. g1 mdr reporting contact email updated. h1-updated to cardiac arrest/death. h6 impact code added 1802. h6 component code 925 added. | B5 additional information received regarding event days before the patient expierenced cardiac arrest and expired .
28288 | 1220648-2024-09482 | 19049573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049573&pc=OZD | 03/28/2024 | 05/10/2024 |  | 04/05/2024 |  |  | 03/28/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025449669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after 5 hours of support the impella cp was explanted due to limb ischemia as evidence by absent pulses and skin mottling. the patient was escalated to impella 5.5 and was reported to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28289 | 1220648-2024-09481 | 19049575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049575&pc=OZD | 06/01/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/01/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199748 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient on impella cp had bleeding around the impella access site. the sheath had significant bleeding and the decision was made to explant the pump due to the excessive bleeding. the impella was weaned and explanted, manual pressure held. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28290 | 1220648-2024-09483 | 19049714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049714&pc=OZD | 06/02/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a concern of poor distal perfusion to right lower extremity was identified. an external femoral to femoral bypass was performed to resolve the limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28291 | 1220648-2024-09484 | 19049729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049729&pc=OZD | 03/28/2024 | 08/06/2024 |  | 04/05/2024 |  |  | 03/28/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025437486 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete | Other |  | Yes | Initial submission | Followup | Followup |  |  | 03/28/2024 |  |  |  |  | The user facility reported an 80 year old male with history of a-fib, coded at the cath lab and was placed on impella for cardiac support. the access site was seen to perforate with the use of the cook dilators and then the 14fr abiomed sheath. there was a hematoma there and later the team had to have the vascular injury repair consultation with vascular surgery. additionally the team explanted the cp and delivered an iabp for support. | [5 empty] | The investigation into the vascular damage - peripheral vessel and the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. it was reported that the bleeding occurred with the combination of the pump insertion and dilation. per the clinical information provided, the patient had significant pad, vessels very narrow due to blockages. therefore, the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related to small/ diseased vessels. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | Med watch report was received and has been attached. | The impella device has not been received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta. assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
28292 | 1220648-2024-09486 | 19049794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049794&pc=OZD | 11/06/2023 | 09/19/2025 |  | 04/05/2024 |  |  | 11/06/2023 | 04/05/2024 |  |  | 05/31/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 was removed and the graft bled out. the graft was folded and over sewn. after sewing was completed, the site was assessed for bleeding. there was no bleeding and the skin was sutured with glue over the top of the incision. | The complainant also reported the impella placement signal in aorta alarm was occurring anytime the patient coughed. team confirmed the correct placement under echocardiography. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the most likely cause of the major access site bleeding was use related. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added- h6 component code 925 d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for an investigation. data log analysis confirmed multiple impella position in aorta alarms eleven hours after beginning support. these were proceeded and followed by suction alarms. additionally, all 5s parameters were within normal bounds except for signal-to-noise ratio (snr), which remained below 1500 throughout the entire case. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b.5. updated to include optical signal issue description. h.6. updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09486. b.1. product problem updated. b.7. other relevant history, including preexisting medical conditions updated. d.10. concomitant medical products and therapy dates updated. g.1. reporting contact email and reporting contact fax number updated.
28293 | 1220648-2024-09487 | 19049811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049811&pc=OZD | 03/30/2024 | 05/16/2024 |  | 04/05/2024 |  |  | 03/30/2024 | 04/05/2024 |  |  | 03/31/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024328796 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 58-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp pump was inadvertently pulled out of the ventricle during patient support. as a result of the poor positioning, the decision was made to replace the device. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was most likely patient movement due to improper technique as the pump came out of position following the transfer. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28294 | 1220648-2024-09488 | 19049839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049839&pc=OZD | 03/30/2024 | 05/16/2024 |  | 04/05/2024 |  |  | 03/30/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025440574 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 46-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while in use with a patient an impella cp failed to flow above 1 liter and experienced consistent suction events despite attempts to reposition with echo. the pump was replaced with a new pump and upon removal, it appeared that a foreign object was present in the impella inflow. no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not returned for investigation. data logs were reviewed which showed suction alarms occurred. consistent suction events while in the unit despite attempts to reposition with echo. the cause of the low pump flow most likely was biomaterial ingestion based on the clinical detail that biomaterial was observed in the outflow after removal. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was discarded by the customer therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28295 | 1220648-2024-09489 | 19049880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049880&pc=OZD | 04/01/2024 | 05/14/2024 |  | 04/05/2024 |  |  | 04/01/2024 | 04/05/2024 |  |  | 11/30/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024428659 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and there were saturated flows after implanting. during ramp up to p-4 team noticed flows of 4.5l/min and left ventricle waveform was completely above aortic placement signal on automated impella controller (aic). all waveforms, including motor current pulsatile. staff discussed possible flow saturation and motor failure. the impella was replaced. the pump was saved for return for investigation. | [5 empty] | The investigation into the saturated flow issue has been completed since the original report was filed. the pump was returned for investigation. upon visual inspection, biomaterial was observed on the impeller blades. per the device investigation and the clinical details, the root cause of the saturated flow was due to biomaterial ingestion. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was received from the customer on (b)(6) 2024 and therefore,an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28296 | 1220648-2024-09491 | 19049962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049962&pc=PYX | 04/01/2024 | 05/16/2024 |  | 04/05/2024 |  |  | 04/01/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025458211 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a rp flex to support a patient admitted with multiple cardiac risk factors. the patient had known mitral valve disease (mvd) coronary artery disease¿(cad), mitral regurgitation¿(mr), tricuspid regurgitation¿(tr), cardiomyopathy and more. the patient was to have a 5.5 and rp placed to support for coronary artery bypass grafting¿¿(cabg). the retroperitoneal bleed (rp) alarmed and stopped within minutes (7 mins) of use and was explanted and replaced by a new rp flex.. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned for investigation. upon investigation of the returned pump, biomaterial was found around the impeller and purge gap. high pressure purge and purge system blocked alarms were reproduced. the root cause of the pump stop issue was determined to be biomaterial. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
28297 | 1220648-2024-09492 | 19049993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19049993&pc=OZD | 04/01/2024 | 05/15/2024 |  | 04/05/2024 |  |  | 04/01/2024 | 04/05/2024 |  |  | 11/30/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434212 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IABP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella implant, there were low pump flows. the flows went from 3.3 l/min to 0.0 l/min. echo did not show a left ventricle thrombus and the patient was on heparin. the impella was explanted after percutaneous coronary intervention. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. cannula kink observed near marker band in middle of the cannula. no other abnormalities found. the cause of the low pump flow was cannula kink due to improper technique. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28298 | 1220648-2024-09493 | 19050018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050018&pc=OZD | 04/01/2024 | 05/13/2024 |  | 04/05/2024 |  |  | 04/01/2024 | 04/05/2024 |  |  | 12/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292107 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that an impella 5.5 could not be implanted in a patient due to the patient's vascular size. instead, an impella cp was successfully implanted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of small vessel size and were not large enough to accommodate 5.5 pump and were successful in implanting impella cp. d4-updated model number. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28299 | 1220648-2024-09494 | 19050043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050043&pc=OZD | 06/02/2022 | 03/10/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 12/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023169893 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was swelling to the extremity which a ct scan revealed a large hematoma to the right upper extremity extending to the chest and a small hematoma to the right groin with a small amount of oozing. the patient became hemodynamically unstable and required the addition of neo and was administered 2 units of packed red blood cells (prbcs). after intubation from respiratory distress, the patient was stabilized and no active bleed was noted. the right upper extremity was stable. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28300 | 1220648-2024-09495 | 19050060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050060&pc=OZD | 04/02/2024 | 06/06/2024 |  | 04/05/2024 |  |  | 04/02/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2025462831 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella cp pump triggered an optical sensor system error alarm during patient support. the aic (automated impella controller) was swapped in an attempt to troubleshoot the issue, but the pump remained inactive after the swap. the pump was replaced and the second impella cp device was started up on the original aic with no noted issues. there was no patient harm. | [5 empty] | The investigation into the optical signal issue and the pump stop issue has been completed since the original report was filed. the device was returned for investigation. the reported issues were not reproducible. no problem was found as the root cause of the optical signal issue based on provided data. the root cause of the pump not detected issue was not determined as relevant logs were not provided and there was limited clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
28301 | 1220648-2024-09497 | 19050074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050074&pc=OZD | 06/04/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/04/2022 | 04/05/2024 |  |  | 03/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Difficult to Remove | [6 empty] | Hematoma | Ischemia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2022 |  |  |  |  | The user facility reported an 80-year-old female was implanted with impella cp for mechanical circulatory support. it was reported that the patient's pedal pulses were difficult to obtain. actions taken to resolve the ischemia is unknown. additionally, it was also reported that the patient had a hematoma. the hematoma was expressed. heparin was reduced and then stopped due to bleeding from the impella site and gi bleeding. the patient received two packed red blood cells, related to the impella implantation. the purge solution was changed to sodium bicarbonate. during explant of the impella cp, the perclose failed and there was damage to the artery. the patient was taken to the operating room for surgical repair. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28302 | 1220648-2024-09496 | 19050076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050076&pc=OZD | 06/02/2022 | 03/10/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023195018 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was some oozing at the access site; however, the patient was given blood for treatment and the oozing was controlled. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28303 | 1220648-2024-09498 | 19050079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050079&pc=OZD | 06/02/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 12/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173691 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that required the patient to receive multiple units of packed red blood cells, platelets, and cryoprecipitate. groin bleeding was addressed and resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | . this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28304 | 1220648-2024-09499 | 19050188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050188&pc=OZD | 06/02/2022 | 03/07/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 05/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022086449 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the patient experienced hematoma which was treated with manual pressure and pressure dressing. 1 unit of packed red blood cells(prbcs) was issued and the device was repositioned. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28305 | 1220648-2024-09500 | 19050223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050223&pc=OZD | 05/31/2022 | 04/05/2024 |  | 04/05/2024 |  |  | 05/31/2022 | 04/05/2024 |  |  | 10/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158676 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, significant oozing noted from the access site. blood products were given. the patient also had ventricular fibrillation and was shocked. also, there was a low pump flow alarm. the impella cp was weaned and removed due to risk of bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28306 | 1220648-2024-09501 | 19050397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050397&pc=OZD | 06/04/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/04/2022 | 04/05/2024 |  |  | 03/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203652 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2022 |  |  |  |  | The user facility reported an 83-year-old male was implanted with impella cp for mechanical circulatory support. it was reported that upon device explanation it was noted that there was only about 5 ml of blood and no clot. there was no bleeding from groin. the patient was taken to the cath lab for angiography and a common femoral full clot was found. the clot was retrieved and flow was restored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28307 | 1220648-2024-09502 | 19050426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050426&pc=OZD | 06/02/2022 | 03/10/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023182547 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was hemolyzing therefore, the pump was removed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28308 | 1220648-2024-09503 | 19050488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050488&pc=OZD | 06/02/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 12/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174588 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a 64-year-old female on impella cp for mechanical circulatory support. it was reported that the patient had a thrombus. this was treated by removing the pump. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
28309 | 1220648-2024-09504 | 19050544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050544&pc=OZD | 06/02/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/02/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023195424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2022 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was bleeding around the peel away sheath which was the same insertion site as the intra-aortic balloon pump ( iabp) and is thought to be the cause. heparin was withheld and the site was redressed with downward pressure held at the site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28310 | 1220648-2024-09505 | 19050552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050552&pc=OZD | 06/03/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/03/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2022 |  |  |  |  | The user facility reported a 52-year-old male on impella cp support due to st elevation myocardial infarction (stemi). while on support, the nurse noted a development of hematuria. the nurse suspected positioning may be an issue. no additional information was provided. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28311 | 1220648-2024-09506 | 19050583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050583&pc=OZD | 05/31/2022 | 05/21/2025 |  | 04/05/2024 |  |  | 05/31/2022 | 04/05/2024 |  |  | 12/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172216 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a minor aorta dissection during guide manipulation and the decision was made to continue with the procedure, which was considered successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28312 | 1220648-2024-09507 | 19050594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050594&pc=OZD | 06/03/2022 | 04/05/2024 |  | 04/05/2024 |  |  | 06/03/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194910 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 72-year-old female patient for mechanical circulatory support. it was reported under echocardiogram, there was insufficient space in the left ventricle for proper positioning of impella. as a result, the patient's experienced hemolysis due to this positioning challenge, therefore, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28313 | 1220648-2024-09508 | 19050663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050663&pc=OZD | 03/23/2024 | 05/13/2024 |  | 04/05/2024 |  |  | 03/23/2024 | 04/05/2024 |  |  | 11/30/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2024 |  |  |  |  | The user facility reported a 62-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed hemolysis. repositioning of the pump, heparin was added to the purge and p level was decreased as treatment. however, the patient continued to have hemolyzed labs and dark urine. the clinical representative stated there were no indications of clots, but there is a possibility of a micro clot causing hemolysis. the pump was explanted and escalated to venoarterial extracorporeal membrane oxygenation and intra-aortic balloon pump. the patient was reported to be stable. | [5 empty] | The investigation for the reported hemolysis and thrombus has been completed. according to the clinical, on the first day of impella cp support, the pump was inserted with some difficulty and the team struggled to optimize the position of the pump. upon arrival to the icu the following day, imaging showed that the pump inflow in mid ventricle and without interaction with any lv structures. later on the same day the patient began to develop signs of hemolysis with dark red urine. the decision was then made to escalate the patient to va ecmo and explant the pump. after the impella device was removed the patient was diagnosed with a platelet clumping hematological condition. no product was returned for a visual inspection. data log analysis confirmed positioning issues with suction alarms at the very beginning of support. position issue resolved immediately. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. the root cause of the thrombus could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
28314 | 1220648-2024-09509 | 19050717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050717&pc=OZD | 06/03/2022 | 05/21/2025 |  | 04/05/2024 |  |  | 06/03/2022 | 04/05/2024 |  |  | 03/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206108 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 69-year-old male patient for mechanical circulatory support. it was reported that during sheath and wire insertion a descending aortic dissection was noted. it was managed with a 25 cm impella sheath. the impella cp was weaned and removed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28315 | 1220648-2024-09510 | 19050726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050726&pc=OZD | 06/01/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/01/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197430 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Ventricular Fibrillation | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient required defibrillation performed three times for ventricular tachycardia and ventricular fibrillation. the user facility noted that the patient's impella leg was mottled and the patient developed compartment syndrome. a fasciotomy performed to the left thigh and left calf to treat the limb complications. the patient's plasma free hemoglobin result was >150. and the urine was tea colored. as a result of the hemolysis, the patient started continuous renal replacement therapy. the patient also experienced bleeding issues from multiple sites. the patient was administered protamine and multiple blood products were transfused to treat the patient. additionally, it was reported the patient coughed, and the impella alarmed for impella in aorta. the decision was made to explant the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28316 | 1220648-2024-09511 | 19050821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050821&pc=OZD | 05/31/2022 | 05/20/2025 |  | 04/05/2024 |  |  | 05/31/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196916 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. while positioning the impella, the patient went back into ventricular fibrillation and the pump pigtail pulled back into position under the papillary muscle. the decision was made to leave the pump without repositioning due to multiple unsuccessful attempts. the positioning issue led to hemolysis which was mitigated by reducing the flow of the pump. additionally, the patient had a run of ventricular tachycardia that resolved with medication and the patient converted himself. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28317 | 1220648-2024-09512 | 19050915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19050915&pc=OZD | 03/28/2024 | 05/14/2024 |  | 04/05/2024 |  |  | 03/28/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025451970 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 59-year-old female with cardiogenic shock and acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at insertion but was stable at assessment. the physician later removed the impella device due to hematoma and blood loss. the doctor believe the intensive care unit staff moved the patient and triggered more bleeding. it was also reported that the patient had an ischemic leg. the patient had blood products transfused. | [5 empty] | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
28318 | 1220648-2024-09513 | 19051039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051039&pc=OZD | 03/28/2024 | 05/14/2024 |  | 04/05/2024 |  |  | 03/28/2024 | 04/05/2024 |  |  | 07/31/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442231 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 72-year-old female with cardiogenic shock and acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had severe bleeding from their groin. a femoral stop was placed, and the leg became molted. the medical staff removed the femoral stop. the medical staff utilized manual pressure, and the old suture was removed and replaced. the bleeding subsided. two units of blood products were provided to the patient. the patient remains on support. | [5 empty] | The investigation for the access site bleed and the limb ischemia has been completed. the impella device was not returned for evaluation. clinical details were sufficient to determine the root cause of the access site bleed. the root cause of the access site bleeding issue was most likely use related, bleeding subsided after old suture was removed and replaced as per the clinical description. the root cause of the limb ischemia could not be determined without additional clinical details. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28319 | 1220648-2024-09514 | 19051174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051174&pc=OZD | 03/29/2024 | 05/15/2024 |  | 04/05/2024 |  |  | 03/29/2024 | 04/05/2024 |  |  | 06/30/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024360347 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 84-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported there was difficulty implanting an impella 5.5 due to calcium seen in the patient's subclavian artery. the patient was instead successfully implanted with an impella cp. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition as the 5.5 was unable to pass through the calcium in the subclavian. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28320 | 1220648-2024-09515 | 19051398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051398&pc=OZD | 06/06/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/06/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198861 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon explanation, it was reported that the device was explanted via surgical cut down/wound closure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28321 | 1220648-2024-09516 | 19051446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051446&pc=OZD | 03/30/2024 | 06/05/2024 |  | 04/05/2024 |  |  | 03/30/2024 | 04/05/2024 |  |  | 06/30/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236901 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient's outcome was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of device with outcome. | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site bleeding. manual pressure was held. the physician added another stitch and heparin was discontinued. the bleeding resolved with no oozing or hematoma after the additional stitch. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, bleeding resolved with no oozing or hematoma after additional stitch as per the clinical description. the instructions for use referenced in the initial report is for the impella cp with smart assist system in sections: section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).", section: general patient care considerations, and section: entering cardiac output. b5-mso review d4-updated model number added- d6a/d6b | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28322 | 1220648-2024-09517 | 19051627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051627&pc=PYX | 03/30/2024 | 05/16/2024 |  | 04/05/2024 |  |  | 03/30/2024 | 04/05/2024 |  |  | 12/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA PUMP 450 |  |  | ABIOMED, INC. |  |  | US |  | 2024320040 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 79-year-old female patient of unknown ethnicity in st-segment elevation myocardial infarction was to be implanted with an impella rp device for mechanical circulatory support. it was reported that despite using a buddy wire to advance into the pulmonary artery, the pump could not advance. posterior atrioventricular was passed. the procedure was aborted. tortuous femoral vein was cited by a fellow along with using the a more difficult left femoral vein access. no patient harm was reported. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because of the tortuous anatomy of the femoral vein and were unable to advance the pump and the implantation was aborted. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28323 | 1220648-2024-09518 | 19051838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051838&pc=OZD | 03/30/2024 | 05/17/2024 |  | 04/05/2024 |  |  | 03/31/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient began to complain of loss of vision followed by disorientation and lethargy during impella cp support. it was noted that a left ventricle¿(lv) thrombus had been present at the time of impella implant. despite a scheduled computed tomography¿(ct) scan coming back as negative for hemorrhaging, the decision was made to immediately explant the pump following reports of the noted symptoms. biomaterial was found on the device following explant. | [5 empty] | The investigation into the stroke/cva has been completed since the original report was submitted. the device was not returned for investigation. it was reported that echo the following day revealed known small left ventricle thrombus. as relevant clinical information was not provided, the root cause of the stroke/cva and the thrombus could not be determined. impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | Serial number, lot number, expiration date, unique identifier (udi) # and device manufacture date are not known. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28324 | 1220648-2024-09519 | 19051868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051868&pc=OZD | 06/06/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/06/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a thrombus that required a thrombectomy to be performed. additionally, the patient had multiple episodes of ventricular tachycardia/ ventricular fibrillation, a hematoma and oozing at the end of the access procedure. manual pressure was applied, successfully managing the impella complication. no blood products or surgical interventions were required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28325 | 1220648-2024-09520 | 19051901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051901&pc=OZD | 06/10/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/10/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183590 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a 71-year-old male was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, there was access site bleeding noted. heparin was withheld due to bleeding resulting in high purge pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28326 | 1220648-2024-09521 | 19051942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051942&pc=OZD | 06/12/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/12/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186856 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 74-year-old female patient was on impella cp support due to non-st elevated myocardial infarction. while on support, oozing around repo sheath was noted. dressing was changed and patient was given blood due to drop in hemoglobin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28327 | 1220648-2024-09522 | 19051977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19051977&pc=OZD | 04/01/2024 | 06/07/2024 |  | 04/05/2024 |  |  | 04/02/2024 | 04/05/2024 |  |  | 11/30/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2025448045 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia. surgical cut down was performed to remove the sheath and subsequent thrombectomy. | [5 empty] | The investigation into the thrombus and the limb ischemia - reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and the limb ischemia was not determined. in accordance with updated procedures, sections d6 and g3 have been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
28328 | 1220648-2024-09523 | 19052037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19052037&pc=OZD | 06/12/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/12/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202108 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 59-year-old female patient was placed onto impella cp support due to percutaneous coronary intervention for lateral wall st elevation. while on support, the pump experienced suction alarms that were resolved by giving patient fluids, 1 unit of blood and decreasing pump flow. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28329 | 1220648-2024-09524 | 19052112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19052112&pc=OZD | 06/07/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 06/07/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186653 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a 55-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, patient developed left ventricle (lv) thrombus. patient was kept on support and continued to be monitored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
28330 | 1220648-2024-09525 | 19052421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19052421&pc=OZD | 08/31/2022 | 05/27/2025 |  | 04/05/2024 |  |  | 08/31/2022 | 04/05/2024 |  |  | 05/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023228732 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the patient developed hematuria and the pump was pulled back under echo. due to elevated partial thromboplastin time (ptt), heparin was withheld. eventually, the device was successfully weaned. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28331 | 1220648-2024-09526 | 19052448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19052448&pc=OZD | 04/02/2024 | 04/22/2024 |  | 04/05/2024 |  |  | 04/02/2024 | 04/05/2024 |  |  |  |  | 04/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1368535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | A united states based field team member was conducting training when the automated impella controller (aic) alarmed with a system self-check error immediately upon boot up. | [5 empty] | The investigation for the reported console (aic) self check error has been completed. the console was returned for evaluation. upon visual inspection of the pbm assembly it¿s been discovered that leaking electrolyte from supercaps c69, c70 caused corrosion of the traces responsible for pbm1 communication. this type of event has been observed before and addressed by fcn 71, however this pbm serial number lays outside of the previously identified range of affected components. during console inspection it¿s been observed that purge door release button was loose (unrelated to complaint). console logs could not be immediately obtained from the device due to file corruption on cf card. after file system repair, data logs were extracted and verified to have been consistent with complaint: multiple error messages referring to sau_powermod_1 (sau_ser10) communication failure were logged. aic self-reboot did not resolve the issue and pbm1 communication loss was permanent, preventing proper console boot-up. the cause of the aic issue was contamination. added- d9 device return date d4-updated model number. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28332 | 1220648-2024-09527 | 19052468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19052468&pc=OZD | 09/16/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 09/16/2022 | 04/05/2024 |  |  | 01/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187403 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The user facility reported a 67-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while on the impella cp the patient experienced ventricular tachycardia. during the exchange, the wire went into papillary and caused sustained ventricular tachycardia. the catheter and wire reinserted without issue and wire pulled back upon impella pigtail entry into lv to avoid papillary muscle. pump started without issue. the pump was weaned and explanted, distal runoff revealed an occluded left leg. left femoral artery stented, clot visualized and multiple passes with penumbra restored distal flow and hemostasis of the impella access. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28333 | 1220648-2024-09529 | 19053628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19053628&pc=OZD | 05/20/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 05/20/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient had no dopplerable pulse in their right lower extremity following impella cp implant. it was noted via angiogram prior to insertion that the right femoral artery/iliac had shown some narrowing due to existing scar tissue. upon observation of the absent pulse, treatment options were discussed, and the decision was ultimately made to remove the pump. following explant, an intra-aortic balloon pump was scheduled for implant for ongoing patient support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28334 | 1220648-2024-09530 | 19053952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19053952&pc=OZD | 04/02/2024 | 05/21/2024 |  | 04/05/2024 |  |  | 04/02/2024 | 04/05/2024 |  |  | 11/30/2025 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024432599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the user facility reported that the impella cp sheath was occlusive in the artery. the physician performed an external fem-fem bypass to resume distal flow. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the limb ischemia ¿ reversible has been completed since the original report was submitted. the pump was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28335 | 1220648-2024-09531 | 19056303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056303&pc=OZD | 10/05/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 10/05/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192069 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported an 86-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patent developed lower extremity ischemia. issue was resolved by insertion of antegrade sheath and solution. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28336 | 1220648-2024-09532 | 19056342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056342&pc=OZD | 11/17/2022 | 03/07/2025 |  | 04/05/2024 |  |  | 11/17/2022 | 04/05/2024 |  |  | 06/30/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023249100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a patient needing an impella cp for mechanical circulatory support. during aortic valvuloplasty the patient's blood pressure dropped once the trans-catheter aortic valve replacement (tavr) valve deployed. the attempt was made to utilize the imeplla cp for support; however, there were issues with the femoral access. due to the patient's vascular anatomy the implant attempt was aborted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28337 | 1220648-2024-09533 | 19056369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056369&pc=OZD | 11/23/2022 | 04/03/2025 |  | 04/05/2024 |  |  | 11/23/2022 | 04/05/2024 |  |  | 02/29/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a (b)(6) female on impella cp for mechanical circulatory support. it was reported there was oozing at the sinoatrial (sa) site. the repositioning sheath was advanced into the peel away sheath to control the bleeding. the patient was weaned from the impella cp without any hemodynamic compromise and explanted. perclose x3 were unsuccessful gaining hemostasis. the physician performed a femoral cutdown and repaired the femoral artery without any difficulty. the patient tolerated procedure and cutdown without any complications. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28338 | 1220648-2024-09534 | 19056378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056378&pc=OZD | 11/23/2022 | 04/05/2024 |  | 04/05/2024 |  |  | 11/23/2022 | 04/05/2024 |  |  | 08/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269796 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 85-year-old male patient for mechanical circulatory support. it was reported that while on impella cp support the patient developed oozing at the access site due to the inability to adequately manage pre-closure sutures placed during insertion of the impella cp. upon removal of the device, an extra perclose was utilized, but it became lodged in the femoral artery and was managed by vascular surgery cut down. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28339 | 1220648-2024-09535 | 19056398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056398&pc=OZD | 12/09/2022 | 04/05/2024 |  | 04/05/2024 |  |  | 12/09/2022 | 04/05/2024 |  |  | 12/31/2023 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179296 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported a 72-year-old male on impella cp for mechanical circulatory support. it was reported the patient was on impella cp and the physician was unable to place device due to patient¿s vasculature. it was decided to remove the impella cp and implant a balloon expanding transcatheter aortic valve replacement (tavr) within the surgical valve. the impella cp was removed without any complications. a balloon expandable tavr was placed across the surgical valve and deployed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28340 | 1220648-2024-09536 | 19056399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19056399&pc=OZD | 12/09/2022 | 05/13/2025 |  | 04/05/2024 |  |  | 12/09/2022 | 04/05/2024 |  |  | 08/31/2024 |  | 04/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261421 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported a 67-year-old male on impella cp for mechanical circulatory support. it was reported that while the patient was on impella cp support, the impella cp pulled back across aorta and ao and lv disassociation and was positioned deep after the 3rd coding episode. the impella cp was later swapped out for the 5.5 device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28341 | 1220648-2024-09375 | 19039886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19039886&pc=OZD | 06/24/2022 | 10/06/2025 |  | 04/04/2024 |  |  | 06/24/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199495 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The complainant reported that the patient encountered limb ischemia and that the pump stopped. the limb ischemia was treated by placing an external bypass. the pump stop was treated by explanting the pump and replacing it with a new one. | The user facility reported a 23-year-old female on impella cp for mechanical circulatory support. it was reported that the pump stopped. this was treated by explanting the pump and replacing it with a new one. no patient harm reported. |  |  |  |  | This report is being filed as part of a retrospective review of historical records. h6: added limb ischemia. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
28342 | 1220648-2024-09376 | 19039942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19039942&pc=OZD | 06/27/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/27/2022 | 04/04/2024 |  |  | 03/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208063 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2022 |  |  |  |  | The user facility reported a 68-year-old male patient implanted with an impella cp device for mechanical circulatory support due to severe left ventricular dysfunction. it was reported while on impella cp support the patient experienced bloody sputum that was resolved by discontinuing systemic heparin. it was reported that the patient experienced septic shock while on impella cp support. no information provided for any intervention provided for sepsis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28343 | 1220648-2024-09377 | 19039972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19039972&pc=OZD | 06/15/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/15/2022 | 04/04/2024 |  |  | 12/31/2023 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164234 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient on impella cp had thrombocytopenia. thrombocytopenia was treated, but no information was provided on treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The root cause of the thrombocytopenia could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
28344 | 19040007 | 19040007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040007&pc=OZD | 03/13/2024 | 03/13/2024 | 03/13/2024 | 04/04/2024 | 03/13/2024 | 04/04/2024 |  | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 1 MO |  | Male |  |  |  | Malfunction | No | Yes | [7 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Device was being used for an impella-assisted pci (percutaneous coronary intervention). when the providers placed the impella catheter into the aic (automated impella controller), it did not recognize the catheter. the abiomed rep was in the room to assist with troubleshooting. decision was made to exchange the catheter. after catheter exchange, the aic did not recognize this one either. call was made to perfusion for assistance. perfusion then called ecmo coordinator for new aic console. perfusion obtained the aic from the cath lab room and the ecmo coordinator met perfusion at the red line doors with new aic to take to cath lab. ecmo coordinator confirmed with the abiomed rep that this new aic that was delivered worked and it did. this caused the patient to receive two catheters and a delay in their procedure.
28345 | 1220648-2024-09378 | 19040062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040062&pc=OZD | 03/22/2024 | 05/08/2024 |  | 04/04/2024 |  |  | 03/22/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025449407 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 72 year-old male with st-segment elevation myocardial infarction and mitral valve disease was implanted with an impella cp device for mechanical circulatory support. during implantation, the pump position was lost, and the team attempted to reimplant the pump to the left ventricle but was unsuccessful. the pump was then explanted and reprimed in a bowl. while repriming, the wire was still retained in the pump and the pump stopped. the impella device was replaced due to concerns of damage. the stop did not occur in the patient, and no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the pump was returned for analysis. the cp pump and the guidewire were visually inspected. a very slight kink near outflow cage observed. no issue or damaged was observed in the returned guidewire and pump. the cause of the delivery issue was use related due to improper technique as they lost lv wire access while advancing the pump. in accordance with updated procedures, section d6 has been revised from the initial submission.
28346 | 1220648-2024-09379 | 19040185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040185&pc=PYX | 03/25/2024 | 05/23/2024 |  | 04/04/2024 |  |  | 03/25/2024 | 04/04/2024 |  |  | 07/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023229830 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 61 year-old male with cardiogenic shock and st-segment elevation myocardial infarction was implanted with an impella rp device for mechanical circulatory support. the impella rp device stopped when the sheath was attempted to be removed. the sheath was then left in place and the impella device was restarted. impella rp device was turned to p5, and the patient was being adequately supported. the patient is stable. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the pump and data logs were returned and investigated. the pump was soaked to get the dried blood off and after that checked with borescope and it was found that the pump had heavy biomaterial wrapped around the impeller blades and the outflow cage. there was also a small catheter kink observed around the 70 centimeters mark. no other physical abnormalities were found. the data logs were investigated and it was seen that there were motor current spikes seen when the pump was implanted suggesting sudden biomaterial ingestion. this was followed by suction and impella stop alarms. the pump was then troubleshooted and the purge pressure was seen to be getting better with change in the p-level and decrease in the motor current spikes. the root cause of the temporary pump stop issue was ingested biomaterial found wrapped around the impeller and outflow cage. instructions for use for the related event are as follows: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 (device not returned) was inadvertently omitted from the previously submitted report and has been added to this report. the product has been returned since the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
28347 | 1220648-2024-09380 | 19040216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040216&pc=OZD | 03/08/2024 | 04/18/2024 |  | 04/04/2024 |  |  | 03/08/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443065 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male was implanted with an impella device as elective placement. however, while closing a graft placed for direct access following impella device removal, the aorta tore requiring a femoral cannulation bypass. a patch was placed, and the chest was left open to manage the bleeding. the patient's outcome was noted as critical. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ cp and 5.5 ifu: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
28348 | 1220648-2024-09382 | 19040283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040283&pc=OZD | 06/19/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/19/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a 81-year-old patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. heparin was withheld and the patient was transfused with multiple units of blood products. the patient also experienced atrial fibrillation while on support which self-corrected. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28349 | 1220648-2024-09383 | 19040289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040289&pc=OZD | 03/26/2024 | 05/02/2024 |  | 04/04/2024 |  |  | 03/26/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1373062 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Alarm System | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 70 year-old female with high risk of percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. a purge pressure high alarm occurred. the staff troubleshot the alarm and replaced the purge cassette. the new purge disc was not detected even though it was securely in place. after multiple attempts, they switched to a new automated impella controller (aic). after the aic switch, the purge disc was detected but the patient continued to alarm for purge pressure high. the decision to remove the impella device was then made. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. to conform to updated procedures, section d2 was revised with updated information.
28350 | 1220648-2024-09384 | 19040295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040295&pc=PYX | 03/22/2024 | 05/10/2024 |  | 04/04/2024 |  |  | 03/22/2024 | 04/04/2024 |  |  | 06/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024360977 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. while on impella rp support, the patient developed access site bleeding. an additional mattress was placed and pressure was held to resolve the bleeding. additionally, the patient was transfused with three units of packed red blood cells. the patient was successfully weaned from the impella rp and survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related, additional mattress sutures were placed to prevent the bleeding as per the clinical description.
28351 | 1220648-2024-09386 | 19040370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040370&pc=OZD | 03/26/2024 | 05/03/2024 |  | 04/04/2024 |  |  | 03/26/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMPTEMPORARY NON-ROLLER TYPE |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455195 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported an 72 year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support after ventricular tachycardia arrest. the patient had a gastrointestinal bleed. the medical staff stopped administering heparin and started administering protonix. the patient¿s hemoglobin continued to drop and one unit of packed red blood cells was given to the patient. the patient continued support. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the vascular damage gi bleed and its relation to impella were not determined since product was not returned and there is a lack of clinical details. the instructions for use for impella cp with smart assist system in section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- d6a/d6b d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28352 | 1220648-2024-09385 | 19040382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040382&pc=OZD | 03/24/2024 | 05/23/2024 |  | 04/04/2024 |  |  | 03/24/2024 | 04/04/2024 |  |  | 06/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431850 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Other | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2024 |  |  |  |  | The user facility reported a 54 year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support due to the patient coding twice and decompensating. while on support, the patient went in and out of atrial fibrillation. additionally, the patient was heparin induced thrombocytopenia (hit) positive, and heparin was discontinued towards the end of support. also, a hand injection revealed a thrombus around the pump. the pump was pulled across the aortic valve, powered down, and removed. a thrombus was noted post-explant. the patient remains on support. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the thrombocytopenia could not be determined due to insufficient clinical details. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and vt/vf was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.1 revised as previously entered incorrectly. h.6 code 925 was previously entered incorrectly. code 4114 was inadvertently omitted from the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
28353 | 1220648-2024-09393 | 19040591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040591&pc=OZD | 03/26/2024 | 05/20/2024 |  | 04/04/2024 |  |  | 03/26/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Intraoperative Cardiac Valve Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 69 year-old male with cardiogenic shock and acute myocardial infarction was implanted with an impella device for mechanical circulatory support. the patient also underwent a mitral valve replacement. on the fourth day of support, the medical staff noticed a thrombus in the patient¿s aorta, surrounding the impella device¿s cannula and catheter. the thrombus was resolved with a penumbra and inari thrombectomy. additionally, there was bleeding that prompted a washout and a possible fasciotomy. the impella device was explanted and replaced with another, escalated impella device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve. if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow. if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4852 and 4775 were reported incorrectly on the previously submitted report. code 4114 was added as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
28354 | 1220648-2024-09394 | 19040616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040616&pc=OZD | 03/27/2024 | 05/20/2024 |  | 04/04/2024 |  |  | 03/27/2024 | 04/04/2024 |  |  | 06/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428516 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 86-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that following impella cp implant, the patient developed a small hematoma and loss pulse in their right foot. as a result of the noted conditions, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and not much information was provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
28355 | 1220648-2024-09395 | 19040637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040637&pc=OZD | 03/27/2024 | 05/09/2024 |  | 04/04/2024 |  |  | 03/27/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025443050 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pain | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump support ongoing for 2 days for a patient admitted in acute hf. the patient was seen to have mvd (mitral valve disease) cad (coronary artery disease) and a cabg (coronary artery bypass grafting) was planned. the team found after the 2 days there was pain in the limb accessed by the 5.5 and the pump was then seen to be thrombosing. the thrombus was venous, a dvt (deep vein thrombosis) rather than arterial clot. the 5.5 was explanted and replaced by a cp pump via the femoral artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus and the pain at insertion site issues has been completed since the original report was submitted. the device was not returned for investigation. limb ischemia also occurred. as the necessary information was not provided and the device was not returned, the root cause of the thrombus and the limb ischemia was not determined. the cause of the pain at insertion site was not determined due to insufficient clinical details. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28356 | 1220648-2024-09396 | 19040693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040693&pc=OZD | 03/25/2024 | 05/21/2024 |  | 04/04/2024 |  |  | 03/27/2024 | 04/04/2024 |  |  | 09/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024405874 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient admitted in with lengthy cardiac history and cardiomyopathy. the support allowed for intervention/pci (percutaneous coronary intervention) of the coronaries. after just under 4 days the pump was explanted. later (2 days post explant) the field team was notified of a need for thrombectomy to the limb after removal there was femoral clot observed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. the correct code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but it was discarded.
28357 | 1220648-2024-09397 | 19040730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040730&pc=OZD | 06/04/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/04/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the clinical staff was unable to locate distal pulses on the left leg of the patient. the patient was removed from pressers except for 10mcg/min of levo. echo was performed at bedside, and the impella positioned well. lower limb extremity perfusion was still an issue. the patient remained on support and was successfully weaned off and the impella was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28358 | 1220648-2024-09398 | 19040766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040766&pc=PYX | 03/27/2024 | 05/23/2024 |  | 04/04/2024 |  |  | 03/27/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025442075 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The complainant reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support. the impella rp flex was placed but due to low flows from the pump, the team explanted the pump after just under 14 hours. there was no patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella pump was returned and was visually inspected. there was biomaterial ingestion noted around the impeller. no broken blades nor cannula kink was observed. the cause of the low pump flow issue was biomaterial ingestion. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report.. the product has been returned since then and the investigation has been completed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28359 | 1220648-2024-09399 | 19040791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040791&pc=OZD | 03/29/2024 | 05/10/2024 |  | 04/04/2024 |  |  | 03/29/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1666954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Output Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old who was a heart transplant rejection was implanted with an impella device. the automated impella controller (aic) had an unexpected controller shut down, controller failure (red alarm). waveforms went flat; however, the impella did not turn off. the team switched to another controller without incident. there was no report of adverse effect to the patient. | [5 empty] | The automated impella controller device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller failure (red alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. unable to reproduce the reported failure mode by running the console with the pump. the root cause for the unexpected shutdown was not determined due to the inability to reproduce.
28360 | 1220648-2024-09400 | 19040812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040812&pc=OZD | 12/04/2023 | 09/26/2025 |  | 04/04/2024 |  |  | 12/04/2023 | 04/04/2024 |  |  | 09/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024400558 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on impella 5.5 support, the patient had gastrointestinal (gi) bleed, therefore heparin had been discontinued. no additional information was provided for the treatment of gi. | Additional information indicated that, during patient support, the pump placement signal stopped working. an echocardiogram was performed to verify correct position. support was continued, and the patient was successfully weaned off the device. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09400. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. the pump was not returned for investigation. data logs were only provided up until 12/01, before patient transfer. no optical signal or placement signal abnormalities were reported in the given logs until 12/01. data logs were not provided for the time during the psnr event. the placement signal was lost two days prior to explant. the pump was kept in place, and no other abnormalities were reported. the root cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-09400.
28361 | 1220648-2024-09401 | 19040871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040871&pc=OZD | 03/30/2024 | 05/13/2024 |  | 04/04/2024 |  |  | 03/30/2024 | 04/04/2024 |  |  | 11/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434339 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation from an impella cp device. the 5.5 pump had a failure to deliver. the 5.5 became kinked and was replaced by a new impella 5.5 pump. no harm or injury from delivery failure through innominate artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was returned for evaluation. upon visual evaluation, the catheter had a kink in it. clinical details also states that the impella pump was removed and the shaft of the impella was bent right before the motor unit. data logs were not relevant to the reported. therefore, the root cause of the delivery issue was mostly patient condition since the impella could not get past the innominate artery into the aorta without success. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission.
28362 | 1220648-2024-09402 | 19040892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040892&pc=OZD | 03/28/2024 | 05/16/2024 |  | 04/04/2024 |  |  | 03/28/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1566931 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during a routine check, the automated impella controller (aic) was turned on and immediately, an alarmed control failure occurred. there was no patient involvement. | [5 empty] | The investigation into the aic ¿ other issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the back housing i/o panel cover and power chord was missing. there was residue build up within the purge unit cabin. the push button wire was broken during the testing. the firmware was found to be corrupted. the root cause for the controller failure was a purge unit driver with corrupted firmware. d9. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28363 | 1220648-2024-09403 | 19040926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040926&pc=OZD | 03/30/2024 | 05/15/2024 |  | 04/04/2024 |  |  | 03/30/2024 | 04/04/2024 |  |  | 10/31/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412180 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. shortly after the implant, there was a high purge alarm. low purge flows were noted. the purge cassette was replaced with no avail. the automated impella controller (aic) was replaced with no resolution. eventually, placement signal showed signs of pump saturation. the impella cp was replaced. | [5 empty] | The investigation for the reported saturated flow has been completed. only the pump was returned for investigation. a significant amount of biomaterial was observed on the impeller and in the purge gap. no physical abnormalities were noted. a high purge pressure alarm appeared during the purge testing. additionally, the pump was unable to start, showing a high motor current alarm during the test run. data logs show a gradual 12 minute rise in purge pressure followed by saturated pump flow after pump was inserted into body. the cause of the high purge pressure issue was biomaterial based on returned product and data log review. the saturation of pump flow was caused by the ingress of biomaterial. added-d9 device return date, h6 med dev prob code 1491 in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28364 | 1220648-2024-09404 | 19040960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19040960&pc=OZD | 06/21/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/21/2022 | 04/04/2024 |  |  | 03/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022107857 |  | 1000115 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The user facility reported a 31-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at the impella access site. the site was redressed with a pressure dressing and the patient was transfused with blood products. it was reported on the ninth day of support with the impella cp, the patient spiked a fever. the medical team decided to wean and explant the device as a result of the fever. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28365 | 1220648-2024-09405 | 19041056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041056&pc=OZD | 03/31/2024 | 05/02/2024 |  | 04/04/2024 |  |  | 03/31/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1156149 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 52-year-old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported that the automated impella controller (aic) displayed a purge disc not detected alarm following a purge cassette exchange. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. logs were reviewed and purge disc not detected alarm issues were confirmed. multiple instances of purge disc not detected were found in the logs on the reported occurrence date. the console was tested. the issue with detecting the purge disc was confirmed, and the purge flag was discovered to be broken. the root cause of this issue was a broken purge disc flag. d.2 and d.6a revised from the initial submission in accordance with updated procedures.d.4 revised as previously entered incorrectly.
28366 | 1220648-2024-09406 | 19041082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041082&pc=OZD | 07/05/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 07/05/2022 | 04/04/2024 |  |  | 04/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214042 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 48-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced access site oozing. manual pressure was held to help stop it and angle matching dressing confirmed. the patient was brought to the operating room for washout. a large clot and one liter of blood were removed. the patient received two units of packed red blood cells and one unit platelets/fresh frozen plasma. the patient did not tolerate weaning extracorporeal membrane oxygenation to impella. the patient had a seizure once back in unit and was taken stat for computed tomography. heparin was withheld in purge. the device was pulled back to 83 centimeters and a double pigtail stent was added to right leg. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28367 | 1220648-2024-09407 | 19041084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041084&pc=OZD | 03/29/2024 | 05/15/2024 |  | 04/04/2024 |  |  | 03/29/2024 | 04/04/2024 |  |  | 09/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444109 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a patient of unknown race with high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the impella implant, the sheath kinked. the patient had calcified and tortuous anatomy so the doctor opted for the longer sheath. during the insertion, the doctor found that the sheath was severely kinked in two places. the doctor removed the sheath and inserted the shorter 14 french sheath. the impella implant was successful. | [5 empty] | The investigation for the reported introducer damage has been completed. the introducer was returned for evaluation. the sheath introducer had 2 kinks as mentioned in the clinical information. patient had diseased vessels. per the clinical information and product investigation, the root cause of the introducer damage - mechanical issue was patient condition related to tortuous vessels. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28368 | 1220648-2024-09408 | 19041101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041101&pc=OZD | 03/31/2024 | 05/02/2024 |  | 04/04/2024 |  |  | 03/31/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025450082 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 66-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the 2 hours after the implant the patient developed a hematoma at the access site. the physician decided to wean and explant the pump. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 & h10 were revised with updated information.
28369 | 1220648-2024-09409 | 19041153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041153&pc=PYX | 03/31/2024 | 08/30/2024 |  | 04/04/2024 |  |  | 03/31/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1158011 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 78-year-old female patient post cardiotomy cardiogenic shock, low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that automated impella controller (aic) had a purge disc failure. the aic was replaced without patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added- d9 device return date, h6 med dev prob code 2994 d2-corrected common device name f6/f8 should have been left blank in the initial report.
28370 | 1220648-2024-09410 | 19041185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041185&pc=OZD | 08/22/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 08/22/2022 | 04/04/2024 |  |  | 06/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092816 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced a cerebrovascular accident and suffered a hemorrhagic stroke. it was reported that the device was was successfully weaned and explanted. no further information is available as the facility declines consent to follow-up. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. -timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient. -neurological examination findings and any focal deficits observed. -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). -laboratory test results, including coagulation profiles and cardiac markers. -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
28371 | 1220648-2024-09411 | 19041208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041208&pc=OZD | 04/01/2024 | 04/19/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1080190 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 68-year-old male patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that automated impella controller (aic) displayed a purge disc not detected alarm following a purge bag exchange. the aic was swapped as to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28372 | 1220648-2024-09412 | 19041263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041263&pc=OZD | 09/19/2022 | 03/11/2025 |  | 04/04/2024 |  |  | 09/19/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023191718 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella cp was removed due to the patient developing limb ischemia. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28373 | 1220648-2024-09413 | 19041277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041277&pc=OZD | 04/01/2024 | 06/05/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC |  |  | US |  | 2025462462 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 58-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that after extended use beyond the intended life cycle of the pump, the patient was noted to have suffered a cerebral hemorrhage due to straining while defecating. the patient was made do not resuscitate (dnr) and subsequently passed away. it was noted that heparin had been present in the purge line during support. a direct connection to the pump was not definitively established and there were no noted deficiencies in the function of the device. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. h1-changed type of reportable event to death. the instructions for use for the impella cp with smart assist system in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28374 | 1220648-2024-09414 | 19041296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041296&pc=OZD | 03/29/2024 | 05/13/2024 |  | 04/04/2024 |  |  | 03/29/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2025455168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The user facility reported a 76-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. following implant, the impella cp pump began to experience a drop in flow rates. via fluoroscopy, it was discovered that the cannula was kinked proximal to the motor. as a result of the noted damage, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was most likely the cannula kink due to improper technique. in accordance with updated procedures, section d6 has been revised from the initial submission.
28375 | 1220648-2024-09415 | 19041320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041320&pc=OZD | 04/01/2024 | 05/15/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  | 11/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024434321 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that while an impella 5.5 pump was in use a purge leak developed at the red hub. there was no patient harm reported. | [5 empty] | The investigation into the purge leak ¿ pump) issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the purge leak - pump issue was not determined since product was not returned for investigation and insufficient clinical information was provided. section: cleaning- ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
28376 | 1220648-2024-09417 | 19041342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041342&pc=OZD | 06/08/2022 | 04/04/2024 |  | 04/04/2024 |  |  | 06/08/2022 | 04/04/2024 |  |  | 03/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Fracture | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, the impella cp had product damage (guidewire), and there was access site bleeding. the impella was inserted through the 14fr x 13 cm peel away sheath and crossed the valve with the pigtail catheter and 0.035 j wire. this was exchanged for the 0.018" abiomed wire, removing the pigtail wire came back across the aortic valve. the 0.018" wire was removed, due to a kink was noted. the patient had bleeding and coagulopathy issues with platelet drop, and the physician removed the impella cp device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28377 | 1220648-2024-09416 | 19041355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041355&pc=OZD | 12/15/2023 | 09/25/2025 |  | 04/04/2024 |  |  | 12/15/2023 | 04/04/2024 |  |  | 09/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396380 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | Thrombocytopenia | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 74-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on impella 5.5 support, the patient had gastrointestinal(gi) bleed and thrombocytopenia, therefore heparin has been discontinued and patient was given blood products. the impella cp was successfully weaned and explanted. | The complainant reported that there were false impella position in aorta and impella position in ventricle alarms. there was normal signal on the left ventricle and normal motor current. the patient had dorsalis pedis and posterior tibial pulsed and a good radial pulse. the placement monitoring was disabled. the patient survived the event. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. upon investigation of the pump, the optical parameters were found to be within specification. placement signal unreliable alarm did not reproduce in the lab. the pump passed the laser continuity test. data logs show that the 5s parameters were also within normal ranges during the case. the aic had software version v12.2. the pump's position was confirmed to be in good position with echocardiography. the root cause of the optical signal issue and false impella in aorta alarms was most likely patient condition related as the 5s parameters were within normal ranges and the pump was confirmed to be in good position. the root cause of the thrombocytopenia issue was not determined since product was not returned and insufficient clinical information was provided. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 f6 and f8 filled out erroneously d6a / d6b h3 device evaluated by manufacturer h6 medical device problem code h6 component code.
28378 | 1220648-2024-09418 | 19041358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041358&pc=OZD | 06/29/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/29/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support, there was bleeding and a hematoma at the impella access site. the site was redressed, and the hematoma was successfully managed. no information was provided on the intervention provided to resolve the bleeding and hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28379 | 1220648-2024-09419 | 19041377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041377&pc=OZD | 04/01/2024 | 04/29/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1666954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) failed to recognize and installed purge cassette during patient support. the sensor was cleaned and the power was cycled, but the issue persisted. the aic was subsequently swapped to resolve the noted failure. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. the issue with detecting the purge disc was reproduced, and the purge flag was confirmed to be broken. the root cause of this issue was a broken purge disc flag. | The automated impella controller device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28380 | 1220648-2024-09421 | 19041413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041413&pc=OZD | 10/22/2023 | 09/19/2025 |  | 04/04/2024 |  |  | 10/22/2023 | 04/04/2024 |  |  | 09/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396392 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/22/2023 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the sheath had split, and subsequently, the hub of the sheath was inserted into the graft and secured in place. | The complainant also reported that there was an impella position unknown alarm due to extremely low pulsatility. |  |  |  |  | The investigation into the optical signal issue and the introducer damage ¿ mechanical issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the mechanical introducer damage and the optical signal issue was not determined because no product was returned, and not enough clinical information was given. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. according to the clinical, on the sixth day of impella 5.5 support an impella position unknown alarm was recorded due to low pulsatility. data log analysis showed that the 5s parameters were within range but the signal-to-noise ratio (snr) varied throughout the case with a "placement signal not reliable" alarm. the cause of the optical signal issue was not determined because no product was returned and not enough clinical information was given. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09421. d10 concomitant medical products and therapy dates updated. f6 and f8 should not have had entries in the initial report. g1 reporting contact email.
28381 | 1220648-2024-09422 | 19041419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041419&pc=OZD | 08/26/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 08/26/2022 | 04/04/2024 |  |  | 04/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222282 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 31-year-old female with implanted with an impella 5.5 for mechanical circulatory support. it was reported a patient¿s anatomy would not allow for axillary access for impella 5.5 implant and the physician decided to take a transaortal approach. the device was placed with no issues during the night. the next morning, it was reported that the patient experienced retrosternal bleeding and the hemodynamics degraded. the plan was to observe and reevaluate in the afternoon. it is noted that the source of the bleed was unclear. the patient¿s condition did not improve, and retrosternal bleeding continued. four (4) units of packed red blood cells (prbc¿s) were administered at night, and three during the day. the patient¿s retrosternal bleeding stopped. the decision was made to escalate the patient to the impella rp. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28382 | 1220648-2024-09424 | 19041433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041433&pc=OZD | 03/30/2024 | 05/21/2024 |  | 04/04/2024 |  |  | 03/30/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1679798 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that shortly after the implant, there was a high purge alarm. low purge flows were noted. the purge cassette was replaced with no avail. the automated impella controller (aic) was replaced with no resolution. eventually, placement signal showed signs of pump saturation. the impella cp was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge issue has been completed. the console was returned for evaluation. the console was tested but functioned as expected. there was no problem found with the console. data logs were reviewed finding that when a known good pump was plugged into the console it did not have the purge errors the failure following the pump across consoles only being resolved after switching pumps points to the console not being the cause of the failure. there was no issue found with console ic10010. the device functioned/operated to specification. added- d9 device return date d2-updated common device name. d4-updated model number.
28383 | 1220648-2024-09420 | 19041434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041434&pc=OZD | 09/22/2022 | 03/11/2025 |  | 04/04/2024 |  |  | 09/22/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported there was a positioning issue with the placement of the impella cp. reportedly the impella was placed on the 0.018 wire and advanced into the left ventricle apex with ease. the wire was removed and the impella was turned on auto, the pump migrated a little deep and started to cause ectopy. the cp was pulled back to a shallower position and the ectopy stopped. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28384 | 1220648-2024-09423 | 19041439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041439&pc=OZD | 04/01/2024 | 05/21/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  | 11/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025436915 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. a user facility reported that upon startup of an impella cp in a patient the pump temporarily stopped and restarted. the pump was replaced to resolve the issue. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the temporary pump stop issue has been completed since the original report was filed. the pump was returned for investigation. no physical abnormalities were found; however, biomaterial was observed in the purge gap. data logs confirmed the temporary pump stop issue. the root cause of the temporary pump stop issue was biomaterial, based on returned product and data log review. in accordance with updated procedures, section d6 has been revised from the initial submission.
28385 | 1220648-2024-09425 | 19041455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041455&pc=OZD | 08/30/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 08/30/2022 | 04/04/2024 |  |  | 10/31/2023 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148570 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there were positioning issues as the pump came into contact with the mitral valve. this was treated by explanting the pump. the patient was sent to unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28386 | 1220648-2024-09426 | 19041482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041482&pc=OZD | 04/01/2024 | 05/17/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  | 11/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438805 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having a 100% occluded left anterior descending artery. while on support, the big sheath was left in place for support and the repo sheath was inserted all the way into the big sheath. staff was unable to flush the sheath. when the impella was weaned, the impella and the sheath were pulled out as one due to the concern for a clot in the big sheath. angio showed no flow down the right leg. balloon angioplasty was performed restoring flow. a clot was noted in at the insertion site and the decision was made to send the patient to vascular surgery for penumbra with possible cut down and clot retrieval. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and thrombus could not be determined. the instructions for use referenced in the initial report is from ifu for the impella cp with smartassist system section: potential adverse events (united states) and now includes cardiac or vascular injury (including ventricular perforation). d4-updated model number g1-updated mfr site name and address in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28387 | 1220648-2024-09427 | 19041499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041499&pc=OZD | 09/27/2022 | 03/11/2025 |  | 04/04/2024 |  |  | 09/27/2022 | 04/04/2024 |  |  | 06/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. reportedly a hematoma was observed after transport to the intensive care unit. the patient was being administered integerlin, heparin and kengreal, and was waking up from sedation. it was reported that integerlin was stopped, and manual pressure was held until the hematoma resolved. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28388 | 1220648-2024-09428 | 19041502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041502&pc=OZD | 06/12/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/12/2022 | 04/04/2024 |  |  | 01/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183570 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, ventricular alarms and suction alarms began. therefore, the pump was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28389 | 1220648-2024-09429 | 19041518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041518&pc=OZD | 09/01/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 09/01/2022 | 04/04/2024 |  |  | 03/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200387 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia in the lower limb where the pump was inserted. subsequently, the decision was made to insert a new pump from the right leg after extracorporeal membrane oxygenation was removed. the patient remained on support and was sent to unit. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28390 | 1220648-2024-09430 | 19041519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041519&pc=OZD | 02/09/2024 | 04/09/2024 |  | 04/04/2024 |  |  | 02/09/2024 | 04/04/2024 |  |  | 10/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024286342 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | [10 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced multiple complications during impella cp support. it was initially noted that two days into support, the patient lost pulse in their left leg. a lower leg fasciotomy, thrombectomy, and fem-fem bypass were performed to counteract the condition. the patient was given four units of rbc (red blood cells) during the procedure. despite the intervention, the ischemia persisted in the left lower extremity and the decision was made to perform a below knee amputation. bleeding from the stump persisted over the next week resulting in seven units of rbc being given. furthermore, intermittent oozing was noted at the site and anticoagulants were subsequently put on hold to manage the condition. the pump was explanted after nineteen days of support. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the limb ischemia ¿ irreversible, the access site bleeding ¿ major, and the vascular damage - peripheral vessel issues has been completed since the original report was filed. the cause of the access site bleeding and the and the vascular damage - peripheral vessel issues was not determined since product was not returned and sufficient clinical information was not provided. as all the necessary information and product was not provided, the root cause of the limb ischemia was not determined.
28391 | 1220648-2024-09431 | 19041581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041581&pc=OZD | 10/03/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 10/03/2022 | 04/04/2024 |  |  | 06/03/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023236947 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was oozing and bleeding at the groin site which was redressed with angle matching and 2 units of blood was given for treatment. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28392 | 1220648-2024-09432 | 19041591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041591&pc=OZD | 09/06/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 09/06/2022 | 04/04/2024 |  |  | 05/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2022 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the patient received two units of blood before decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28393 | 1220648-2024-09433 | 19041628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041628&pc=OZD | 06/30/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/30/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200617 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support the patient experienced limb ischemia. an external femoral to femoral bypass was performed but it did not resolve the distal limb ischemia. therefore, the device was explanted, | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28394 | 1220648-2024-09434 | 19041641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041641&pc=OZD | 10/05/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 10/05/2022 | 04/04/2024 |  |  | 06/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023248344 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported a 60-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient's right groin had a hematoma at the device site which was treated with manual pressure, causing the oozing at the site to stop; once the pump's dressing was changed. it was unclear of the source of bleeding and low volume status; however, the patient was given a unit of blood for persistent thrombocytopenia. therefore, the decision was made to have the device removed. due to the left ventricle (lv) was severely impaired, the decision was made to exchange the device through the re-access port of the right femoral artery; which the pump was inserted with an alternate sheath in hopes of providing superior hemostasis. after the device was pulled into the aorta, the patient decompensated further and the insertion of second impella cp was initiated quickly. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28395 | 1220648-2024-09436 | 19041648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041648&pc=OZD | 09/08/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 09/08/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202612 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced occluded blood flow to the limb. the decision was made to explant the impella device. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28396 | 1220648-2024-09435 | 19041649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041649&pc=OZD | 03/30/2024 | 05/16/2024 |  | 04/04/2024 |  |  | 03/30/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025456625 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a 26-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a cannula kink occurred while attempting to implant an impella cp in a patient of small size. the user facility was clear that this issue occurred due to patient anatomy and acute angle of aortic arch, and not because of a product defect. the impella was removed, and the patient was scheduled for extracorporeal membrane oxygenation (ecmo) and cutdown for the impella device. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition as the patient has a small distance from the ventricle and an acute angle in the arch resulting in resistance during delivery. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28397 | 1220648-2024-09437 | 19041670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041670&pc=OZD | 10/07/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 10/07/2022 | 04/04/2024 |  |  | 03/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209738 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 63-year-old male patient in acute myocardial infarction/cardiogenic shock was being placed onto impella cp support. while on support, the patient experienced access site bleeding / hematoma. sutures were placed and heparin reduced to resolve hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28398 | 1220648-2024-09438 | 19041689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041689&pc=OZD | 09/11/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 09/11/2022 | 04/04/2024 |  |  | 06/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that no doppler pulses were found in the patients left leg. a fluid bolus was administered, the patient was weaned off levophed, and pulses were reassessed. due to the patient size and vascular condition the impella cp was replaced with the impella 5.5. the impella 5.5 was inserted without reported complications and good distal pulses were found. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28399 | 1220648-2024-09439 | 19041730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041730&pc=OZD | 09/24/2023 | 09/25/2025 |  | 04/04/2024 |  |  | 09/24/2023 | 04/04/2024 |  |  | 05/31/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388874 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/25/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 82-year-old male patient for mechanical circulatory support. the patient was experiencing oozing at all access sites. the patient was not given systemic heparin and the purge was switched to sodium bicarbonate. | The complainant also reported that as soon as the impella cp was inserted and support was initiated, the placement signal was noted to be significantly different than the patients actual blood pressure reading from the arterial line - roughly around a 30 to 40 point difference in readings. this also led to false alarms on the device, and the placement monitoring had to be disabled. pump positioning was confirmed. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the most likely cause of the major access site bleeding was use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09439. d6a revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. according to the clinical, the placement signal was noted to be significantly different than the patient's actual blood pressure reading from the arterial line. echo imaging confirmed the pump was in the correct position. clinical notes indicated there was around a 30 to 40 point difference in these pressure readings. this led to false alarms being recorded and resulted in the decision being made to disable placement signal monitoring. later, the next day an attempt was made to reposition the pump. once the repositioning sheath was in place, bleeding was observed coming from the distal end of the blue suture hub. upon further investigation, it was noted that the blue suture hub was not sealed appropriately around the repositioning unit near the wire re-access port. the decision was then made to replace it with a new impella device. no product was returned for an investigation. data log analysis confirmed the placement signal was low with impella position in the aorta alarms. no other abnormalities were found as the 5s parameters were all stable. the root cause of the optical signal issue is most likely due to the patients condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal issue. b7 updated with additional information. g1 reporting contact email updated.
28400 | 1220648-2024-09440 | 19041745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041745&pc=OZD | 06/18/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/18/2022 | 04/04/2024 |  |  | 07/31/2023 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116668 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced loss of pulsatility in both legs. the patient was escalated to an impella 5.5 and the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. b1 adverse event: changed from no to serious injury. d4 model number. g2 health professional: changed from no to yes.
28401 | 1220648-2024-09441 | 19041753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041753&pc=OZD | 06/12/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/12/2022 | 04/04/2024 |  |  | 01/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189337 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia that was resolved with external fem. the procedure was completed without complications, flow to distal limb restored. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28402 | 1220648-2024-09442 | 19041822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041822&pc=OZD | 10/11/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 10/11/2022 | 04/04/2024 |  |  | 06/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, device exhibited a purge leak coming through pump motor area. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28403 | 1220648-2024-09443 | 19041865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041865&pc=OZD | 03/30/2024 | 05/15/2024 |  | 04/04/2024 |  |  | 03/30/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025463324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2024 |  |  |  |  | The user facility reported a female patient of an unknown age, race, and ethnicity was to be implanted with an impella cp device for mechanical circulatory support. inability to cross aortic valve while attempting to implant the impella cp was reported. implantation was attempted three times with an aortic valvuloplasty, but was unsuccessful. impella insertion was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported difficulty inserting/removal issues has been completed. no product or logs were returned. clinical details state that the patient had stenotic aortic valve and was able to cross the valve with the guidewire but not with the impella. attempted 3 times and the aortic valvuloplasty done was still unsuccessful. the impella insertion was aborted. the root of the delivery issue was most likely patient condition related. d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28404 | 1220648-2024-09444 | 19041909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19041909&pc=OZD | 10/11/2022 | 03/11/2025 |  | 04/04/2024 |  |  | 10/11/2022 | 04/04/2024 |  |  | 07/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023240714 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella insertion site was changed to the right femoral artery due to when inserting from the left femoral artery the impella diverging during the first 8fr sheath insertion. insertion was achieved using a 16fr sheath after difficulty with a 14fr dilator. contrast imaging revealed bleeding from the left femoral that was punctured, resulting in coil embolization being performed. hemostasis was confirmed through contrast enhanced computer tomography, hand compression was used for 1 hour as a precautionary measure. reportedly there was no bleeding at the impella insertion site. however, the bleeding continued from the left femoral puncture, the impella was removed early the following day to end the heparin administration. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28405 | 1220648-2024-09445 | 19042015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042015&pc=OZD | 03/31/2024 | 05/13/2024 |  | 04/04/2024 |  |  | 03/31/2024 | 04/04/2024 |  |  | 07/31/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 381 PUMP SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2025440514 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Insufficient Flow or Under Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported an 82-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after 38 hours of support, the pump stopped repeatedly and was restarted. upon restart, the pump had low flows/suction that prompted the team to explant the pump. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and low pump flow issues has been completed since the original report was filed. the pump was returned for analysis. the data logs show a motor current spike and confirm pump stop due to 'motor current high' alarms, as well as a continuous suction after restarting the pump. upon product evaluation, biomaterial was found wrapped around the impeller at the inflow cage. the root cause of the temporary pump stop was determined to be biomaterial.
28406 | 1220648-2024-09446 | 19042070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042070&pc=OZD | 06/08/2022 | 09/30/2025 |  | 04/04/2024 |  |  | 06/08/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported that a 67-year old male patient noted as high risk percutaneous coronary insertion was being placed onto impella cp support. while on support, the pump experienced positioning issues (proximity to mitral valve), leading to hemolysis. as a result, device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number.
28407 | 1220648-2024-09447 | 19042270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042270&pc=OZD | 03/31/2024 | 05/14/2024 |  | 04/04/2024 |  |  | 03/31/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116347 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the purge cassette had repeated leaks. in order to rectify the leaking, the automated impella controller (aic) had to be replaced. no patient harm was reported. | [5 empty] | The investigation for the reported aic disc/flag issues has been completed. the console was returned for evaluation. during testing, the flag was stuck and the dextrose buildup inside the retainer also saw the purge fluid deposit behind the purge system. data logs were reviewed which confirmed that during patient support, the purge disc was removed, and the pressurized disc was immediately reinserted (this could have made the user have difficulty inserting the disc). most likely, the pressurized disc could have punctured and leaked the purge fluid into the retainer, which might have stuck the flag. the console was turned on, and a different cassette was inserted. within a minute, the cassette was removed and reinserted. most likely, the stuck flag could have punctured the purge disc. then, the console was shut down and restarted, and the same cassette was inserted and removed several times. every time the disc was inserted, the console received "purge disc not detected ¿ alarm". this symptom likely aligns with the reported failure mode. the root cause of the leakage was most likely the punctured purge disc membrane caused by the immovable flag, which was caused by the purge fluid ingress. added: d9 device return date, h6 component code 765 in accordance with updated procedures, sections d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir."
28408 | 1220648-2024-09448 | 19042294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042294&pc=OZD | 06/13/2022 | 04/04/2024 |  | 04/04/2024 |  |  | 06/13/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201279 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing from all sites. the patient received four units of packed red blood cells, 955mls cell saver, a unit of platelets, 4 units fresh frozen plasma and tranexamic acid. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28409 | 1220648-2024-09449 | 19042408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042408&pc=OZD | 03/31/2024 | 05/16/2024 |  | 04/04/2024 |  |  | 03/31/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2025454668 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Insufficient Flow or Under Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2024 |  |  |  |  | The user facility reported a 68-year-old male patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that following an extended priming sequence, an implanted impella cp pump began to experience low pump flows and low purge volume alarms. positioning was verified via echocardiogram and the decision was subsequently made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow and the high purge pressure issues has been completed since the original report was submitted. the device was not returned for investigation. data logs did not show any high purge pressure. however, priming issue was identified. there were suction and positioning alarms, but the pump was confirmed on echo to be in correct position. the root cause of the priming issue was most likely use related since they were able to resolve the priming alarms without replacing the pump. the root cause of the low pump flow issue was not determined since product was not returned and limited clinical information was provided. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28410 | 1220648-2024-09450 | 19042549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042549&pc=OZD | 06/13/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/13/2022 | 04/04/2024 |  |  | 02/28/2023 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055636 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the implantation of the impella cp implant was aborted due to the patient¿s vascular anatomy. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28411 | 1220648-2024-09451 | 19042657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042657&pc=OZD | 06/14/2022 | 04/04/2024 |  | 04/04/2024 |  |  | 06/14/2022 | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191482 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had hemolysis upon insertion, which was treated by explanting the impella cp device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records
28412 | 1220648-2024-09452 | 19042749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042749&pc=OZD | 04/01/2024 | 06/07/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  | 11/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024435423 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant from the study reported that during impella cp support for male patient, the patient experienced heparin induced thrombocytopenia that was moderate in severity. the event was determined to not be related to the impella but probably related to the impella procedure. heparin was stopped and argatroban was initiated. additionally, the patient was noted to have an infection at the impella access site. the event was determined to be definitely related to the impella procedure but not related to the impella device. the patient received targeted antibiotic therapy and vacuum assisted closure therapy. | [5 empty] | The investigation into the thrombocytopenia and the infection/sepsis issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the thrombocytopenia was not determined due to insufficient clinical details. because evidence related to infection/sepsis has not been provided, the root cause cannot be determined. d6a, d6b. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
28413 | 1220648-2024-09453 | 19042844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042844&pc=OZD | 11/05/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 11/05/2022 | 04/04/2024 |  |  | 04/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023224108 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/05/2022 |  |  |  |  | The user facility reported a 38-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the patient's urine output tinged with blood. the signs and symptoms of hemolysis were treated with switching back to heparinized purge and continuous renal replacement therapy (crrt ). during explanting of the device, a clot formed in the iliac. penumbra was used allowing the access point to be stented and hemostasis achieved. the closure method was completed with preclose w perclose proglide x2. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28414 | 1220648-2024-09454 | 19042889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042889&pc=OZD | 04/01/2023 | 06/07/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  | 12/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303752 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility in japan reported a patient of unknown age and gender was mentioned in a literature report to have been implanted with an impella cp device for mechanical circulatory support. it was reported in the publication that the patient sustained a left ventricle (lv) perforation after the pump was placed to the lv and along the posterior wall. the posterior wall perforated, and the team placed a drain for pericardiocentesis, and surgery was performed. the patient survived and was transferred out to a rehabilitation hospital after 66 days. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical information provided, the root cause of ventricle perforation was most likely improper positioning due to improper technique. in accordance with updated procedures, sections d6 and g3 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
28415 | 1220648-2024-09455 | 19042955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19042955&pc=OZD | 11/14/2022 | 03/07/2025 |  | 04/04/2024 |  |  | 11/14/2022 | 04/04/2024 |  |  | 08/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 71-year-old male was implanted with an impella cp device for mechanical circulatory support. during implantation, the insertion of the short 14fr sheath was noted to be cracked which was removed and replaced with the long sheath. eventually the device the was successfully weaned. no patient harm reported. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28416 | 1220648-2024-09456 | 19043091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043091&pc=OZD | 11/17/2022 | 05/28/2025 |  | 04/04/2024 |  |  | 11/17/2022 | 04/04/2024 |  |  | 08/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265393 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 72-year-old patient was implanted with an impella cp for mechanical circulatory support. it was reported that while on support the patient experienced hemolysis. the patient remains on continuous renal replacement therapy (crrt) and received 1 unit packed red blood cells (prbc). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28417 | 1220648-2024-09457 | 19043163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043163&pc=OZD | 04/01/2024 | 05/09/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168283 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 59-year-old female patient in st-segment elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the automated impella controller (aic) displayed a purge disc not detected alarm during a cassette change. the aic was swapped as to resolve the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. added- d9 device return date in accordance with updated procedures, sections d2 has been revised from the initial submission.
28418 | 1220648-2024-09458 | 19043292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043292&pc=OZD | 11/18/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 11/18/2022 | 04/04/2024 |  |  | 04/30/2023 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022081102 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device the patient developed hematoma post pump implant. the device was explanted and a intra-aortic balloon pump (iabp) was placed for cardiac therapy. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28419 | 1220648-2024-09459 | 19043329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043329&pc=OZD | 11/18/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 11/18/2022 | 04/04/2024 |  |  | 07/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253348 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 82-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on support it was noted that the patient encountered a perforation of the left anterior descending coronary artery (lad) and developed pericardial effusion. the physician accessed the pericardial and drained 350cc of blood. the patient was stabilized and the device was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28420 | 1220648-2024-09460 | 19043464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043464&pc=OZD | 11/18/2022 | 03/10/2025 |  | 04/04/2024 |  |  | 11/18/2022 | 04/04/2024 |  |  | 08/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023265619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. distal pulses were difficult to doppler bilaterally and the right lower extremity(rle) was appearing more discolored than the left lower extremity (lle). the patient's condition became stable; however, the lower extremity was still cool and dusky and there was no pulse. the right lower extremity had a dopplerable post tibial pulse, but no pedal and the leg appeared decently perfused. eventually, a stent was placed in the common femoral artery (cfa) with a successful closure. the device was successfully weaned. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28421 | 1220648-2024-09462 | 19043507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043507&pc=OZD | 04/01/2024 | 05/17/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  | 11/30/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024435030 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, there were suction events. there was a spike in the motor current and left ventricle signals. it was suspected that biomaterial was ingested. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. the returned pump was visually inspected, and biomaterial was observed around impeller, and the cannula shape appeared to be deformed. the cause of the low pump flow was biomaterial ingestion. in accordance with updated procedures, section d6 has been revised from the initial submission.
28422 | 1220648-2024-09461 | 19043525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043525&pc=OZD | 11/18/2022 | 09/17/2025 |  | 04/04/2024 |  |  | 11/18/2022 | 04/04/2024 |  |  | 07/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259232 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 78-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient experienced an internal bleed from femoral artery from initial stick. the patient¿s hemoglobin dropped from 10 to 6.8 from the time of insertion requiring blood transfusion. the patient received 2 units of packed red blood cells in the operating room. hemodynamics improved allowing for a successful weaning of the impella cp. the vascular surgeon performed vascular surgery repair by cutdown and the impella was removed. reportedly the patient remained hemodynamically stable throughout the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported for the access site bleeding has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleeding could not be determined. g1 mdr reporting contact email updated.
28423 | 1220648-2024-09463 | 19043601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043601&pc=OZD | 11/23/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 11/23/2022 | 04/04/2024 |  |  | 06/30/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248061 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 52-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the staff observed a thrombus formation in the left lower extremity (lle) within 2 hours of explant. the clot was removed from the left superficial femoral artery and the flow was restored. the patient was taken to unit in stable condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. limb ischemia: when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. thrombus: when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
28424 | 1220648-2024-09464 | 19043647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043647&pc=OZD | 04/01/2024 | 04/19/2024 |  | 04/04/2024 |  |  | 04/01/2024 | 04/04/2024 |  |  |  |  | 04/04/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137163 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility biomedical engineer reported the automated impella controller's (aic) purge disc was not connecting. the aic was replaced. no patient was involved. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28425 | 1220648-2024-09465 | 19043905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19043905&pc=OZD | 11/24/2022 | 04/04/2024 |  | 04/04/2024 |  |  | 11/24/2022 | 04/04/2024 |  |  | 07/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231476 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 73-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. oozing was noted at the impella access site with the peel away sheath in place and the blue hub sutured to the patient¿s leg. the touhy-borst valve was not tight but the catheter was still at 100 cm which was appropriate per cardiac cath lab staff. aggrastat was discontinued and the groin site was redressed. moreover, the stents to the left anterior descending (lad) and circumflex arteries were occluded and the right coronary artery was closed. the impella remained in the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28426 | 1220648-2024-09466 | 19044358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19044358&pc=OZD | 11/25/2022 | 04/04/2024 |  | 04/04/2024 |  |  | 11/25/2022 | 04/04/2024 |  |  | 08/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267574 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper or Incorrect Procedure or Method | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 72-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced continuous oozing at the impella site which was dressed heavily for transport. oozing continued after transport from one hospital to another. upon removal of dressing, it was discovered that the blue site hub was under the skin as the incision was too large. the patient was taken to the lab to address the bleeding. however, the patient's symptoms had subsided and filling pressures were back down to normal. it was determined that the impella would no longer be needed when a swan was used to reveal satisfactory filling pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28427 | 1220648-2024-09467 | 19044557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19044557&pc=OZD | 08/31/2023 | 09/19/2025 |  | 04/04/2024 |  |  | 08/31/2023 | 04/04/2024 |  |  | 07/31/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377579 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/31/2023 |  |  |  |  | The user facility reported a 41-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient began to experience frequent ectopy once back in the intensive care unit. echocardiogram revealed catheter measuring deep into left ventricle (lv). the physician repositioned the catheter to address the issue. the patient is hemodynamically stable. | Additionally, the complainant also reported that during repositioning, placement signal and left ventricle waveform became flat and began to alarm placement signal not reliable. troubleshooting was done and everything was working accordingly. alarm audio was disabled. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia could not be determined. added- h6 component code 925 d4-corrected model number. added- d6a/d6b. f6/f8 should have been left blank in the initial report. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs about 2 hours & 15 minutes into support, the placement signal (ps) spiked to 3000 mmhg, which triggered ps not reliable alarms. this event remained for the rest of the case. logs also showed signal-to-noise ratio, lamp & light were out of specification during the case indicated a fiber break. product was not returned for review. the root cause of the optical signal issue was most likely a fiber break. the root cause of the fiber break was unable to be determined as limited clinical details were provided and no product was returned for review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b.1. product problem updated. b.5. updated to include optical signal issue description. h.6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09467. b.7. other relevant history, including preexisting medical conditions, updated. d.10. concomitant medical products and therapy dates updated. g.1. reporting contact email and reporting contact fax number updated.
28428 | 1220648-2024-09468 | 19044692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19044692&pc=OZD | 10/15/2022 | 04/04/2024 |  | 04/04/2024 |  |  | 10/15/2022 | 04/04/2024 |  |  | 07/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023253096 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/15/2022 |  |  |  |  | The user facility reported a 58-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support and the patient experienced a thrombus formation. thrombus was formed in the original 14x13cm sheath so was exchanged for the 14x25cm sheath. when the sheath was being peeled away, a massive amount of blood was lost. once repo sheath was in place, bleeding ceased but the patient required drips and a rapid transfusion of 1l of whole blood. the bleeding ceased once the repo sheath was advanced and in place. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28429 | 1220648-2024-09469 | 19044930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19044930&pc=OZD | 08/18/2023 | 09/19/2025 |  | 04/04/2024 |  |  | 09/23/2023 | 04/04/2024 |  |  | 05/31/2025 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353391 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/25/2023 |  |  |  |  | The complainant also reported that there were frequent impella position unknown alarms. the placement signal (ps) is not reliable. the nurse was advised to ignore the impella position unknown alarm at the time and to continue to monitor cuff pressures. the patient was stable without any changes in condition. it was further reported that a preventing retrograde flow alarm appeared. the motor current increased with the alarm, and the ps continued to remain high and would even disappear. the nurse was advised to disable ps monitoring as well as retrograde flow alarms so the impella would not run higher than p5 with a possible failing sensor. the nurse reiterated that the patient¿s condition had not changed. | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the heparin drip was on hold due to endotracheal versus airway oozing causing large clots to form. the impella was repositioned, without difficulty from 3.4 centimeters (cm) to 5.2 cm mid-inlet to atrioventicular. the patient had no signs of hemolysis but was having arrhythmias upon turning. |  |  |  |  | The investigation for the reported ventricular fibrillation (vf) and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf and thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ added- h6 component code 925 corrections to the initial mfr report: d4-corrected model number added-d6a/d6b f6/f8 should have been left blank. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for an investigation. data log analysis confirmed multiple impella in aorta alarms and placement signal spikes throughout support. the logs also confirmed a decreasing trend in signal-to-noise ratio (snr) and contrast. upon further inspection of the imc logs, the console software version was also found to be v12.1. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with optical signal issue description. h6 updated with optical signal coding. corrections that were made from the initial manufacturer device report number. b1 product problem updated. b7 other relevant history, including preexisting medical conditions, updated. d10 concomitant medical products updated. g1 reporting contact email and reporting contact fax number were updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28430 | 1220648-2024-09470 | 19046710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19046710&pc=OZD | 06/05/2021 | 05/20/2025 |  | 04/04/2024 |  |  | 06/05/2021 | 04/04/2024 |  |  | 12/31/2023 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165796 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2021 |  |  |  |  | The user facility reported an impella cp was implanted in a 81-year-old male patient for mechanical circulatory support. it was reported during insertion, that the patient experienced hemolysis, which resolved by refining the placement position through echocardiography and adjusting the level of assistance accordingly. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28431 | 1220648-2024-09471 | 19046761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19046761&pc=OZD | 05/31/2022 | 04/04/2024 |  | 04/04/2024 |  |  |  | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195196 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support gi bleeding, as well as bleeding from groin sites was noted. decision made to explant and decannulate, then escalate to a impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28432 | 1220648-2024-09472 | 19046985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19046985&pc=OZD | 05/31/2022 | 05/19/2025 |  | 04/04/2024 |  |  |  | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200446 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the position was confirmed via echo but before leaving, the patient started having suction and decoupled left ventricle waveform. the impella was repositioned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28433 | 1220648-2024-09473 | 19047175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19047175&pc=OZD | 05/31/2022 | 04/04/2024 |  | 04/04/2024 |  |  |  | 04/04/2024 |  |  | 01/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183634 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, there was no heparin in purge due to right groin bleeding complication. the patient was diagnosed with hematuria and the physician was unable to advance wire beyond sheath. contrast injection through impella sheath revealed what initially appeared to be a dissection. access obtained on left femoral artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28434 | 1220648-2024-09474 | 19047919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19047919&pc=OZD | 06/03/2022 | 04/03/2025 |  | 04/04/2024 |  |  |  | 04/04/2024 |  |  | 02/29/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202173 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2022 |  |  |  |  | The user facility reported a 31-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced bleeding at hemostatic valve of peel away, vt (ventricular tachycardia), some septic component to shock, and positioning issues. pigtail noted caught under pap, dr lee able to free with pullback. there were some challenges finding sweet spot for position as pigtail of cp noted to be causing irritation/stable vt. eventually able to torque impella cp, inflow about 3 centimers from aov (aortic velocity), pointed mid apex, vt (ventricular tachycardia) calmed. it was reported that the patient was in stable condition. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28435 | 1220648-2024-09475 | 19047973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19047973&pc=OZD | 06/03/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/03/2022 | 04/04/2024 |  |  | 01/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183530 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2022 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella continously presented suction alarms and low filing pressure. a large volume of normal saline and albumin was given overnight. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28436 | 1220648-2024-09476 | 19047992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19047992&pc=OZD | 06/04/2022 | 04/03/2025 |  | 04/04/2024 |  |  | 06/04/2022 | 04/04/2024 |  |  | 03/31/2024 |  | 04/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204989 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 60-year-old male patient for mechanical circulatory support. it was reported that the pump got caught under the papillary muscle and as a result the patient had hemolysis. the patient received two unit of packed red blood cells (rbcs). the impella cp device was explanted, and the patient was escalated to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | . this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28437 | 1220648-2024-09289 | 19030822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19030822&pc=OZD | 06/22/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/22/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178885 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing from the impella access site at the groin. heparin was discontinued to stop the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28438 | 1220648-2024-09290 | 19030863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19030863&pc=OZD | 09/16/2022 | 03/07/2025 |  | 04/03/2024 |  |  | 09/16/2022 | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237582 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient started feeling numbness in left foot; foot was noticeably colder than right foot; faint dopplerable pulses. cardiac index (ci) 3.0, co 5.1, central venous pressure (cvp) 6, pa 24/12, left ventricular end diastolic pressure (lvedp) on automated impella controller (aic) 4, no drips. percutaneous coronary intervention (pci) to be performed when patient is stronger. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28439 | 1220648-2024-09291 | 19030918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19030918&pc=OZD | 09/17/2022 | 03/12/2025 |  | 04/03/2024 |  |  | 09/17/2022 | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The complainant reported a 63-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, impella repositioning sheath was fully occluded. therefore, impella was explanted, and circulation was restored. md believes incident was from vaso spasm. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28440 | 1220648-2024-09292 | 19030920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19030920&pc=OZD | 07/01/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 07/01/2022 | 04/03/2024 |  |  | 02/29/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200105 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced heparin induced thrombocytopenia and bleeding at the impella access site. heparin was discontinued and was switched to bicarbonate. the perclose sutures at the impella access site failed so patient was taken to operating room for cutdown and surgical repair. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28441 | 1220648-2024-09293 | 19030974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19030974&pc=OZD | 09/17/2022 | 03/07/2025 |  | 04/03/2024 |  |  | 09/17/2022 | 04/03/2024 |  |  | 01/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225239 |  | 1000115 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The complainant reported a 41-year-old female patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on impella cp patient had oozing from site. the patient had manual compression with frequent sutures to contain the bleeding. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28442 | 1220648-2024-09294 | 19031054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031054&pc=OZD | 09/27/2022 | 05/20/2025 |  | 04/03/2024 |  |  | 09/27/2022 | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235293 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a 74-year-old patient needing mechanical circulatory support was implanted with impella cp. it was reported that the patient suffered from hematuria earlier in the day and the patient¿s urine was dark red in color. the health care professional (hcp) reduced the pump speed level from p9 to p7 and the urine was said to be pink and clearing and ultimately was amber in color with slight pink hue. the patient was to be weaned from the impella cp in the cardiac catheterization laboratory (ccl). the device was withdrawn from the aorta and the repo sheath side access port was wired. upon removal, an occlusion was noted of the external iliac artery distal to the sheath site with collateral flow and organized thrombus. medical staff were consulted, and a decision was made to intervene via extremity. it was reported that the patient was successfully weaned and survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28443 | 1220648-2024-09295 | 19031063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031063&pc=OZD | 08/19/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 08/19/2022 | 04/03/2024 |  |  | 05/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, patient developed hematoma and swelling above the impella access site was noted. the physician was at the bedside and noted that the hematoma was growing in size at the impella access site. the decision was made to rapidly wean the impella due to suspected bleed from impella insertion site. two units of packed red blood cells were given to the patient. continuous ooze noted from impella site despite forward tension dressing and the impella was explanted and the previously deployed perclose sutures were tightened and manual pressure was held for 15 minutes. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28444 | 1220648-2024-09296 | 19031133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031133&pc=OZD | 09/28/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 09/28/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204505 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 82-year-old patient needing mechanical circulatory support. it was reported that while on impella cp support, the patient experienced oozing at the groin site. heparin drip was stopped. the site was redressed, and heparin was adjusted, the bleeding stopped. heparin was added back to the purge solution. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28445 | 1220648-2024-09297 | 19031222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031222&pc=OZD | 08/19/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 08/19/2022 | 04/03/2024 |  |  | 05/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230857 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was placed on impella cp support after presenting with chest pain and non-st elevated myocardial infarction and left heart catheterization revealed multi vessel disease. it was reported while on impella cp support the patient experienced a cold foot and the patient was taken to interventional radiology for penumbra and percutaneous transluminal angioplasty (pta) of the central femoral artery (cfa). it was reported while on impella cp support, the patient experienced hematuria. transesophageal echocardiography (tte) imaging was attempted for repositioning of the impella, however, it provided limited view. the physician decided to explant the impella due to worsening hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28446 | 1220648-2024-09298 | 19031309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031309&pc=OZD | 09/12/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 09/12/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202186 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported that a 55-year-old male patient was placed on impella cp smart assist support for venting purposes post percutaneous coronary intervention (pci). the patient developed bleeding and a hematoma requiring blood transfusions. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28447 | 1220648-2024-09299 | 19031405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031405&pc=OZD | 06/09/2022 | 03/28/2025 |  | 04/03/2024 |  |  | 06/09/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170438 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2022 |  |  |  |  | The user facility reported a 57-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was found to be positive for heparin induced thrombocytopenia. heparin was removed from the purge. the patient remained on impella cp support and was transferred to the cardiac care unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28448 | 1220648-2024-09300 | 19031417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031417&pc=OZD | 03/12/2024 | 04/29/2024 |  | 04/03/2024 |  |  | 03/12/2024 | 04/03/2024 |  |  | 11/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024428127 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 37-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. a transthoracic echocardiogram was completed and revealed a large left ventricle thrombus near the inlet of the impella pump. hemodynamically the patient had improved and due to the noted thrombus, a plan to wean the patient off the pump was made and completed. an intra-aortic balloon pump (iabp) was placed following removal of the impella device for on-going patient support. | [5 empty] | The investigation into the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28449 | 1220648-2024-09301 | 19031450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031450&pc=OZD | 03/18/2024 | 05/17/2024 |  | 04/03/2024 |  |  | 03/18/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025456866 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support for male patient, the pump had low volume despite volume substitution >2l. the pump position was checked to be correct via tte. prior to impella placement, heparin had been injected and act was 261. the pump was ultimately explanted and a new impella was placed without issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. the returned pump was visually inspected. broken blades were found in the impeller. the cause of the low pump flow issue was damaged impeller due to interaction with another devise.
28450 | 1220648-2024-09302 | 19031476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031476&pc=OZD | 03/15/2024 | 05/13/2024 |  | 04/03/2024 |  |  | 03/18/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025455394 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for male patient, the patient experienced superficial femoral artery occlusion leading to pump explant. the patient is noted to be stable following explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation."
28451 | 19031596 | 19031596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031596&pc=OZD | 09/12/2023 | 10/31/2023 | 10/31/2023 | 04/03/2024 | 10/31/2023 | 04/03/2024 |  | 04/03/2024 |  |  |  |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 444125 |  | 0048-0003 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | [7 empty] | Failure of Implant | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella device would not start due to wire entrapment. catheter placed in patient, then staff tried to pull out sheath, which got stuck and would not come out properly. they removed the device and used a new impella to finish the procedure.
28452 | 1220648-2024-09312 | 19031847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031847&pc=OZD | 03/27/2024 | 05/01/2024 |  | 04/03/2024 |  |  | 03/27/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025455802 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user reported that while on impella support the patient developed access site bleeding and hematoma. the physician decided to remove the impella, and manual pressure and a fem stop was applied. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ .¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding was most likely due to patient movement since the acute bleeding was noted in the groin after the patient had begun moving. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
28453 | 1220648-2024-09313 | 19031873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031873&pc=OZD | 11/23/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 11/23/2022 | 04/03/2024 |  |  | 08/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient underwent a computed tomography (ct) scan of brain to assess for bleed. pedal pulses were unable to be obtained, however cap refill was present. it was recommended to wean off pressors and reassess flow down the leg. patient hemoglobin dropped from a previous blunt force head trauma and bleeding at the impella access site which resulted in persistent suction alarms on the impella. the patient received two units of packet red blood cell (prbc)s with moderate resolution to the suction alarms. the impella access site was redressed and bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28454 | 1220648-2024-09314 | 19031887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031887&pc=OZD | 06/06/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/06/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177180 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had access site bleeding and limb ischemia noted by the patient's cold and blue right leg. the access site bleeding was treated by placing a pressure dressing, removing heparin, and giving the patient two units of blood and vitamin k. the limb ischemia was treated by removing the peel-away sheath and advancing the repositioning sheath. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28455 | 1220648-2024-09315 | 19031930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031930&pc=OZD | 04/01/2024 | 06/24/2024 |  | 04/03/2024 |  |  | 03/28/2024 | 04/03/2024 |  |  | 10/31/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024420608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Pressure Problem | [6 empty] | Renal Impairment | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that during placement of the impella device a high purge pressure alarm was generated. the pump placement was successful and no harm to the patient was reported. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used: ¿required temporary dialysis catheter. renal function improved and returned to baseline..¿ patient is stable . |  |  |  |  | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | Instructions for use for the related event are as follows: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the high purge pressure and the renal failure issues has been completed since the original report was submitted. the product was not returned for investigation. data log revealed that there was a corresponding motor current spike followed by saturated flows suggesting that there was a biomaterial buildup. adding tpa resolved it. therefore, the root cause of the high purge pressure issue was most likely due to biomaterial. the root cause of the renal failure issue was not determined since product was not returned and unknown relation to impella. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28456 | 1220648-2024-09316 | 19031969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031969&pc=OZD | 03/29/2024 | 04/03/2024 |  | 04/03/2024 |  |  | 03/29/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1346474 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/29/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a red battery alarm during power up. the aic was removed from service and a backup aic was put into place. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28457 | 1220648-2024-09317 | 19031997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19031997&pc=OZD | 11/24/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 11/24/2022 | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226992 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 60-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due acute chest pain and triple vessel disease. at the beginning of the percutaneous coronary intervention (pci), the patient became hypotensive and an impella cp was placed emergently. while on support, the patient experienced red urine and the groin access site was actively bleeding. the access site was successfully redressed with assistance from the physician assistant. two days later, the access site dressing was noted to be slightly saturated. the clinical support walked the rn through adequate angle matching and dressing technique. 2 perclose devices were placed. there were no further reports of oozing. hemolysis remained unresolved until pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28458 | 1220648-2024-09318 | 19032001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032001&pc=OZD | 03/31/2024 | 06/04/2024 |  | 04/03/2024 |  |  | 04/01/2024 | 04/03/2024 |  |  | 10/31/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024419243 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility in japan reported a patient of unknown age, gender, and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after 55 hours of support pump was explanted due to persistent suction despite all methods of troubleshooting. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned for investigation visual inspection found biomaterial inside the pump housing & wrapped around the impeller. the root cause of the low flow was biomaterial ingestion. in accordance with updated procedures, section d6 has been revised from the initial submission.
28459 | 1220648-2024-09319 | 19032131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032131&pc=OZD | 03/23/2024 | 06/20/2024 |  | 04/03/2024 |  |  | 04/01/2024 | 04/03/2024 |  |  | 09/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024406351 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Migration or Expulsion of Device | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility in japan reported a female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the physician attempted to adjust the position of the impella to a shallower position it slipped out of the aorta. the impella was subsequently removed from the patient. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning has been completed. the impella device was not returned for evaluation, however, clinical details were sufficient to determine the root cause. the cause of the positioning issue most likely use related due to improper technique as the pump slipped out to aorta while while adjusting it to a shallower position. f6/f8-should have been left blank.
28460 | 1220648-2024-09320 | 19032180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032180&pc=OZD | 03/16/2024 | 05/17/2024 |  | 04/03/2024 |  |  | 03/16/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446833 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support for 43-year-old male patient, the automated impella controller (aic) displayed a low purge pressure alarm. the purge cassette was exchanged but the alarm was ongoing. therefore, the aic was also exchanged but the issue continued. following unsuccessful troubleshooting, the decision was made to exchange the impella for a new one. the issue resolved and the patient was noted to be stable. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella cp pump was returned for evaluation was visually inspected. dried blood was observed along the cannula and inlet/outlet cages. dried blood was dissolved to clear observation. biomaterial was found around impella. no other abnormalities found. the cause of the low pump flow was biomaterial ingest.
28461 | 1220648-2024-09321 | 19032196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032196&pc=OZD | 06/08/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/08/2022 | 04/03/2024 |  |  | 01/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184723 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a 68-year-old patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. they attempted implant was not successful due to patient's difficult anatomy with peripheral artery disease. as a result, procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28462 | 1220648-2024-09322 | 19032267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032267&pc=OZD | 03/18/2024 | 05/13/2024 |  | 04/03/2024 |  |  | 03/18/2024 | 04/03/2024 |  |  | 11/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424515 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient had leg ischemia leading to pump explant. the ischemia began after repositioning unit insertion. the leg is noted to be well perfused after explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28463 | 1220648-2024-09323 | 19032534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032534&pc=OZD | 03/18/2024 | 05/06/2024 |  | 04/03/2024 |  |  | 03/18/2024 | 04/03/2024 |  |  | 11/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433342 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a hematoma developed at the access site, and two units of packed red blood cells were given to the patient. the patient was successfully escalated to an impella 5.5 device. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28464 | 1220648-2024-09324 | 19032643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032643&pc=OZD | 06/07/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 09/30/2022 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2021007632 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia due to lost pulse to the right foot and heparin was administered. the patient also experienced bleeding/hematoma to the impella site and hemoglobin dropped to 4. surgical exploration was completed and the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28465 | 1220648-2024-09325 | 19032766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032766&pc=OZD | 06/07/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172117 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon arrival at the hospital, the patient had no pulse in the foot. during and after the introducer was inserted, the patient lost pulses at the limb. medical intervention for the ischemia was not provided and is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28466 | 1220648-2024-09326 | 19032970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032970&pc=OZD | 06/07/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171981 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ventricular fibrillation / ventricular tachycardia with return of spontaneous circulation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28467 | 1220648-2024-09327 | 19032974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19032974&pc=OZD | 11/07/2023 | 09/23/2024 |  | 04/03/2024 |  |  | 11/07/2023 | 04/03/2024 |  |  | 09/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402039 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician was unhappy with the placement of the pump and decided to remove and rewire for entry. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. according to the clinical, the pump was placed into position under fluoroscopy and the guidewire was removed. however, after starting the pump the physician was not happy with the placement and decided to remove the pump and rewire for entry. no product was returned for an investigation. the most likely cause of the positioning issues was use related.
28468 | 1220648-2024-09328 | 19033194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033194&pc=OZD | 06/08/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 06/08/2022 | 04/03/2024 |  |  | 10/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149019 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient presented hemolyzed urine and labs. the impella cp device was repositioned. the repositioning of the device did not fully correct the hemolyzed urine. care was withdrawn and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28469 | 1220648-2024-09329 | 19033266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033266&pc=OZD | 11/07/2023 | 09/23/2024 |  | 04/03/2024 |  |  | 11/07/2023 | 04/03/2024 |  |  | 09/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had hemolysis and it was reported that the staff had issues with obtaining lab results due to the lab stating specimens were hemolyzed, so heparin infusion had never been started. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. according to the clinical, heparin was never started as the labs also stated specimens were hemolyzed. the decision was then made to explant the pump. no product was returned for an investigation. data log analysis confirmed an initial suction alarm at the beginning of the case with rising purge pressure and saturated flows. a bag change was performed 3 hours later after "purge pressure high" alarms occurred but that did not reduce the purge pressure. there was no anticoagulants being utilized. the root cause of the hemolysis was most likely due to ingested biomaterial. the root cause of the high purge pressure is most likely due to biomaterial. added- h6 component code 925 corrections for the initial mfr report: added- d6a/d6b.
28470 | 1220648-2024-09331 | 19033465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033465&pc=OZD | 06/12/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/12/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206499 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 28-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient was successfully unloaded and placed on venoarterial extracorporeal membrane oxygenation (va ecmo). the patient suffered left lower extremity ischemia where the va ecmo was cannulated and diminished pulses in the right lower extremity. the decision was made to remove the impella cp and cannulate to the centrally placed va ecmo. the device was explanted and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition, any concomitant medication, vascular size or variations thereof, detailed description of the symptoms experienced by the patient, physical examination findings, including pulse assessment and visual examination of the affected limb, diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography, laboratory test results, including blood tests for markers of inflammation or clotting disorders, any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension, or patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28471 | 1220648-2024-09330 | 19033471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033471&pc=OZD | 06/08/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/08/2022 | 04/03/2024 |  |  | 02/29/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had a massive hematoma that required manual pressure and c-clamp pressure for 3 hours. the hematoma continued to grow and was not improving, with bleeding still occurring around the venous site on the same groin. the patient had cold extremities, only doppler pulses, and was pale. the access site was gained on the opposite side and the team was able to do dry closure technique with perclose. the groin was then held, and the blood milked for the tissue, the hematoma had better result and wasn¿t as big, bruising had already developed on the pelvic area / groin and right upper leg. the impella was weaned and removed due to the bleeding/oozing and unresolved hematoma. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28472 | 1220648-2024-09332 | 19033669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033669&pc=OZD | 06/28/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/28/2022 | 04/03/2024 |  |  | 09/30/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129772 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 40-year-old female was implanted with an impella cp for mechanical circulatory support. it was reported the patient was hospitalized with fulminant myocarditis and impella cp and extracorporeal membrane oxygenation (ecmo) support were required. after device insertion, the dorsalis pedis artery was weakly palpable with doppler, so a reverse sheath was inserted as a preventive measure. the blood flow on the impella side was observed to be poor and the plan was to remove the device once the patient¿s heart function improved. after the impella was explanted, the color in the lower extremities returned and there were no signs of ischemia. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28473 | 1220648-2024-09333 | 19033716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033716&pc=OZD | 09/29/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 09/29/2022 | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219980 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 61-year-old female patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to mid right coronary artery lesion. a kink was noticed in the cannula on x-ray and the pump was replaced for another impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28474 | 1220648-2024-09334 | 19033719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033719&pc=OZD | 09/29/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 09/29/2022 | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patent had vascular damage. during explant, patient suffered continued bleeding and unsuccessful closure due to patient having peripheral artery disease (pad). the vascular surgery team made the decision to cut and surgically close the affected area. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28475 | 1220648-2024-09335 | 19033737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033737&pc=OZD | 09/18/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 09/18/2022 | 04/03/2024 |  |  | 05/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225221 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 96 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the patient¿s vitals were not stable and on trying to insert the impella cp device as a replacement from the intra-aortic balloon pump the catheter could not be inserted due to blood vessels in the lower limbs and tortuosity. therefore, the insertion of the impella was stopped. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28476 | 1220648-2024-09336 | 19033825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033825&pc=OZD | 06/13/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/13/2022 | 04/03/2024 |  |  | 02/29/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported a 65-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced frequent ectopy and runs of ventricular tachycardia due to deep pump placement. the pump was pulled back and ectopy stopped. the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
28477 | 1220648-2024-09337 | 19033903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19033903&pc=OZD | 06/13/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/13/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205602 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patients' platelets dropped. one unit of blood was administered. the patient continued on impella cp support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. the cause of the thrombocytopenia was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
28478 | 1220648-2024-09338 | 19034013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034013&pc=OZD | 06/18/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/18/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171887 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced an access site bleed during support. heparin was withheld to resolve the bleeding. the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28479 | 1220648-2024-09339 | 19034095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034095&pc=OZD | 06/24/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/24/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 62-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced persistent bleeding at the impella access site. two perclose devices were deployed, an angioseal was attempted but the arteriotomy was too large for deployment, and manual pressure was held for 15 minutes until hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28480 | 1220648-2024-09340 | 19034135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034135&pc=OZD | 11/14/2023 | 09/24/2025 |  | 04/03/2024 |  |  | 11/16/2023 | 04/03/2024 |  |  | 09/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405063 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 43-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at the impella site. the plan was to evacuate. the patient received packed red blood cells. patient is stable post issue. | The complainant also reported that the optical signal was lost. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the access site bleed has been completed. pump was not returned for investigation. information related to hematoma at access site is not provided in the clinical details. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information provided. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: g1 mdr reporting contact email. h6 component code 925. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were investigated, and a placement signal not reliable (psnr) alarm was observed. the placement signal (ps) goes to 3000 and never recovers, signal-to-noise ratio (snr) drops to 0, lamp goes to 100, contrast is a spread, gain increases and light decreases. these symptoms are similar to that of a fiber break since the ps doesn't recover. therefore, the root cause of the optical signal issue was most likely a damaged optical fiber line due to excessive force. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number . d10 concomitant medical products and therapy dates updated. g1 reporting contact fax number updated.
28481 | 1220648-2024-09341 | 19034167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034167&pc=OZD | 09/18/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 09/18/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023225691 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 56-year-old male was implanted with the impella cp device for mechanical circulatory support. the patient went into ventricular tachycardia during removal of the impella sheath. the patient was shocked, drips were administered and a temporary pacer was placed for treatment. due to the circumstances the patient was escalated to the impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
28482 | 1220648-2024-09342 | 19034255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034255&pc=OZD | 09/27/2022 | 03/11/2025 |  | 04/03/2024 |  |  | 09/27/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212549 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was a repeat echo for repositioning vs explant due to continuous suction. swan was placed with appropriate co/ci/svo2. the impella was successfully explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28483 | 1220648-2024-09343 | 19034357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034357&pc=OZD | 09/27/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 09/27/2022 | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced oozing at the groin site and hematuria was noted. it was not noted how the oozing and hematuria were resolved, however, the patient was successfully weaned and the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28484 | 1220648-2024-09344 | 19034374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034374&pc=OZD | 12/01/2023 | 10/27/2025 |  | 04/03/2024 |  |  | 12/06/2023 | 04/03/2024 |  |  | 10/31/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408655 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 77-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the physician pulled the impella back to 3.6 centimeters in the intensive care unit due to one event of ventricle tachycardia. the pump remained on support till planned removal in the operating room during the ventricular septal defect surgery. the patient is stable post issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28485 | 1220648-2024-09345 | 19034448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034448&pc=OZD | 09/27/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 09/27/2022 | 04/03/2024 |  |  | 01/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187149 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was continuous suction (low flow). swan was placed with appropriate co/ci/svo2. the patient was weaned and impella successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28486 | 1220648-2024-09346 | 19034564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034564&pc=OZD | 03/20/2024 | 04/10/2024 |  | 04/03/2024 |  |  | 03/20/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153939 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported that the automated impella controller (aic) displayed a red battery failure alarm. the battery was at 0%. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. batteries were confirmed to be unable to power the console and could not charge. the cause of the aic issue was a depleted battery. the cause of the depletion is unknown since logs do not go further back in time.
28487 | 1220648-2024-09347 | 19034592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19034592&pc=OZD | 06/18/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/18/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was bleeding at the groin access site. manual pressure and the use of femostop controlled the bleeding. heparin was also withheld. the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28488 | 1220648-2024-09349 | 19035041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035041&pc=OZD | 11/02/2022 | 03/12/2025 |  | 04/03/2024 |  |  | 11/02/2022 | 04/03/2024 |  |  | 03/23/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127127 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported the patient experienced oozing at the access site and the pump was placed too deep. to resolve the access site oozing, the clinical team recommended more forward tension sutures. the pump was found on echogram to be positioned at 4.4cm, while urinary output was clear yellow and waveforms looked correct on the console, so it was not repositioned. the impella was left in place at power setting p-7, 3.3lpm. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28489 | 1220648-2024-09348 | 19035096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035096&pc=OZD | 06/29/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 06/29/2022 | 04/03/2024 |  |  | 12/31/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171661 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that hemolysis was confirmed from chronic diastolic heart failure (cdhf) drainage when the auxiliary level was changed from p-4 to p-6 during icu management. after that, when it was lowered to p-4 and hemolysis improved. additionally, the patient experienced access site bleeding; a hematoma was observed. a blood transfusion was required to resolve the problem. the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28490 | 1220648-2024-09350 | 19035169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035169&pc=OZD | 08/10/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 08/10/2022 | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023214018 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the device was weaned and explanted, due to the patient having limb ischemia. once the pump was explanted, the patient improved with no needed inotropes and had a successful closure with a stent. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28491 | 1220648-2024-09351 | 19035199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035199&pc=OZD | 03/20/2024 | 05/06/2024 |  | 04/03/2024 |  |  | 03/20/2024 | 04/03/2024 |  |  | 09/30/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398242 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Difficult to Advance | [6 empty] | Aortic Valve Stenosis | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that difficulty was encountered during attempted delivery of an impella 5.5 pump. during imaging, it was observed that the patient had critical aortic stenosis. after multiple unsuccessful attempts were made to cross the valve, the decision was made to transfer the patient for better imaging. however, upon arrival, the patient coded and was placed on extracorporeal membrane oxygenation for emergency support. the implant was subsequently aborted. | [5 empty] | The investigation into the delivery issues and the vf/vt/af/at has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the delivery issue was most likely due to patient condition since the patient had critical aortic stenosis. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28492 | 1220648-2024-09352 | 19035305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035305&pc=OZD | 11/04/2022 | 03/12/2025 |  | 04/03/2024 |  |  | 11/04/2022 | 04/03/2024 |  |  | 04/03/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221727 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The complainant reported a 70-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the impella insertion, the distal left main artery and proximal left circumflex artery were ballooned. pulsatility was lost during the balloon inflation. the patient remained stable despite loss of pulsatility during the stent deployment. the impella sheath was peeled away and the reposition sheath was advanced. towards the end of support, the patient had amber/red urine. the pump was removed shortly after, and it is unknown if the signs of hemolysis improved. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28493 | 1220648-2024-09353 | 19035572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035572&pc=OZD | 09/07/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 09/07/2022 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288837 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 45-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a controller error alarm. at the time of the alarm, it was observed that the placement signal was not showing up on the display. the unit was removed and replaced by a new aic. there was no harm to the patient. | [5 empty] | The investigation has been completed. the device was returned. console logs from the reported day of event were consistent with the complaint and showed that at some point during case start, optical bench data acquisition was stopped which resulted in controller error alarm and perceived placement signal issues (the two reported failure modes are directly related). this was followed by malfunction of one of pump detection mechanisms that relies on signal-to-noise ratio values, and in order to shut down the console, the power had to be withdrawn. the automated impella controller (aic) was tested and the issue was not reproduced, the console's optical system operated within specification. at that time, the root cause could not be determined. it was recommended to replace the optical bench as a precaution prior to returning the console back to customer. further analysis of console data, performed later by the software team, revealed a software defect, where a process erroneously stops optical bench data acquisition. this issue manifests if pump is unplugged unexpectedly at a certain time during case start, without prior prompt from the case start wizard. the root cause for the aic issue was a software issue.
28494 | 1220648-2024-09354 | 19035575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035575&pc=OZD | 08/29/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 08/29/2022 | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023224179 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 48-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced an ischemic stroke during implantation and dropped into a coma. a cerebral computed tomography (ct) scan showed no major occlusion in the major cerebral vessel; however, an infarction was found. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | To make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
28495 | 1220648-2024-09355 | 19035817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035817&pc=OZD | 09/03/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 09/03/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023225689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2022 |  |  |  |  | The user facility reported a 21-year-old female in acute myocardial infarction/ cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, thrombosis developed in the left main coronary artery which was aspirated with penumbra. the patient had lost pulses in both legs and a ct scan of both legs was a pending decision. the device's catheter was replaced with a new catheter, due to the pump was positioned in the aorta. procedure outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28496 | 1220648-2024-09356 | 19035869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19035869&pc=OZD | 09/09/2023 | 07/31/2024 |  | 04/03/2024 |  |  | 09/09/2023 | 04/03/2024 |  |  | 05/31/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373531 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2023 |  |  |  |  | The user facility reported a 74-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella in aorta alarmed and measured just slightly across the valve. repositioning was suggested but the physician did not want to do it due to non sterility of the impella catheter. the pump was shallow. impella in aorta alarm continued with pulsatile motor current. impella position was verified but not removed. the patient is stable post issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow and the positioning issue has been completed. the impella was not returned for evaluation. according to the clinical, "impella in aorta" alarms were reported and the pump was measured just slightly across the valve. data logs were reviewed which did not find any impella in aorta alarms, but only "position unknown." the physician did not want to reposition due to non-sterility of the catheter and imaging was used to confirm the pump was in good position but data logs show there were low flows under 1 l/min for the duration of the case along with persistent suction alarms. additionally, pulsatile motor current was reported. there was a trend in the placement signal. the most likely cause of the low flow was patient condition related. added- d6a/d6b f6/f8 should have been left blank in the initial report.
28497 | 1220648-2024-09358 | 19036792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19036792&pc=OZD | 11/11/2022 | 03/10/2025 |  | 04/03/2024 |  |  | 11/11/2022 | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023219334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported repositioning of the device occurred, due to the urine output showed pink-red and hematuria was present which cleared up. the patient had a hematoma extending down to the scrotum; however, it was stabilized. the patient was stable and successfully weaned off the device. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma." this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding and positioning issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
28498 | 1220648-2024-09357 | 19036796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19036796&pc=OZD | 11/02/2022 | 04/03/2024 |  | 04/03/2024 |  |  | 11/02/2022 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274736 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 60-year-old male was implanted with an impella cp and was escalated an impella 5.5. the 5.5 was successfully placed and then automated impella controller (aic) alarms occurred that prompted the team to troubleshoot. the impella stopped, air was in the purge, and the impella stopped retrograde flow. in an attempt to resolve the issue, the console was reportedly turned off and then back on, but the alarms reportedly persisted. the console was then exchanged for a backup. there were no adverse events. instructions for use were followed and the aic was replaced. no lasting harm or injury to the patient was reported. | [5 empty] | The investigation has been completed. the product was returned. console logs from the reported day of event were mostly consistent with the complaint and confirmed presence of alarms. however, they do not support the report of alarms being unresolved after automated impella controller (aic) shutdown. in fact, the aic shutdown could not be performed due to false pump detection - console was powered down after alternating current (ac) power was disconnected and batteries were disengaged. after the aic powered back up, the original alarms were resolved, and only ¿unexpected controller shutdown¿ alarm was shown. further analysis of the console logs revealed loss of data from the optical bench that was caused by a combination of a software (sw) defect that presented itself after the pump was unplugged during case start, and a most likely malfunction of 4-in-1 mis-reading signals. because the data from the optical bench was not being received/processed, pump detection mechanism malfunctioned, and after the pump was disconnected from the console, the decrease in motor current (to 0ma) was misinterpreted as ¿pump stop¿. the console was tested, but the issue could not be reproduced. inspection of the optical system revealed minor damage to the optical pump connector fiber, that was unrelated to this failure mode. the root cause of the aic issue was a software issue.
28499 | 1220648-2024-09359 | 19037065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19037065&pc=OZD | 11/11/2022 | 03/10/2025 |  | 04/03/2024 |  |  | 11/11/2022 | 04/03/2024 |  |  | 07/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023255789 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 79-year-old female in post cardiotomy cardiogenic shock low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was reported after the implantation of the device, the patient experienced bleeding at the access site and the heparin in the purge solution was replaced with sodium bicarbonate and heparin-free. however, venous bleeding and cardiac tamponade occurred and required for the patient's chest to be reopened twice for stabilization. due to a sudden drop in the platelet count, heparin-induced thrombocytopenia (hit) was suspected. therefore, the device was successfully weaned based on the judgement that the left ventricular support was currently sufficient. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28500 | 1220648-2024-09360 | 19037413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19037413&pc=OZD | 03/24/2024 | 05/08/2024 |  | 04/03/2024 |  |  | 03/24/2024 | 04/03/2024 |  |  | 07/31/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 381 PUMP SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024432507 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2024 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, concerns were had regarding limb ischemia and compartment syndrome. bilateral fasciotomies, a thrombectomy, and peripheral stenting were performed in the leg in response to the condition. the pump was removed and the patient was escalated to impella 5.5 support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
28501 | 1220648-2024-09361 | 19037633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19037633&pc=OZD | 03/25/2024 | 05/16/2024 |  | 04/03/2024 |  |  | 03/25/2024 | 04/03/2024 |  |  |  |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025456066 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the same day of impella device insertion, the physician brought the patient back to the catheterization lab and removed the device due to suction alarms. further information noted that the impella device was repositioned multiple times with echocardiogram at bedside. the right ventricle was evaluated, and no issue was indicated. intravenous fluid (ivf) bolus and albumin were administered. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for analysis, and no issue found with the pump. however, according to the data logs review, many positioning alarms in aorta and suction alarms occurred throughout the case. the cause of the low pump flow was most likely improper positioning due to improper technique. d6a, d6b, d9. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28502 | 1220648-2024-09362 | 19037815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19037815&pc=OZD | 03/24/2024 | 05/15/2024 |  | 04/03/2024 |  |  | 03/25/2024 | 04/03/2024 |  |  | 08/31/2025 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 381 PUMP SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2025443757 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Insufficient Flow or Under Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 41-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the pump was only able to run at p1 due to persistent suction above p1. additionally, the patient had a computerized tomography scan and the catheter was pulled back in the aorta. the impella was replaced and no patient harm was reported. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. data logs were reviewed, and suction alarms observed with low flow after pressure level dropped. no motor current spike or other abnormalities were found. the cause of the low pump flow issue was not determined since product was not returned and due to inconclusive evidence per log review. the cause of the positioning issue was most likely use related per clinical. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28503 | 1220648-2024-09363 | 19038184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038184&pc=OZD | 11/24/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 11/24/2022 | 04/03/2024 |  |  | 08/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265744 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | Pumping Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 71-year-old female patient was implanted with an impella cp for mechanical circulatory support. during support, it was reported that the physician was to restart heparin in the purge once bleeding subsided. in addition, disseminated intravascular coagulation (dic) panel for suspected heparin induced thrombocytopenia (hit) was noted. patient's hemoglobin dropped to 5, therefore, blood was given to the patient. there were also concerns with device suction and alarm impella in aorta. the patient was successfully weaned and the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury .this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28504 | 1220648-2024-09364 | 19038316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038316&pc=OZD | 11/25/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 11/25/2022 | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 50-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced some access site bleeding. the patient was given blood transfusions, and it stopped the volume of bleeding. suction alarms sounded, however, were resolved once the bleeding stopped. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28505 | 1220648-2024-09365 | 19038359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038359&pc=OZD | 11/25/2022 | 04/04/2025 |  | 04/03/2024 |  |  | 11/25/2022 | 04/03/2024 |  |  | 05/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231178 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 64-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient was not perfusing distal to the impella site. the decision was made to leave the device in for unloading and slowly wean and explant the following day. during the explant procedure, when the physician removed the peel away sheath and advanced the repositioning sheath, the patient¿s iliac was noted to still be occluded so an external bypass was performed to ensure perfusion to the right leg. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28506 | 1220648-2024-09366 | 19038400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038400&pc=OZD | 11/26/2022 | 04/03/2025 |  | 04/03/2024 |  |  | 11/26/2022 | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236593 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Tachycardia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2022 |  |  |  |  | The user facility reported a 35-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced ventricular tachycardia (vt). positioning was difficult to view on trans-thoracic echo and therefore trans-thoracic echo probe was used. left ventricle (lv) measurement under trans-thoracic echo appeared to be 5 cm and the doctor was able to reposition to 3.8 cm. after the impella pulled into the descending aorta, the patient was noted to have severe gastro intestinal bleeding. the impella was removed and the patient stabilized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28507 | 1220648-2024-09367 | 19038546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038546&pc=OZD | 10/07/2022 | 10/06/2025 |  | 04/03/2024 |  |  |  | 04/03/2024 |  |  | 06/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248575 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | Positioning Problem |  |  |  |  | Cardiac Arrest | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 63-year-old patient implanted with an impella cp for mechanical circulatory support as a high-risk percutaneous coronary intervention. five minutes after device insertion, the patient sat straight up on the cath table causing the impella to go back into the aorta. the impella was taken out and re-inserted. the patient coded and cardiopulmonary resuscitation (cpr) was started. while pulling the impella back to the 14fr sheath, the patient raised his right knee into the air. after pulling the device out of the body it was noted that the device was sheared and the inlet/pigtail was still in the common iliac artery. the patient was then put on venoarterial extracorporeal membrane oxygenation (va ecmo). the vascular surgery team was brought in to help snare the inlet/pigtail but was unsuccessful. it was decided that once va ecmo was weaned the patient would go to the operating room (or) and a cutdown would be done to retrieve to inlet/pigtail. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28508 | 1220648-2024-09368 | 19038603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038603&pc=OZD | 11/03/2022 | 04/03/2024 |  | 04/03/2024 |  |  |  | 04/03/2024 |  |  | 04/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216059 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported a 73-year-old male patient implanted with an impella cp for mechanical circulatory support. while on support the patient experienced a cerebral hemorrhage. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28509 | 1220648-2024-09369 | 19038781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038781&pc=OZD | 11/16/2022 | 05/13/2025 |  | 04/03/2024 |  |  |  | 04/03/2024 |  |  | 08/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251394 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Migration or Expulsion of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 74-year-old male patient was implanted with an impella cp for mechanical circulatory support. it was reported that while on support the patient experienced hemolysis and there were positioning issues. standard troubleshooting did not resolve the hemolysis and the pump accidentally slipped into the aorta. as the patient was stable, the pump was able to be removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction section h6 : medical device problem code was inadvertently reported incorrectly in the initial report which has been corrected in this report.
28510 | 1220648-2024-09370 | 19038807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038807&pc=OZD | 06/07/2022 | 04/03/2025 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204866 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute st-elevated myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia (vt) that was resolved by placement onto veno-arterial (va) ecmo. additionally, the pump experienced positioning issues that were resolved by removal of the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28511 | 1220648-2024-09371 | 19038867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038867&pc=OZD | 06/07/2022 | 04/03/2025 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 11/30/2023 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a 72-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. access site bleeding was observed. as a result, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28512 | 1220648-2024-09372 | 19038887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038887&pc=OZD | 06/07/2022 | 04/03/2025 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 02/29/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202040 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a 70-year-old female noted as high risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. during the procedure, the patient went into ventricular tachycardia (vt) requiring 5 shocks. the patient had no positive pedal pulses via doppler on the left leg. the patient was also bleeding from the site. heparin was lowered and manual pressure held. bleeding eventually stopped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition, any concomitant medication , vascular size or variations thereof, detailed description of the symptoms experienced by the patient , physical examination findings, including pulse assessment and visual examination of the affected limb , diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography , laboratory test results, including blood tests for markers of inflammation or clotting disorders , any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension , patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28513 | 1220648-2024-09373 | 19038902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038902&pc=OZD | 06/07/2022 | 04/03/2025 |  | 04/03/2024 |  |  | 06/07/2022 | 04/03/2024 |  |  | 03/31/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209597 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the pump stopped working. therefore, device was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28514 | 1220648-2024-09374 | 19038935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19038935&pc=OZD | 11/08/2023 | 04/30/2024 |  | 04/03/2024 |  |  | 11/08/2023 | 04/03/2024 |  |  | 11/30/2024 |  | 04/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295149 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 72-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support the patient developed oozing at the impella access site. manual pressure was applied, angle match addressed and repositioned, but oozing continued. fem stop was placed above access site to achieve hemostasis. additionally, the aic began having ¿purge flow blocked¿ alarms. multiple trouble shooting attempts made, but the purge flow continued to be greater than 1100 and the pump stopped shortly after and was subsequently removed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the purge leak ¿ cassette, the pump stop, and the access site bleeding ¿ major issues has been completed since the original report was filed. the pump was returned and evaluated. there was no leak to be found. the root cause of the purge leak - cassette is most likely due to use. there were no abnormalities seen from the introducer sheath or reposition unit. the root cause of the access site bleeding cannot be determined due to limited clinical information provided. there was biomaterial found in the purge gap reproducing the high purge pressure and on the magnet of the pump. the root cause of the pump stop that led to the high purge pressure is due to biomaterial. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to conform with updated procedures, section d6 and h10 were revised with updated information.
28515 | 1220648-2024-09173 | 19018124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018124&pc=OZD | 06/20/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/20/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209843 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a loss pulse in right foot overnight. echocardiogram (echo) showed heavily calcified arteries and small arteries the down the leg. no information was provided on any intervention performed for the limb ischemia. it was also reported the patient had an episode of atrial fibrillation while on impella cp support. the patient was administered amiodarone, and the pump position was confirmed under transesophageal echocardiography. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28516 | 1220648-2024-09174 | 19018157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018157&pc=OZD | 06/22/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 06/22/2022 | 04/02/2024 |  |  | 11/30/2023 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149102 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had signs of hemolysis. heparin was discontinued. the impella position was assessed via transthoracic echocardiography (tte), indicating a somewhat suboptimal placement with the inlet appearing slightly close to the mitral valve, but the physician optimized the positioning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28517 | 1220648-2024-09175 | 19018296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018296&pc=OZD | 09/24/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 09/24/2022 | 04/02/2024 |  |  | 05/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230305 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 68-year-old female patient experienced v-fib arrest requiring cv. decision was made to place impella cp. it was reported that during support, the patient required 3 shocks due to ventricular tachycardia. impella was on p7 with good flows and position, however there was oozing at the access site, but good distal pulses. the impella was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28518 | 1220648-2024-09176 | 19018598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018598&pc=OZD | 10/13/2023 | 06/21/2024 |  | 04/02/2024 |  |  | 10/13/2023 | 04/02/2024 |  |  | 08/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed a hematoma at the impella cp access site. the reposition sheath was re-sutured to match entry angle and a femstop was placed with light pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the root cause of access site bleeding was determined to be use issue because the bleeding was resolved by re-suturing and matching the entry angle. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 code 4114 (device not returned) was inadvertently omitted from the previously submitted report and has been added to this report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
28519 | 1220648-2024-09177 | 19018741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018741&pc=OZD | 10/13/2023 | 05/14/2024 |  | 04/02/2024 |  |  | 10/13/2023 | 04/02/2024 |  |  | 04/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355586 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 86-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient on impella support experienced bleeding after the patient was rolled. forward tension sutures were replaced and the site was angle matched. additionally, heparin was not in the purge but bicarb was. groin site stable without bleeding. additionally, the 14 french sheath split and repositioning sheath was advanced. | [5 empty] | The investigation into the access site bleeding ¿ major and the introducer damage ¿ mechanical issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical information, shortly after beginning impella cp support the groin site started bleeding after the patient was rolled. the bleed was effectively treated, and the device remained in the patient after the bleed was resolved. the most likely cause of the major access site bleeding was patient movement. potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28520 | 1220648-2024-09178 | 19018840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018840&pc=OZD | 09/26/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 09/26/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 65-year-old male on impella cp support for an unplanned higher-risk percutaneous coronary intervention (hrpci). the physician made multiple attempts made to wire the left circumflex artery (lcx). unfortunately, the wire was diving under the stent. after failing to cross the lcx stent, the physician consulted computed tomography (ct) scan. the decision was made to get patient more in depth ct scan consult. the impella cp was removed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28521 | 1220648-2024-09179 | 19018849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018849&pc=OZD | 03/11/2024 | 04/02/2024 |  | 04/02/2024 |  |  | 03/11/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1216752 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Component Missing | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had cp support was being prepped when the aic (automated impella controller) was found to be missing the purge disc/disc retainer. the aic was removed and replaced per the ifu (instructions for use) recommendations. the new aic was used for support. | [5 empty] | The impella device was received from the customer on 18-mar-2024 data review: logs are irrelevant to this complaint as the clinical staff only reported that the consoles purge disc retainer was found to be broken during prep. device analysis: console was tested after arriving in fs. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken (see attached image). before sending this console back to the customer, fs was instructed to replace the blue purge disc retainer [3001195], purge disc detect flag [3001198], and purge flag spring [3001196], validate sensor accuracy via section 12 of the pm procedure [0042-7309], and perform a full functional check on the console and optical system [0042-7316].
28522 | 1220648-2024-09180 | 19018874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018874&pc=OZD | 06/06/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/06/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200191 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, patient access site began bleeding and was given blood products. still oozing persist and heparin was held from the purge system. the patient¿s bleeding stopped overnight. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28523 | 1220648-2024-09182 | 19018929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018929&pc=OZD | 09/02/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 09/02/2022 | 04/02/2024 |  |  | 05/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228090 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device presented active suction alarm. it was reported that the device appeared to be either out of position or obstructed potentially. the pump had constant suction, flat motor current and very low flows. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28524 | 1220648-2024-09181 | 19018937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018937&pc=OZD | 03/13/2024 | 04/30/2024 |  | 04/02/2024 |  |  | 03/13/2024 | 04/02/2024 |  |  | 10/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024417739 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | Ischemia | Tachycardia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support for emergent circulation, airway, breathing. while on support, the patient did not have dopplerable pulses. the patient was taken to the cath lab and distal perfusion catheter was placed and distal flow was restored. additionally, the cp measured 2 cm across the valve. pump on echo - echocardiogram (ultrasound)- looked caught in papillary muscle and the doctor wanted to be sure that the impella was not causing the patient's ventricular tachycardia (vt). patient shocked multiple times to convert out of vt. impella eventually removed to see if removal of pump mitigated the vt, shocked out of vt after pump removed and vt started again. the doctor discussed possibly placing an impella 5.5, but thought due to rhythm issues and possible left ventricle irritation to just do ecmo at this time. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. as all the necessary information was not provided, the root cause of the limb ischemia and the vf/vt/af/at was not determined. potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ to conform with updated procedures, sections d6 and h10 have been revised with updated information.
28525 | 1220648-2024-09183 | 19018948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19018948&pc=OZD | 09/29/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 09/29/2022 | 04/02/2024 |  |  | 05/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229795 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Stenosis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 88-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, distal wire perforation occurred. user facility attempted coiling and ended up injecting fat for embolization. an emergency pericardiocentesis (-300ml) procedure was needed. patient was hemodynamically stable at the end of the case. perforation required covered stent and/or balloon tamponade. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28526 | 1220648-2024-09184 | 19019020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019020&pc=PYX | 03/14/2024 | 05/01/2024 |  | 04/02/2024 |  |  | 03/14/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025458202 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a thrombus was discovered during impella rp support. coinciding with the noted thrombus, the optical signal became unreliable and flows in the pump decreased. a tpa protocol was initiated to treat the thrombus. support with the pump remained ongoing following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical issue and thrombus has been completed. data logs were reviewed which showed that there is a motor current spike followed by snr drops to 0, contrast decreases, and lamp and gain increase. light remains relatively stable. 5.5 hours into the case, data logs show a potential ingestion event (motor current spike, slight ps upwards shift, and speed deviation) followed by a drop in flows which eventually recover. the device was not returned for evaluation. the root cause of the optical signal issue was not determined since product was not returned. the root cause of the low flow was most likely biomaterial. the instructions for use for impella rp with smartassist system-section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ added- h6 impact code 4644, h6 component code 925 added- d6a/d6b. d4-updated model number. added h6 med dev prob code 2917.
28527 | 1220648-2024-09185 | 19019083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019083&pc=OZD | 09/08/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 09/08/2022 | 04/02/2024 |  |  | 04/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 75-year-old female was implanted with an impella cp device for mechanical circulatory support. soon after implant, return of spontaneous circulation (rosc) was attained. device flows were lower than expected with suction alarms. a kink was seen in the impella cannula before the outlet. the health care professional (hcp) was advised to replace the device in lieu of a peel away introducer on the left side. the arterial sheath on the left was upsized to an abiomed fourteen (14) french introducer. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28528 | 1220648-2024-09186 | 19019108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019108&pc=OZD | 09/30/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 09/30/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236847 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2022 |  |  |  |  | The user facility reported a 86-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, patient developed tamponade. therefore, the impella was turned off and after a minute of no pulse, the patient showed signs of a pulsatile a-line. impella was quickly turned back on and up to p-9. patient remained stable and surgeon performed sternotomy with washout. the clot was removed, and bleeding was repaired around the heart. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28529 | 1220648-2024-09187 | 19019112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019112&pc=OZD | 03/18/2024 | 05/04/2024 |  | 04/02/2024 |  |  | 03/18/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426050 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed to support a patient admitted with multiple cardiac and systemic comorbidities. the team had the pump supporting for over 18 days when the pump came out of position and was unable to be redelivered across the valve. in the attempts the pump kinked and so was replaced. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
28530 | 1220648-2024-09188 | 19019130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019130&pc=PYX | 03/26/2024 | 05/23/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  | 08/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404269 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a rp placed to support a patient in rv (right ventricle) failure s/p aaa (abdominal aortic aneurysm) repair and cabg (coronary artery bypass grafting). the rp was placed but explanted after 20hours due to hemolysis concerns. the team instead placed iabp (intra-aortic balloon pump) and ecmo (extracorporeal membrane oxygenation). the hemolysis was diagnosed by labs hemolyzing, urine color change and rising ldh (lactate dehydrogenase) levels. | [5 empty] | The investigation has been completed. based on the clinical details, it is unclear how the hemolysis was resolved. the complaint device was not returned for investigation. no data logs were available to review. insufficient clinical data was provided. the cause of the hemolysis issue could not be determined. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.4 revised pma/510(k) number as previously entered incorrectly. h.3 evaluation summary attached was previously marked and should not have been. h.6 health effect - impact code was inadvertently left off the previously submitted report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
28531 | 1220648-2024-09189 | 19019152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019152&pc=OZD | 03/19/2024 | 05/06/2024 |  | 04/02/2024 |  |  | 03/19/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435798 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The complainant reported a patient undergoing aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient did not receive enough blood flow and the patient¿s left leg was cool to the touch with intermittent doppler pulse. an antegrade sheath was in place in an attempt to assist with the condition, but the decision was ultimately made to remove the pump. initial closure of the site was unsuccessful and the patient was taken to the or for a femoral cutdown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28532 | 1220648-2024-09191 | 19019164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019164&pc=OZD | 03/26/2024 | 05/16/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  | 06/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427229 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 75-year-old male of unknown ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was being defibrillated upon arrival and due to patient vasculature and status the initial attempt to implant the impella cp for support was unsuccessful. the physician was able to successfully place a new impella cp pump. the patient remained in critical condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely use related, the impella was not able to cross valve but tip of impella pig tail/inlet kinked as per the clinical description. impella cp with smartassist. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ d6a, d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28533 | 1220648-2024-09190 | 19019169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019169&pc=OZD | 07/04/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 07/04/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204923 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a loss of pulses. the physician decided to explant the device as a result of the limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28534 | 1220648-2024-09192 | 19019190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019190&pc=OZD | 03/05/2024 | 09/19/2025 |  | 04/02/2024 |  |  | 03/05/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427505 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/09/2024 |  |  |  |  | The user facility reported that a 35-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that pump support was ongoing when there was a right axillary hematoma, but was resolving. it was reported that a washout was planned. the support was for 34 days which is beyond indication. | The complainant also reported that the pump additionally lost optical signal but remained on for support until explant at the time of open heart transplant. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for a visual inspection. data log analysis confirmed all 5s parameters were stable. however, the placement signal drop, plateau, and downward trend were consistent with sensor drift. the cause of the optical signal issue was sensor wear. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-09192. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
28535 | 1220648-2024-09193 | 19019208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019208&pc=OZD | 03/26/2024 | 06/17/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  | 12/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABOIMED, INC |  |  | US |  | 2024300677 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Improper or Incorrect Procedure or Method | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The patient expired. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device with the patient's outcome. | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was discovered that due to the user failing to tighten the fixation ring, the impella cp became dislodge from the left ventricle. a second pump was placed, and support continued. |  |  |  |  | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. based on the clinical description, the root cause of the positioning issue was most likely use related. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 last name was revised as previously entered incorrectly. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 codes 2017 and 4114 were reported incorrectly on the previously submitted report. new codes has been added. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28536 | 1220648-2024-09194 | 19019212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019212&pc=OZD | 03/21/2024 | 05/06/2024 |  | 04/02/2024 |  |  | 03/21/2024 | 04/02/2024 |  |  | 09/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024409946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp (and ecmo - extracorporeal membrane oxygenation-) placed for support of a cgs/ami (cardiogenic shock/acute myocardial infarction) patient with vt (ventricular tachycardia) and aicd (automatic implantable cardioverter defibrillator). the pump supported for 3 days but during that time hemolysis was present. pfhg (plasma-free hemoglobin) lab values were elevated and the patient was placed on crrt (continuous renal replacement therapy). the team attempted to reposition, but the pump was caught under the papillary. the cp was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the hemolysis and the positioning issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the cause of the hemolysis issue was most likely improper positioning due to improper technique. the cause of the positioning issue most likely was use related due to improper technique. adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28537 | 1220648-2024-09195 | 19019261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019261&pc=OZD | 03/27/2024 | 05/01/2024 |  | 04/02/2024 |  |  | 03/27/2024 | 04/02/2024 |  |  | 09/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025444919 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to the patient going into arrest and when they arrived at the hospital they coded. while on support, the patient had no pulsatility on the impella side. additionally, the patient had access site bleeding and hematoma. one unit of packed red blood cells were provided, manual pressure was held, and the impella was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time of bleed was 366 which is above the range limit. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28538 | 1220648-2024-09196 | 19019267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019267&pc=OZD | 09/13/2022 | 04/07/2025 |  | 04/02/2024 |  |  | 09/13/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023236591 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient developed scant oozing/bleeding at the access site, which was discovered when the patient was turned overnight. the care team were walked through impella site re-dressing. the site was found to be otherwise clean and intact with doppler left pedal pulses. the patient was also receiving platelets (blood products) due to thrombocytopenia and placement of temporary dialysis catheter. heparin induced thrombocytopenia antibody was positive and heparin was removed from purge solution and replaced with 25u/ml of bicarbonate, argatroban drip was placed for anticoagulation. the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28539 | 1220648-2024-09197 | 19019275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019275&pc=OZD | 03/21/2024 | 05/01/2024 |  | 04/02/2024 |  |  | 03/21/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430615 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the transfer from community to the tertiary center the team observed an access site ooze. the team remedied with pressure and angle matching however, the day after the team gave two units of red blood cells (rbc)s during percutaneous coronary intervention. the team later achieved hemostasis with one perclose and three units of rbcs. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use as the anticoagulation management since the coagulation was supratherapeutic and since the hemostasis was achieved by adding a perclose. added-d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28540 | 1220648-2024-09198 | 19019319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019319&pc=OZD | 03/22/2024 | 05/07/2024 |  | 04/02/2024 |  |  | 03/22/2024 | 04/02/2024 |  |  | 04/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348638 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. the flow was checked distal to sheath and none were noted. the patient had no pulses in the right lower extremity. the decision was made to place a fem-fem bypass which was successfully performed. after, the pulses were dopplerable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
28541 | 1220648-2024-09199 | 19019331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019331&pc=OZD | 03/27/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/27/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025439480 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | It was further reported that the impella pump was stuck in the papillary since implant. there had been no cpr performed and no patient anatomy or movement issues observed. repositioning attempts were made with no success in resolving position issue. also, patient had ongoing hemolysis since implant necessitating pump removal. | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the cp was placed despite knowledge of the delivery would cause limb ischemia. the team left in the 14 french (fr) sheath so that distal perfusion could be placed. the limb perfusion was placed, and sheath flushed every hour. the patient remained on support. |  |  |  |  | The investigation for the reported limb ischemia, positioning issues, and hemolysis has been completed. no product was returned for a visual inspection. data log was reviewed, suction alarms and few impella in aorta alarms observed. according to the clinical, on the first day of impella cp support the patient began to develop signs of ischemia in both distal limbs. the physician noted that they had been aware the pump would be occlusive to the leg before they placed it. an attempt was then made to resolve the limb ischemia successfully. pump was stuck in the papillary muscle of the mitral apparatus since implant and repositioning was unsuccessful. the cause of the positioning issues was most likely use related per clinical. the cause of the hemolysis issue was most likely improper positioning of the device. the root cause of the limb ischemia could not determined without additional clinical details. the instructions for use for the impella cp with smart assist system in section: potential adverse events (united states) states, "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". added: b1: checked product problem, b5: added additional narrative., h6 clinical code 1886, h6 med dev prob code 3009 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28542 | 1220648-2024-09201 | 19019379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019379&pc=OZD | 03/28/2024 | 09/26/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025458635 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/28/2024 |  |  |  |  | The user facility reported an 82-year-old male in st elevated myocardial infarction and respiratory distress was implanted with an impella cp device for mechanical circulatory support. it was reported that shortly after insertion there was a steady stream of blood oozing from the access site and from the patient¿s mouth. the observation of hematoma forming at access site was also noted. manual pressure was held which seemed to slow the oozing, and pressure dressing was applied, however blood was appearing under the dressings. manual pressure was resumed and held for over 2 hours. during this time the experienced rhythm issues and a drop in mean arterial pressure. tranexamic acid (antifibrinolytic agent) and fresh frozen plasma was administered. the patient began shooting blood form the impella site. the staff mentioned the pump needed to be removed or replaced to control the bleeding. the family decided to make the patient a do not resuscitate (dnr) and the patient expired after 4.08 hours of support. it is unknown if the impella contributed to the patient's expiration. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections h1 and h6. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding was most likely due to patient condition since the patient had bleeding from multiple sites including both the fem sites and mouth. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 & h10 were revised with updated information.
28543 | 1220648-2024-09202 | 19019388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019388&pc=OZD | 03/28/2024 | 04/29/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1251270 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 58-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when pump being prepped, placement signal not reliable alarmed. after troubleshooting alarm resolved so decision made to continue inserting current cp. once inserted optimal flows never achieved with constant ventricle alarm when impella was in proper place. after 5 minutes spent repositioning and making adjustments without success, decision made to switch pump and automated impella controller (aic). no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - loss of opm data has been completed since the initial report was submitted. the console was returned and tested. inspection of the console revealed contamination of the optical pump connector fiber. the reported placement signal issues were caused by defective optical bench, but contamination of the optical pump connector might have contributed to the symptoms. the cause of the issue was a defective component.
28544 | 1220648-2024-09203 | 19019422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019422&pc=OZD | 03/22/2024 | 05/23/2024 |  | 04/02/2024 |  |  | 03/23/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024434431 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Material Fragmentation | Detachment of Device or Device Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2024 |  |  |  |  | Additional information states that the patient expired on support of the impella cp. the patient had developed dic and bleeding could not be controlled from the removal issue. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. when removing the impella 5.5, the blood vessel diameter, which was 5.5 mm at the time of insertion, was smaller than that, so the 5.5 could not be removed. the shaft broke when it was pulled by the physician. removal of the impella from the femoral artery (fa) using a snare catheter was also attempted but was unsuccessful. the catheter shaft was held down with forceps and pulled, the catheter shaft came off from the cannula. part of the device remained in the patient's body. the parts that remained inside the patient¿s body are from the cannula to the tip of the inhaler. the motor part was removed. an impella cp was inserted to provide circulatory support. surgery was done to repair the vessel and blood products were given. it was further reported that bleeding occurred during the 5.5 removal and additional treatments were given due to the difficulty in removal. additional information received stating that the remaining impella 5.5 cannula part was surgically removed. the tip of the inlet was located at the origin of the brachiocephalic artery, and the distal end of the cannula was located proximal to the anastomosis of the subclavian artery. the blood vessel was used to anastomose and bypass the healthy blood vessels on the central and peripheral sides of the subclavian artery. the bleeding was caused by the cannula that was stuck and had torn the subclavian artery. although the subclavian artery was torn, there was no serious bleeding as it was covered with a cannula. after the cannula was removed, an artificial blood vessel was used to anastomose and bypass the healthy blood vessels on the central and peripheral sides of the subclavian artery. artificial blood vessel replacement surgery was performed on the damaged area. bleeding is treated with blood transfusions, but the exact number of units is unknown. |  |  |  |  | The investigation for the reported vascular damage and removal issues has been completed. no product was returned for investigation. field image confirms the torn cannula. the patient's vessel size had decreased since insertion and made removal difficult. there was bleeding and damage to the subclavian artery due to this and blood products had to be given and vascular repair. the root cause of the cannula detachment issue was most likely use related due to the pump requiring excessive force to be removed. the root cause of the vascular damage peripheral vessel issue was most likely patient condition related because the vessels had decreased in size since implant making removal difficult and excessive force was required, ultimately leading to vessel damage. the instructions for use referenced in the initial report is from ifu impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states), which now includes "(including ventricular perforation)¿ b2 is updated with box for death checked. b5-updated outcome h1-changed type of reportable event to death. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28545 | 1220648-2024-09204 | 19019437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019437&pc=OZD | 09/16/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 09/16/2022 | 04/02/2024 |  |  | 04/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218246 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding and was administered blood products. the patient also experienced fluctuating hemolysis. the patients lasix was turned off and lactate dehydrogenase was sent to monitor blood in urine. urinalysis confirmed red blood cells in urine while blood in urine was 2. the reported issue was continuously monitored while the patient was on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28546 | 1220648-2024-09200 | 19019439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019439&pc=OZD | 12/15/2023 | 09/25/2024 |  | 04/02/2024 |  |  | 12/15/2023 | 04/02/2024 |  |  | 03/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345305 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2023 |  |  |  |  | The user facility reported a 71-year-old male patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on support when access site bleeding was experienced. best practices for groin management were discussed. additionally, heparin was removed from the purge solution and replaced with bicarb due to bleeding issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the activated clotting time (act) value was 170 which is in the therapeutic range and root cause is not determined. added- h6 component code 925. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mdr mfr report: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank.
28547 | 1220648-2024-09205 | 19019497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019497&pc=PYX | 03/23/2024 | 05/15/2024 |  | 04/02/2024 |  |  | 03/23/2024 | 04/02/2024 |  |  | 09/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024435710 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2024 |  |  |  |  | The complainant reported a patient (age and race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support during a percutaneous coronary intervention (pci). after the pci procedure with impella rp support and during the removal of the 23fr sheath, the impella rp fell back into the patient's right ventricle. after about an hour of trying to put the impella back into place the physician elected to remove the impella rp. the patient was sent to icu for observation and recovery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. no product was returned. the positioning issue was a pulled-out type issue with the impella rp falling back into right ventricle during removal of the 23 french sheath. the cause of the positioning issue was most likely use related per clinical review. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 2199 and 3099 were reported incorrectly on the previously submitted report. new codes has been added to health effect - impact codes and medical device problem codes. code 3233 was added to investigation findings as it was inadvertently left off the previously submitted report.
28548 | 1220648-2024-09206 | 19019498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019498&pc=OZD | 06/06/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/06/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197634 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had access site bleeding. the patient was given five units of blood and dressing changed. the impella site is dry, and the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28549 | 1220648-2024-09207 | 19019530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019530&pc=OZD | 07/07/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 07/07/2022 | 04/02/2024 |  |  | 12/31/2023 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161747 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported after impella cp was explanted, the patient developed bleeding at the impella access site. manual pressure was applied and did not resolve the bleeding. a balloon was inserted from the contralateral side, and as the balloon expanded from within the blood vessel at the access site, pressure was applied to stop the bleeding. contrast imaging was used to confirm hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28550 | 1220648-2024-09208 | 19019541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019541&pc=OZD | 03/26/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438082 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during support, a purge pressure high alarm was noted on impella 5.5. in trouble shooting alarm the purge cassette was changed. the new purge disc was not detected although securely in place. after multiple attempts a new aic was turned on and the pump switched to the new aic. after switching to a new aic purge disc was detected but the patient continued to experienced purge pressure high alarms. decision was made to remove the impella. it was further reported that at the beginning of impella support there was a small hematoma that formed above the impella site and heparin was stopped. hematoma is controlled. the patient is stable. | [5 empty] | The investigation into the high purge pressure and the access site bleeding ¿ major issues has been completed since the original report was filed. the pump was returned for analysis. high purge pressure and purge blocked alarms were confirmed, which did not resolve during the case. during product review, kink was found along the catheter, biomaterial was found in the purge gap, the inflow cage, and above the impeller. the purge pressure did not reproduce after removing the when the biomaterial. the root cause of access site bleeding - major, was determined to be anticoagulation management related as reducing heparin resolved the issue. the root cause of high purge pressure was most likely biomaterial as pump ran successfully after clearing. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28551 | 1220648-2024-09209 | 19019582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019582&pc=OZD | 10/02/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 10/02/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2022 |  |  |  |  | The user facility reported 57-year-old patient was being placed onto impella cp support due to st-elevated myocardial infarction. while on support, the patient experienced access site bleeding that was resolved with antihypertensives and pressure dressing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28552 | 1220648-2024-09210 | 19019597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019597&pc=OZD | 03/26/2024 | 05/15/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443304 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support due to suspected takotsubo cardiomyopathy. after the insertion, peripheral runoff showed no distal flow. right femoral, antegrade sheath was placed for external bypass. additionally, the patient had access site bleeding. dressing was changed every hour, perclose sutures were tightened, surgical foam was used, and one unit of packed red blood cells were provided. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding issue was most likely use related, perclose sutures were tightened as per clinical description provided. d6a, d6b.
28553 | 1220648-2024-09211 | 19019636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019636&pc=OZD | 09/25/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 09/25/2022 | 04/02/2024 |  |  | 05/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228836 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2022 |  |  |  |  | The user facility reported a 49-year-old male on impella cp after arresting in er. while on support, patient presented signs of hypovolemic shock with hemoglobin noted 6.5. four units of blood products were received. impella cp flows of 2.2 achieved and still not tolerating over p4 due to suction alarms. the impella cp was repositioned several times via cath lab and bed side. echo showed very hyper dynamic ventricle and difficult positioning. during the icu groin was bleeding and a small hematoma was noted. the attending physician added a purse-string suture to help oozing. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28554 | 1220648-2024-09212 | 19019646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019646&pc=PYX | 03/24/2024 | 05/07/2024 |  | 04/02/2024 |  |  | 03/24/2024 | 04/02/2024 |  |  | 03/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024397286 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2024 |  |  |  |  | The user facility reported a 60-year-old male was implanted with an impella 5.5 as a bridge to transplant. after the patient was transplanted with left ventricle assist device, an impella rp flex was implanted for right heart support. the impella rp flex was removed after four days due to low flows. the patient was implanted with a centrimag right ventricle assist device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the pump was returned for evaluation. during inspection, the impeller was found to be clean. there were no kinks noted. the low flow was not reproduced while running in a water bucket at p-9. no other abnormalities were observed. data logs show that 78 hours into support there is a motor current spike, ps shift, and speed deviation while at p-9 with immediate drop in flows. this is indicative of ingestion. the biomaterial was most likely kicked out as it was continued to run in the patient. the root cause of the low pump flow issue was most likely biomaterial. the instructions for use for impella rp with smartassist system in section: using the automated impella controller with the impella catheter states "achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup".
28555 | 1220648-2024-09213 | 19019669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019669&pc=OZD | 08/28/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 08/28/2022 | 04/02/2024 |  |  | 04/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023224077 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a 53-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, continuous suction alarms were received. the pump was positioned across the atrioventricular and there was a kink noted within the cannula near the outflow cage which caused low flows. repositioning of the pump was attempted multiple times; however, the device could not be positioned properly. therefore, the decision was made to explant the device and implant a new impella cp which was prepped and successfully inserted without issues. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28556 | 1220648-2024-09214 | 19019727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019727&pc=OZD | 03/26/2024 | 05/13/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433309 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support the patient experienced high purge pressure without alarms. the cassette was changed without resolution, and tissue plasminogen activator was added to the purge to treat for thrombosis. additionally, a hematoma developed and heparin was subsequently withheld. impella was removed the next day and patient did not require pharmacological support, remains vented, and the presentation of compromised neurostatus. doctor will discuss next steps for plan of care with family. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported high purge pressure (hpp), thrombus, and access site bleed has been completed. the pump and data logs were returned and evaluated. there was biomaterial found wrapped around the impeller blades and the outflow cage. although no hpp alarms were seen during investigation, the purge pressure was seen rise up to 1000mmhg. the root cause of the high purge pressure issue was biomaterial found wrapped around the impeller blades and purge gap reducing the purge flow as a result increasing the purge pressure. the root cause of the access site bleeding issue was not determined since insufficient clinical information was provided. the instructions for use for impella cp with smartassist for use during cardiogenic shock and high risk cpi in section: using the automated impella controller with the impella rp with smartassist system catheter states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ added-d9 device return date, h6 med dev prob code 3012, h6 component code 925 in accordance with updated procedures, section d6 has been revised from the initial submission.
28557 | 1220648-2024-09215 | 19019752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019752&pc=OZD | 06/30/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 06/30/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213255 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 67-year-old male patient in acute myocardial infarction (ami)/ cardiogenic shock (cgs) was implanted with an impella cp device for mechanical circulatory support for a percutaneous coronary intervention (pci) to the left anterior descending artery (lad). it was reported while on impella cp support, the patient experienced bleeding at the impella access site and gastrointestinal bleeding. systemic heparin was discontinued to treat the bleeding. it was also reported while on support the patient experienced episodes of ventricular tachycardia and the patient was administered amiodarone. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28558 | 1220648-2024-09216 | 19019759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019759&pc=OZD | 10/06/2022 | 03/11/2025 |  | 04/02/2024 |  |  | 10/06/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248063 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and the staff attempted to place the device in the right axillary artery but was unsuccessful due to the size of the artery. the impella 5.5 was removed and replaced with an impella cp without issues. the patient remained on impella cp support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28559 | 1220648-2024-09217 | 19019825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019825&pc=OZD | 10/06/2022 | 10/10/2024 |  | 04/02/2024 |  |  | 10/06/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194972 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2022 |  |  |  |  | The complainant reported the patient was being placed onto impella cp support due to decompensation while awaiting impella 5.5 scheduled for the next day. while on support, the patient became bacteremic, which was resolved with antibiotics. additionally, the pump pulled into the aorta, which was resolved by advancing with echocardiogram guidance. the impella cp was explanted, and the patient escalated to the impella 5.5. the patient remained on impella 5.5 support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | Manufacturer device report 1220648-2024-09217 erroneously submitted. the reported complaint against infection was against the subsequently implanted 5.5, not the subject impella cp device of this record. positioning issues were resolved by repositioning with echo and, as such, do not constitute a complaint.
28560 | 1220648-2024-09218 | 19019835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019835&pc=OZD | 03/25/2024 | 05/01/2024 |  | 04/02/2024 |  |  | 03/25/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, bleeding at the impella access site occurred. manual pressure and another forward suture were added. hemoglobin dropped and one unit of red blood cells were given. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28561 | 1220648-2024-09219 | 19019879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019879&pc=OZD | 10/11/2022 | 03/11/2025 |  | 04/02/2024 |  |  | 10/11/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253018 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and the patient experienced a loss of distal pulses while using a doppler. it was also observed that the pump exhibited low flow. the physician decided to explant the device. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28562 | 1220648-2024-09220 | 19019895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19019895&pc=OZD | 06/06/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/06/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206126 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The complainant reported the patient was implanted with a impella cp support due to cardiac arrhythmia. while on support, the patient experienced slight access site bleeding. additionally, the pump presented suction alarms that were resolved with repositioning and fluids. planned escalation of patient from impella cp to impella 5.5. patient transferred from outside facility. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28563 | 1220648-2024-09226 | 19020115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020115&pc=OZD | 03/27/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/27/2024 | 04/02/2024 |  |  | 08/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268448 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Human-Device Interface Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that perclose done prior to impella cp dilation when impella initially implanted failed. post close x4, failed. 14fr sheath placed and surgical cutdown was done. hemostasis achieved. the patient is stable. | [5 empty] | The investigation for the reported removal issue has been completed. pump was returned for evaluation. upon visual inspection, no abnormalities were found with the pump or repositioning unit. clinical details noted that perclose x2 sutures were placed prior to cp dilation and insertion. after pump was explanted, perclose sutures failed, post close sutures were attempted and also failed. a 14fr sheath was placed, surgical cutdown was performed and hemostasis was achieved. no issues were found with leg under doppler. the root cause of the removal issue which led to the access site bleed was due to a use issue as the perclose sutures failed and no issue was found with the pump or repositioning sheath. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28564 | 1220648-2024-09227 | 19020187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020187&pc=OZD | 10/24/2022 | 03/11/2025 |  | 04/02/2024 |  |  | 10/24/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023231737 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | A japan complainant reported that the patient was placed on impella cp smart assist support and developed lower limb ischemia. the impella was explanted and replaced with an intra-aortic balloon pump (iabp) after closure surgery. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. (explant date unknown). the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28565 | 1220648-2024-09228 | 19020217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020217&pc=OZD | 03/27/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/27/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443359 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during priming the impella was attached to the aic (automated impella controller) when a purge system open alarm occurred. the purge fluid was not moving into the purge side arm. due to concern for the patient a new impella cp and aic was used successfully. the patient is stable. | [5 empty] | The investigation for the reported priming issue has been completed. "purge system open" alarm reported. it was observed that the purge fluid was not moving into the purge side arm. data logs were not returned. upon evaluation of the product, there were no visual abnormalities and the priming issue did not reproduce. there was no problem found as the root cause of the priming issue. added- d9 device return date corrected d1/d2 brand name, common device name d4-updated model number in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28566 | 1220648-2024-09229 | 19020223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020223&pc=OZD | 03/26/2024 | 04/29/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1341199 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility biomedical department reported that the automated impella controller (aic) had a battery failure. the aic was removed from service. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - battery comm. failure (yellow alarm) has been completed since the initial report was submitted. the console was returned and tested. battery failure issue was able to be reproduced. the root cause of this issue was battery internal cell imbalance.
28567 | 1220648-2024-09230 | 19020246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020246&pc=OZD | 09/10/2023 | 10/14/2024 |  | 04/02/2024 |  |  | 09/10/2023 | 04/02/2024 |  |  | 03/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341959 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2023 |  |  |  |  | The user facility reported a 77-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, the patient experienced a hematoma at the access site. manual pressure was applied and preclose was used. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the cause of the access site bleeding was patient movement related as the patient was restless and hematoma developed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added-d6a/d6b f6/f8 should have been left blank.
28568 | 1220648-2024-09231 | 19020276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020276&pc=OZD | 03/28/2024 | 05/15/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025451906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 32-year-old male of unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support the impella cp had low flows, despite confirmed proper placement and adequate volumes. the decision was made to replace pump with new impella cp which resolved the reported issue. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella cp was not returned for investigation. data logs were reviewed and position in ventricle alarm noted before low flow issue. suction alarms were observed with low flow after positioning issue resolved. the cause of the low pump flow issue was not determined since product was not returned and due to inconclusive evidence per log review. the cause of the positioning issue was most likely use related. added-h6 med dev prob code 3009 added- d6a/d6b. in accordance with updated procedures, sections d6, h6, and h10 have been revised from the initial submission.
28569 | 1220648-2024-09233 | 19020324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020324&pc=OZD | 03/28/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025448689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 58-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when pump being prepped, placement signal not reliable alarmed. after troubleshooting alarm resolved so decision made to continue inserting current cp. once inserted optimal flows never achieved with constant ventricle alarm when impella was in proper place. after 5 minutes spent repositioning and making adjustments without success, decision made to switch pump and console. no patient harm was reported. | [5 empty] | The investigation into the optical signal issue and the low pump flow issues has been completed since the original report was submitted. it was determined that the console was the potential cause of the optical signal issue, per the investigation into the console in complaint #(b)(4). the root cause of the low flow issue was not determined since product was not returned and data logs were inconclusive. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28570 | 1220648-2024-09232 | 19020325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020325&pc=OZD | 03/15/2024 | 05/02/2024 |  | 04/02/2024 |  |  | 03/15/2024 | 04/02/2024 |  |  | 09/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 52-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported during impella implant, the patient went into ventricular fibrillation and was shocked once. patient was stabilized and impella cp was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation (vf) has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf could not be determined as the device was not returned nor sufficient clinical details. the instructions for use states impella cp with smart assist system in the section: potential adverse events (united states) states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ added-h6 impact code 4627 to conform to updated procedures, sections d6 & h10 were revised with updated information.
28571 | 1220648-2024-09234 | 19020326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020326&pc=OZD | 06/06/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/06/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195986 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site bleeding and vascular damage. the access site bleeding was treated by deploying three perclose devices. the vascular damage was treated by doing a percutaneous transluminal angioplasty to the common femoral artery that had been narrowed or undergone stenosis followed by vascular surgery to complete the repair. the impella cp was weaned and explanted in the cath lab. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
28572 | 1220648-2024-09235 | 19020365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020365&pc=OZD | 03/25/2024 | 06/04/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456569 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. when the impella cp was placed to the left ventricle (lv), the patient went into a ventricular fibrillation storm. the impella cp was then explanted. extracorporeal membrane oxygenation was instead placed and percutaneous coronary intervention was performed. at the conclusion, the team attempted to deliver the same pump to the lv and delivery failed. the impella cp was aborted. | [5 empty] | The investigation has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vf was not determined. the root cause of the delivery issue was most likely due to patient condition since the pump could not pass due to the large size of the patient. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ i d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 first name revised as previously entered incorrectly. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28573 | 1220648-2024-09236 | 19020437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020437&pc=OZD | 08/03/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 08/03/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the patient had known peripheral vascular disease, a highly calcified right iliac artery, and complete occlusion of the left iliac artery. vascular surgery placed a covered stent to the right iliac and a short impella sheath was placed. however, they were unable to pass the impella through calcium proximal to the stent. they then attempted placing a long sheath and were still unable to pass the calcium, after which they attempted peripheral balloon inflation and deflation to place the long sheath, but they still could not cross the lesion. the decision was made to explant the sheath and abort the insertion procedure. the device was not placed, and the entrance site was closed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28574 | 1220648-2024-09237 | 19020441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020441&pc=OZD | 11/12/2023 | 10/24/2024 |  | 04/02/2024 |  |  | 11/12/2023 | 04/02/2024 |  |  | 07/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388644 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed oozing and hematoma at the impella access site. the oozing was addressed, and manual pressure was applied to press out the hematoma. additionally, when the patient was being pulled up in the bed, the catheter was caught on the foot board, which caused the blue suture hub to disconnect from the reposition sheath. therefore, the reposition sheath was pulled out of the arteriotomy and the impella was pulled into aorta. to address this issue, the impella device was explanted at the bedside and patient remains hemodynamically stable. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the access site bleed and the positioning issues has been completed. no product was returned for analysis. the clinical details stated that the catheter got caught on the bed and pulled the impella into the aorta. based on the clinical details the root cause of the positioning issues is most likely due to patient movement. the root cause of the access site bleeding cannot be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added-d6a/d6b f6/f8 should have been left blank.
28575 | 1220648-2024-09238 | 19020553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020553&pc=OZD | 06/07/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/07/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194823 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on initial start-up of the impella cp device, it had a high motor current and the device shut down. the staff had to perform one restart. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records
28576 | 1220648-2024-09239 | 19020622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020622&pc=OZD | 09/17/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 09/17/2022 | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235270 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella site. the sheath was repositioned and sutured in a new position. oozing ceased. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28577 | 1220648-2024-09240 | 19020733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020733&pc=OZD | 06/07/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/07/2022 | 04/02/2024 |  |  | 12/31/2023 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181042 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient bent their leg drastically causing pulsatile bleeding from around orange lip and valve of impella sheath. the 7fr was removed and the bleeding stopped. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28578 | 1220648-2024-09241 | 19020770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020770&pc=OZD | 08/04/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 08/04/2022 | 04/02/2024 |  |  | 04/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224870 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. as the patient became febrile, and developed worsening hypotension and confusion, with lactate rising, the shock team was consulted and decision made to place patient on extracorporeal membrane oxygenation (ecmo) with impella cp support for left ventricle unloading. it was reported that while the patient was on impella cp support, the impella alarmed for suctioning intermittently and the medical observed chugging and chattering from ecmo. echocardiography (echo) was performed at bedside and showed the impella placement too deep at 5cm, and was repositioned under echo guidance. there were still suctioning issue occurring, and the medical team made the decision to lower the p level and observe. it was reported the patient experienced hemolysis while on impella cp support. the pump was lowered from p3 to p2, and was the impella was explanted due to worsening hemolysis symptoms. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28579 | 1220648-2024-09242 | 19020855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020855&pc=OZD | 08/03/2023 | 04/30/2024 |  | 04/02/2024 |  |  | 08/03/2023 | 04/02/2024 |  |  | 04/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350737 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2023 |  |  |  |  | The user facility reported a 65-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding that prompted the team to re-suture to the skin after the initial two percloses did not take. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the reported access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the bleeding was improper angle matching since the repo hub was sutured flat to the patient instead of propped up at an angle. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28580 | 1220648-2024-09244 | 19020995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19020995&pc=OZD | 06/08/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 06/08/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump slipped into the aorta and it was unable to be re-crossed. consequently, the pump was removed and replaced with a new one. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28581 | 1220648-2024-09243 | 19021051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021051&pc=OZD | 06/27/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/27/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191497 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 74-year-old male female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella site and developed a hematoma. the patient was transfused with packed red blood cells (prbcs). it was also reported that the patient experienced hemolysis while on impella cp support. it was reported the patient experienced the hemolysis through impella explant, but no further information was provided if there was any additional intervention. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28582 | 1220648-2024-09246 | 19021169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021169&pc=OZD | 03/26/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438687 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old female noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. upon insertion of the impella device through the 14 french abiomed sheath, the physician used a cook needle at ten o'clock to gain pci access via single access technique. the abiomed companion sheath was inserted over a wire establishing single access but resulted in the 14 french abiomed sheath splitting just proximal to the hemostatic valve. this resulted in bleeding noted from the sheath. when the bleeding was noted, the coronary vessels were already wired, so the physician worked quickly to finish the pci and successfully explanted the impella device and sheath without any issue. the patient is stable. | [5 empty] | The investigation for the introducer damage has been completed. the returned 7fr lp sheath was visually inspected and tip damage was observed. data log review not applicable. if the needle stick in the 14fr introducer is very far out towards the edge of the 14fr valve, the user is most likely to run into the sharp edge of the inner geometry of the peel away hub when doing single access. the cause of the introducer damage was most likely user related since the needle stick was likely too lateral. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28583 | 1220648-2024-09247 | 19021335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021335&pc=OZD | 08/15/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 08/15/2022 | 04/02/2024 |  |  | 04/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. heparin was discontinued and a suture was placed around the blue hub to stop the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28584 | 1220648-2024-09249 | 19021359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021359&pc=OZD | 11/13/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 11/13/2022 | 04/02/2024 |  |  | 08/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263402 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Pumping Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/13/2022 |  |  |  |  | The user facility reported a 57-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, device had suction issue due to kink into the cannula. the impella device was explanted, and patient was hemodynamically stable. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28585 | 1220648-2024-09250 | 19021383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021383&pc=OZD | 06/08/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/08/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203049 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced heparin-induced thrombocytopenia positive during support. additionally, gi bleed was found as there was bleeding in the rectum as well as access site bleeding noted by the oozing. no further information is known on what intervention was done for the issues. a purge leak was also noted by the lead at the connection and was resolved by conducting a purge bypass as changing the purge cassette was unsuccessful. the patient remained on support and was transferred to unit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." this report is being filed as part of a retrospective review of historical records.
28586 | 1220648-2024-09251 | 19021672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021672&pc=OZD | 03/27/2024 | 05/16/2024 |  | 04/02/2024 |  |  | 03/27/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137162 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The biomedical engineer reported the automated impella controller (aic) had issues with the purge disk not connecting. the console was replaced. no patient was involved. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure was unreproducible. there was no issue found during the investigation of the complaint. the device functioned/operated to specification.
28587 | 1220648-2024-09253 | 19021793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19021793&pc=OZD | 03/27/2024 | 05/15/2024 |  | 04/02/2024 |  |  | 03/27/2024 | 04/02/2024 |  |  | 09/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445432 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. it was reported that a patient was being prepped for impella pump insertion. the sheath was placed when the patient lost distal pulses. in response, the pump insertion was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details.
28588 | 1220648-2024-09254 | 19022123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19022123&pc=OZD | 06/10/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/10/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207141 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a patient needing mechanical circulatory support due to st-elevated myocardial infarction/cardiogenic shock. while on support, the patient experienced access site bleeding. the physician decided to wean the device at completion of the case. the patient tolerated well and the impella cp device was explanted | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28589 | 1220648-2024-09255 | 19022300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19022300&pc=OZD | 09/26/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 09/26/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195035 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The complainant reported the patient was on impella cp support due to a stemi (st-segment elevation myocardial infarction) requiring cardiopulmonary resuscitation (cpr) and defibrillation in the field before achieved rosc (return of spontaneous circulation) and brought to cardiac catheterization laboratory (ccl) for insertion. it was reported that while on impella cp support, the patient had an unobtainable doppler pulse in the impella leg, and hemolysis was noted. the p level was lowered to treat the hemolysis, but the hemolysis did not resolve while on impella cp support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28590 | 1220648-2024-09245 | 19022326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19022326&pc=OZD | 08/11/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 08/11/2022 | 04/02/2024 |  |  | 04/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209988 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 73-year-old male with heart failure and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a hematoma at the impella access site. surgery was performed to remove the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28591 | 1220648-2024-09256 | 19022466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19022466&pc=OZD | 03/28/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1607645 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility biomedical department reported that an automated impella controller (aic) had a purge cassette door that was broken. no patient involvement was reported. | [5 empty] | The investigation for the console (aic) purge disc/flag issues has been completed. the console was returned for evaluation. upon inspection, there was no damage visible to the console¿s purge door. data logs were not relevant given the logs don't include information regarding functionality of the purge door. therefore the root cause of the reported issue could not be determined since the issue could not be reproduced. in accordance with updated procedures, section d2 has been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28592 | 1220648-2024-09257 | 19023210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19023210&pc=OZD | 06/10/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/10/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. during support there was access site bleeding noted by hematoma formation, which was treated by applying manual pressure, a femstop device, and two units of blood. although the hematoma formation subsided, bleeding returned at the site again and this was treated by providing the patient with another two units of blood. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28593 | 1220648-2024-09258 | 19023389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19023389&pc=OZD | 03/28/2024 | 05/02/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  |  |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025440649 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 46-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed access site bleeding. as treatment, a femstop was place and the patient received two units of red blood cells. the patient had improved hemodynamically, and the pump was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to patient condition since the patient was coagulopathic from shock liver. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 & h10 were revised with updated information.
28594 | 1220648-2024-09259 | 19023394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19023394&pc=OZD | 08/23/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 08/23/2022 | 04/02/2024 |  |  | 05/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231150 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The user facility reported a 86-year-old male patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented with aortic stenosis and decreased ejection fraction. it was reported that the impella cp was initially placed but a kink formed in the aorta, and poor flow was noted. the impella was removed and successfully placed a second time with no flow issues reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28595 | 1220648-2024-09252 | 19023978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19023978&pc=OZD | 08/18/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 08/18/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208911 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 58-year-old male patient was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support the patient became septic. the sepsis was managed through discussions of alternate axillary approaches for weaning and mobilization. the decision was made to explant the impella, treat patient for sepsis and reevaluate for a left ventricular assist device. it was also reported that the impella had suction alarms above p6. transesophageal echocardiography (tee) was used to confirm impella position. impella was found to be too far into the left ventricle and was pulled back under tee guidance to an optimal position. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28596 | 1220648-2024-09260 | 19024108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19024108&pc=OZD | 03/28/2024 | 05/14/2024 |  | 04/02/2024 |  |  | 03/28/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024430985 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the physician might have gone too far when lifting to suction pleural effusion. the doctor was unable to pull back as the patient was completely dependent on the impella 5.5 prior to opening the chest. the left ventricle was repaired with stitches without any complications. the patient was stable. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the ventricle perforation issue was use related to improper technique. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ "dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve."
28597 | 1220648-2024-09261 | 19024195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19024195&pc=OZD | 06/11/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/11/2022 | 04/02/2024 |  |  | 01/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178473 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/11/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support had access site bleeding. the patient was given protamine reversal, and the heparin purge was stopped, also a blood transfusion was given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28598 | 1220648-2024-09262 | 19024382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19024382&pc=OZD | 08/25/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 08/25/2022 | 04/02/2024 |  |  | 12/31/2023 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170358 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a patient with suspected corona myocarditis was implanted with an impella cp for mechanical circulatory support. it was reported that while on impella cp support, the patient was bleeding and was given a blood transfusion. bleeding was not resolved until the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28599 | 1220648-2024-09248 | 19024888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19024888&pc=OZD | 03/26/2024 | 05/15/2024 |  | 04/02/2024 |  |  | 03/26/2024 | 04/02/2024 |  |  | 10/31/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418854 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that they were unable to advance an impella 5.5 device into the patient's ascending aorta. an impella cp was deployed instead. there was no patient harm reported. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of delivery issue is determined to be patient condition because acute angle of the patient innominate artery and the pump was unable to pass into ascending aorta. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28600 | 1220648-2024-09263 | 19026201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19026201&pc=OZD | 06/16/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 06/16/2022 | 04/02/2024 |  |  | 02/29/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200634 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the patient had limb ischemia. manual compression was held for one hour. the impella was explanted at the patient's bedside due to occlusion of the right leg. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28601 | 1220648-2024-09265 | 19026713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19026713&pc=OZD | 09/20/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 09/20/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212162 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the pump exhibited low pump flow. the placement was verified under echo and later explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28602 | 1220648-2024-09264 | 19026921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19026921&pc=OZD | 09/17/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 09/17/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206051 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported a 76-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had signs of hemolysis due to the catheter being pulled back in the aorta. the impella cp was repositioned, and the patient received one unit of packed red blood cells (prbc). the patient was then hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28603 | 1220648-2024-09266 | 19027813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19027813&pc=OZD | 06/18/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 01/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187504 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/18/2022 |  |  |  |  | The user facility reported a patient needing mechanical circulatory support due to shortness of breath and anterior myocardial infarction. the patient developed bleeding and oozing at the impella cp access site. the bleeding was resolved after adding a stitch to the site. purge concentration decreased by half. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28604 | 1220648-2024-09267 | 19028149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028149&pc=OZD | 09/20/2022 | 10/08/2025 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/20/2022 |  |  |  |  | The user facility reported a 52-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the distal reperfusion sheath/cannula clotted resulting in limb ischemia. the patient was sent to the operating for a fasciotomy and impella removal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28605 | 1220648-2024-09269 | 19028337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028337&pc=OZD | 10/27/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 81-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient underwent bypass surgery and experienced atrial fibrillation. it was reported while explanting the impella cp device, the physician had trouble with access and bleeding at site due to the mesh from a recent hernia surgery. the physician was able to cutdown and remove the impella cp and used distal digital pressure to force any clot to bleed out and not down leg. patient has not been on heparin since surgery, so sheath was full of clot. the impella cp was explanted and was clear of any build up. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28606 | 1220648-2024-09270 | 19028515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028515&pc=OZD | 11/03/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 01/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173668 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported a 78-year-old male was implanted with an impella cp device for mechanical circulatory support. while on support, a thrombus had formed due to heparin-induced thrombocytopenia (hit) and was able to be aspirated. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28607 | 1220648-2024-09268 | 19028569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028569&pc=OZD | 09/20/2022 | 04/04/2025 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198322 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2022 |  |  |  |  | The user facility reported a 65-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient underwent a femorofemoral (ff) bypass, but a blood clot was noted to have formed in the route. the clot was aspirated and opened, but it occluded again. the patient had absent pulses and mottling in their lower limb. once the impella cp was explanted, the dorsalis pedis artery was stated to be palpable and the color in the lower extremity returned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition , any concomitant medication , vascular size or variations thereof, detailed description of the symptoms experienced by the patient , physical examination findings, including pulse assessment and visual examination of the affected limb , diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography , laboratory test results, including blood tests for markers of inflammation or clotting disorders , any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension , patient's response to previous treatments or interventions for vascular conditions . with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions, description of the location and characteristics of the thrombus (e.g., size, shape, consistency), relevant laboratory test results, such as coagulation profiles and platelet counts, imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus, any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states), medications or treatments that the patient was receiving at the time of thrombus formation, surgical or procedural details if applicable (e.g, cardiac catheterization), any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events), presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
28608 | 1220648-2024-09272 | 19028573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028573&pc=OZD | 09/21/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238064 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 91-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a pericardial bleed that had to be removed. 200 milliliters of fluid was removed. advanced cardiac life support protocol was initiated. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28609 | 1220648-2024-09275 | 19028715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028715&pc=OZD | 09/23/2022 | 05/20/2025 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235817 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support thrombocytopenia and hemolysis were noted. hemolysis was treated under echo guidance. the patient received packed red blood cells for the reported issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28610 | 1220648-2024-09273 | 19028731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028731&pc=OZD | 11/11/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 07/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258988 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 29-year-old male was implanted with an impella cp with a history of non-ischaemic cardiomyopathy. it was report that during support, the patient starting to decompensate despite impella cp support. patients cardiac power output and pulmonary artery pulse pressure were very low. in addition, the patient spiked a fever requiring increasing pressor. impella position was evaluated due to hematuria. it looked to be slightly out of position however the decision was made not to adjust the position. the impella was removed due to unresolved hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28611 | 1220648-2024-09271 | 19028733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028733&pc=OZD | 09/21/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234918 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experience access site bleeding and a hematoma. the patient was administered t two units of blood and systemic heparin was discontinued for a few hours. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28612 | 1220648-2024-09274 | 19028756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028756&pc=OZD | 09/22/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 12/31/2023 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176038 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient's urine was red and labs hemolyzed. impella turned to p5 and intra-aortic balloon pump 1:2 at 30 percent per doctor. there was also bicarb in purge. the patients urine cleared. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28613 | 1220648-2024-09277 | 19028845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028845&pc=OZD | 09/24/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 05/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230128 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella device for mechanical circulatory support. while on support, there was slight dried ooze at the access site; intervention is unknown. additionally, the patient had runs of non-sustained ventricular tachycardia on support. the impella was angled towards the mitral valve and was shallow. the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28614 | 1220648-2024-09276 | 19028866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19028866&pc=OZD | 09/23/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235012 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced oozing at access site. the impella was weaned and removed. the 14 fr sheath was left in place temporarily to see if patient recovered; it was removed later. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28615 | 1220648-2024-09278 | 19029313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029313&pc=OZD | 11/15/2022 | 04/02/2024 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 07/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258975 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 65-year-old male patient for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed ischemia as noted by sluggish distal flow. the impella cp pump was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28616 | 1220648-2024-09279 | 19029495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029495&pc=OZD | 11/17/2022 | 05/22/2025 |  | 04/02/2024 |  |  |  | 04/02/2024 |  |  | 06/30/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235757 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 64-year-old male patient implanted with an impella cp for mechanical circulatory support due to an onset of acute myocardial infarction (ami). it was reported that the oscor was not able to be removed due to compromised sterility, subsequently, the repo sheath suture wings were cut off. distal pulses were ok and the impella was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28617 | 1220648-2024-09280 | 19029713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029713&pc=OZD | 11/17/2022 | 05/23/2025 |  | 04/02/2024 |  |  | 11/17/2022 | 04/02/2024 |  |  | 08/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263732 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 42-year-old female patient, noted as high risk percutaneous coronary intervention, was implanted with an impella cp device for mechanical circulatory support. the medical team attempted post-closing the site with 2 percloses, however was unsuccessful. the cardiac physician did not want to attempt a dry closure with a balloon tamponade, so he called vascular to do a surgical cut down and repair. he placed the 14fr sheath back into the arteriotomy and sent the patient to the or. the patient tolerated the removal of the sheath and the procedure went well. estimated amount of blood loss is unknown. the patient was sent to the icu post procedure and doing very well. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28618 | 1220648-2024-09281 | 19029722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029722&pc=OZD | 11/16/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 11/16/2022 | 04/02/2024 |  |  | 03/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207136 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had gastrointestinal bleed (gi bleed) and was treated by giving the patient two units of blood and one unit of platelets. gi bleed improved and later did not require a transfusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28619 | 1220648-2024-09282 | 19029742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029742&pc=OZD | 11/16/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 11/16/2022 | 04/02/2024 |  |  | 07/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023256341 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Decrease in Pressure | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 86-year-old female patient implanted with an impella cp for mechanical circulatory support due to a st-elevated myocardial infarction. while on support, the pump presented low purge flow that was resolved by tissue plasminogen activator protocol. additionally, the pump experienced positioning issues that led to hemolysis, which was resolved by explanting the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28620 | 1220648-2024-09283 | 19029752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029752&pc=OZD | 11/17/2022 | 04/04/2025 |  | 04/02/2024 |  |  | 11/17/2022 | 04/02/2024 |  |  | 08/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266451 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. additionally, the pump experienced suction alarms due to volume issues. there was no information provided related to the resolution of these events. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28621 | 1220648-2024-09284 | 19029757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029757&pc=OZD | 11/18/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 11/18/2022 | 04/02/2024 |  |  | 07/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257029 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 52-year-old male patient implanted with an impella cp for mechanical circulatory support due to a left anterior descending occluded artery. it was observed that the groin site had oozing. partial thromboplastin time came back supra-therapeutic although act 178.the patient was weaned from the impella cp with a successful closure in the cath lab. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28622 | 1220648-2024-09285 | 19029792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029792&pc=OZD | 11/22/2022 | 05/27/2025 |  | 04/02/2024 |  |  | 11/22/2022 | 04/02/2024 |  |  | 08/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262852 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Pumping Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported a (b)(6) female was placed on an impella cp for mechanical circulatory support after having a hemodynamic collapse during an elective rotoblador of left anterior descending artery (lad). because of the extravasation from the first access site and low pump flows, it was determined to explant the impella. the impella was removed using a two 8 fr sheath technique. hemostasis at the groin was achieved with manual compression of the initial site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28623 | 1220648-2024-09286 | 19029828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029828&pc=OZD | 11/22/2022 | 04/02/2024 |  | 04/02/2024 |  |  | 11/22/2022 | 04/02/2024 |  |  | 07/31/2024 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023259502 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported a 69-year-old female patient was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced limb ischemia. patient had diminished pedal pulses to the left leg. the surgeon decided to implant the impella 5.5, however, due to the patient's medical condition, the impella cp was successfully placed instead and the patient was able to come off bypass. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28624 | 1220648-2024-09287 | 19029847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029847&pc=OZD | 12/17/2023 | 05/22/2024 |  | 04/02/2024 |  |  | 12/17/2023 | 04/02/2024 |  |  | 09/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401811 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, the patient had experienced oozing at the access site. the purge was transitioned from heparin to bicarbonate. additionally, the sheath was repositioned, fem-stop was applied, the patient was transfused with two units of pack red blood cells (prbc), and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was high patient act values, the patient had act value above 180 and heparin was also stopped as per the clinical description.
28625 | 1220648-2024-09288 | 19029853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19029853&pc=OZD | 02/29/2024 | 04/17/2024 |  | 04/02/2024 |  |  | 02/29/2024 | 04/02/2024 |  |  | 11/30/2025 |  | 04/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438248 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 53-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient had signs of hemolysis with pink urine. the patient was receiving blood products and patient¿s signs of hemolysis were improving. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the access site bleeding ¿ minor and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding and the hemolysis was not determined because sufficient clinical details and the data was not returned to determine the cause and the product was not returned for analysis.
28626 | 1220648-2024-09089 | 19010430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010430&pc=OZD | 07/05/2022 | 03/11/2025 |  | 04/01/2024 |  |  | 07/05/2022 | 04/01/2024 |  |  | 01/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported that the decision for impella cp was made after the patient went into cardiac arrest. the impella was placed and removed due to inability to recross vasculature. a new impella was placed using the same access as the prior impella. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28627 | 1220648-2024-09090 | 19010510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010510&pc=OZD | 03/07/2024 | 04/24/2024 |  | 04/01/2024 |  |  | 03/07/2024 | 04/01/2024 |  |  | 06/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438751 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 58-year-old female was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed limb ischemia and the staff could not detect signals in the left leg. the patient was escalated to an impella 5.5 due to vascular issues with the left leg per vascular recommendations. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
28628 | 1220648-2024-09091 | 19010536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010536&pc=OZD | 09/10/2022 | 04/08/2025 |  | 04/01/2024 |  |  | 09/10/2022 | 04/01/2024 |  |  | 05/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228105 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the impella device was alarming for suction and diastolic flows were decreased. echocardiography was performed to confirm impella placement, and the impella was found to be too deep into the ventricle. under echocardiography guidance, the physician attempted to pull the pump back for optimal placement, but pulled the pump too far back and it exited the left ventricle. the physician attempted to recross the aortic valve but was unsuccessful. the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28629 | 1220648-2024-09092 | 19010562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010562&pc=OZD | 03/14/2024 | 04/23/2024 |  | 04/01/2024 |  |  | 04/01/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a prolonged purge system blocked alarm. tissue plasminogen activator protocol was unsuccessful in clearing the blockage. purge flow was less than two cubic centimeters per hour for over eight hours. the impella 5.5 and the automated impella controller (aic) were exchanged. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge issue. the console was returned for evaluation. the console was tested and was unable to reproduce the event #408 and #409 by running the test pump and the cassette. purge pressure sensor accuracy. the fringe pattern looks good, which confirms no issue with the optical bench. data logs were reviewed which showed the pump was re-plugged three times, and it was most likely to unwind if the white catheter was coiled (note: there was no pump detection issue or impella defective alarm). after the pump was re-plugged for the third time, in 5 hours and 40 minutes, the console received the first purge pressure high (pph) alarm, event #408, and the purge system blocked alarm, event #409. after 12 minutes of the purge pressure high alarm but events #408 and #409 were never resolved. the purge flow was less than 2ml/hr. then the user performed a couple of 'purge fluid bag change' and one 'cassette and bag change' procedure but the event #408 and #409 was persistent. then, the pump was unplugged, and, through the case start wizard, a new cassette and pump were plugged in. after the new pump connection, in 2 minutes, the console received a placement signal not reliable (psnr) alarm, which was resolved immediately after the pump was started. still, events #408 and #409 were persistent. the purge flow was less than 2ml/hr for more than 8 hours, confirmed in the data logs, which confirms the reported failure mode. no explicit evidence was found in the logs to point out the cause for the high purge pressure and the purge system block alarm. there was also no evidence of a piston block in the logs. the root cause of the high purge pressure and the blocked purge system was not determined due to the inability to reproduce. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28630 | 1220648-2024-09093 | 19010564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010564&pc=OZD | 09/15/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 09/15/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was successfully placed, however the placement signal was not working and the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28631 | 1220648-2024-09094 | 19010636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010636&pc=OZD | 03/15/2024 | 05/02/2024 |  | 04/01/2024 |  |  | 03/15/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025458539 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, the patient developed an access site hematoma. as treatment, anticoagulation therapy was held and evacuation at the cut down site was performed. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. to conform to updated procedures, section d6 was revised with updated information.
28632 | 1220648-2024-09095 | 19010647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010647&pc=OZD | 03/27/2024 | 05/14/2024 |  | 04/01/2024 |  |  | 03/27/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025439386 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | Pressure Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the physician attempted to reposition the pump away from the mitral valve and the purge flows decreased, and the purge pressure was above 1050. the physician decided to leave the cp in place with the blocked purge alarm. the pump eventually stopped and was briefly restarted before stopping again. the decision was made to explant the pump and the patient was reported to be stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported issues has been completed since the original report was filed. the pump and purge cassette were returned for investigation. according to the clinical report, during repositioning, the md torqued the catheter, and the purge system blocked alarm appeared without resolution. the cause of the pump stop issue was a kinked purge line near motor housing, resulting in high purge pressure. this kink caused blood and biomaterial to enter the motor housing. the root cause of the positioning issue was most likely patient movement as this occurred after hospital transfer. in accordance with updated procedures, section d6 has been revised from the initial submission.
28633 | 1220648-2024-09096 | 19010678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010678&pc=OZD | 09/18/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 09/18/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at the groin impella access site. the bleeding was resolved with additional sutures. it was also reported while on support the patient experienced heparin induced thrombocytopenia (hit) and as a result systemic heparin was discontinued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28634 | 1220648-2024-09097 | 19010679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010679&pc=OZD | 03/27/2024 | 05/09/2024 |  | 04/01/2024 |  |  | 03/27/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1746775 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that before the device entered the patient that during priming the impella was attached to the automated impella controller (aic) when a purge system open alarm occurred. the purge fluid was not moving into the purge side arm. due to concern for the patient a new impella cp and aic was used successfully. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - pump issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. there was no issue found during the investigation of the reported complaint. the console functioned/operated to specification during the clinical procedure, and no issues were reproduced during testing.
28635 | 1220648-2024-09098 | 19010738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010738&pc=OZD | 06/10/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/10/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200613 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding. to resolve the bleeding, perclose was tightened, a femostop was applied, manual pressure was applied, the site was redressed and heparin was discontinued and replaced with dextrose. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28636 | 1220648-2024-09099 | 19010786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010786&pc=OZD | 06/10/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/10/2022 | 04/01/2024 |  |  | 01/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189502 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a 78-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing and a hematoma at the impella access site. the impella cp was explanted to help with the bleeding issues; hemostasis was achieved at the end of closure procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28637 | 1220648-2024-09100 | 19010817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010817&pc=OZD | 03/20/2024 | 05/21/2024 |  | 04/01/2024 |  |  | 03/20/2024 | 04/01/2024 |  |  | 06/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377458 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 24-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient had access site bleeding. blood products were provided to the patient and a stitch was applied. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, a stitch was applied to achieve hemostasis as per the clinical description. the instructions for use referenced in the initial is from ifu for impella cp with smart assist system in sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28638 | 1220648-2024-09101 | 19010876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010876&pc=OZD | 06/10/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/10/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207858 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing and a hematoma at the impella access site. to resolve the issue, femstop and manual pressure were applied and the impella cp was subsequently explanted. the patient's outcome was reportedly hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28639 | 1220648-2024-09102 | 19010904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010904&pc=OZD | 03/21/2024 | 05/13/2024 |  | 04/01/2024 |  |  | 03/21/2024 | 04/01/2024 |  |  | 07/31/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440661 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 72-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant, the patient suffered a vascular perforation at the access site. the perforation and its associated hematoma were treated in the operating room while undergoing coronary artery bypass grafting. the patient was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ | The impella device was discarded by the customer. the investigation has been completed since the original report was submitted. the cause of the vascular damage issue was most likely use related per clinical review. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report. additional manufacturer narrative on the previously submitted report stated that the device was not returned but it was discarded.
28640 | 1220648-2024-09103 | 19010918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010918&pc=OZD | 09/19/2022 | 10/08/2025 |  | 04/01/2024 |  |  | 09/19/2022 | 04/01/2024 |  |  | 06/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was was being placed on an impella cp device for mechanical circulatory support. it was reported during insertion of the impella device the introducer was damaged due to the patient having heavy calcium in the iliacs. the introducer was removed and a new sheath was opened and inserted with success. it was also reported during insertion, the pump was unable to across the aortic valve due to aortic stenosis. the medical team decided to abort the insertion and the pump was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28641 | 1220648-2024-09104 | 19010973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19010973&pc=OZD | 06/11/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/11/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198894 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was out of position and had low flows. the pump repositioned from 1.8 to 3.7. it was reported that the patient had left ventricle recovery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28642 | 1220648-2024-09105 | 19011046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011046&pc=OZD | 06/12/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/12/2022 | 04/01/2024 |  |  | 12/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162488 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The user facility reported a 81-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during intubation in the icu, the motor waveform became flat, the flow rate decreased due to double pulse and the patient had an episode of ventricular tachycardia. additionally, it was determined under echocardiogram that the pigtail was found in the left ventricle, and it was very close to coming out. therefore, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28643 | 1220648-2024-09106 | 19011089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011089&pc=OZD | 03/21/2024 | 09/26/2025 |  | 04/01/2024 |  |  | 03/21/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443569 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/21/2024 |  |  |  |  | The user facility reported an unspecified male patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that following impella cp placement, the automated impella controller (aic) displayed motor current was damped pulsatile with 366/29. the doctor was unsure if the pump was damaged as a 16 french cook sheath was used during the implant procedure. there was no reported patient harm. | The complainant reported that the device exhibited optical signal issues. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow issue has been completed since the original report was submitted. the pump was not returned for investigation. clinical data analysis discovered optical signal issue and temporary pump stop issue. the root cause of the optical signal issue was not determined since product was not returned. the root cause of the temporary pump stop was not determined since product was not returned. in accordance with updated procedures, section d6 has been revised from the initial submission. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was not returned. there is a screenshot of the automated impella controller (aic) showing suction and placement signal not reliable (psnr) alarms. data logs show consistent placement signal (ps) low alarms throughout the case. there are also psnr alarms. there are also suction alarms throughout the case. there are position in ventricle alarms at some of the instances of psnr. the 5s parameters are relatively stable. the pump prior to this one had no optical issues. the root cause of the optical signal issue was not determined since product was not returned. data logs also show a hmc pump stop 1 hour into the case which restarts and continues to run for the rest of the case. the root cause of the temporary pump stop was not determined since product was not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issues "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
28644 | 1220648-2024-09107 | 19011110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011110&pc=OZD | 06/13/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/13/2022 | 04/01/2024 |  |  | 04/30/2023 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the groin access site. to resolve the bleeding, the dressing was changed multiple times and manual pressure was applied. additionally, the patient received three units of packed red blood cells. patient outcome was reported as improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28645 | 1220648-2024-09108 | 19011150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011150&pc=OZD | 06/15/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/15/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the groin access site. the impella was repositioned, manual pressure was applied, and heparin was withheld. no new oozing was noted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28646 | 1220648-2024-09109 | 19011162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011162&pc=OZD | 07/07/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 07/07/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023213988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. durin implantation, multiple attempts to position the device occurred. the physician tried to pull back and the pump either crossed the atrioventricular (av) and was unable to recross or it immediately went under the pap muscle; which would not release without pushing back across the (av). hemolysis developed due to the device was placed too deep and caused the urine output to be red. volume was given and the p-level was reduced. the patient's urine was still bloody and there were no plans for repositioning the device. however, the patient continued to improve overnight and the pump was explanted without any issues. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28647 | 1220648-2024-09110 | 19011190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011190&pc=OZD | 06/19/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/19/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump was unable to advance into iliac proximal to aorta. therefore, the impella cp implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28648 | 1220648-2024-09111 | 19011193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011193&pc=OZD | 07/06/2022 | 03/10/2025 |  | 04/01/2024 |  |  | 07/06/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023194722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that impella in aorta alarms were being received. the physician echoed and the pump was pulled back with the pigtail through the valve which advanced and confirmed the echo. the position looked good; however, the patient developed hemolysis. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28649 | 1220648-2024-09112 | 19011269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011269&pc=OZD | 07/11/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 07/11/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211491 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Migration | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 79-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to hemodynamically unstable. while on support, the echo confirmed migration of the device. moreover, the patient¿s lab was abnormal leading to hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28650 | 1220648-2024-09113 | 19011348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011348&pc=OZD | 03/22/2024 | 05/10/2024 |  | 04/01/2024 |  |  | 03/22/2024 | 04/01/2024 |  |  | 03/31/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024320106 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient pulled the impella back and the impella stopped. the surgeon kept the pump at p0 and the impella 5.5 was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the positioning issue and the pump stop has been completed. no product was returned for evaluation. data logs were reviewed and the right logs at end of case show the pump was pulled into the aorta when a motor current spike of 30000 occurred and alarm #115, "impella stopped (rpm deviations)" was triggered. clinical details noted that the patient pulled the pump back 5-7cm on day of explant and pump stop occurred at that time. the root cause of the pump stop was most likely due to an electrical open as pump stopped when patient pulled impella.
28651 | 1220648-2024-09115 | 19011424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011424&pc=OZD | 06/27/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/27/2022 | 04/01/2024 |  |  | 01/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a hematoma was noted. to resolve the issue, heparin was discontinued with the addition of bicarb to the purge. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28652 | 1220648-2024-09116 | 19011451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011451&pc=OZD | 10/02/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 10/02/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226079 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at the impella access site. forward tension sutures, angle matching, and sterile dressing was applied to the site. the patient was than transferred to a hospital bed and a hematoma was observed above the impella site. the physician made the decision to explant the impella due to the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28653 | 1220648-2024-09117 | 19011521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011521&pc=OZD | 07/13/2022 | 03/12/2025 |  | 04/01/2024 |  |  | 07/13/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the left ventricle (lv) alarm alerted that the device positioning was too deep. the impella cp was initially placed 4.0cm from the valve, but eventually repositioned at 3.5cm from the aortic valve. the health care provider observed that the catheter pigtail was caught in the papillary (pap) muscle, and the device was repositioned to release the catheter from the pap. it was reported that while in the critical care unit, the patient experienced hemolysis and oozing from the access site overnight. the patient was given plasma-free hemoglobin every six hours until hemolysis cleared. it was reported that two (2) units of packed red blood cells (prbc¿s) were administered, site dressings were changed, and the purge solution was switched to sodium bicarbonate in response to the access site bleed. additional support was provided as a second impella cp was implanted for ongoing support for the left ventricular ejection fraction (lvef). the impella 5.5 was placed across the aortic valve and the imeplla cp was explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: access site bleeding ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ hemolysis ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28654 | 1220648-2024-09118 | 19011604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011604&pc=OZD | 08/02/2022 | 04/03/2025 |  | 04/01/2024 |  |  | 08/02/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215726 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The complainant reported a female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced paroxysmal atrial tachycardia due to small vessel size. post procedure, there was an instant rise of total hemodialysis, and the impella cp weaning was accelerated due to a substantial decrease in limb perfusion. the device was explanted in the operating room the following day. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28655 | 1220648-2024-09119 | 19011608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011608&pc=OZD | 06/28/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/28/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023210565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the aic alarm optical sensor failure/controller failure was displayed. the aic was replaced with another controller and the same alarm happened. the impella was replaced with a new impella device successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d5 operator of device: health professional added d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28656 | 1220648-2024-09120 | 19011622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011622&pc=OZD | 11/14/2022 | 03/12/2025 |  | 04/01/2024 |  |  | 11/14/2022 | 04/01/2024 |  |  | 07/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient then developed a significant arterial groin bleed. the dressing was changed and the previous sutures were removed and replaced with forward tension sutures. the impella position was confirmed with echocardiogram. the patient also experienced 4-5 beat intervals of ventricular tachycardia that would resolve on its own. the impella cp was later replaced with an impella 5.5. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28657 | 1220648-2024-09121 | 19011691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011691&pc=OZD | 11/15/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/15/2022 | 04/01/2024 |  |  | 08/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260502 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella cp insertion site (left groin). fempstop was applied to the left groin and the dressing was changed. patient was removed from integrilin and placed on a cangrelor drip. bicarb in purge. cardiology placed new mattress sutures that stopped the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28658 | 1220648-2024-09122 | 19011747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011747&pc=OZD | 11/16/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/16/2022 | 04/01/2024 |  |  | 08/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hematoma at the site. therefore, heparin was withheld. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28659 | 1220648-2024-09123 | 19011860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011860&pc=OZD | 08/04/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 08/04/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209238 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 28-year-old female was being implanted with an impella cp device into the left heart chamber for mechanical circulatory support. due to the position of the patient¿s aortic block interfering with the device motor, the impella cp was temporarily removed, and a mitral valve prolapse was performed. an attempt was made to re-implant the device without the use of a guidewire. however, the impella cp would not pass through the patient¿s vasculature. a new device was implanted, and post-operative course progressed well. the patient was successfully weaned from the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28660 | 1220648-2024-09124 | 19011913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011913&pc=OZD | 09/01/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 09/01/2022 | 04/01/2024 |  |  | 05/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228752 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 77-year-old male patient for mechanical circulatory support. while the patient was on impella cp support, it was determined under echocardiogram while verifying impella position, the impella cp was found to be caught in the papillary muscle. multiple attempts were made to manipulate the impella back without success. the impella device was removed, cleaned, and reinserted successfully to middle of left ventricle. echocardiography was used again to verify position, and the device was again found to be placed behind a papillary muscle. due to the patient's condition, the medical team made the decision to not attempt to reposition again as the inlet was away from the mitral valve and papillary muscle. it was also reported a small hematoma was noticed and addressed, but there was no information provided on what intervention was used to resolve the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28661 | 1220648-2024-09125 | 19011961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19011961&pc=OZD | 10/27/2023 | 08/21/2024 |  | 04/01/2024 |  |  | 10/27/2023 | 04/01/2024 |  |  | 01/31/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328128 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 52-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during transport, suction alarms were noted and the p level was decreased to p2 to break suction. blood was administered while in the ambulance. upon arrival to the medical center, it was noticed that the automated impella controller showed decoupling with lower than expected flows. after an echocardiogram was performed, the impella cp was repositioned. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the pump was not returned for investigation. data logs show that the pump flow was lower than expected, with active suction and 'impella position unknown' alarms. this issue was seen resolved after lowering the p-level to p3. no motor current spikes were observed during the case run. the placement signal decreased and flattened over time during the reported low pump flow event. according to the clinical report, bleeding was reported after transfer to icu and was managed after placing femstop. also, low pump flow was noted after the patient transfer. the cause of the low pump flow issue was most likely patient condition related, based on given clinical and data log review. added- b5 mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
28662 | 1220648-2024-09126 | 19012067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012067&pc=OZD | 11/16/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/16/2022 | 04/01/2024 |  |  | 11/30/2023 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 82-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. there were difficulties attempting closure of the impella cp device. the physician gained access on the right side to go up and across to use balloon for hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28663 | 1220648-2024-09127 | 19012089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012089&pc=OZD | 08/08/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 08/08/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213661 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support. during support a purge pressure high alarm which was resolved with repositioning the patient. in addition, there was a low pump flow. the patient was successfully weaned and the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
28664 | 1220648-2024-09128 | 19012142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012142&pc=OZD | 11/26/2022 | 03/12/2025 |  | 04/01/2024 |  |  | 11/26/2022 | 04/01/2024 |  |  | 06/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2022 |  |  |  |  | The user facility reported a 68-year-old male on impella cp for mechanical circulatory support. it was reported the physician experienced difficulty advancing the impella cp across the atrioventricular nove (av). the device was explanted and a dobutamine drip was started at a low dose. the patient experienced hematuria. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28665 | 1220648-2024-09129 | 19012354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012354&pc=OZD | 08/11/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 08/11/2022 | 04/01/2024 |  |  | 05/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219587 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 11 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | Mechanical circulatory support and had positioning issues. there was a suction alarm due to impella in the aorta in too deep. stable hemodynamics was achieved. impella explanted and implanted heartmate 3 (hm3). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28666 | 1220648-2024-09130 | 19012605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012605&pc=OZD | 03/16/2024 | 05/28/2024 |  | 04/01/2024 |  |  | 03/21/2024 | 04/01/2024 |  |  | 11/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427048 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 74-year-old male in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion, when the guidewire was removed, the impella also came out. the impella was removed, and a new guidewire and pump were used to replace the impella. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the positioning issue was most likely use issue since the impella came out of left ventricle during guidewire removal. d6a, d6b.
28667 | 1220648-2024-09131 | 19012667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012667&pc=OZD | 08/13/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 08/13/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 67-year-old female was implanted with an impella cp device for mechanical circulatory support due to critical left main coronary artery calcification, chest pain and shortness of breath. while on support, high purge pressure alarm appeared to be due to kink; alarm resolved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28668 | 1220648-2024-09132 | 19012783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012783&pc=OZD | 03/18/2024 | 05/21/2024 |  | 04/01/2024 |  |  | 03/21/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025441895 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 61-year-old female post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 had moved out of position and into the aorta. the pump was removed and replaced with an impella cp device to resolve the issue. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was patient movement because the impella 5.5 had protruded into the ascending aorta right after patient moved to icu. in accordance with updated procedures, section d6 has been revised from the initial submission.
28669 | 1220648-2024-09133 | 19012985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19012985&pc=OZD | 08/13/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 08/13/2022 | 04/01/2024 |  |  | 03/31/2023 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063621 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 77-year-old patient needing mechanical circulatory support due to st segment elevation myocardial infarction. while on support, the patient had hemolysis. possible heparin-induced thrombocytopenia with lactate dehydrogenase at 2000 and platelets dropped to 17. p-level dropped to p-5, with plan to drop to p-3 later to improve hemolysis. transthoracic echocardiogram ordered to confirm position. started mil and increased norepinephrine due to having to turn down impella power. 2 units of platelets given. patient desatted to 85% and suctioned with numerous clots and bleeding noted. impella was down to p3 and drips increased overnight. due to low p level, the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28670 | 1220648-2024-09134 | 19013050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013050&pc=OZD | 03/22/2024 | 05/15/2024 |  | 04/01/2024 |  |  | 03/22/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025463783 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | Material Split, Cut or Torn | [5 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was noted that the pump¿s placement was very lateral and interacting with mitral valve leaflets. repositioning attempting was made. however, the pump always migrated back laterally. the pump was left out of position. while on support, the patient developed a gastrointestinal bleed and the heparin was held due to the bleed. it was also noted that the sterile sleeve was torn. the patient remained on support and was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the sterile sleeve damage was not determined since product was not returned. the root cause of the positioning issues was most likely patient condition related. the pump kept going out of position despite being repositioned. the root cause of the vascular damage gi bleed and its relation to impella were not determined since product was not returned and there is a lack of clinical details.
28671 | 1220648-2024-09135 | 19013073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013073&pc=OZD | 10/08/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 10/08/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202128 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had oozing at the groin. to assist in treating the bleeding at the groin the medical team adjusted anticoagulants by temporarily removing systemic heparin from the purge solution and administering a bolus of integrilin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28672 | 1220648-2024-09136 | 19013125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013125&pc=OZD | 06/24/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 06/24/2022 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 56-year-old male patient needing an impella cp for mechanical circulatory support. it was reported that the impella insertion was aborted due to the patient's anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28673 | 1220648-2024-09137 | 19013134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013134&pc=OZD | 03/22/2024 | 05/07/2024 |  | 04/01/2024 |  |  | 03/22/2024 | 04/01/2024 |  |  | 11/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024424842 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient experienced a nosebleed and developed a hematoma at the access site. heparin was stopped as treatment and the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
28674 | 1220648-2024-09138 | 19013366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013366&pc=OZD | 11/24/2023 | 09/23/2024 |  | 04/01/2024 |  |  | 11/24/2023 | 04/01/2024 |  |  | 06/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379710 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 30-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed some oozing at their impella access site during support. femostop was placed to treat the condition. access site improved and the air from the fem stop has been removed completely but it is still in place for support on the impella sheath. there was some concern for heparin induced thrombocytopenia, so anticoagulant had not been started yet. however, the team was planning on starting a bivalrudin drip. platelets were low but trending up. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and thrombocytopenia could not be determined. added- h6 component code 925. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
28675 | 1220648-2024-09139 | 19013398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013398&pc=OZD | 08/31/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 08/31/2022 | 04/01/2024 |  |  | 05/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231179 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 26-year-old female was implanted with an impella cp device for mechanical circulatory support. during support hemolysis was noted. the patient also reported pain in the right leg. the impella was repositioned due to the hemolysis and the patient reported that the pain lessened. the impella was removed via cutdown by vascular surgery and was replaced with a heartmate 3. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28676 | 1220648-2024-09140 | 19013480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013480&pc=OZD | 10/12/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 10/12/2022 | 04/01/2024 |  |  | 04/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076881 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2022 |  |  |  |  | The complainant reported a patient, race unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the impella 5.5 was placed for more support due to a decreased ejection fraction (ef) function and needing more pressors. it was reported that after the peel-away sheath was removed there was bleeding from the impella site. manual pressure, femostop, and sutures were used to resolve the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28677 | 1220648-2024-09141 | 19013541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013541&pc=OZD | 06/25/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 06/25/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195374 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/25/2022 |  |  |  |  | The user facility reported a 75-year-old male patient on impella cp support due to a st-elevation myocardial infarction (stemi). while on support, the patient experienced a couple of ventricular tachycardia (v-tach) runs through the night and was shocked back. amiodorone and lidocaine drips were started which to resolved the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28678 | 1220648-2024-09142 | 19013569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013569&pc=OZD | 11/06/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/06/2022 | 04/01/2024 |  |  | 07/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253593 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2022 |  |  |  |  | The complainant reported a patient, race unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient was on impella cp support after suffering a vf (ventricular fibrillation) arrest at home. while on support a hematoma was noted. the hematoma was managed by manual pressure; however, heparin was not administered due to bleeding concerns. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28679 | 1220648-2024-09143 | 19013580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013580&pc=OZD | 03/22/2024 | 05/08/2024 |  | 04/01/2024 |  |  | 03/22/2024 | 04/01/2024 |  |  | 09/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025444737 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 70-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp was out of position. the physician attempted unsuccessfully to reposition the device. the device was explanted and a new impella cp was implanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue was most likely use related, per clinical review. in accordance with updated procedures, section d6 has been revised from the initial submission.
28680 | 1220648-2024-09144 | 19013655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013655&pc=OZD | 09/02/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 09/02/2022 | 04/01/2024 |  |  | 06/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235248 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that required placement of a bulb drain. additionally, the pump experienced positioning issues that were causing vigorous coughing, intermittent suction, and large decrease in pulsatility. this was resolved by advancing the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28681 | 1220648-2024-09145 | 19013734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013734&pc=OZD | 11/07/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/07/2022 | 04/01/2024 |  |  | 06/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240833 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The complainant reported a patient, age and race unknown, noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion of the impella a peroration occurred when accessing the right coronary artery (rca), which requiring covered stent and balloon tamponade. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. h3 evaluation summary attached: should be no.
28682 | 1220648-2024-09146 | 19013818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013818&pc=OZD | 03/23/2024 | 05/08/2024 |  | 04/01/2024 |  |  | 03/23/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025443132 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2024 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day seven of support, the patient developed a hematoma at the access site that prompted wound evacuation. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | The investigation into access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28683 | 1220648-2024-09147 | 19013887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013887&pc=OZD | 10/12/2023 | 05/15/2024 |  | 04/01/2024 |  |  | 10/12/2023 | 04/01/2024 |  |  | 08/31/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393769 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 53-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported during impella support, the patient experienced a small hematoma at the groin site and signs of hemolysis with dark red urine. p levels were reduced to p6 and the patient was given iv fluids. manual pressure was applied, and the groin and dressing were clean, dry, and intact. the patient was stable and the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the access site bleeding and the hemolysis issues was not determined since product was not returned and limited clinical information was given.
28684 | 1220648-2024-09148 | 19013912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013912&pc=OZD | 09/02/2022 | 05/28/2025 |  | 04/01/2024 |  |  | 09/02/2022 | 04/01/2024 |  |  | 05/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231666 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support there were hemolysis concerns, and the patient was given volume and switched to bicarb due to address bleeding concerns. the patient was later successfully weaned from impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28685 | 1220648-2024-09149 | 19013946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19013946&pc=OZD | 03/24/2024 | 05/08/2024 |  | 04/01/2024 |  |  | 03/24/2024 | 04/01/2024 |  |  | 06/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427691 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Dislodged or Dislocated | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2024 |  |  |  |  | The user facility reported a 58-year-old male in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was receiving cardiopulmonary resuscitation (cpr), the impella was inserted. during the delivery of cpr, the impella became dislodged, and the decision was made to implant with a new impella cp device. the second impella cp was successfully implanted, and the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was use related with pump being dislodged due to cpr during support. the instructions for use for the impella cp with smartassist for use during cardiogenic shock in the cautions sections states: "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device. check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance".
28686 | 1220648-2024-09150 | 19014030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014030&pc=PYX | 12/10/2023 | 06/17/2024 |  | 04/01/2024 |  |  | 12/10/2023 | 04/01/2024 |  |  | 02/28/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023194525 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 72-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported during impella support, the patient experienced some oozing from their left femoral artery retrograde puncture (lfa rp) access. to address the bleed, the tourniquet band (tb) was tightened, and manual pressure as well as a femostop device were applied. due to coagulopathy, the patient required multiple transfusions. ultimately the team decided to withdraw care due to low likelihood of survival. no allegations were made towards the impella for cause of death. | Upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding is determined to be patient condition because of bleeding from multiple sites and it was also mentioned that patient was coagulopathic. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
28687 | 1220648-2024-09151 | 19014031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014031&pc=OZD | 11/08/2022 | 10/08/2025 |  | 04/01/2024 |  |  | 11/08/2022 | 04/01/2024 |  |  | 06/30/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248182 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Crack | Leak/Splash | Activation, Positioning or Separation Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while patient was on impella cp support, a crack was found after the patient bent their leg and a leak formed at the introducer's hub and valve. to correct the leak, the medical team decided to peel away and insert a repo sheath. during procedure, the impella moved forward into the left ventricle and after the repo sheath was inserted the impella was drawn back into the aorta. multiple attempts were made to reposition the impella back into the left ventricle, but impella could not cross back into the left ventricle. echocardiography was used to verify impella placement, and the impella was shown to be in the aorta and an apical left ventricle thrombus was noted. the medical team decided to leave the pump in the aorta at a p1 and monitor patient. patient remained stable, and was able to be successfully weaned and explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-09151. d6a and d6b revised from the initial manufacturer device report 1220648-2024-09151 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-09151 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-09151 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-09151. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28688 | 1220648-2024-09152 | 19014068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014068&pc=OZD | 03/23/2024 | 05/08/2024 |  | 04/01/2024 |  |  | 03/23/2024 | 04/01/2024 |  |  | 11/30/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024429247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed limb ischemia. there was a loss of pedal pulses and the foot became cool to the touch. the decision was made to wean and remove the impella, which resolved the reported issue. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28689 | 1220648-2024-09153 | 19014109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014109&pc=OZD | 09/08/2022 | 04/01/2024 |  | 04/01/2024 |  |  | 09/08/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210543 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 75-year-old female on impella cp for mechanical circulatory support due to high-risk percutaneous coronary intervention. while on support, at the end of the case the patient was having issues maintaining pulsatility. the patient experienced access site bleeding and required more flow. the impella cp was explanted and implanted with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28690 | 1220648-2024-09154 | 19014186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014186&pc=OZD | 03/24/2024 | 12/31/2024 |  | 04/01/2024 |  |  | 03/24/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025454932 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/24/2024 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the impella cp was unable to run any higher than p1 and the patient had weak distal pulses noted on the right lower extremity. vascular surgery was consulted and the impella was removed. the patient was reported to be stable. | Additional information received that an above knee amputation was performed due to the patient having rhabdomyolysis. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported low pump flow and limb ischemia has been completed. no product was returned. data logs were reviewed and several suction alarms observed with low flow. no motor current spike or other abnormalities were found. the cause of the low pump flow issue was not determined since product was not returned and due to inconclusive evidence per log review and insufficient clinical information. the cause of the low pump flow could not be determined. in accordance with updated procedures, section d6 has been revised from the initial submission. | B2 additional outcome added. b5 updated with additional patient information. h6 health effect - impact code 4626 added to reflect patient outcome the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09154. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing.
28691 | 1220648-2024-09155 | 19014219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014219&pc=OZD | 11/10/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/10/2022 | 04/01/2024 |  |  | 07/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253644 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed access site bleeding when the preclose failed after the peel away was removed. the patient was subsequently taken for surgical removal and exploration. additionally, it was reported that there had been a pump stop, however, no further details had been provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28692 | 1220648-2024-09156 | 19014303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014303&pc=OZD | 11/10/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/10/2022 | 04/01/2024 |  |  | 08/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263940 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The complainant reported a patient, race unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient experienced a loss of pulses in the left lower extremity. the patient was escalated to the impella 5.5 and the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28693 | 1220648-2024-09157 | 19014364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014364&pc=OZD | 03/19/2024 | 06/04/2024 |  | 04/01/2024 |  |  | 03/25/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455811 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 88-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on impella cp support expired due to multiple organ failure secondary to septic shock due to infection. the route of infection is unknown, and there were signs before the impella insertion but impella relationship to the infection is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use referenced in the initial report is for the cp with smartassist in section: table 3.2 impella catheter components. in accordance with updated procedures, sections h10 have been revised from the initial submission.
28694 | 1220648-2024-09158 | 19014480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014480&pc=OZD | 11/14/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/14/2022 | 04/01/2024 |  |  | 07/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 72-year-old female on impella cp for mechanical circulatory support. it was reported the patient was on impella cp support and the patient experienced some cardiac arrhythmias during insertion. the pump was turned off of auto and left at p6 to allow the heart to settle. no further reports of arrhythmia reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28695 | 1220648-2024-09159 | 19014519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014519&pc=OZD | 03/25/2024 | 05/14/2024 |  | 04/01/2024 |  |  | 03/25/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025448769 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 72-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp placement, resistance was met when attempting to advance the companion sheath through the impella sheath valve. imperfection was noted at the distal end of the sheath. this sheath was ultimately pulled off the table in favor of a new companion sheath that did pass successfully. there was no patient harm. | [5 empty] | The investigation for the reported introducer damage has been completed. the introducer was returned for evaluation and was visually inspected. crumpling of the sheath tip was observed while on the dilator. data log review not applicable. if the needle stick in the 14fr introducer is very far out towards the edge of the 14fr valve, the user is most likely to run into the sharp edge of the inner geometry of the peel away hub when doing single access. therefore, the root cause of the 7fr companion sheath tip damage was most likely user related since the needle stick was likely too lateral. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28696 | 1220648-2024-09160 | 19014540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014540&pc=OZD | 09/09/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 09/09/2022 | 04/01/2024 |  |  | 09/30/2023 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137014 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 75-year-old female was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella cp demonstrated low pump flow and suctioning issues and cannula kink was found. the physician was advised to replace the impella and use the recommended peelaway introducer in the left side. the second impella was prepped. a pigtail and wire were passed through the left side introducer but dissected. the sheath was removed and the physician decided to abort replacing due to the dissection and to place the patient on extracorporeal membrane oxygenation (ecmo). ther impella was removed through gore sheath and the right side gore sheath was exchanged for the ecmo cannula and venous access was obtained on the left side for venous ecmo cannula. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28697 | 1220648-2024-09161 | 19014791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014791&pc=OZD | 03/23/2024 | 05/29/2024 |  | 04/01/2024 |  |  | 03/25/2024 | 04/01/2024 |  |  |  |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450482 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 25-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was put on impella cp support following worsening right heart failure condition. it was noted the patient had bleeding complications due to liver failure and hemolysis. there were no impella related complications, but the pump was noted to have been in suction. impella was in optimal position. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the renal failure issues has been completed since the original report was submitted. the pump was not returned for investigation. the data logs showed many suction alarms around the time that hemolysis and renal failure was reported. there are no indicative motor current spikes or a placement signal trend. the root cause of the hemolysis and renal failure cannot be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d6a, d6b.
28698 | 1220648-2024-09162 | 19014961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19014961&pc=OZD | 06/29/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/29/2022 | 04/01/2024 |  |  | 02/29/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202660 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Extravasation | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was a fraction deep and the patient experienced hemolysis as indicated by pink/red urine and lab results. percutaneous coronary intervention of the circumflex artery was attempted. after first balloon inflation, extravasation was noted. the procedure was aborted, but the impella was left in the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: cp and 5.5 ifu: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28699 | 1220648-2024-09163 | 19015495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19015495&pc=OZD | 06/29/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 06/29/2022 | 04/01/2024 |  |  | 07/31/2023 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022119003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that multiple suction alarms occurred and were not resolved with iv fluid bolus. the patient's p-level was decreased to relieve suction and chemical support was increased to maintain mean arterial pressure (map). however, the transesophageal echocardiogram (tee) confirmed the device was in the appropriate placement. subsequently, the patient developed right ventricle (rv) dysfunction and the decision was made to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28700 | 1220648-2024-09164 | 19015649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19015649&pc=OZD | 06/30/2022 | 05/16/2025 |  | 04/01/2024 |  |  | 06/30/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Insufficient Flow or Under Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 89-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed access site bleeding which was resolved by re-suturing and dressing to align with angle of insertion. additionally, the pump presented suction alarms which were resolved by transfusion of blood products and repositioning of the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28701 | 1220648-2024-09165 | 19015689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19015689&pc=OZD | 07/01/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 07/01/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211254 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 84-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during transfer of the patient, the impella cp pulled out, rendering the device unsecured and out of position. the responding physician turned the impella to p0 and removed it from the body. the patient developed a hematoma at the access site, which was addressed by applying pressure to the area. the patient remained stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28702 | 1220648-2024-09166 | 19015830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19015830&pc=OZD | 07/06/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 07/06/2022 | 04/01/2024 |  |  | 03/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022065679 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 63-year-old female in post cardiotomy cardiogenic shock low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was reported that the device alarmed impella in aorta alarm with small but visible motor current. the device was only pulling about .4l of flow. the nurse was instructed on how to disable placement alarms and was advised to use the echo to confirm the placement of the device. additionally, placement signal low alarm appeared, but resolved on its own with volume. pulmonary stenosis low alarms was persistent, due to the patient's pressure not being in correlation. the device was discontinued after 10 days after implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28703 | 1220648-2024-09167 | 19016002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19016002&pc=OZD | 11/07/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/07/2022 | 04/01/2024 |  |  | 11/30/2023 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023162366 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 44-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp was weaned and removed, a post angiogram revealed a large thrombus in the common femoral artery. the patient had absent pulses and mottling in their right foot due to the thrombus. the medical team decided to use a jeti thrombectomy system to successfully remove the thrombus and pulses returned to the patient's foot. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28704 | 1220648-2024-09168 | 19016243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19016243&pc=OZD | 11/11/2022 | 04/07/2025 |  | 04/01/2024 |  |  |  | 04/01/2024 |  |  | 04/24/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204941 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 56-year-old male patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the procedure, the tissue plasminogen activator was infusing through purge when they experienced low purge flow high purge pressure. the pressure improved. no patient harm or injury reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28705 | 1220648-2024-09169 | 19016519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19016519&pc=OZD | 11/12/2022 | 04/01/2024 |  | 04/01/2024 |  |  |  | 04/01/2024 |  |  | 05/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230401 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported during impella support, that the patient experienced limb ischemia in the left leg. the physician placed new antegrade dps to the art line and was flushed every hour. pulses could be heard with the doppler and the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28706 | 1220648-2024-09170 | 19016855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19016855&pc=OZD | 11/12/2022 | 04/07/2025 |  | 04/01/2024 |  |  | 11/12/2022 | 04/01/2024 |  |  | 08/31/2024 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263858 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Positioning Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 36-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced pump positioning issues where the device was almost pulled out. subsequently, the device was repositioned and pulled back one cm the next morning due to decreased irritability with short runs of ventricular tachycardia (vt) patient was successfully weaned from the impella and the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28707 | 1220648-2024-09171 | 19016874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19016874&pc=OZD | 10/25/2023 | 06/24/2024 |  | 04/01/2024 |  |  | 10/25/2023 | 04/01/2024 |  |  | 08/31/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393273 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 51-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported during the exchange of the reposition sheath for the existing 14fr, the impella was inadvertently pulled into the aorta. subsequently, the impella was removed, and the left ventricle was rewired. no further information was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed since the original report was submitted. no product was returned for a visual inspection. no data logs were returned for analysis. the most likely cause of the positioning issues was use related. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 health effect - impact code and medical device problem code were revised in accordance with updated procedures. new codes have been added.
28708 | 1220648-2024-09172 | 19016882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19016882&pc=OZD | 09/11/2023 | 10/14/2024 |  | 04/01/2024 |  |  | 09/11/2023 | 04/01/2024 |  |  | 05/31/2025 |  | 04/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348495 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/11/2023 |  |  |  |  | The user facility reported a 81-year-old male patient implanted with an impella cp device for mechanical circulatory support. it was reported that during the impella implant, the patient experienced bleeding at the impella access site. the patient was administered units of packed red blood cells and fluid to resolve the issue. | The patient was on impella support for 93.1 hours before expiring. the relationship between death and impella is unknown, the impella operated without any problems until death. per the medical safety officer review there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09172: upon further review the cause of death was unknown, requiring a regulatory report. b2 outcomes attributed to adverse event; changed to death b5 patient outcome and mso statement d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h1 type of reportable event. h6 health effect - impact code 1802 missing. h6 component code revised from the initial submission in accordance with updated procedures. | Upon further review the access site bleeding was due to 4fr sheath insertion prior to impella insertion. there is currently no allegation of a malfunction or serious injury associated with the implanted pump. a regulatory report is no longer necessary.
28709 | 1220648-2024-09083 | 19009405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19009405&pc=OZD | 03/08/2024 | 04/26/2024 |  | 03/31/2024 |  |  |  | 03/31/2024 |  |  | 07/31/2025 |  | 03/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377011 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced ectopy and right arm tingling. an echocardiogram was performed which revealed deep placement, but no repositioning was performed at this time. if ectopy persists/increases, or suction events/hemolysis occur then another echocardiogram will be obtained and probable repositioning of the impella. the patient continued to have ectopy and it was reported that it was the patient¿s baseline. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation (vf) has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf could not be determined as the device was not returned nor sufficient clinical details. ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
28710 | 1220648-2024-09084 | 19009409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19009409&pc=PYX | 07/11/2023 | 10/03/2025 |  | 03/31/2024 |  |  |  | 03/31/2024 |  |  | 07/31/2024 |  | 03/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023232096 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/11/2023 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The us complainant had a rp placed for a patient already on the cp with biventricular failure issues. the rp lost placement signal (ps) but remained on till care was withdrawn 2 days later. no harm from the ps loss. | The user facility reported a 72 year old male with a impella rp for mechanical circulatory support. it was reported that the impella rp had lost placement signal (ps) but remained on until care was withdrawn two days later. there was no patient harm reported. |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-09084 was provided in sections b2, b5, g3, g4, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data review: data showed placement signal (ps) was negative & low (24 mmhg) when pump started. signal not reliable and cvp began trending down with suction triggering at 3:31 pm on (b)(6) 2023. lamp, gain, and light were in specification and steady throughout support indicating no defects with the optical line. an alarm 115 (impella stopped - rpm deviations) was observed related to unplugging the pump. about 32 hours & 15 minutes later, ps not reliable alarms were triggered & contrast decreased below 10%. the alarm resolved but the contrast value remained low. no motor current spikes unrelated to p-level changes were observed throughout support. pump flow remained within specification throughout support. product was not returned for review. based on clinical description & data review, the cause of the optical signal issue was most likely patient condition related. clinical reported rv on echo showing very little movement and care was eventually withdrawn. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation has been completed. data showed placement signal (ps) was negative and low (24 mmhg) when the pump started and back to normal range after the pump was re-zeroed. all 5s parameters were in specification. about three hours and 40 minutes into the case, the motor current (mc) spiked and pump stop occurred that triggered an alarm. the pump could be restarted. about 32 hours and 15 minutes later, ps not reliable alarms were triggered and contrast decreased below 10%. re-zeroing resolved the alarm but the contrast value still was out of specification. this event repeated two more times and was resolved with re-zeroing of the pump. about ten hours later, this event reappeared. no re-zeroing was performed at this time. four hours later, the pump was stopped manually and disconnected from the console. the product was not returned for review. based on clinical description and data review, a spike in mc, ps, and purge pressure (pp) indicated that the cause of the low pump flow was biomaterial ingestion.
28711 | 1220648-2024-09085 | 19009413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19009413&pc=OZD | 07/13/2023 | 03/31/2024 |  | 03/31/2024 |  |  |  | 03/31/2024 |  |  | 12/31/2024 |  | 03/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307521 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported there was difficulty placing an impella 5.5 in a 76 year old female due to severe peripheral vascular disease (pvd). a balloon was placed in the axillary/subclavian as a second attempt but also failed. a buddy wire was then placed and the impella 5.5 eventually crossed the aortic valve. however, catheter shaft kinking was noted where the motor and catheter meet. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation has been completed. no product was returned. as per clinical review, the patient had severe peripheral vascular disease at the axillary artery making it difficult to place the impella on first attempt and catheter shaft kinking was noted below the motor housing after multiple attempts. the impella advanced further and was inserted after several attempts with the buddy wire used. the cause of the delivery issue was patient condition related with catheter kink noted during insertion with resistance at axillary artery due to peripheral vascular disease.
28712 | 1220648-2024-09086 | 19009417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19009417&pc=OZD | 08/21/2023 | 03/31/2024 |  | 03/31/2024 |  |  |  | 03/31/2024 |  |  | 07/31/2025 |  | 03/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024372374 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/21/2023 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was access site was bleeding at the anastamosis. the bleed was not treated with any blood products or vascular repair. heparin was held. no lasting harm or injury. | [5 empty] | The investigation has been completed. no product was returned. as per clinical, the patient had bleeding at the access site with activated clotting time (act) values at 208 and heparin was stopped. the cause of the access site bleeding issue was most likely anticoagulation management related due to high patient act values per clinical review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28713 | 1220648-2024-09087 | 19009419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19009419&pc=OZD | 09/27/2023 | 03/31/2024 |  | 03/31/2024 |  |  |  | 03/31/2024 |  |  | 05/31/2025 |  | 03/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368682 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. in the catheterization lab, as soon as the physician advanced the impella over .18 wire in the left ventricle, it was noticed that the mean arterial pressure (map) dropped to 35 mmhg. when the physician did a 2d echocardiogram, there was a massive ar. the impella was not yet started at this point. the physician pulled the impella into the aorta and the pressures were stabilized. the physician decided to complete percutaneous transluminal coronary angioplasty (ptca) while the impella was still hanging in the arch of the aorta. the physician thought that if the impella started, it was not going to be effective given the massive ar (aortic regurgitation). the ptca procedure was completed and after, the impella was explanted. | [5 empty] | The investigation has been completed. no product was returned. the root cause of the delivery issue was due to the massive aortic regurgitation condition of the patient and the doctor decided to explant the impella. hence, the root cause was determined to be patient condition.
28714 | 1220648-2024-09088 | 19009423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19009423&pc=OZD | 10/12/2023 | 03/31/2024 |  | 03/31/2024 |  |  | 10/12/2023 | 03/31/2024 |  |  | 04/30/2025 |  | 03/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340020 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient had bleeding from the impella access site. the site was injected with epinephrine and manual compression was applied. furthermore, the patient developed hemolysis. two units of blood products were administered to the patient to assist with the access site bleeding and hemolysis. | [5 empty] | The investigation has been completed. the purge cassette was returned for evaluation, but the impella cp pump was not returned. as per clinical, the patient developed hemolysis due to suction alarms and was also having bleeding at the impella groin side with partial thromboplastin time (ptt) around 147 seconds. the patient was assessed with two liters of volume to resolve the hemolysis. during bleeding, ptt was found to be high at 147 seconds. the cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical information per clinical. the cause of the hemolysis issue was most likely patient condition related with additional volume given resolving complication per clinical and data log review. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis,¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
28715 | 1220648-2024-08991 | 19002621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002621&pc=OZD | 09/12/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 09/12/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023212727 |  | 550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 70-year-old female female in st elevation myocardial infarction was being placed on an impella 5.5 for mechanical circulatory support. it was reported the cardiac surgeon attempted to implant impella 5.5 for heart disfunction. the attempted implant was not successful due to vascular disease. after several unsuccessful attempts, the cardiac surgeon made the decision to implant the impella cp instead. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28716 | 1220648-2024-08992 | 19002716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002716&pc=OZD | 08/21/2022 | 04/17/2025 |  | 03/29/2024 |  |  | 08/21/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023222969 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a 70-year-old patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the impella alarmed for "impella in aorta". the impella was attempted to be repositioned by medical staff with echocardiography but the pigtail was unable to be advanced through the aortic valve. the patient was hemodynamically stable and already weaned from catecholamines, so the decision was made to explant the impella device. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28717 | 1220648-2024-08993 | 19002775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002775&pc=OZD | 01/30/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 01/30/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1085507 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 01/30/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a console with a battery failure on start up. the aic (automated impella controller) was removed from use and the ifu followed and a backup controller used. no harm or injury. | [5 empty] | The automated impella controller device was received from the customer on (b)(6)2023. data analysis: aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) and a battery level low alarm due to low pack voltage. see ic1211 aic logs complaint date in attachments section. the ltpowerdata from the complaint date showed a cell voltage decline in only one of battery 2¿s cells which occurred as the battery failure alarm triggered (see ic1211 battery analysis according to ltpowerdata.png in attachments). device analysis: the failure mode was reproduced on an engineering bench. batttest further confirmed that battery 2 had an imbalanced cell. the battery failure alarm was due to a defective battery 2. the root cause of the defective battery was due to single cell failure. before sending this console back to the customer. fs replaced the battery set (6120mah, 3900-0032), validated charging, and performed a full functional check on the console and optical system (0042-7316).
28718 | 1220648-2024-08994 | 19002803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002803&pc=OZD | 03/13/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 03/13/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1152882 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/13/2023 |  |  |  |  | No patient involvement: the complainant reported that their aic (automated impella controller) displayed a battery failure alarm upon power up. the device was swapped with a new aic as a result of the failure. | [5 empty] | The impella device was received from the customer on 21-mar-2023. data analysis: aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360). see ic2949 aic logs complaint date in attachments section. the ltpowerdata from the complaint date showed a cell voltage decline in only one of battery 1¿s cells which occurred as the battery failure alarm triggered (see ic2949 battery analysis according to ltpowerdata.png in attachments). device analysis: the failure mode was reproduced on an engineering bench. the battery lcd indicator was blank (fully discharged). battery 1 was replaced to resolve the failure alarm. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. before sending this console back to the customer. fs replaced the battery set (6120mah, 3900-0032), validated charging, and performed a full functional check on the console and optical system (0042-7316).
28719 | 1220648-2024-08995 | 19002818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002818&pc=OZD | 05/17/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 05/17/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1576134 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/17/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant in prep of the console for a cp support found the purge disc to be broken. the aic (automated impella controller) was replaced per ifu (instructions for use) recommendation with a backup and support proceeded without any noted harm or injury from delay. support was successful for 94 hours before patient expired on support. no allegation impella caused or contributed to the patient outcome. | [5 empty] | The impella device was received from the customer on 07-jul-2023. data analysis: on the complaint date, console ic8809 received ¿purge fluid not detected error¿ upon inserting the purge cassette (s/n: (b)(6) lot 1663628) after case start step 3, which denotes that at this point the purge retainer and flag might be broken/missing. it confirms the reported failure. device analysis: the console ic8809 was tested after arriving in fs. upon visual inspection, it was confirmed that the purge blue retainer and the flag were missing. no dextrose residue was found. see the attachments in the complaint folder 23-15016 for the missing blue retainer and the flag. the root cause for the missing purge retainer and the flag was due to the defective component ((b)(6)purge pressure sensor). before returning this console back to the customer, fs was instructed to perform the following, install new purge pressure sensor ((b)(6)). validate sensor accuracy via section 12 of the pm procedure ((b)(6)). perform a full functional check on the console and optical system ((b)(6)).
28720 | 1220648-2024-08996 | 19002852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002852&pc=OZD | 06/07/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 06/08/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1327081 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/08/2023 |  |  |  |  | No patient involvement: the complainant reported that a (white) alarm on the console displayed ¿console running on battery power¿. at that time the console was plugged into an outlet. the second alert on the aic (automated impella controller) indicated in its display a ¿cord icon¿ with a red ¿x¿. lastly, a (white) alarm alerted the user of an ¿unexpected controller shutdown¿. the decision was made to use a backup aic for the procedure. no patient was involved. no harm or injury was reported. | [5 empty] | The impella device was received from the customer on (b)(6)2023. data analysis: aic logs confirmed the reported issues. the console did not recognize the ac mains power had been connected. additionally, the console unexpected shutdown which was likely related to the user toggling the transport connection switch during device troubleshooting in the field. device analysis: upon arrival, the failure mode was confirmed. the connection of ac power was not recognized by the console. there was 120v input into the ac/dc power supply, but there was not the expected 24.0 v output to the pbm, rather 3.2 v. the root cause of the ac power issues was due to a defective ac/dc power supply. before sending this console back to the customer, fs was instructed to replace the ac/dc power supply (4050-0041) and perform sections 20 and 25 of the pm procedure. fs also routinely performs a full functional check on the console and optical system (0042-7316).
28721 | 1220648-2024-08998 | 19002860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002860&pc=OZD | 03/28/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 03/28/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1574757 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/28/2023 |  |  |  |  | The complainant reported an automated impella controller (aic) battery failure. no patient was involved. a loaner was provided. | [5 empty] | The investigation has been completed. after analyzing the automated impella controller (aic), it was confirmed that several events of battery failure occurred on the event date. the ltpowerdata before the complaint date showed cell imbalance with battery one cell four which caused battery pack internal electronics to shut down that battery. the console was tested after arriving. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the cause of the aic issue was a defective battery.
28722 | 1220648-2024-08997 | 19002865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002865&pc=OZD | 08/22/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/22/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Pressure Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, there were alarms for high pressure. the medical team reports this was likely do to the cessation of systemic heparin due to oozing from the impella access site. systemic heparin was resumed to resolve the alarms. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records
28723 | 1220648-2024-08999 | 19002869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002869&pc=OZD | 06/22/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 06/22/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1174507 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/22/2023 |  |  |  |  | The user facility reported an automated impella controller (aic) was being returned for a battery failure, red alarm. a back up console had been used for the patient. there was no patient harm reported. | [5 empty] | The investigation for the reported aic alarm issue has been completed. the aic was returned for evaluation. during functional testing, the alarm produced suggests that the battery transport switch had been engaged. data logs were reviewed do not confirm the reported alarm and show that the console was operating within specification. the root cause was determined to be disengaged/reengaged battery transport switch.
28724 | 1220648-2024-09000 | 19002921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002921&pc=OZD | 06/23/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 06/23/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119395 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/23/2023 |  |  |  |  | The user facility reported a 77-year-old male patient needing mechanical circulatory support. it was reported that the automated impella controller (aic) door does not open without having to manually pull it. there was no patient harm reported. | [5 empty] | The investigation into the purge door issue has been completed. the console was returned for evaluation. visual evaluation observed build up of purge solution near the purge door latch assembly. data logs were not relevant to the door issue. therefore the root cause of the purge door issue was due to observed build up of purge solution near the purge door latch assembly
28725 | 1220648-2024-09001 | 19002966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002966&pc=OZD | 06/27/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 06/27/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1309666 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported an automated impella controller (aic) "failure to boot". it was taken to biomed for troubleshooting but was unsuccessful. the console was taken out of service. there was no patient harm as a result of the issue. | [5 empty] | The investigation for the reported aic "system self check" error has been completed. the aic was returned for evaluation and visual inspection found that pbm board had corroded copper traces near the super capacitors. logs were reviewed revealed multiple communication issues between the pbm and the epc. the root cause for the aic not powering on was due to loss of communication from the pbm to the 4 in 1 caused by corrosion of the copper traces by leaking electrolyte of the super capacitors.
28726 | 1220648-2024-09002 | 19002993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19002993&pc=OZD | 04/26/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 04/26/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133096 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/26/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an automated impella controller (aic) with a battery failure. the team attempted to troubleshoot and found the aic needed replaced. instructions for use were followed and a backup was utilized for patient use. no harm or injury was reported. | [5 empty] | The investigation has been completed. logs confirmed the reported failure. there was a battery failure alarm transport connection blown/ missing and a battery level low alarm due to low pack voltage. the ltpowerdata from the complaint date showed a cell voltage decline in only one of battery one¿s cells which occurred as the battery failure alarm triggered. the battery failure alarm transport connection blown/ missing and battery level low alarm were caused by failure of battery one due to internal cell imbalance. the cause of the automated impella controller (aic) issue was by failure of battery one (defective battery) due to internal cell imbalance.
28727 | 1220648-2024-09003 | 19003025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003025&pc=OZD | 07/01/2023 | 09/26/2024 |  | 03/29/2024 |  |  | 07/01/2023 | 03/29/2024 |  |  | 03/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333635 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2023 |  |  |  |  | The user facility reported an impella cp in a 68yo male patient for mechanical circulatory support. it was reported that the patient experienced suction problems and a thrombus was observed at the pump outlet. the team made the choice not to explant or replace the pump as the patient condition was poor. the patient subsequently expired from multiple organ failure unrelated to use of the device. | [5 empty] | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09003. b1 updated to adverse event. b2 outcomes: death and date of death. f6/f8-should have been left blank. g1 mdr reporting contact name and address missing. h1 type of reportable event: death. h.6 code 4611 was reported incorrectly. new code was added to health effect - impact code. | The investigation for the reported low pump flow and thrombus has been completed. the impella device nor the data logs were returned for evaluation. however, clinical details provided states that an echo revealed a thrombus at the pump outlet. therefore, the most likely root cause of the low pump flow was due to patient condition, as biomaterial was found in the outlet, as per clinical details. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ it also states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28728 | 1220648-2024-09015 | 19003420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003420&pc=OZD | 06/09/2023 | 09/11/2024 |  | 03/29/2024 |  |  | 06/09/2023 | 03/29/2024 |  |  | 11/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311193 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Pressure Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2023 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella support for 57 hours when there was high purge pressure. the complainant reached out to unit after seeing pump stoppage on impella connect. they attempted to restart twice with no success. advised to pull pump past valve to prevent aortic insufficiency and to treat according to team management with goals of care if not replacing. patient expired on support. the death is unrelated to the pump malfunction. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation has been completed. no product was returned. data logs were reviewed and found that there was a rise in purge pressure with purge flow decreasing and purge flow ticks were stalling. during unkinking of the purge line, there was a sudden decrease in purge pressure observed. the cause of the high purge pressure issue was most likely a kinked purge line in the catheter based on data log review. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-09015 was provided in sections b2, b5, h1, and h6.
28729 | 1220648-2024-09017 | 19003540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003540&pc=OZD | 06/15/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 06/15/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196892 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 63-year-old female on impella cp for mechanical circulatory support. while on support, the patient experienced limb ischemia. ecmo (extracorporeal membrane oxygenation) flow was reduced to 3.5l/min. the impella cp was still at p-2. va (venoarterial) emco was the most appropriate therapy and used impella cp as vent. emco was placed and both l (left) and r (right) distal antegrade catheters placed for bilateral perfusion through emco circuit. ultimately, the patient was taken to or to have the impella cp and emco removed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
28730 | 1220648-2024-09018 | 19003561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003561&pc=OZD | 09/11/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 09/11/2022 | 03/29/2024 |  |  | 11/30/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168785 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a 29-year-old male was implanted with an impella cp device for mechanical circulatory support. the impella cp encountered low pump flow issues. the pump had constant suction and multiple attempts were made to reposition the pump despite being in the appropriate location per the echocardiography. subsequently, the decision was made to explant the pump due to the constant suction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28731 | 1220648-2024-09019 | 19003587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003587&pc=OZD | 07/12/2023 | 10/14/2024 |  | 03/29/2024 |  |  | 07/13/2023 | 03/29/2024 |  |  | 08/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124695 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 33 year old male on an impella cp for mechanical circulatory support. it was reported that there was bleeding around repositioning. staff performed an angle match, manual pressure was held, and femstop was applied. during insertion, the impella was initially deep and slack was removed. during this time, the patient had no urine output. as a result, lasix was administered and the urine appeared pink and slowly grew darker red. angiomax had been used to anticoagulate and the activated clotting time (act) was higher than 400. | The patient experienced intermittent ventricular tachycardia, which resolved following the physician's adjustments to the pump position and the removal of slack. |  |  |  |  | The investigation for the hemolysis and the access site bleed has been completed. the impella device was not returned for evaluation. data logs were reviewed and no positioning alarms, motor current spike or suction issue were found. placement signals were trending down during the event of hemolysis. clinical details state that urine output appeared to be darker red and red frothy fluid accumulated in lungs and congested, lasix was administered. bleeding at access site was around repositioning sheath unit and angiomax had been used to anticoagulate and act was higher than 400 during bleeding. the cause of the access site bleeding issue was due to anticoagulation management related with high patient act values greater than 400 observed during bleeding per clinical. the cause of the hemolysis issue was most likely patient condition related per clinical and log review. the instructions for use states ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected low s or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. it also states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Additional information for the initial mfr 1220648-2024-09019 was provided in sections b5, b7, d6a, d6b, h1, h6. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28732 | 1220648-2024-09021 | 19003598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003598&pc=OZD | 08/25/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/25/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231114 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. patient experienced ventricular tachycardia, but the impella device maintained a mean arterial pressure of 70. the patient was cardioverted a number of times and decision made to take patient to the cardiac catheterization lab and image. the impella was repositioned under fluoroscope. upon removal of impella cp, the angioseal failed and the arteriotomy began bleeding significantly. a balloon tamponade was used to cover a deployed stent and hemostats was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28733 | 1220648-2024-09024 | 19003707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003707&pc=OZD | 06/19/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 06/19/2022 | 03/29/2024 |  |  | 01/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a 63-year-old male on impella cp for mechanical circulatory support. it was reported that the patient on impella cp had ooze from sites. the registered nurse also stated that systemic heparin is off. the concentration of purge was reduced to resolve the ooze. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28734 | 1220648-2024-09025 | 19003740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003740&pc=OZD | 06/21/2022 | 03/11/2025 |  | 03/29/2024 |  |  | 06/21/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206515 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient experienced slow oozing and was provided 3 units of packed red blood cells. the patient was reported to be stable and remained on impella cp support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ impella cp: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28735 | 1220648-2024-09026 | 19003768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003768&pc=OZD | 09/15/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 09/15/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199892 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/15/2022 |  |  |  |  | The user facility reported a 39-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. en route to the hospital, the groin started bleeding and pressure was applied and mattress suture applied to groin site to obtain hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. this report is being filed as part of a retrospective review of historical records.
28736 | 1220648-2024-09027 | 19003809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003809&pc=OZD | 06/22/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 06/22/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438552 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/22/2023 |  |  |  |  | The user facility reported an automated impella controller (aic) had a "purge cassette not detected" alarm. a loaner was sent for use. there was no patient harm reported. | [5 empty] | The investigation has been completed. the impella has been returned for evaluation. logs confirmed the reported complaint of the purge fluid not being detected as evidenced by multiple logs of purge pressure errors. even though there were a couple of errors, there was no evidence of purge cassette not detected errors. visual inspection revealed dextrose residue within the purge cassette cabin as well as within the purge pressure sensor. the blue purge retainer was observed to not be flush with the pressure foil membrane. this resulted in the displacement of the purge flag. the root cause of the purge fluid not detected error was due to a broken purge retainer, which impacted the accuracy of the pressure reading. the cause of the issue was a defective component.
28737 | 1220648-2024-09030 | 19003911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003911&pc=OZD | 09/18/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 09/18/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 72-year-old female was implanted with an impella cp device for mechanical circulatory support. after the impella cp was weaned and removed, a post angiogram revealed a small thrombus burden at bifurcation of right superficial femoral artery and profunda. it was successfully aspirated with a thrombectomy device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined this report is being filed as part of a retrospective review of historical records.
28738 | 1220648-2024-09028 | 19003950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19003950&pc=OZD | 09/05/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 09/05/2023 | 03/29/2024 |  |  | 06/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024376314 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2023 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp insertion that was aborted due to native anatomy. the team had inserted the 14fr short and tried to advance the pump but, due to stenosis was unable to do so. the insertion was aborted and there was no patient injury or pump damage. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. per the clinical information provided, the root cause of the delivery issue was patient condition related to stenosis.
28739 | 1220648-2024-09031 | 19004037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004037&pc=OZD | 06/23/2022 | 03/11/2025 |  | 03/29/2024 |  |  | 06/23/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211625 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during support, the patient failed to meet index and had little to no pulsatility on the arterial line. the decision to escalate to peripheral venoarterial extracorporeal membrane oxygenation was made and the patient taken was taken to ccl where venoarterial extracorporeal membrane oxygenation (ecmo) was placed. the patient started to regain some pulsatility on arterial line. heparin was removed from purge due to partial thromboplastin times consistently coming back at over 200. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28740 | 1220648-2024-09032 | 19004182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004182&pc=OZD | 06/28/2022 | 03/07/2025 |  | 03/29/2024 |  |  | 06/28/2022 | 03/29/2024 |  |  | 01/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023187702 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the physician repositioned the pump with the echo during the night, due to the patient having increased ectopy which was resolved. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
28741 | 1220648-2024-09033 | 19004208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004208&pc=OZD | 09/18/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 09/18/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203576 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 83-year-old male was implanted with an impella cp device for mechanical circulatory support and that the patient had access site bleeding noted by the hematoma and oozing at the groin site and limb ischemia as patient lost pulsatility in their distal leg. the access site bleeding was treated by applying manual pressure and giving the patient two units of blood. the decision was made to explant the impella cp device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28742 | 1220648-2024-09034 | 19004282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004282&pc=OZD | 05/31/2023 | 08/24/2024 |  | 03/29/2024 |  |  | 06/24/2023 | 03/29/2024 |  |  | 10/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276769 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2023 |  |  |  |  | The user facility reported a 53-year-old male patient on an impella cp for high risk surgery. it was reported that the flow rate on the impella suddenly dropped to zero unexpectedly during patient support. it was noted that spike like waveforms had appeared occasionally around the time of the event. during observation, the flow rate on the display returned and no further complications were encountered. the patient died of multiple organ failure due to circulatory failure, unrelated to device. there was no patient harm resulting from the notes issue. | [5 empty] | The investigation has been completed. the impella has been returned for evaluation. visual inspection found a cannula kink near the outlet cage. data logs were reviewed which revealed multiple suction alarms and low motor currents. the patient died due to circulatory failure unrelated to the impella. therefore, the cause of the low pump flow most likely was use related issue due to cannula kink. | D6a and d6b added on manufacturer device report 1220648-2024-09034. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-09034. b1 updated to adverse event. b2 outcomes: death and date of death. g1 mdr reporting contact facility and fax number missing. h1 type of reportable event: death. h.6 code 2199 was reported incorrectly. new code was added to health effect - impact code.
28743 | 1220648-2024-09035 | 19004320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004320&pc=OZD | 06/30/2022 | 03/11/2025 |  | 03/29/2024 |  |  | 06/30/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202142 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient later developed a small hematoma at the access site. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28744 | 1220648-2024-09036 | 19004361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004361&pc=OZD | 09/01/2022 | 03/10/2025 |  | 03/29/2024 |  |  | 09/01/2022 | 03/29/2024 |  |  | 12/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023178332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a low platelet count, and systemic heparin was discontinued as a result. the impella cp had suction and placement signal alarms. transesophageal echocardiography (tee) was used to observe the impella had been pulled back into the aorta. the impella was repositioned under tee and advanced into the left ventricle. there were no further reports of positioning alarms or concerns. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning and thrombocytopenia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
28745 | 1220648-2024-09037 | 19004401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004401&pc=OZD | 06/29/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 06/29/2023 | 03/29/2024 |  |  | 05/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024343784 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2023 |  |  |  |  | The user facility reported an impella 5.5 was implanted in a 52 year old female patient for mechanical circulatory support. it was reported that there was a high purge pressure during impella support. the purge cassette was replaced to resolve the issue. there was no patient harm reported. | [5 empty] | The investigation has been completed. the impella device was not returned for evaluation. logs were not returned for review. the root cause of the high purge pressure could not be determined due to insufficient clinical information and no logs or product returned. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28746 | 1220648-2024-09038 | 19004474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004474&pc=OZD | 09/19/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 09/19/2022 | 03/29/2024 |  |  | 06/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237599 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the peel away sheath was removed and peeled away outside of the patient, per guidance, but hemostasis was not achieved with advancing repositioning sheath. manual pressure was held until access was achieved on the contralateral side. peripheral balloon was used to tamponade impella access site after impella was weaned and removed to properly manage bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records
28747 | 1220648-2024-09039 | 19004513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004513&pc=OZD | 09/06/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 09/06/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234881 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while attempting to place the impella, there were difficulties advancing the sheath, which were resolved by using a shorter sheath with the same dilator. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28748 | 1220648-2024-09041 | 19004652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004652&pc=OZD | 09/26/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 09/26/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209783 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 73-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support hemolysis and limb ischemia was noted. it was difficult to assess right femoral pedal pulses. very faint pulse was found. warming device applied to bilateral lower extremities for better circulation. bair hugger monitoring system used. patient remained on support for 5 days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28749 | 1220648-2024-09042 | 19004683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004683&pc=OZD | 09/06/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 09/06/2022 | 03/29/2024 |  |  | 09/30/2022 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1504708 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk coronary artery bypass graft was being implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support, there was bleeding at the impella access site at the right groin. pressure was applied to the site and a suture was applied to resolve the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28750 | 1220648-2024-08880 | 19004696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004696&pc=OZD | 12/05/2022 | 03/27/2024 |  | 03/29/2024 |  |  | 12/05/2022 | 03/29/2024 |  |  | 07/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259465 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient exhibited hemolysis as recognized by bloody urine output. the device was turned down to p6 and eventually replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28751 | 1220648-2024-08975 | 19004873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19004873&pc=OZD | 08/28/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/28/2022 | 03/29/2024 |  |  | 07/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump encountered positioning issues as it was getting pulled into the aorta. as a result, the pump was explanted as there was no possibility for repositioning of the pump as it was not visible with echo. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28752 | 1220648-2024-08915 | 19005007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005007&pc=OZD | 03/19/2024 | 05/13/2024 |  | 03/29/2024 |  |  | 03/20/2024 | 03/29/2024 |  |  | 09/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025444637 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that an implanted impella cp pump ingested a clot while coming off cardiopulmonary bypass (cpb). the motor current subsequently spiked and the pump experienced continuous suction despite good positioning and volume. as a result of the noted issue, the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. data logs confirmed the failure mode. the logs showed after the pump was started and ran about two hours 22 minutes, multiple motor currents spiked followed the low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 name prefix/title, first name and last name were revised as information previously reported was incorrect. e.3 revised occupation to go along with the corrections in e.1.
28753 | 1220648-2024-08914 | 19005032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005032&pc=OZD | 03/20/2024 | 05/09/2024 |  | 03/29/2024 |  |  | 03/20/2024 | 03/29/2024 |  |  | 11/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2025436856 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 61-year-old female with cardiogenic shock and with a compromised left anterior descending artery and left circumflex artery was implanted with an impella device for mechanical circulatory support. during the impella device explantation, a clot was found in the right common femoral artery. a balloon was inflated in the area of the clot. an angiogram was then taken, and the balloon was removed. the patient remained on support in stable condition. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number was revised from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 name prefix/title was inadvertently left off the previously submitted report and was added. g.1 revised manufacturing site name and address section as previously entered incorrectly. h.6 codes 4414 and 4627 were reported incorrectly on the previously submitted report. the correct odes have been added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report.
28754 | 1220648-2024-09043 | 19005195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005195&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/29/2024 |  |  | 03/15/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025442086 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that high purge pressure alarms occurred on the impella 5.5. the clinician tried all troubleshooting techniques without resolution of alarm. the clinician will try to change the purge cassette and the tpa protocol. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the high purge pressure cannot be determined with insufficient data logs and product to investigate. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ to conform to updated procedures, sections d6 & h10 were revised with updated information.
28755 | 1220648-2024-09046 | 19005259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005259&pc=OZD | 09/07/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 09/07/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 64-year-old female noted as high risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion of the impella, the physician encountered delivery issues due to the patient's anatomy. as a result, the impella cp insertion was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28756 | 1220648-2024-08976 | 19005284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005284&pc=OZD | 06/23/2022 | 03/28/2024 |  | 03/29/2024 |  |  | 06/23/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198113 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Detachment of Device or Device Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a 70-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. maintaining systolic blood pressure (sbp) was an issue but was achieved with medicine. when the physician went in to balloon the left anterior descending (lad) artery, the patient went flat on the impella. perforation was noted in the lad. after the stents were able to be placed, pressures increased and pulsatility came back. the impella was left in overnight since the patient wasn¿t tolerating the wean. the peel away sheath was removed; however, broke apart without peeling away. once repo sheath was in place, groin was bleeding. anesthesia started the patient on dobutamine. sbp improved and the impella was removed. hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28757 | 1220648-2024-09047 | 19005287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005287&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/29/2024 |  |  | 03/15/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450274 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had cool legs and no doppler pulses were located. the impella was removed due to the ischemia. the staff knew this would happen due to the patient's anatomy. they were hopeful that the reperfusion sheath would allow for flow, however the patient still had ischemia and the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ to conform to updated procedures, sections d6 & h10 were revised with updated information.
28758 | 1220648-2024-09049 | 19005294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005294&pc=OZD | 03/20/2024 | 06/11/2024 |  | 03/29/2024 |  |  | 03/20/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, UK |  |  | ABIOMED, INC. |  |  | US |  | 1502178 |  | 0042-0010-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 82-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during prep, the automated impella controller (aic) was turned on and displayed 1 red alarm (battery failure) and 1 yellow alarm (battery communication failure). the site confirmed the console had not been turned on in months. the procedure was still performed successfully and there was no reported patient harm. there were no issues reported during implant, during support, or during explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issue has been completed since the original report was submitted. the device was not returned for investigation. based on problem description, it is most likely that console batteries were fully depleted, however we could not verify this because neither the console nor its data was returned for investigation. the root cause could not be determined. in accordance with updated procedures, section d2 has been revised from the initial submission.
28759 | 1220648-2024-09048 | 19005326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005326&pc=OZD | 11/02/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 11/02/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201249 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during placement the impella would not make the turn on the aortic arch and seemed to be stuck on something even when it was attempted to be removed. the physician was going to attempt to start over with the insertion when the patient went into complete heart block, asystole. the patient coded and cpr started, epi given and a pacer was placed left groin. after return of spontaneous circulation was achieved, the decision was made to place an intra-aortic balloon pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28760 | 1220648-2024-09050 | 19005352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005352&pc=OZD | 03/21/2024 | 05/10/2024 |  | 03/29/2024 |  |  | 03/21/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025454677 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during planned explant of the impella cp device, the introducer buckled as the pump was being removed and began to split. the pump had to be removed concurrently with the 14 fr sheath as a result of the complication. there was no patient harm. | [5 empty] | The impella device was received from the customer on (b)(6)2024 and therefore,an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the difficulty inserting/removing pump through introducer issue has been completed since the original report was filed. the pump was returned for analysis. the pump was stuck in the peel away sheath along with a stiff buddy wire. the returned product shows that correct insertion practices were not followed. the root cause of the difficulty removing pump through introducer issue was use related to improper technique.
28761 | 1220648-2024-09051 | 19005356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005356&pc=OZD | 06/14/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/05/2023 | 03/29/2024 |  |  | 02/28/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322884 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fracture | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported a 67 year old male was on impella cp support for mechanical circulatory support. it was reported that while the patient was undergoing impella sheath insertion, the hemostatic valve was broken and there was blood leakage reported from the hemostatic valve of the introducer. additionally, bleeding was noted from the access site. the procedure was stopped and a new sheath was inserted from a different route. there was no additional treatment other than repositioning sheath replacement. | [5 empty] | The investigation has been completed. the pump was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical information was provided. additionally, the introducer was not returned for evaluation. the root cause of the introducer damage - mechanical could not be determined. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28762 | 1220648-2024-09053 | 19005442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005442&pc=OZD | 03/21/2024 | 05/02/2024 |  | 03/29/2024 |  |  | 03/21/2024 | 03/29/2024 |  |  | 10/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025467849 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump began to experience high purge pressure during patient support. to counteract the noted issue, a tpa protocol was initiated. the alarm was disabled and support with the implanted pump continued. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the high purge pressure was not determined since product was not returned. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ to conform to updated procedures, section d6 & h10 were revised with updated information.
28763 | 1220648-2024-09054 | 19005498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005498&pc=OZD | 03/22/2024 | 05/10/2024 |  | 03/29/2024 |  |  | 03/22/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433484 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Break | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during placement of the impella cp, the short 14f peel away sheath broke. part of the clear diaphragm came out of the part it¿s secured in. the impella was successfully placed. it was further reported that after insertion of the cp in the cath lab, distal pulses checked and no doppler on either foot. md declined fem-fem bypass due to patient¿s distal disease in the superficial femoral artery and the patient¿s vessel being small. the patient remained on support for three additional days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and introducer damage has been completed. the device was not returned for evaluation. clinical details were not sufficient to determine root cause. therefore the root cause of the limb ischemia and introducer damage has been completed. added- d6a/d6b.
28764 | 1220648-2024-09056 | 19005599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005599&pc=OZD | 11/07/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 11/07/2022 | 03/29/2024 |  |  | 06/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238874 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 37-year-old male patient was being implanted with an impella cp device for mechanical circulatory support and that the pump had difficulty in advancing due to the femoral artery being too small for the impella sheath. therefore, the medical team planned for a reperfusion catheter. while the patient was on impella cp support, the patient experienced limb ischemia in both extremities. the pump was repositioned, during repositioning the patient went into ventricular tachycardia and a left atrial appendage thrombus was found. no additional information was provided for the treatment of the vt and thrombus. patient was monitored until impella cp device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
28765 | 1220648-2024-09057 | 19005606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005606&pc=OZD | 07/14/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 07/14/2022 | 03/29/2024 |  |  | 09/30/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144138 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the impella cp required repositioning as it was noted to have been pushed under the papillary muscle and was unable to be freed. the physician opted to explant the impella. closure was successful utilizing the preclose with perclose proglide x2 method. the patient survived and received post-case support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28766 | 1220648-2024-09058 | 19005619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005619&pc=OZD | 03/25/2024 | 05/16/2024 |  | 03/29/2024 |  |  | 03/25/2024 | 03/29/2024 |  |  | 07/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443622 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 24-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that flows are lower than expected on the impella cp. the clinician was not able to turn p-level up without suction alarms. on impella connect a motor spike was observed, lower than expected flows, and copious suction alarms despite good position and volume status. the decision was made to explant the cp. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported low pump flow has been completed. the impella pump was returned and visually inspected. biomaterial observed around the impeller. no cannula kink, broken blades, nor other abnormalities were found. the cause of the low pump flow and suction alarms was biomaterial ingestion. added-d9 device return date d4-updated model number added-d6a/d6b.
28767 | 1220648-2024-09059 | 19005657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005657&pc=OZD | 03/26/2024 | 09/23/2025 |  | 03/29/2024 |  |  | 03/26/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447076 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/26/2024 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed an access site hematoma and heparin was put on hold. the patient was reported as stable. | The complainant also reported that there were placement signal alarms, and the staff disabled the audio and monitored. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the optical signal issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the optical signal issue and the access site bleeding was not determined since product was not returned, some data logs are missing, and there are limited clinical details. d6a | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs show that about 110 hours into support, the placement signal spikes to 3000 mmhg, the signal-to-noise ratio drops to 0, light decreases, lamp goes to 100. contrast increases in altitude, and gain increases. the root cause of the optical signal issue was not determined since product was not returned and some data logs are missing. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number. b7 other relevant history, including preexisting medical conditions updated. d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 have erroneous entries in the initial report - should have been blank. g1 reporting contact email and reporting contact fax number updated.
28768 | 1220648-2024-09061 | 19005682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005682&pc=OZD | 03/26/2024 | 05/09/2024 |  | 03/29/2024 |  |  | 03/26/2024 | 03/29/2024 |  |  | 11/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024428521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The user facility reported an 80-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient developed hemolysis and the pump was removed. upon explant, the patient had limited flows down leg after impella removal. a thrombectomy was performed and two large clots were found. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported limb ischemia, thrombus, and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. data logs were reviewed finding there was no evidence of positioning issues, suction alarms or motor currents (mc) spikes that were recorded in the logs. therefore, the root cause of the limb ischemia, thrombus, and hemolysis could not be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient condition including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28769 | 1220648-2024-09062 | 19005687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005687&pc=OZD | 07/05/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/05/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300416 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/05/2023 |  |  |  |  | It was reported by the user facility's biomedical engineer that the console had a cut in the power cable. there was no patient involvement. a backup console was used for patient use. | [5 empty] | The investigation has been completed. the logs were not necessary for this type of complaint. the console was visually inspected and a cut in the power cable was confirmed. the cause of the issue was a defective power cable.
28770 | 1220648-2024-09064 | 19005708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005708&pc=OZD | 08/02/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 08/02/2022 | 03/29/2024 |  |  | 12/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164216 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced lower extremity ischemia, cold sensation at the access site, and was noted to be slightly purple in color. the patient was successfully weaned from the impella cp and the device was explanted leaving the sheath in place. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28771 | 1220648-2024-09063 | 19005709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005709&pc=OZD | 03/27/2024 | 04/23/2024 |  | 03/29/2024 |  |  | 03/27/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1109482 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 68-year-old female undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support there was an unexpected automated impella controller (aic) shutdown. the shutdown prompted the team to reboot. the patient had lost pressure momentarily during the reboot. aic is again running, and team will return the console for investigation. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - pump stop issue has been completed since the initial report was submitted. the console was returned and tested. the failure mode was not reproduced naturally while running a pump during the investigation. the cause of the aic issue was the 3rd party hardware/software that runs io-graphics.
28772 | 1220648-2024-09065 | 19005855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005855&pc=PYX | 07/09/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/09/2023 | 03/29/2024 |  |  | 05/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024355189 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/09/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. while on impella rp support, the p level slowly increased from p-2 to p-6. pump stop alarm was note. the pump defaulted to p-0 with no flow. the pump was immediately placed back to the previous p level and restored without complication. the staff continued to monitor the patient and impella. once in the intensive care unit, another pump stop occurred. the pump was explanted and the patient was transferred for higher level care. no patient harm was reported due to the issue. | [5 empty] | The investigation has been completed. the product was returned for evaluation. data logs were reviewed, and motor current high alarms were observed with pump stop. multiple motor current spikes were observed initially after the case start within five minutes but after was able run at p-6 at low flows. low flows were observed at p-level higher than p-5. again, pump stop was observed at the end after two hours. the pump was visually inspected and found biomaterial ingested on the impeller. no other abnormalities were found during the case review. the cause of the pump stop issue was due to ingested biomaterial over the impeller per field investigation and log review. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
28773 | 1220648-2024-09066 | 19005878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005878&pc=OZD | 08/05/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/05/2022 | 03/29/2024 |  |  | 01/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma. the antegrade sheath was removed and subsequently the impella access site was noted to be clean/dry/soft. additionally, it was reported that no doppler pulses were found on either of the patients feet. the patient was taken to operating room to decannulate. the impella site was clean, the patient received 1 unit of blood, there were no alarms, urine was clear, and distal pulses present. the pump was successfully explanted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28774 | 1220648-2024-09067 | 19005937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19005937&pc=OZD | 08/30/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/30/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213228 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support the patient experienced oozing at the impella access site. the medical staff noted the knee immobilizer was pulling on the site and was adjusted to alleviate the oozing. it was also noted during this time that the patient's activated partial thromboplastin clotting time (aptt) was >100 seconds with a target of 60 seconds, the heparin was discontinued temporarily as a result. the bleeding resolved with no further intervention. the patient also experienced red urine, hemolyzed labs, and evidence of acute kidney failure with elevated kidney lab values . the pump was repositioned and urine cleared, but later on red urine was noted. the pump was turned down, and urine began to clear. there was insufficient information provided to determine if the patient was treated further for the hemolysis experienced while on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28775 | 1220648-2024-09055 | 19006018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006018&pc=OZD | 03/24/2024 | 05/10/2024 |  | 03/29/2024 |  |  | 03/24/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025451032 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that an impella cp device stopped functioning after being successfully placed in a patient. the device performed for 24 hours with lower than expected flow rate but then the purge rate dropped dramatically and the pump alarmed intermittently, and then stopped. several unsuccessful attempts to restart the pump were made. the pump was successfully removed from the patient and no harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was submitted. the device was not returned for investigation. the data logs show multiple impella position alarms and a gradual purge pressure rise followed by a motor current spike. the high purge pressure did not resolve. the root cause of pump stop was high purge pressure, the cause of high purge pressure was not determined as no product was returned. d6a, d6b, g3.
28776 | 1220648-2024-09069 | 19006070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006070&pc=OZD | 08/11/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/11/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 26-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the physician was adjusting the position, and the pump was pulled out of the ventricle. the physician was unable to recross with the device. the patient was deemed stable at that time, so the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28777 | 1220648-2024-09016 | 19006125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006125&pc=OZD | 09/10/2022 | 03/10/2025 |  | 03/29/2024 |  |  | 09/10/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023230750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support there was oozing at the groin site and platelets were trending down. therefore, the pump was successfully weaned and explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28778 | 1220648-2024-09070 | 19006158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006158&pc=OZD | 08/12/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/12/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221597 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced oozing at the access site. to resolve the oozing, the clinical team stopped heparin and applied additional sutures. the patient remained on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28779 | 1220648-2024-09071 | 19006198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006198&pc=OZD | 08/12/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/12/2022 | 03/29/2024 |  |  | 05/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228760 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 90-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the pump encountered delivery issues and was unable to be implanted in the first access site; wires were unable to be passed across the stents found in the patient's aorta. subsequently, the pump was implanted at a second site that was a little higher on the right femoral site. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28780 | 1220648-2024-09072 | 19006236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006236&pc=OZD | 08/12/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/12/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217264 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that while on support, the impella cp demonstrated low pump flow and suctioning issues. p-levels were reduced to resolve the alarms. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28781 | 1220648-2024-09052 | 19006262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006262&pc=OZD | 03/21/2024 | 05/06/2024 |  | 03/29/2024 |  |  | 03/21/2024 | 03/29/2024 |  |  | 07/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443818 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Puncture/Hole | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, there were suction alarms present. albumin and and blood were provided. additionally, there was oozing around the impella site. two units of packed red blood cells were provided and surgical repair was performed with the cp in place. staff noticed that blood was leaking from the red impella plug. the impella was turned off, explanted and replaced. the patient remained on support with the replaced pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed, low pump flow, and product damage has been completed. the impella device was returned for evaluation. upon inspection of the returned product, a hole in the catheter was found at the 70cm mark and on the reposition sheath. blood was also found in the red handle. clinical details state that the patient had a drop in hemoglobin and had suction alarms. oozing was noted at the impella cp access site about 12 hours into support and a blood leak from red handle was noticed after surgical repair. upon further inspection, clinical noted a puncture/knick in catheter after impella explant. the root cause of the access site bleeding was most likely use related. the root cause of the product damage (hole in catheter) was determined to be use related, which is most likely contributed to the suction alarms (low pump flow). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28782 | 1220648-2024-09073 | 19006312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006312&pc=OZD | 08/13/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 08/13/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223621 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced reoccurring bleeding at the access site. vascular surgery added sutures which was successful in stopping the bleeding. the patient was also administered one unit of pack red blood cells, and systemic heparin was held. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28783 | 1220648-2024-09074 | 19006441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006441&pc=OZD | 08/16/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/16/2022 | 03/29/2024 |  |  | 01/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185801 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 43-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. limb ischemia was noted with absent pulses. distal perfuser added to extracorporeal membrane oxygenation and impella overnight. patient remained on support and was sent to unit. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28784 | 1220648-2024-09040 | 19006561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006561&pc=OZD | 07/03/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/03/2023 | 03/29/2024 |  |  | 11/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295099 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2023 |  |  |  |  | The user facility reported an impella cp was implanted in a 42 year old male patient for mechanical circulatory support. it was reported that the pump dislodged the pump from the left ventricle (lv) when the patient became restless. the team was unable to re-advance the pump across the valve. the patient was stable and the pump was explanted. no further hemodynamic support was needed. | [5 empty] | The investigation has been completed. the product was not returned for investigation. per clinical details, the impella was pulled by the patient, and the doctor failed to reposition it. the root cause of the positioning issue was patient movement based on provided clinical information.
28785 | 1220648-2024-09045 | 19006737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006737&pc=OZD | 03/15/2024 | 04/19/2024 |  | 03/29/2024 |  |  | 03/15/2024 | 03/29/2024 |  |  | 11/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024314511 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported an 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed access site bleeding. a femstop was placed above the insertion site and one unit of packed red blood cells transfused as treatment. the bleeding resolved and the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
28786 | 1220648-2024-09060 | 19006772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006772&pc=OZD | 07/29/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 07/29/2022 | 03/29/2024 |  |  | 12/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported 86-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. fifteen minutes after the impella cp was placed, the purge pressure (pp) increased into the thousands and the pump flow (pf) was 1.7 milliliters per hour. the patient was successfully weaned with no hemodynamic compromise. the device was explanted and hemostasis was achieved. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28787 | 1220648-2024-09044 | 19006832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006832&pc=OZD | 06/30/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/04/2023 | 03/29/2024 |  |  | 02/28/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333071 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/04/2023 |  |  |  |  | The user facility reported a 85-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had acute myocardial infarction/multivessel lesion. at a later date, the culprit and remnant branch that had not been fully opened were treated. although the percutaneous coronary intervention was successful, the pupillary immobility was confirmed, so a computed tomography scan of the head confirmed a cerebral hemorrhage and the impella was removed. | [5 empty] | The investigation has been completed. no product and data logs were returned. as per clinical, the patient experienced a cerebral hemorrhage and later the pump was explanted. it was unknown whether the timing of heparin administration was causing cerebrovascular accident concern during percutaneous coronary intervention. the cause of the stroke issue was most likely patient condition related and unknown relation to impella due to insufficient clinical/log review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28788 | 1220648-2024-09075 | 19006852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006852&pc=OZD | 08/20/2022 | 10/27/2025 |  | 03/29/2024 |  |  | 08/20/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210845 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2022 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support, it was noted that impella catheter needed to be repositioned. the catheter was advanced under echo guidance. the physician was satisfied with the positioning. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | D6a and d6b revised from the initial manufacturer device report 1220648-2024-09075 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-09075 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-09075 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-09075. h6 investigation conclusions revised from the initial submission in accordance with updated procedures.
28789 | 1220648-2024-09068 | 19006867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006867&pc=OZD | 08/08/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/08/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 380011 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 84-year-old male patient was attempting to be implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was unable to advance due to patient¿s vascular anatomy and high calcification within the patient's limbs. therefore, the impella cp implant was aborted. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. the root cause of the delivery issue was determined to be patient condition as the patient had highly calcified vasculature preventing the pump delivery. the root cause of the introducer damage-mechanical was determined to be patient condition as the patient had highly calcified vasculature. this report is being filed as part of a retrospective review of historical records.
28790 | 1220648-2024-09029 | 19006978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19006978&pc=OZD | 10/10/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 10/10/2023 | 03/29/2024 |  |  | 08/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024381532 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 61 year old female on an impella 5.5 for mechanical circulatory support. it was reported that due to complex vascular anatomy, the impella 5.5 was unable to be delivered in to ascending aorta. despite multiple attempts at troubleshooting implant, case was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
28791 | 1220648-2024-09023 | 19007049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19007049&pc=OZD | 06/18/2022 | 03/11/2025 |  | 03/29/2024 |  |  | 06/18/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that the patient on impella cp had access site bleeding and cassette leak. the purge cassette was leaking, it was changed and resolved. the patient was oozing from all invasive lines, the purge solution was changed to sodium bicarbonate. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. impella cp ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28792 | 1220648-2024-09022 | 19007058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19007058&pc=OZD | 07/19/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/19/2023 | 03/29/2024 |  |  | 12/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024304081 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2023 |  |  |  |  | The user facility reported a 75 year old male on an impella cp for mechanical circulatory support. it was reported that during insertion, everything except the motor was able to pass through the stents in the iliacs. after multiple attempts the impella was aborted for the case. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
28793 | 1220648-2024-09014 | 19007200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19007200&pc=OZD | 08/24/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/24/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199916 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2022 |  |  |  |  | The user facility reported an 80-year-old patient was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, the impella cp was found to be too deep in the ventricle with echocardiography. the pump was found under the chordae tendineae against the patient's lateral wall. due to the concern over how small the ventricle was in size, the team was not comfortable pulling back the pump and decided to discontinue pump support instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28794 | 1220648-2024-08887 | 19007234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19007234&pc=OZD | 07/09/2022 | 03/27/2024 |  | 03/29/2024 |  |  | 07/09/2022 | 03/29/2024 |  |  | 11/30/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162440 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/09/2022 |  |  |  |  | The user facility reported a 66-year-old male patient implanted with an impella cp for mechanical circulatory support. it was reported, that while on support, the patient developed an ischemic leg on the impella side. the pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28795 | 1220648-2024-09020 | 19007489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19007489&pc=OZD | 06/16/2022 | 03/10/2025 |  | 03/29/2024 |  |  | 06/16/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023200612 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a 78-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported when the pump was started, the high motor current alarm was on and the pump stopped. the device was explanted and revealed the easy guide lumen was still within the cannula and outlet which was left during implantation. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28796 | 1220648-2024-09076 | 19008407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008407&pc=OZD | 12/04/2023 | 05/24/2024 |  | 03/29/2024 |  |  | 12/04/2023 | 03/29/2024 |  |  | 05/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 79 year old male on an impella cp for mechanical circulatory support. it was reported that the patient had oozing from the access site. the heparin drip was stopped and the patient was switched to bicarbonate. later, the team made the choice to abandon the coronary artery bypass graft (cabg) and allow the patient to expire under care/comfort measures. | The patient was withdrawn from care and expired after 70 hours on support. upon review by abiomed's medical safety officer, there is not enough information to associate use of the device with patient outcome. |  |  |  |  | The impella device has not been returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, as the access site was treated by dressing change and the angle was matched as per the clinical description. the instructions for use referenced in the initial report are from ifu for impella cp with smartassist for use during cardiogenic shock and high risk cp, sections: general patient care considerations, entering cardiac output, and potential adverse events (united states) which now includes statement "cardiac or vascular injury (including ventricular perforation).¿ b5-added patient outcome from medical safety officer in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28797 | 1220648-2024-09077 | 19008760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008760&pc=OZD | 02/11/2024 | 06/11/2024 |  | 03/29/2024 |  |  | 02/28/2024 | 03/29/2024 |  |  | 06/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024369998 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Other | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/28/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 70 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured thrombocytopenia with a "possible" relationship to the impella device. there was no noted treatment required, and the patient recovered approximate 4 days later. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured hemolysis with a "definite" relationship to the impella device: the patient's hospitalization was prolonged, as a result. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombocytopenia and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. it was reported that the data log does not indicate any events related to motor current spikes or positioning issues during the case run, which could cause hemolysis due to mechanical failure of the device. the cause of the thrombocytopenia and the hemolysis issues was not determined since product was not returned and sufficient clinical information was not provided. in accordance with updated procedures, sections d6 and g3 have been revised from the initial submission. | Additional information was received and entered into b5 (hemolysis). the investigation has been reopened. a supplemental mdr will be filed upon completion of our analysis.
28798 | 1220648-2024-09078 | 19008793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008793&pc=OZD | 01/29/2024 | 04/12/2024 |  | 03/29/2024 |  |  | 02/28/2024 | 03/29/2024 |  |  | 11/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424849 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding an (b)(6) female patient presenting for high rick percutaneous coronary intervention for impella support. during support, the patient endured an access site pseudoaneurysm, 10 days post explant, requiring 2u prbc delivery. the allegation had a "possible" relationship to the impella device. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported vessel damage has been completed. the pump was not returned for investigation. as per the given clinical details, patient developed an access site pseudoaneurysm 10 days post-explant, requiring 2 units of prbcs. the patient underwent hrpci with impella on 19jan2024. an ultrasound on 20jan2024 showed no pseudoaneurysm, but a repeat ultrasound on 29jan2024 revealed one measuring 0.4 x 0.3 cm at the impella site. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. the vascular damage's relation to the impella is unknown, as there was no sign of the pseudoaneurysm at the access site the day after explant. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ d4-updated model number.
28799 | 1220648-2024-08769 | 19008796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008796&pc=OZD | 07/15/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 07/15/2022 | 03/29/2024 |  |  | 11/30/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023159611 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced hemolysis and lower limb ischemia during treatment. it was also reported that the hemolysis resolved, however, the lower extremity ischemia continued from insertion of the device until its removal. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28800 | 1220648-2024-08694 | 19008797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008797&pc=OZD | 03/01/2024 | 04/19/2024 |  | 03/29/2024 |  |  | 03/01/2024 | 03/29/2024 |  |  | 11/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025436758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 70-year-old male patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced non-sustained pulseless ventricular tachycardia (vt). staff repositioned the pump and the patient seemed to have improved. however, the patient was transferred and experienced ectopy again. additionally, the staff assumed there was possible debris ingested into the pump. the pump was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (non-sustained pulseless vt) and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vf/vt (non-sustained pulseless vt) and the thrombus was not determined.
28801 | 1220648-2024-08852 | 19008798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008798&pc=OZD | 03/18/2024 | 05/16/2024 |  | 03/29/2024 |  |  | 03/18/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2025441767 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a 5.5 support ongoing for approximately 1 month for a patient admitted in cgs (cardiogenic shock). the pump stopped around a month after the pump was thought to have ingested thrombus. the pump was explanted and not replaced. the patient was managed by medical/pharma support only. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported thrombus and pump stop has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus nor the pump stop could not be determined.' the instructions for use in the initial report is from section: direct aortic insertion, section: using the automated impella controller with the impella catheter, and section: impella stopped for the impella 5.5 with smart assist for use during cardiogenic shock. d4-updated model number h5-changed to yes in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
28802 | 1220648-2024-08724 | 19008799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008799&pc=OZD | 08/07/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/07/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208902 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a 75-year-old male patient needing mechanical circulatory support. it was reported that the pump encountered positioning issues. the pump had slipped out of the left ventricle and after several attempts, the pump was able to be correctly repositioned. however, another positioning issue was later encountered, and it was then decided to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the positioning issue was unable to be determined due to limited clinical details. this report is being filed as part of a retrospective review of historical records.
28803 | 1220648-2024-09012 | 19008800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008800&pc=OZD | 06/25/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/06/2023 | 03/29/2024 |  |  | 01/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333667 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported a 75 year old female with an impella cp for acute myocardial infarction. it was reported that after the pump was explanted, the experienced a drop in hemoglobin that required 1 unit of packed red blood cells (prbcs). the team confirmed the issue was due to iron deficiency anemia. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were not provided. however, clinical details states the patient was diagnosed with iron-deficiency anemia during support which resulted in low hemoglobin. therefore, the root cause of hemolysis was most likely due to patient condition & unrelated to impella.
28804 | 1220648-2024-09013 | 19008801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008801&pc=OZD | 07/08/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/10/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1429248 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/10/2023 |  |  |  |  | The user facility reported a console (aic) being used on a 47 year old male for mechanical circulatory support. it was reported that the aic failed to recognize an installed purge disc during preparation for impella support. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag was missing and there was a crack in the container. the root cause of this issue was a missing flag and crack in the retainer.
28805 | 1220648-2024-09008 | 19008802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008802&pc=OZD | 07/04/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/04/2023 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/04/2023 |  |  |  |  | The user facility reported a console (aic) in use on a 49 year old female requiring mechanical circulatory support. it was reported that the softkeys were not responsive, the menus could not be opened, and the grey knob also had no effect. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported aic soft key/knob issues has been completed. the aic was returned for evaluation. visual and functional evaluation found that the device functioned as expected. the reported error could not be reproduced. there were no signs of damage to the device. logs were reviewed which revealed imc_kbd errors at different times during use but not consistently. therefore the root cause of the reported issue could not be determined since the issue could not be reproduced.
28806 | 1220648-2024-08797 | 19008803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008803&pc=OZD | 06/21/2022 | 03/27/2024 |  | 03/29/2024 |  |  | 06/21/2022 | 03/29/2024 |  |  | 10/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158821 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 06/21/2022 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from impella site. manual pressure was held, tuohy locked and blood products were administered. during support the patient decompensated and was having frequent suction alarms after being transferred to intensive care unit. the pulse pressure narrowed. there was no right pedal pulse after femstop was placed. vascular surgery was consulted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28807 | 1220648-2024-08850 | 19008804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008804&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/29/2024 |  |  | 03/15/2024 | 03/29/2024 |  |  | 04/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353571 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old female noted as high risk percutaneous coronary intervention (pci) was to be implanted with an impella cp device for mechanical circulatory support. when inserting the impella cp, there was resistance where the iliac and descending aorta met. the physician tried shock waving that area. the pump became about halfway up but became stuck right by the motor housing. after two to three more unsuccessful attempts, the impella cp device placement was aborted and the physician continued pci without support. there was no report of an adverse event to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the delivery issue is determined to be patient condition because the pump was unable to pass through vasculature, where the pump was unable to pass from the iliac into descending aorta where the pump got stuck by motor housing even after multiple attempts. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. to conform to updated procedures, sections d6 & h10 were revised with updated information.
28808 | 1220648-2024-08962 | 19008805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008805&pc=OZD | 08/10/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/10/2022 | 03/29/2024 |  |  | 04/30/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224757 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. access site bleeding was noticed at the site around the repositioning sheath. due to the patient's increased risk of bleeding, the decision was made to explant the impella cp device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28809 | 1220648-2024-08798 | 19008806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008806&pc=OZD | 03/13/2024 | 03/27/2024 |  | 03/29/2024 |  |  | 03/13/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1250838 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during prep, the automated impella controller (aic) did not detect the purge cassette. the cassette was exchanged but the issue did not resolve. therefore, the aic was exchanged to resolve the issue. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported "aic - purge disc/flag issues" has been completed since the original report was filed. product was returned for investigation. the console was tested after arriving. the ic4500 was missing the purge flag. the console repeatedly was unable to detect purge discs (alarm #402) at the start of the case. the cause of the issue was a missing purge flag.
28810 | 1220648-2024-08851 | 19008807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008807&pc=OZD | 03/16/2024 | 04/29/2024 |  | 03/29/2024 |  |  | 03/18/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1430362 |  | 0042-0000=JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when the automated impella controller (aic) had a purge disc not detected alarm. staff tried to reconnect the purge disc to no avail. a new purge disc was attempted and it was unsuccessful. the staff replaced the aic and the purge disc was installed without an issue. | [5 empty] | The impella device was received from the customer on 21-mar-2024 and therefore,¿an evaluation of the device was is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. console was inspected after arriving in fs and the purge flag was confirmed to be broken. the root cause of this issue was a broken/missing purge flag.
28811 | 1220648-2024-08770 | 19008808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008808&pc=OZD | 07/01/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 07/01/2022 | 03/29/2024 |  |  | 12/31/2023 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 75-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. after insertion of the impella cp, the cheater flag got stuck in the impella and the staff did not notice. the red cord of the cheater flag got tangled in the rotor when pump was started, and pump failure occurred. the wire went through the pigtail of the impella and did feed the cheater flag through, but when pulled apart, it stayed behind unnoticed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28812 | 1220648-2024-08891 | 19008809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008809&pc=OZD | 06/30/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 06/30/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201676 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 76-year-old female patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a patient presented with a new onset of atrial fibrillation (af), rapid ventricular response (rvr), chronic total occlusion (cto), and right coronary artery (rca). it was reported that during impella cp support the patient developed a hematoma at the impella access site and treated with covering the femoral artery with a stent. it was also reported during support, the medical team was unable to pass a 7f terumo destination sheath through the introducer, and a 7f 10cm slender sheath was used instead with success. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28813 | 1220648-2024-08796 | 19008810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008810&pc=OZD | 03/13/2024 | 04/09/2024 |  | 03/29/2024 |  |  | 03/13/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1273595 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prep, automated impella controller (aic) did not detect the purge cassette. the issue was attempted to be resolved by rebooting console power cycle, but was unsuccessful. therefore, the aic was exchanged and the problem resolved. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge flag issue has been completed. data logs were reviewed which revealed that while attempting to start a case, the device could not get past step 4 due to "piston blocked". the console continuously tried to prime but since the piston would not move, priming could never be completed and triggered alarm 417 for purge system failure (piston blocked). reviewing the right logs, the purge flow ticks never exceeded 23 which supports the lack of piston range seen in the imc logs. the aic was returned for evaluation and confirmed that the shaft was blocked for the purge with signs of glucose around it. therefore, the cause of the piston block was a glucose ingress. added- h6 impact code 4629.
28814 | 1220648-2024-08856 | 19008811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008811&pc=OZD | 03/20/2024 | 05/10/2024 |  | 03/29/2024 |  |  | 03/20/2024 | 03/29/2024 |  |  | 09/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024403925 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 60-year-old male was implanted with an impella 5.5 device as a bridge to transplant. it was reported that after 7 days of support the patient developed an access site bleed, which prompted the discontinuation of the heparin. the pump remained on for support with sodium bicarb running in the purge. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned for analysis. no abnormalities were found. the root cause of the access site bleeding issue was most likely anticoagulation management as stopping heparin was able to resolve the hematoma.
28815 | 1220648-2024-08855 | 19008812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008812&pc=OZD | 03/19/2024 | 04/23/2024 |  | 03/29/2024 |  |  | 03/19/2024 | 03/29/2024 |  |  | 07/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441409 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a (b)(6)-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed to support a patient admitted with fibrillation and long term coumadin therapy. the team could not control the access site bleed from the high inr despite a 2hour hold and so the team explanted the pump, placed a perclose and an angioseal and ballooned for hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the bleeding was due to a high bleeding time. therefore, the root cause of the access site bleed was most likely due to anticoagulation management. d4-updated model number.
28816 | 1220648-2024-08853 | 19008813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008813&pc=OZD | 03/18/2024 | 04/19/2024 |  | 03/29/2024 |  |  | 03/18/2024 | 03/29/2024 |  |  | 06/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364877 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the md inserted 5.5 into the graft over the guidewire and noticed bleeding at the anastomosis site. md fixed the bleeding area and advanced the 5.5 catheter from the graft to the axillary artery and noticed more bleeding. the md removed the impella 5.5 and used pledgets, sutures, thrombin, and bioglue at the bleeding site until hemostasis was achieved. one unit of packed red blood cells was given. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by pledgets, sutures, thrombin and bioglue.
28817 | 1220648-2024-08857 | 19008814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008814&pc=OZD | 06/23/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 06/23/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204773 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, hemolysis was detected in the urine and patient began bleeding from the insertion site. the medical team decided to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28818 | 1220648-2024-09006 | 19008815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008815&pc=OZD | 06/04/2023 | 09/11/2024 |  | 03/29/2024 |  |  | 06/05/2023 | 03/29/2024 |  |  | 03/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335499 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female |  |  |  | Death | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 49-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported low pump flow after initiation. the decision was made to remove the device. a pulmonary embolism was suspected and the patient ultimately expired. no involvement was noted between the device and adverse event/death. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation has been completed. data confirmed low flow when pump started and remained to the rest of the case that triggered auto suction response and impella flow reduced alarms. the product was not returned for review. based on clinical description and data, the cause of the issue was patient condition related. | B2: the date of the patient's death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-09015 was provided in sections b2, b5, h1, and h6.
28819 | 1220648-2024-08854 | 19008816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008816&pc=OZD | 06/23/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 06/23/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023209605 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Unraveled Material | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was loaded per usual fashion. the 0.018 wire was pulled out and the end of it uncoiled and expanded. the doctor confirmed the wire did not break off anywhere. however, the patient experienced a hematoma and bleeding. there was no medical or surgical intervention reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28820 | 1220648-2024-08858 | 19008817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008817&pc=OZD | 03/15/2024 | 05/29/2024 |  | 03/29/2024 |  |  | 03/15/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443191 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that multiple attempts were made to pass the impella 5.5 through the innominate artery. the device would not pass. the decision was made to transition to an open direct insertion with a new impella 5.5 due to kinking at the catheter. there was no harm to the patient. | The patient expired on support. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through axillary/innominate/aorta as per the clinical description provided. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.3 revised as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation code. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report.
28821 | 1220648-2024-09007 | 19008818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008818&pc=OZD | 06/14/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 06/14/2022 | 03/29/2024 |  |  | 01/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023183155 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported a 88-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while implanting the device, left groin access was obtained and a long pump sheath was inserted under the fluoro; however, the pump's sheath kinked due to the patient's anatomy. therefore, the decision was made to access the right groin with the short pump's sheath. however, there were difficulties with the guide advancing and removing the guide caused the pump to pull back into the aorta which was noted on the fluoro that the pump's catheter was coiled in the descending aorta. the physician decided to remove the device and stage the patient for right coronary artery lesion. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28822 | 1220648-2024-09004 | 19008819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008819&pc=OZD | 06/13/2022 | 03/29/2024 |  | 03/29/2024 |  |  | 06/13/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197518 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported a 67-year-old female was implanted with an impella cp device for mechanical circulatory support. during support the patient developed a large hematoma, about the size of cantaloupe at the impella access site. the patient was given 4 units of blood. the patient¿s hemoglobin dropped from a baseline 12.1 to 6.6. hemostasis was obtained using manual pressure and femstop. | [5 empty] | The impella cp device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28823 | 1220648-2024-08913 | 19008820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008820&pc=OZD | 03/19/2024 | 05/09/2024 |  | 03/29/2024 |  |  | 03/19/2024 | 03/29/2024 |  |  |  |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025441198 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 67-year-old female of an unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had observed small femoral arteries. the pump was placed despite this. the pump caused limb ischemia and explanted after just 19 hours and was escalated to the 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported limb ischemia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 and 3233 were inadvertently left off the previously submitted report and were added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
28824 | 1220648-2024-09009 | 19008821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008821&pc=OZD | 07/04/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/04/2023 | 03/29/2024 |  |  | 04/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350705 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/04/2023 |  |  |  |  | The user facility reported an impella cp on a 63 year old female patient with an acute myocardial infarction. it was reported that the patient was decompensating and the decision was made to place the impella. the doctor attempted to place the impella, but was unable to get it across the valve. the implant was aborted and the patient expired shortly after. the patient's death was not attributed to use of the impella device. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
28825 | 1220648-2024-09005 | 19008822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008822&pc=OZD | 08/22/2022 | 09/30/2025 |  | 03/29/2024 |  |  | 08/22/2022 | 03/29/2024 |  |  | 03/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205171 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The complainant reported that a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had positioning issues due to anatomy which included a short aorta ascendens and a small cavum. it was reported after being implanted successfully, it was observed the flow rate was reduced due to poor position of the device. the patient was hemodynamically stable after a successful high risk percutaneous coronary intervention, and the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
28826 | 1220648-2024-09010 | 19008823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008823&pc=OZD | 08/31/2022 | 04/08/2025 |  | 03/29/2024 |  |  | 08/31/2022 | 03/29/2024 |  |  | 02/29/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201670 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a(b)(6)female was implanted with an impella cp device for mechanical circulatory support. while on support, the impella flows reduced to sub-optimal performance which required lowering of the p-level to break suction. patient continued to show improvement. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28827 | 1220648-2024-09011 | 19008826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008826&pc=OZD | 07/05/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 07/05/2023 | 03/29/2024 |  |  | 04/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338014 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported an 83 year old male on impella cp for mechanical circulatory support. it was reported that the pump triggered suction alarms when support levels were increased above p2. multiple factors were investigated in an attempt to determine a cause, however, there was no improvement in the noted issue. flows remained low throughout the course of support. there was no patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was returned for evaluation. visual inspection found a cannula kink near the inlet and there was biomaterial in the top of the impella and a thrombus in the outlet. the root cause of the low pump flow was cannula kink due to patient condition related as patient's blood vessel tortuosity was severe that kinked the cannula and impacted the flow path, suction occurred and biomaterial accumulated. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. the ifu further states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
28828 | 1220648-2024-09079 | 19008833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008833&pc=OZD | 09/27/2023 | 03/29/2024 |  | 03/29/2024 |  |  | 02/05/2024 | 03/29/2024 |  |  | 05/31/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376994 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/05/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient was found to have a axillary hematoma and systemic anticoagulation had to be held. the patient remained on impella support. | [5 empty] | The investigation into the patient's access site bleeding has been completed. no product was returned for investigation. the root cause of the access site bleeding - major was not determined due to a lack of clinical details. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28829 | 1220648-2024-09080 | 19008842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008842&pc=OZD | 02/05/2024 | 04/09/2024 |  | 03/29/2024 |  |  | 02/05/2024 | 03/29/2024 |  |  | 02/28/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333470 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following placement of the repositioning sheath during impella cp support, the access site continued to bleed and hemostasis could not be achieved. as a result, the patient was transferred to another facility and the pump was replaced with an impella 5.5 device. a surgical wound closure was successful in obtaining hemostasis at the femoral site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
28830 | 1220648-2024-09081 | 19008861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008861&pc=OZD | 02/06/2024 | 03/29/2024 |  | 03/29/2024 |  |  | 02/08/2024 | 03/29/2024 |  |  | 12/31/2024 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024321403 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/08/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the placement of the impella cp the patient had five episodes of ventricular fibrillation after placement of the pig-tail catheter in the left ventricle. medications were given. the safety team determined that ventricular fibrillation is possibly related to the impella procedure. the patient recovered. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. vt/vf - cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
28831 | 1220648-2024-09082 | 19008901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19008901&pc=OZD | 02/18/2024 | 04/18/2024 |  | 03/29/2024 |  |  | 02/29/2024 | 03/29/2024 |  |  | 06/30/2025 |  | 03/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369996 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 53 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. approximately 18 days after support, the patient presented to the emergency department with symptoms of right calf tenderness and swelling. deep vein thrombosis was diagnosed and "apixaban" was administered. the allegation came with a "probable" relationship to the impella device. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
28832 | 1220648-2024-08928 | 18990078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990078&pc=OZD | 06/15/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/15/2022 | 03/28/2024 |  |  | 11/30/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162164 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella cp device for mechanical circulatory support. during the procedure there was some mild concern for limb ischemia. the medical team decided to remove peel away sheath and insert repo sheath. there was slight difficulty advancing repo sheath into arteriotomy. the impella cp was explanted in the procedural area. no additional information was provided for the suspected ischemia. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28833 | 1220648-2024-08924 | 18990259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990259&pc=OZD | 08/03/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 08/03/2022 | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224427 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 62-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced a perclose failure during closure where the new perclose was not passing over the wire. the patient received two stents in the right coronary artery (percutaneous transluminal angioplasty). the surgeon opted to use 8fr angioseal advanced into the vessel to obtain good hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28834 | 1220648-2024-08926 | 18990276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990276&pc=OZD | 03/19/2024 | 05/10/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  | 06/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024431035 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced a hematoma with oozing. manual pressure was held for 60 minutes. the physician unsuccessfully attempted to adjust the perclose, so a femstop was applied and skin sutures were added to that access site. the patient was reported to be in stable condition. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was determined to be anticoagulation management as the patient had a high act value of 420. heparin was held until it could be brought down. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28835 | 1220648-2024-08876 | 18990287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990287&pc=OZD | 08/14/2022 | 09/21/2024 |  | 03/28/2024 |  |  | 08/14/2022 | 03/28/2024 |  |  | 09/30/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Deliver | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2022 |  |  |  |  | The patient had a superficial femoral artery (sfa) occlusion preventing flow through the introducer. consequently, a fem-fem bypass was performed to bypass the occlusion. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 58-year-old female patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there were delivery issues. the patient had a superficial femoral artery (sfa) occlusion preventing flow through the introducer. consequently, a fem-fem bypass was performed to bypass the occlusion. additionally, the complainant reported that there was minor access site bleeding noted with a slight ooze at the insertion site. this was resolved on its own as the bleeding eventually slowed down. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08876 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28836 | 1220648-2024-08878 | 18990306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990306&pc=OZD | 09/07/2022 | 05/16/2025 |  | 03/28/2024 |  |  | 09/07/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194996 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Septic Shock | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient had mild ooze at groin access site. the blue hub was advanced into the skin tract and angle matching was performed to resolve oozing, the patient also experienced loss of pulses in both legs and distributive shock. the medical team decided to explant the impella cp due to the unresolved limb ischemia and distributive shock. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. b1: corrected from ¿product problem¿ to ¿adverse event¿ and ¿product problem¿ g1: corrected ¿reporting contact first name¿ and ¿reporting contact last name¿.
28837 | 1220648-2024-08879 | 18990369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990369&pc=OZD | 07/01/2022 | 03/27/2024 |  | 03/28/2024 |  |  | 07/01/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202724 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 77-year-old female patient needing an impella cp for mechanical circulatory support. it was reported that upon placement of the impella, the patient went into ventricular tachycardia (vt) and was shocked multiple times. the patient was monitored until the impella cp device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28838 | 1220648-2024-08930 | 18990406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990406&pc=OZD | 07/11/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 07/11/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023201609 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported an 80-year-old female patient needing an impella cp for mechanical circulatory support. while positioning the pump, it was noted that the action was causing ventricular tachycardia (vt) due to the anatomical structure of the patient's heart. after finding a comfortable position, support was able to continue with the implanted device. no intervention was required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28839 | 1220648-2024-08927 | 18990494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990494&pc=OZD | 08/10/2022 | 05/16/2025 |  | 03/28/2024 |  |  | 08/10/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214344 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 75-year-old male was implanted with an impella cp device for mechanical circulatory support. access site bleeding was observed. attempts to insert two manta vascular closure devices were unsuccessful. subsequently, a balloon tamponade with rotating manual pressure was used and the site was injected with flowable d-stat to assist with clotting. additionally, a fem-stop was also placed, and the bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28840 | 1220648-2024-08931 | 18990529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990529&pc=OZD | 03/19/2024 | 05/06/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025453879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 66-year-old male was to be implanted with an impella 5.5 device due to being escalated from an impella cp. it was reported that the implant was aborted due to the patient¿s anatomy. no patient harm was reported. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of delivery issue is determined to be patient condition because of the calcified condition of the vessels and the insertion was aborted. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28841 | 1220648-2024-08929 | 18990538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18990538&pc=OZD | 06/22/2022 | 09/17/2025 |  | 03/28/2024 |  |  | 06/22/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207385 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | Life Threatening | Disability | Required Intervention | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 06/22/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when explanting the pump, the doctor disconnected an a line hooked up to the side port on the repositioning sheath, then there was no blood return. a small angiogram was taken which revealed minimal flow in the left iliac. the doctor advanced a 5 french pigtail into the descending aorta and performed an angiogram which revealed a large thrombus burden in the distal aorta at the bifurcation of the iliacs. a thrombectomy was planned to be performed. patient started to hemodynamically decompensate and anesthesia was called to intubate patient prior to going to surgery. upon anesthesia induction, patient went into cardiac arrest. the patient's family decided to the withdraw care after 70.85 hours of support. medical safety review of the event noted there is not have enough information to exclude impella device as an associated factor in the patient's demise. | The user facility reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump presented suction alarms and the patient presented with hemolysis, both of which were found to be caused by the pump being too shallow. these issues were resolved upon repositioning of the pump. additionally, the patient experienced a thrombus while on impella support, which was to be resolved with thrombectomy. |  |  |  |  | Corrected data: b2, b5, d7a, f6, f8, h1. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: access site bleeding. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ hemolysis ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
28842 | 1220648-2024-08937 | 18991021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991021&pc=OZD | 06/18/2022 | 05/16/2025 |  | 03/28/2024 |  |  | 06/18/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023188341 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/18/2022 |  |  |  |  | The user facility reported a 75-year-old female in post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was reported the patient experienced severe right distal limb ischemia and the physician decided to have the device removed. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28843 | 1220648-2024-08890 | 18991241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991241&pc=OZD | 06/29/2022 | 05/16/2025 |  | 03/28/2024 |  |  | 06/29/2022 | 03/28/2024 |  |  | 09/30/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159503 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The complainant reported an 81-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support, the patient experienced oozing at the groin access site. the impella was weaned and explanted early to resolve bleeding. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28844 | 1220648-2024-08940 | 18991272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991272&pc=OZD | 06/19/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/19/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207130 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma while on support. the issue was resolved by explanting the pump followed by manual compression. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28845 | 1220648-2024-08938 | 18991276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991276&pc=OZD | 03/20/2024 | 05/07/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a (b)(6)-year-old female noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. the patient suffered a perforation in the left ventricle during impella device support. it was noted leading up to the event, the patient had suffered from three circumflex artery related infarctions during that week. a pci procedure was initiated. however, the patient became hypotensive with unstable pressure. cardiopulmonary resuscitation was initiated. a pericardial effusion was identified. ventriculography was utilized to note a small perforation in the anterolateral wall of the left ventricle. surgical intervention was undertaken, but the patient's left ventricle was noted to be extremely friable. the original perforation was managed, but additional leaks began to appear due to the poor condition of the left ventricle. the patient subsequently experienced heavy blood loss for which transfusions were initiated. the decision was ultimately made to stop the procedure, and the patient passed away. a definitive cause for the perforation was not established. | [5 empty] | The investigation for the reported vessel damage has been completed. no product was returned for evaluation. data logs were reviewed and lt log shows decreases in flows and motor speed with suction alarms occurring due to patient¿s loss of volume. "impella position unknown" alarms also noted later in the case as patient lost volume and pulsatility. clinical details noted that patient's left ventricle wall was very friable and diseased/damaged prior to impella placement. additionally, once one leak was repaired, another would occur without interaction with the pump. the root cause of the ventricle perforation was due to patient condition as ventricle was very friable and diseased. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28846 | 1220648-2024-08939 | 18991278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991278&pc=OZD | 06/15/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/15/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204116 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 73-year-old male was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient had developed hemolysis. no additional information was provided for the treatment of the hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28847 | 1220648-2024-08941 | 18991307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991307&pc=OZD | 06/23/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/23/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported an 63-year-old male patient noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported flow was checked down the leg following perclose placement and there was no flow distal to the stick site and a dissection was noted. an arterial sheath in the other leg to go up and over was used and ballooned the dissection and good flow was noted down leg on fluoro. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28848 | 1220648-2024-08943 | 18991515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991515&pc=OZD | 06/19/2022 | 05/21/2025 |  | 03/28/2024 |  |  | 06/19/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the impella catheter kinked. multiple attempts to manipulate the catheter failed, and the device was removed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28849 | 1220648-2024-08945 | 18991642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991642&pc=OZD | 06/15/2022 | 03/10/2025 |  | 03/28/2024 |  |  | 06/15/2022 | 03/28/2024 |  |  | 07/31/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022111896 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and a hematoma at the access site which was resolved through redressing and angle matching. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28850 | 1220648-2024-08942 | 18991680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991680&pc=OZD | 06/25/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/25/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205052 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/25/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 52-year-old female patient for mechanical circulatory support. the patient was experiencing bleeding from the leg fasciotomy. the purge was switched to sodium bicarbonate and heparin was discontinued. the patient received three unit of packed red blood cells (prbc). | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28851 | 1220648-2024-08944 | 18991776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991776&pc=OZD | 06/26/2022 | 03/12/2025 |  | 03/28/2024 |  |  | 06/26/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193753 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while patient was on impella cp they had some hemolysis and thrombus. impella was removed due to hemolysis and during the removal there was thrombus. fogarty was performed and was able to get the thrombus resolved. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28852 | 1220648-2024-08925 | 18991815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991815&pc=OZD | 06/15/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/15/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187794 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 91-year-old female was implanted with an impella cp device for mechanical circulatory support. blood was noted in the catheter and the physician also discovered a perforation on the left anterior descending artery, which required fat embolism and micro-coils and the patient underwent a pericardiocentesis. the patient was administered four units of packed red blood cells. a balloon tamponade and covered stent were also placed. after treatment and revascularization, the physician elected to wean and explant the impella device. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28853 | 1220648-2024-08947 | 18991973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18991973&pc=OZD | 03/20/2024 | 06/24/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502178 |  | 0042-0010-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. during set-up for the procedure, the automated impella controller (aic) ) was turned on and displayed one red alarm (battery failure) and one yellow alarm (battery communication failure). the site confirmed the console had not been turned on in months. the procedure was still performed successfully and there was no reported patient harm. there were no issues reported during implant, during support, or during explant. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. based on problem description, it is most likely that console batteries were fully depleted. however, we could not verify this because neither the console nor its data was returned for investigation. the root cause could not be determined. d.2 common device name and g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28854 | 1220648-2024-08886 | 18992096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992096&pc=OZD | 07/08/2022 | 03/10/2025 |  | 03/28/2024 |  |  | 07/08/2022 | 03/28/2024 |  |  | 12/31/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023178569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a small hematoma at the access site and was sent to the or for cutdown removal as treatment. hemostasis was achieved and the device successfully weaned and explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28855 | 1220648-2024-08883 | 18992143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992143&pc=OZD | 02/03/2024 | 04/16/2024 |  | 03/28/2024 |  |  | 02/26/2024 | 03/28/2024 |  |  | 07/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376651 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured recurring ventricular tachycardia with a "possible" relationship to the impella device. multiple cardioversions were required. impella device support was continued for the patient for 4 more days after the incident. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28856 | 1220648-2024-08918 | 18992196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992196&pc=OZD | 03/19/2024 | 05/04/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024436205 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 26-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after seven hours of support, the impella cp was explanted and the patient was to be escalated to impella 5.5, but due to presence of a clot in the left ventricle. the decision was made to escalate to extracorporeal membrane oxygenation (ecmo) instead. the patient was reported as stable. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28857 | 1220648-2024-08946 | 18992232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992232&pc=OZD | 06/19/2022 | 05/21/2025 |  | 03/28/2024 |  |  | 06/19/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205307 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Difficult to Advance | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2022 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced ischemia in the left leg that resolved without any intervention specified. additionally, it was reported that there was a pump positioning issue which was resolved by manipulation of the catheter. the pump also experienced a suction issue that was resolved with administering a neo bolus. a clot was noted at the left superficial femoral artery, which was aspirated with a pnenumbra catheter. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28858 | 1220648-2024-08920 | 18992315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992315&pc=OZD | 03/22/2024 | 05/14/2024 |  | 03/28/2024 |  |  | 03/22/2024 | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435582 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having non-st elevated myocardial infarction. after the implant, the impella kinked after it was placed. the flows then cut in half. the impella was replaced. | [5 empty] | The investigation has been completed. the impella pump was returned for evaluation and was visually inspected. a cannula kink and catheter kink were observed. no biomaterial was observed. no broken blades were found. no other abnormalities were found. the cause of the low pump flow was cannula kink due to improper technique. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 2199 was reported incorrectly on the previously submitted report. the correct code has been added to health effect - impact codes. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28859 | 1220648-2024-08922 | 18992400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992400&pc=OZD | 06/22/2022 | 03/12/2025 |  | 03/28/2024 |  |  | 06/22/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, there was bleeding from groin site due to extreme calcification of the vessel. therefore, surgical wound closure was used. moreover, the left ventricle (lv) collapsed; therefore, there was hemolysis. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28860 | 1220648-2024-08932 | 18992401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992401&pc=OZD | 06/15/2022 | 05/16/2025 |  | 03/28/2024 |  |  | 06/15/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203708 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Difficult to Advance | [5 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 56-year-old male was implanted with an impella cp device for mechanical circulatory support. while attempting to place the impella, it would not pass the aortoiliac bifurcation. upon second attempt, the impella was removed and successfully placed without difficulty. also, the pump had decreased flows, and the patient went into ventricular fibrillation and was shocked one time. additionally, the patient¿s left leg was cold and mottled, an external femoral bypass was performed and subsequently the leg returned to normal color. the family decided to withdraw care and remove the impella. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28861 | 1220648-2024-08923 | 18992489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992489&pc=PYX | 03/19/2024 | 05/31/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  | 08/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024396629 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | There is not enough information to associate use of device with outcome. | The user facility reported a 76-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that during the impella rp implant, the physician attempted to advance the sheath into the patient¿s vasculature, but both attempts resulted in cracks at the distal tip of the dilator. the dilators were removed without injury to the patient. the physician then used a gore sheath without difficulty and the impella rp was implanted successfully. there was no patient harm reported. |  |  |  |  | The investigation into the introducer damage - mechanical issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinician information provided, the root cause of the introducer damage - mechanical was most likely a sheath tip split as confirmed by the image provided showing a damaged dilator. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28862 | 1220648-2024-08948 | 18992506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992506&pc=OZD | 06/20/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/20/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190641 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, a thrombus was found in the left ventricle and was addressed by pulling the pump into the aorta. the patient remained hemodynamically stable and was brought to operating room by vascular surgeon. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28863 | 1220648-2024-08888 | 18992678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992678&pc=OZD | 07/10/2022 | 03/27/2024 |  | 03/28/2024 |  |  | 07/10/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214188 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/10/2022 |  |  |  |  | The user facility reported a 32-year-old male patient was implanted with an impella cp for mechanical circulatory support. the patient experienced access site bleeding that required a femstop and surgicel. heparin was stopped due to low platelets. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28864 | 1220648-2024-08889 | 18992693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992693&pc=OZD | 07/11/2022 | 05/19/2025 |  | 03/28/2024 |  |  | 07/11/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181577 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a male was implanted with an impella cp for mechanical circulatory support. the patient experienced oozing at the access site of the femoral artery. blood products were provided and a reduction in systemic heparin was required to manage the complication. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28865 | 1220648-2024-08905 | 18992754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992754&pc=OZD | 03/19/2024 | 05/07/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025443062 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 62-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that while on support, the patient developed a hematoma and some swelling around their access site. coinciding with the presence of the hematoma, the patient reported numbness/tingling in their fingers in the affected extremity. the patient was given two units of blood, and a washout was performed to treat the hematoma. following the intervention, neurological sensation returned to the extremity. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned, evaluated, and visually inspected. data logs showed gradual increase in the purge pressure which was resolved by adding tpa. from the clinical information and the data logs, the root cause of the high purge pressure issue was gradual biomaterial buildup. the root cause of the access site bleeding cannot be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
28866 | 1220648-2024-08895 | 18992842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992842&pc=OZD | 03/18/2024 | 05/09/2024 |  | 03/28/2024 |  |  | 03/18/2024 | 03/28/2024 |  |  | 06/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Pressure Problem | [6 empty] | Ischemia | Perforation of Vessels | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 68-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during support there was high purge pressure. tissue plasminogen activator was added, and the purge cassette was replaced which did not resolve the issue. the decision was made to wean and explant the impella cp. following the explant the patient had no flows down their right lower extremity. it was found that the perclose was pinched at the closure site and the patient had a clot in their iliac artery. vascular surgery was consulted, and exploratory vascular surgical repair was conducted. the patient was reported as stable. | [5 empty] | The investigation has been completed. no product was returned for evaluation. clinical details noted that after removal of impella cp, the patient experienced no flow to right lower extremity and it was found that the perclose closure was pinched and a clot was present above the perclose closure in the right external iliac artery. a vascular repair was performed to excise clot from iliac artery. as the necessary information was not provided, the root cause of the thrombus and limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4627 was inadvertently added twice on the previously submitted report. code 4582 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28867 | 1220648-2024-08892 | 18992850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992850&pc=OZD | 03/06/2024 | 04/23/2024 |  | 03/28/2024 |  |  | 03/06/2024 | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024412192 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after 29 days of support the patient developed hemolysis as evidence by a recent spike in lactate dehydrogenase (ldh) and plasma free hemoglobin (pfhgb). the impella 5.5 was exchanged for impella cp, and the patient was transitioned to comfort care. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed multiple motor current (mc) spikes during the case along with auto suction and suction alarms were triggered. logs also showed there were many intermittent alarms pump position was in aorta but resolved quickly. based on clinical description and data analysis, the root cause of hemolysis was most likely biomaterial ingestion since there was a spike in mc prior to the hemolysis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis
28868 | 1220648-2024-08894 | 18992853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992853&pc=OZD | 03/18/2024 | 08/14/2024 |  | 03/28/2024 |  |  | 03/18/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1747385 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The complainant reported that when turning on the automated impella controler, an error message occurred - system self check failed, change console immediately. the aic was exchanged. there was no patient involvement. | [5 empty] | The investigation for the console (aic) self check error has been completed. the console was returned for evaluation. during inspection, on initial bootup, the console displayed the ¿system self check failed¿ screen. replacing the original pbm with a known-good pbm on resolved the issue. upon connecting the original defective pbm to the known good test fixture, the reported failure mode was reproduced. data logs were reviewed confirmed that upon initial boot up, a pbm communication issue. this resulted in an automatic console reboot after the log ¿post: watchdog reset. after reboot, the communication issue was persistent. the root cause of the system self-check failure was a pbm malfunction. added- d9 device return date d2-corrected common device name d4-corrected model number f6/f8 should have been left blank in the initial report. h6 med dev prob code 2898 changed to 2880. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
28869 | 1220648-2024-08900 | 18992899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992899&pc=OZD | 03/18/2024 | 05/04/2024 |  | 03/28/2024 |  |  | 03/18/2024 | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2024415935 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 52-year-old male was implanted with an impella 5.5 device for pre-operative placement support. the patient required high risk surgery/coronary artery bypass graft. two days post implant of the impella device, the team observed the patient to have suffered a cerebral vascular accident. the team removed heparin and sent the patient for imaging. the patient survived and remained on support. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury".
28870 | 1220648-2024-08896 | 18992900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992900&pc=OZD | 06/30/2022 | 04/09/2025 |  | 03/28/2024 |  |  | 06/30/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192876 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 85-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient developed a hematoma at the impella insertion site and heparin was withheld. it was also reported that during support, the patient would experience ectopy when the pump p-level was above p5. it was reported the patient had an episode of atrial fibrillation and became hypotensive. echocardiography was recommended to the medical team to verify impella placement and reposition as needed. no further actions were taken by the medical team, as the patient was already in the process of being weaned off the impella device and no further alarms or complications were observed with a lower p level. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28871 | 1220648-2024-08901 | 18992920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992920&pc=OZD | 03/19/2024 | 05/17/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  | 08/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397019 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient suffered a cerebral hemorrhage during impella cp support and passed away. it was noted that the purge solution had heparin at the time of the hemorrhage. per the physician, it was believed that the cerebral infarction had occurred prior to impella implant when the patient was initially unstable. it was further discussed that the improved perfusion provided by the pump may have led to a cerebral hemorrhage from the damaged area. a definitive cause for the cerebral hemorrhage could not be established. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. the instructions for use referenced in the initial report is from the ifu for impella cp with smart assist system. section: potential adverse events (united states)". d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28872 | 1220648-2024-08903 | 18992953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992953&pc=OZD | 03/18/2024 | 04/23/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406505 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Improper Flow or Infusion | [6 empty] | Pulmonary Edema | [12 empty] | Required Intervention | [5 empty] | PROTECK DUO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the field representative was notified of suction and "impella in aorta" alarms. the team did not acknowledge the alarms initially and the patient's condition deteriorated. the patient went into pulmonary edema and was re-intubated. upon recognizing decreased flows, the p level was adjusted and the issues with flow resolved. the patient's condition improved but when the patient coughed, the alarms continued. the impella support was not interrupted nor replaced. | [5 empty] | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues low pump flow could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28873 | 1220648-2024-08933 | 18992971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992971&pc=OZD | 06/22/2022 | 03/12/2025 |  | 03/28/2024 |  |  | 06/22/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced access site bleeding that required stoppage of heparin and addition of blood products to resolve. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28874 | 1220648-2024-08934 | 18992977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18992977&pc=OZD | 10/09/2022 | 04/07/2025 |  | 03/28/2024 |  |  | 10/09/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212532 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 58-year-old female in st segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient developed slight oozing and heparin was withheld. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28875 | 1220648-2024-08910 | 18993013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993013&pc=OZD | 03/15/2024 | 05/29/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  | 05/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352084 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Migration or Expulsion of Device | Patient-Device Incompatibility | Unintended Movement |  |  |  | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The patient's outcome was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of device with outcome. | The user facility reported a 76-year-old male with non-st elevated acute myocardial infarction and cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient had the impella device successfully implanted during cardiopulmonary resuscitation (cpr) for cardiac arrest. during the cpr compressions, the impella device became dislodged. due the patient having aortic stenosis, the device was unable to be reinserted. the patient was do-not-resuscitate, therefore therapy was discontinued. the patient eventually expired. |  |  |  |  | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. added-b5 medical safety officer review d4-updated model number added- d6a/d6b. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
28876 | 1220648-2024-08911 | 18993034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993034&pc=OZD | 03/19/2024 | 04/18/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  | 09/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024402258 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 59-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced excessive bleeding from their access site during impella cp support. fem-stop was placed to manage the condition, but the bleeding continued. the fem-stop was subsequently removed, the per-close was tightened, and pressure was held to manage the condition. the physician was advised to match the insertion angle at the site and a mattress suture was placed to achieve hemostasis. support continued with the implanted pump following the intervention. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since the bleeding occurred due to use issue since hemostasis was achieved by tightening the per-close, matching the insertion angle at the site and adding mattress suture. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28877 | 1220648-2024-08912 | 18993052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993052&pc=OZD | 03/20/2024 | 05/08/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451867 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 65-year-old male with non-st elevated myocardial infarction was implanted with an impella device for mechanical circulatory support. during support, access site bleeding occurred. the patient received four units of packed red blood cells and was stable. the impella was weaned and explanted, with plans for replacement. | [5 empty] | The investigation for the reported access site bleed has been completed. the pump was returned for investigation along with the introducer. there was no damage found with the pump and the introducer. the root cause of the access site bleeding was determined to be use issue because peel-away sheath was left in place during the duration of the runtime and was not removed and even upon recommendation to resolve the bleeding issues. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site name and address as previously entered incorrectly. h.6 code 4641 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report. additional manufacturer narrative on the previously submitted report stated that the device was discarded but the introducer and pump were returned. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
28878 | 1220648-2024-08897 | 18993057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993057&pc=PYX | 03/18/2024 | 10/15/2025 |  | 03/28/2024 |  |  | 03/18/2024 | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIMED, INC. |  |  | US |  | 2024428105 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 39-year-old male with cardiomyopathy was implanted with an impella rp flex device for mechanical circulatory support. the impella device became entangled with another device and was unable to be repositioned without removing the impella device. the procedure was started over and the impella device was successfully placed. | The complainant also reported that the patient was transported to the intensive care unit for recovery. during transport, a placement signal not reliable alarm occurred with a message reading flow calculations disabled. the alarm self-resolved. |  |  |  |  | The investigation into the delivery issues and the optical signal issue issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. case image confirms the swan-pump entanglement. the root cause of the positioning issue was most likely use related as the pump fell out of position as the sheath was being peeled away and the repo advanced; the pump then became tangled in the leave in swan. the pump was unable to be repositioned because it was tangled in the swan and had to be fully explanted and reinserted. it was reported that the cannula was kinked which likely occurred during the positioning issues. the root cause of the placement signal issue was not determined since the product and the data logs were not returned. "do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation." | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email and reporting contact fax number updated.
28879 | 1220648-2024-08951 | 18993169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993169&pc=PYX | 03/20/2024 | 05/10/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025448783 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | Pressure Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 52-year-old male with post cardiotomy cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support. an ongoing purge blocked alarm occurred on the impella device. low flows and erratic motor current were reported. the impella device was explanted, with the patient put on nitric oxide for right ventricle support. the patient was not on systemic anticoagulation. upon removal of the pump, a large clot was found on the distal end. there was no patient harm. | [5 empty] | The investigation into the high purge pressure and the low pump flow issues has been completed since the original report was filed. the pump was returned for analysis. upon visual inspection biomaterial was found on outflow cage and wrapped around impeller blade. the root cause of the low pump flow was due to biomaterial on the impeller blade as no anticoagulation was used. the root cause of the high purge pressure was sudden biomaterial ingestion. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
28880 | 1220648-2024-08907 | 18993175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993175&pc=OZD | 03/19/2024 | 05/03/2024 |  | 03/28/2024 |  |  | 03/19/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025448838 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that patient was in severe shock and when the pump was delivered to the left ventricle (lv) and placed deep. the patient had ventricular fibrillation (vfib) from pump placement and was shocked multiple times. the patient also had signs of hemolysis from the deep positioning. the team later gave blood and kept the pump on for support despite this until support was escalated to the 5.5 pump. | [5 empty] | The investigation for the reported hemolysis and ventricular tachycardia (vt) has been completed. the device was not returned for evaluation. data logs confirmed the pump was in aortic area (about 11.25 hours into the case). there was no alarm impella position in ventricle triggered. pump then was stopped manually and disconnected from the console. additionally, based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of the vt could not be determined without additional clinical details. added- d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.
28881 | 1220648-2024-08950 | 18993180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993180&pc=OZD | 06/20/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/20/2022 | 03/28/2024 |  |  | 12/31/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175012 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced limb ischemia. an external bypass was performed. it was also reported that initial echo showed that the device was too deep and caught on cordae. the impella was repositioned to prevent further hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28882 | 1220648-2024-08949 | 18993206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993206&pc=OZD | 06/16/2022 | 03/10/2025 |  | 03/28/2024 |  |  | 06/16/2022 | 03/28/2024 |  |  | 09/30/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022136570 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a 93-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that there was oozing at the groin site (access site) which was treated with manual pressure, withholding heparin, and stitching. however, the groin site continued to have some bleeding during the night, nevertheless, the next morning the site was soft without hematoma. eventually, the device was successfully removed. one of the two perclose sutures failed to complete the closure and the third preclose was unable to be placed. therefore, perclose x2 with manual pressure was applied and hemostasis was achieved. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28883 | 1220648-2024-08953 | 18993336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993336&pc=OZD | 10/22/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 10/22/2022 | 03/28/2024 |  |  | 07/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255609 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2022 |  |  |  |  | The user facility reported a 39-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed significant bleeding from the impella sheath. the physician placed 2 additional 2-0 silk sutures to the impella sheath. the dressing was changed and gauze was placed to elevate the sheath. manual pressure was held and femstop was currently in place. additionally, the patient went to complete heart block which required a temporary transvenous pacermaker (ttp) and then the patient went to the cardiac catheterization lab for pacer placement. the medical team made the decision to escalate the patient to an impella 5.5 for additional mechanical support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records
28884 | 1220648-2024-08952 | 18993502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993502&pc=OZD | 06/16/2022 | 03/10/2025 |  | 03/28/2024 |  |  | 06/16/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023194897 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the device, the sheath was repositioned causing oozing from the site. tighten the preclose and angle the entry site higher stopped the oozing. the patient experienced bleeding from the nose and the foley, which showed bloody urine. heparin was stopped being distributed and the low partial thromboplastin time required the heparin to be restarted. the patient's urine improved and was clear and there was some oozing from the pump site and there were no nose bleeds. eventually the patient was successfully weaned off the device. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28885 | 1220648-2024-08955 | 18993583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993583&pc=OZD | 07/11/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 07/11/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183104 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient was on impella support due to non-st elevated myocardial infarction, chest pain, hypertension, and acute myocardial infarction and cardiogenic shock. while on support, the patient experienced ventricular fibrillation. the patient expired after being on support for 380.18 hours. although the patient expired while on impella support, no allegations were made towards the impella for cause of death. | [5 empty] | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28886 | 1220648-2024-08956 | 18993686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993686&pc=OZD | 06/16/2022 | 10/01/2025 |  | 03/28/2024 |  |  | 06/16/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023205508 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that there were multiple attempts to reposition the device which were unsuccessful, due to the left ventricle was small with suspected interaction of the pump with the mitral valve. eventually, the device was successfully weaned and explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-08956 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-08956 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-08956 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08956. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-08956.
28887 | 1220648-2024-08958 | 18993697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993697&pc=OZD | 06/24/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 06/24/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182808 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 79-year-old male in non-st segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, there was no distal flow observed after the introducer was placed. a repo sheath was placed and distal flow was achieved. | [5 empty] | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28888 | 1220648-2024-08959 | 18993856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993856&pc=OZD | 03/20/2024 | 05/06/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426093 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Failure to Sense | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy and decompensation was implanted with an impella 5.5 device for mechanical circulatory support. during ongoing support, there were ongoing placement signal alarms. the staff turned the alarm audio off. the patient reported pain at the insertion site, and a hematoma and oozing developed at the insertion site. intervention for the hematoma and access site oozing remains unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. no product was returned. upon review of the data logs, oscillating contrast issue observed with psnr alarm relevant to definitive failure pattern suspecting the console. no problem was found with the pump. please refer to complaint # (b)(4) for an investigation into the console ic 2532. the cause of the access site bleeding issue was not determined due to insufficient clinical information and since product was not returned. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28889 | 1220648-2024-08957 | 18993894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18993894&pc=OZD | 06/27/2022 | 04/17/2025 |  | 03/28/2024 |  |  | 06/27/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the peel-away sheath was removed completely from the body, broken and peeled away, and the repositioning unit was advanced into the arteriotomy with good hemostasis. the patient remained on support. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
28890 | 1220648-2024-08960 | 18994067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994067&pc=OZD | 08/10/2022 | 04/17/2025 |  | 03/28/2024 |  |  | 08/10/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185904 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 42-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when attempting to peel away the peel away sheath, one of the wings snapped off and the peeling away process could not be completed. subsequently, the peel away sheath was cut using scissors. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28891 | 1220648-2024-08963 | 18994163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994163&pc=OZD | 03/22/2024 | 05/08/2024 |  | 03/28/2024 |  |  | 03/22/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445466 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | As a result of the purge leak, the pump was explanted the same evening the leak was found. no harm to the patient was reported. | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella 5.5 was placed to escalate care from the intra-aortic balloon pump (iabp) for a patient admitted with a abdominal aortic aneurysm and angina/shortness of breath. the impella 5.5 support was placed but when the purge pressure rose, the team used tissue plasminogen activator in the purge. the purge sidearm then began to leak. |  |  |  |  | The investigation for the high purge pressure and purge leak has been completed. pump was returned for evaluation. upon visual inspection of the purge sidearm within the retainer, no abnormalities were observed. additionally, no kinks in catheter or biomaterial covering purge gap upon visual inspection. pump was connected to console and purged with colored fluid. leak in purge sidearm was present almost immediately. sidearm retainer was removed and purge filter was found to be leaking and a crack in the filter was observed. sidearm bypass was performed and high purge pressure was reproduced. no ingress was found in the purge line leading up to the motor. the catheter was cut just below the motor and high purge pressure resolved. a teardown of the motor was performed, and when pulling the impeller blade, biomaterial was found inside the purge gap. the motor was cut at the purge gap and was able to purge successfully. data logs were reviewed and left logs showed a build-up of purge pressure over a period of 2 hours before high purge pressure alarms were triggered. bag change was performed after ¿high purge pressure¿ alarms were triggered to add heparin to purge. high purge pressure resolved itself suddenly and purge flow increased, triggering ¿purge pressure low¿ alarms that persisted throughout case. clinical detailed noted that purge flow was 2.5 and purge pressure was 925 and tpa protocol was started. the root cause of the purge leak was due to damaged sidearm filter. the root cause of the high purge pressure was due to biomaterial buildup inside the purge gap. added-b5 addl info to narrative, d9 device return date, h6 impact code 4644. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28892 | 1220648-2024-08936 | 18994430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994430&pc=OZD | 06/17/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 06/17/2022 | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023202662 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was bleeding at the insertion site. systemic heparin was held. subsequently, the site was reportedly dry, clean and intact. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
28893 | 1220648-2024-08966 | 18994497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994497&pc=OZD | 03/25/2024 | 05/07/2024 |  | 03/28/2024 |  |  | 03/25/2024 | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434163 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that an impella device was successfully placed in the patient. during impella use the patient developed a new hematoma and subsequently, heparin was turned off. the hematoma decreased and heparin was restarted the next day. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be anti coagulation management since the hematoma was decreased upon turning off the heparin. added-d6a/d6b, d4 updated model number | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28894 | 1220648-2024-08961 | 18994733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994733&pc=OZD | 03/21/2024 | 05/06/2024 |  | 03/28/2024 |  |  | 03/21/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324934 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The complainant reported that upon start-up, the automated impella controller (aic) displayed a system self-check error. the aic was exchanged. no patient was involved. | [5 empty] | The investigation for the console (aic) self check error has been completed. the console was returned for evaluation. the failure was not reproducible upon start up of the console. the console's internal circuitry inspected visually for any non conformance with no observable anomaly. there was nominal power being supplied from the pbm to the impellatronic board as well as the epc. a known good pump and purge line was connected to the console and run for more than four hours at the highest p-level. there was no observable anomaly. the data logs from the console confirm that there was a system self check failure event that was issued on the date of occurrence. from the logs, its was observed that the failure occurred when the console rebooted itself after initially booting up with cpld errors and then unexpectedly shutting down. the root cause of the system self check failure could not be determined due to the inability to reproduce the failure. however, an anomaly characterized by errors of communication between the cpld and the epc led to the unexpected shutdown and subsequent system self check failure of the console on the date of occurrence. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28895 | 1220648-2024-08965 | 18994799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994799&pc=OZD | 03/25/2024 | 05/16/2024 |  | 03/28/2024 |  |  | 03/25/2024 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025448037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 80-year-old male noted as high risk percutaneous coronary intervention (pci) was implanted with an impella device for mechanical circulatory support. the limb's access site was noted to have hematoma and bleeding followed by loss of distal pulses. angiogram showed some distal flow to impella device; however, the decision was made to insert fem-fem bypass to ensure perfusion at recommendation of vascular surgeon. once this was completed and pulses obtained, site dressed again as was previously done. due to oozing at site, the physician decided to place femstop without inflating which resulted in good hemostasis. the pump remained on despite event after the high-risk pci. | [5 empty] | The investigation into the limb ischemia ¿ reversible and the minor access site bleeding issues has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. d6a, d6b. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28896 | 1220648-2024-08971 | 18994850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18994850&pc=OZD | 10/30/2022 | 11/02/2024 |  | 03/28/2024 |  |  | 10/30/2022 | 03/28/2024 |  |  |  |  | 03/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105810 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Software Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup |  |  |  | 10/30/2022 |  |  |  |  | The user facility reported automated impella controller (aic) screen freezing during cassette change. the issue was resolved, however, a new cassette was required. | [5 empty] | Add information d4 lot number 1105810 section d1/d2 was incorrectly reported as impella 2.5 and common device name was incorrectly reported as temporary non-roller type left heart support blood pump in the initial mdr 1220648-2024-08971. section d4 was incorrectly reported as impella 2.5, and serial number incorrectly reported as (b)(6). section h5 was incorrectly reported as yes labeled for single use. section h8 usage of device incorrectly reported as initial use of device. | The investigation for the reported console ¿froze¿ during purge cassette change has been completed. the console was returned for evaluation. the logs were reviewed and relevant to the reported issue. the console was evaluated, and it was confirmed that the cause of the aic screen frozen during the purge change wizard was due to a software issue. this report is being filed as part of a retrospective review of historical records.
28897 | 1220648-2024-08954 | 18995322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18995322&pc=OZD | 03/20/2024 | 05/10/2024 |  | 03/28/2024 |  |  | 03/20/2024 | 03/28/2024 |  |  | 08/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443556 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2024 |  |  |  |  | The user facility reported an 85-year-old female was implanted with an impella cp device for mechanical circulatory support. during the implantation, the introducer became occlusive following insertion. a fem-fem external bypass was performed. support with the pump continued following the intervention. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28898 | 1220648-2024-08972 | 18995369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18995369&pc=OZD | 11/10/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 11/10/2022 | 03/28/2024 |  |  | 06/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249021 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 56-year-old female on impella cp support due to ventricular fibrillation arrest with 1 shock. while on support, the staff were unable to obtain pulses in the right groin. the impella cp was subsequently explanted. | [5 empty] | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. corrected h6 from 1888 hemorrhage/blood loss/bleeding and 3160 vascular dissections to 1942 ischemia. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
28899 | 1220648-2024-08964 | 18995488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18995488&pc=OZD | 03/25/2024 | 05/14/2024 |  | 03/28/2024 |  |  | 03/25/2024 | 03/28/2024 |  |  | 06/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024367262 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during patient transport via flight the nurse performed a console to console exchange and a yellow alarm was generated. the pump stopped and was unable to be restarted. the physician successfully replaced the pump with another impella 5.5 device. | [5 empty] | The investigation for the pump stop has been completed. the impella was returned for evaluation. upon inspection of the product, blood was found in the red pump plug and a hole at the 35cm mark was found. the impella defective alarm triggered in the field was because the hole in the catheter allowed blood ingress into red pump plug which shorted the electrical components preventing the console from reading the contents of the pump's double eeprom. biomaterial was also found in the cannula and on the impeller. the pump was run after clearing and pump stop was not reproduced. no logs were returned and there is no imc and journal data for the case dates. the root cause of the pump stop was determined to be blood ingress due to a hole in the catheter resulting from improper use. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
28900 | 1220648-2024-08898 | 18995506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18995506&pc=OZD | 03/18/2024 | 05/06/2024 |  | 03/28/2024 |  |  | 03/18/2024 | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024422932 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 78-year-old male of unknown race and ethnicity post cardiotomy cardiogenic shock, low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the pump was unable to make the turn into the ascending aorta during the implant. after multiple attempts, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through due to tortuosity of vessels as per the clinical description. "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter." "to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28901 | 1220648-2024-08969 | 18995515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18995515&pc=OZD | 09/24/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 09/24/2022 | 03/28/2024 |  |  | 07/31/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124654 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 57-year-old male was implanted with an impella cp device for mechanical circulatory support due to becoming unstable and coding after being placed on an intra-aortic balloon pump prior to percutaneous coronary intervention. while on support, a left ventricle thrombus was noted. the impella cp was explanted and the patient was in stable condition. | [5 empty] | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions, description of the location and characteristics of the thrombus (e.g., size, shape, consistency), relevant laboratory test results, such as coagulation profiles and platelet counts, imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus, any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states), medications or treatments that the patient was receiving at the time of thrombus formation, surgical or procedural details if applicable (e.g, cardiac catheterization), any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events), presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28902 | 1220648-2024-08935 | 18995749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18995749&pc=OZD | 06/23/2022 | 03/28/2024 |  | 03/28/2024 |  |  | 06/23/2022 | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188746 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a 55-year-old male was implanted with an impella cp and the patient developed a large hematoma at the left insertion site that was noted prior to the completion of the coronary artery bypass grafting. the physician removed the impella post cabg completion and a surgical closure was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28903 | 1220648-2024-08902 | 18996044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996044&pc=OZD | 07/01/2022 | 04/08/2025 |  | 03/28/2024 |  |  | 07/01/2022 | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023208231 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient developed a hematoma and bleeding at the impella access site. the patient was transfused with multiple units of blood and a purse string suture was placed to resolve the bleed. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28904 | 1220648-2024-08884 | 18996116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996116&pc=PYX | 12/08/2023 | 04/17/2024 |  | 03/28/2024 |  |  | 02/26/2024 | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024418640 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/26/2024 |  |  |  |  | In addition to the originally reported event, the patient also endured thrombocytopenia with a "possible" relationship to the device. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured acute blood loss with a "possible" relationship to the impella device. multiple blood products were required, and device was explanted as a result. |  |  |  |  | Included new information received in b(5) (thrombocytopenia) and additional codes in h(6). the investigation is still in progress. a supplemental report will be filed at the conclusion of our investigation. | The investigation into the access site bleeding ¿ major and the thrombocytopenia issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding and the thrombocytopenia issues was unable to be determined as no product and insufficient clinical details were provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.
28905 | 1220648-2024-08904 | 18996468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996468&pc=OZD | 03/18/2024 | 05/02/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024409296 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that following implant the patient exhibited bleeding at the access site. perclose sutures were adjusted to tamponade the site and manage the condition. the following day, the patient developed hemolysis and anuresis. the pump was explanted later that morning following the diagnosis. | [5 empty] | The investigation into the hemolysis and the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. it was reported there were intermittent suction alarms after pump running for 7 hours. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. based on clinical description, the root cause of access site bleeding was most likely due to anticoagulation management related. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 & h10 were revised with updated information. | The impella device was not received from the customer and therefore,m an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
28906 | 1220648-2024-08967 | 18996496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996496&pc=OZD | 03/22/2024 | 05/06/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423322 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 47-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was on impella cp support and there were low pump flows noted. after the implant, there were immediate suction events. blood products were given to the patient. a kink was noted in the cannula and the device was removed and replaced. | [5 empty] | The investigation for the reported low pump flow and product damage has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the cause of the low pump flow issue most likely was cannula kink due to improper technique. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28907 | 1220648-2024-08919 | 18996578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996578&pc=OZD | 03/22/2024 | 05/31/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 02/28/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2024310481 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There is not enough information to associate use of device with outcome. | The user facility reported a 77-year-old female was implanted with an impella device and high purge pressure was experienced during patient support. tpa was added to the purge line to manage the issue. there was no patient harm. |  |  |  |  | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. however, the data stick was returned. the root cause of high purge pressure was most likely biomaterial as it was able to be resolved by adding tpa. b5, d6a, d6b. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28908 | 1220648-2024-08874 | 18996683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996683&pc=OZD | 06/30/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 10/31/2023 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 60-year-old male patient was on impella cp support due to 15% ejection fraction. while on support, the patient developed a large hematoma and external bleeding which was resolved with manual compression and vascular surgery. device was replaced with an impella 5.5. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28909 | 1220648-2024-08974 | 18996832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996832&pc=PYX | 03/25/2024 | 05/16/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 09/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404240 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during preparation, the impella rp device would not prime. purge fluid was exiting the purge tubing, but when connected to the purge side arm several failures to prime notifications were displayed. no kinks or blocked areas were seen. new impella device pump was opened for use. there was no harm to the patient. | [5 empty] | The investigation for the priming issue has been completed. the impella device was returned for evaluation. upon evaluation of the product, adhesive was found cured within the filter neck. the priming issue was reproduced. data logs were not returned. the root cause of the priming issue was determined to be improper manufacturing operator technique resulting in adhesive in the filter neck. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28910 | 1220648-2024-08916 | 18996979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18996979&pc=OZD | 03/13/2024 | 05/16/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024409974 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced bleeding from the needle hole in the anastomosis. re-stuturing was performed and the bleeding was stabilized. the patient is stable and remained on impella 5.5 support. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation. the failure was identified as access site bleeding during clinical data review. the root cause of the access site bleeding major issue was most likely use related as there was a needle hole observed in the anastomosis and re-suturing it resolved the issue. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
28911 | 1220648-2024-08908 | 18997065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997065&pc=OZD | 03/19/2024 | 03/04/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1739661 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 29-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that the automated impella controller (aic) had a purge leak. the purge cassette (pc) and the aic had leakage at the rear of the purge pressure transmitter and instructions for use were followed. the aic and pc were removed from service. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the console (aic) purge disc/flag issues has been completed. data logs were reviewed, and no relevant information was found in the logs. console was returned fore evaluation. the reported complaint was confirmed; leaked purge fluid was observed contaminating the area around the purge disc retainer. the purge fluid was causing the purge disc detect flag to malfunction; when a purge disc was inserted, the console was unable to detect it. the purge motor and purge pressure sensor were tested and confirmed to be working correctly. the cause of the aic issue was contamination of the purge disc flag due to purge fluid ingress. added d9 device return date corrections to the initial mfr report: f6/f8 should have been left blank. g1 mdr reporting contact email.
28912 | 1220648-2024-08875 | 18997099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997099&pc=OZD | 10/10/2022 | 03/27/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215965 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 57-year-old female was implanted with an impella cp. while patient was on impella cp, the patient had site oozing. manual pressure was applied but persisted. the doctor ordered no systemic heparin currently due to oozing, impella was removed and the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28913 | 1220648-2024-08909 | 18997260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997260&pc=OZD | 06/15/2022 | 05/22/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206528 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/15/2022 |  |  |  |  | The user facility reported a 80-year-old female was implanted with an impella cp device for mechanical circulatory support and that it was determined through angiogram that the patient had a dissection that required a balloon tamponade via the left femoral artery access site. the patient was successfully weaned off the impella cp and hemodynamics was normalized. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
28914 | 1220648-2024-08970 | 18997266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997266&pc=OZD | 03/24/2024 | 05/29/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422719 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2024 |  |  |  |  | The patient's outcome was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of device with outcome. | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 would not advance through the introducer and graft freely. the impella 5.5 appeared damaged. a new 5.5 was used and it did not have any difficulty advancing through the new introducer and was inserted successfully. |  |  |  |  | The investigation for the reported difficult insertion has been completed. the device was not returned for evaluation. clinical details stated that impella 5.5 catheter would not advance through introducer and graft freely. clinical team reported a catheter kink. a new 5.5 and introducer was used and there was no difficulty implanting. the root cause of difficulty inserting the pump through introducer was not determined as no product was returned. the instructions for use referenced in the initial report is from ifu for the impella 5.5 with smart assist for use during cardiogenic shock, section potential adverse events (united states) , which now includes statement "cardiac or vascular injury (including ventricular perforation)¿ added- b5 medical safety officer review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28915 | 1220648-2024-08977 | 18997271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997271&pc=OZD | 07/07/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198689 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention and st elevated myocardial infarction, was implanted with an impella cp device for mechanical circulatory support. while on support, the patient showed signs of hemolysis and was started on dialysis. the patient also experienced episodes of dysrhythmia during the night. the patient was reported to be stable and the impella position was verified via echocardiogram. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 mdr reporting contact name and address g1 manufacturing site name and address | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28916 | 1220648-2024-08885 | 18997279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997279&pc=OZD | 07/03/2022 | 03/27/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023180592 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2022 |  |  |  |  | The user facility reported a 57-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced a hematoma and oozing at the insertion site when moving around a lot. the hematoma and oozing was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28917 | 1220648-2024-08899 | 18997522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997522&pc=OZD | 03/19/2024 | 05/07/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 10/31/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419177 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 26-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was unable to advance through the impella sheath as the sheath had a kink. a new kit was opened and a new 14 fr sheath placed. a new impella cp was also inserted because of concerns of damaged to the impella when attempting to push through the kink. the second cp was successfully placed. | [5 empty] | The investigation for the reported introducer kink has been completed. the device was not returned for evaluation. clinical details were provided stated that the impella cp was unable to advance through 14 fr. 25cm sheath. once the device was removed, an injection through the sheath revealed a kink due to stick angle and body habitas. the root cause of introducer damage - mechanical was most likely a sheath kink. added- d6a/d6b, d4 updated model number | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28918 | 1220648-2024-08893 | 18997630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997630&pc=OZD | 03/14/2024 | 05/15/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025447615 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The complainant reported a patient (race unknown) post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during implant, the cannula kinked after several attempts to get the pump over a narrow area at the transition from the axial artery to the brachial artery. the pump was not removed as the kink in the cannula did not impact patient support. | [5 empty] | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical information provided, the problem was identified as a delivery issue. the cause of the delivery issue was patient condition due to patient tortuosity vessels condition that kinked the cannula. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28919 | 1220648-2024-08917 | 18997633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18997633&pc=OZD | 03/21/2024 | 04/29/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261785 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | This complaint was not related to a clinical procedure, and there were no adverse effects related to this event. | No patient involvement:us complainant reported the automated impella controller (aic) had an issue with the purge disc holder was broken off. |  |  |  |  | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added- b5 additional narrative d4- updated model number. | The aic was received from the customer, the investigation has begun, once the investigation is complete, a supplemental mdr will be filed.
28920 | 1220648-2024-08982 | 18998136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998136&pc=OZD | 04/04/2023 | 03/28/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 1209416 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | Crack | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/05/2023 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a purge disc cracked when placing purge disc. the pump was switched to another console and the procedure was completed successfully. no patient harm. | [5 empty] | The impella device was received from the customer on (b)(6)2023. data analysis: after analyzing the console's data (aic logs), it was found there a couple different alarms on the complaint date: purge pressure low alarms and purge disc not detected alarms. see logs in the attachments section. device analysis: the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken (broken blue disc retainer - see attached image). before returning this console back to the customer, fs was instructed to replace the purge disc retainer/ flag, validate sensor accuracy via section 12 of the pm procedure (0042-7309), and perform a full functional check on the console and optical system (0042-7316).
28921 | 1220648-2024-08979 | 18998153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998153&pc=OZD | 03/22/2023 | 03/28/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1414254 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Self-Activation or Keying | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/22/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a controller failure that occurred during use of the device. there pump was switched to a new console with no harm to the patient. | [5 empty] | The impella device was received from the customer on (b)(6)2023 console logs from the reported day of event are consistent with complaint and show that pump 417339 failed to start, resulting in ¿impella stopped¿ alarm (#13) due to motor current high. repeated attempts to start the pump were unsuccessful, each time resulting in ¿impella stopped¿ alarm, but also a controller error alarm #138 ¿pump speed deviation detected by pmd¿. the pump was eventually transferred to a backup console (ic3733) where it was able to be started. console ic7151 was tested in danvers but the issue was not reproduced, and an engineering pump and a pump simulator were able to be started without issues. there are known scenarios where pump might be unable to start for example due to clot forming during insertion, which gets dislodged during pump manipulation and/or reinsertion, but we were unable to verify this theory due to pump not having been returned for investigation. the root cause of pump being unable to start remains undetermined, although we did not find any irregularities with console operation. repair: perform functional check of the console.
28922 | 1220648-2024-08973 | 18998165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998165&pc=OZD | 03/25/2024 | 05/09/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461215 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old female with cardiomyopathy was to be implanted with an impella 5.5 device due to needing a bridge to heart transplant. the device was unable to pass though the vessel. a decision was made to abort implant of impella 5.5 and place an impella cp through graft. there was no report of adverse effect to the patient. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. the implant was unsuccessful due to being unable to pass though the vessel as per the clinical description. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28923 | 1220648-2024-08882 | 18998207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998207&pc=OZD | 07/01/2022 | 05/28/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212140 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 31-year-old female patient implanted with an impella cp for mechanical circulatory support. while on support the patient experienced excessive bleeding at the impella site and gastrointestinal (gi) bleeding. bicarb was added to the purge. additionally, it was reported the patient had pulseless electrical activity (pea) arrest. one round of cardiopulmonary resuscitation (cpr) was given and a return of spontaneous circulation occurred, however the patient expired on support. no allegations were made towards the impella for cause of death. | [5 empty] | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28924 | 1220648-2024-08881 | 18998392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998392&pc=OZD | 07/01/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214034 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Suction Problem | Pumping Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 60-year-old female patient was on an impella cp support due to st-elevated myocardial infarction. while on support, the patient experienced large amounts of bleeding from the right groin. additionally, a low pump flow indicated by suction alarms, were present on the device. both issues were resolved with five units of red blood cells (rbc¿s). | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28925 | 1220648-2024-08921 | 18998515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998515&pc=OZD | 03/19/2024 | 05/08/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an echocardiogram was performed and a large thrombus along the anteroseptal wall was seen. the impella cp was removed. the patient is stable. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 4114 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28926 | 1220648-2024-08984 | 18998785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998785&pc=OZD | 06/29/2022 | 03/07/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193821 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient's illiac size was too small so after the pci procedure the cp was removed to due concerns for limb ischemia. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28927 | 1220648-2024-08987 | 18998793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998793&pc=OZD | 06/20/2022 | 03/28/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188117 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient¿s feet were cold and pulses could not be found. subsequently, the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28928 | 1220648-2024-08877 | 18998831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998831&pc=OZD | 02/29/2024 | 04/18/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 11/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434621 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. heparin was discontinued and purse string sutures were placed. the patient received multiple blood products while on support. the patient remained on impella support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, purse string sutures were placed to attain hemostasis as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28929 | 1220648-2024-08986 | 18998874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998874&pc=OZD | 10/19/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189825 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 52-year-old male on impella cp for mechanical circulatory support. the patient experienced bleeding at insertion site and ventricular fibrillation. the patient was given an additional (b)(4) units heparin. the impella cp was removed due to suboptimal pump position and concern for clot along cannula. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28930 | 1220648-2024-08983 | 18998931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998931&pc=OZD | 06/27/2022 | 03/07/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 02/29/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195272 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The complainant reported the patient was on impella cp support due to st-elevated myocardial infarction. while on support, the patient experienced ischemia in the left lower extremity that was resolved with external fem-fem- bypass. additionally, a large clot was noted on the end of the pigail as the pump was being explanted. explant of the pump proceeded successfully without any noted complications to the procedure. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Thrombus: analysis summary the failure mode of "thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. limb ischemia: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication vascular size or variations thereof detailed description of the symptoms experienced by the patient physical examination findings, including pulse assessment and visual examination of the affected limb diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography laboratory test results, including blood tests for markers of inflammation or clotting disorders any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
28931 | 1220648-2024-08978 | 18998971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18998971&pc=OZD | 02/24/2023 | 03/28/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  |  |  | 03/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1390491 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/24/2023 |  |  |  |  | No patient involvement: the complainant reported that their aic (automated impella controller) had a broken purge clip. the device was removed from service as a result of the damage. | [5 empty] | The automated impella controller device was received from the customer on 03-mar-2023. data analysis: imc logs revealed that there were purge pressure low and open alarms which could have been related to the reported issue. see imc logs attached. device analysis: the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken (broken blue disc retainer - see attached image). before returning this console back to the customer, fs was instructed to replace the purge disc retainer, validate sensor accuracy via section 12 of the pm procedure (0042-7309), and perform a full functional check on the console and optical system (0042-7316).
28932 | 1220648-2024-08906 | 18999565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18999565&pc=OZD | 03/19/2024 | 05/07/2024 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 09/30/2025 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABOIMED, INC |  |  | US |  | 2024410050 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported an 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed red tinged urine requiring one unit of packed red blood cells and it was also noted upon placement check via echocardiogram the device was noted too deep and was pulled back. the patient was reported to be in stable condition. | [5 empty] | The investigation for the hemolysis has been completed. the complaint device not returned for investigation. data loges were reviewed finding no motor current spikes observed. impella in aorta positioning alarms were confirmed on (b)(6). the cause of the hemolysis issue cannot be determined since insufficient clinical data provided. positioning issue resolved but we don't know if that solved hemolysis issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.
28933 | 1220648-2024-08985 | 18999620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18999620&pc=OZD | 10/13/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218279 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported a 89-year-old female on impella cp for mechanical circulatory support. it was reported that the patient developed a hematoma on the groin and femstop was applied to reduce the hematoma expansion. the physician noted a pseudo aneurysm that required a covered stent. the physician believed this occurred during implant because of the patient¿s small vessel size and tortuosity. the patient's purge solution was switched d5w to sodium bicarb. the plan was also to start heparin drip at 500 u/hr once the ptt comes down and currently, the ptt was over 200. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows:
28934 | 1220648-2024-08980 | 18999707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18999707&pc=OZD | 08/04/2022 | 05/20/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217069 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 67-year-old patient needing mechanical circulatory support. it was reported that a. hematuria was observed so an echo was performed and the device was measuring 6cm. the device was pulled back but inadvertently was pulled completely out of the ventricle. the patient was then taken to the cath lab where the impella was removed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28935 | 1220648-2024-08988 | 18999739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18999739&pc=OZD | 06/21/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 03/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206507 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced hemolysis. the complication was successfully managed at bedside. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records
28936 | 1220648-2024-08981 | 18999892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18999892&pc=OZD | 07/08/2022 | 03/12/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215578 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 59-year-old male patient needing mechanical circulatory support. while on support, the patient had left distal perfusion catheter in place from subclavian to radial. the patient had lost pulses in left arm. impella successfully weaned and explanted. no clot noted on impella catheter. clot visualized in sheath. images revealed clot formation. the inari flow triever retrieval/aspiration system deployed and clot was successfully removed | [5 empty] | This report is being filed as part of a retrospective review of historical records. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. analysis summary the failure mode of "thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28937 | 1220648-2024-08989 | 19000246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19000246&pc=OZD | 06/21/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis presented. the patient remained on intra-aortic balloon pump 1:1 with full augmentation. echocardiogram was done to verify position. it was reported that the hemolysis resolved. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
28938 | 1220648-2024-08968 | 19000249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19000249&pc=OZD | 08/11/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 04/30/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 46-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. access site bleeding was observed. the surgeon placed covered stent in common femoral artery to obtain hemostasis. patient developed stent thrombosis while in the operating room. angioplasty and thrombectomy procedures were performed to resolve stent thrombosis. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28939 | 1220648-2024-08990 | 19000330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19000330&pc=OZD | 06/30/2022 | 04/08/2025 |  | 03/28/2024 |  |  |  | 03/28/2024 |  |  | 01/31/2024 |  | 03/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced bilaterial limb ischemia. the physician added dobutamine to increase pulsatility and distal limb perfusion. the patient was successfully weaned off the impella cp and hemodynamics was normalized. | [5 empty] | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28940 | 1220648-2024-08781 | 18985262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985262&pc=OZD | 03/02/2024 | 05/08/2024 |  | 03/27/2024 |  |  | 03/08/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454423 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Female |  |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant from the study reported a patient on impella 5.5 support experienced tachyarrhythmia absoluta, up to 150 beats per minute on admission. amiodarone saturation was in progress and two external electrical cardioversione (ecvs) with conversion to sinus rhythm occurred during the day. the impella had suction alarms and volume was administered. following ecv, there were numerous polytopic ventricular extra systoles with long pauses, and then asystole. the patient had a pacemaker that was not taking over unless stimulation was given. the impella's position was checked via transesophageal echocardiogram (tee), which showed the impella was twisted and was pointing towards the septum and not the apex. the device was obstructing the opening of the mitral valve. the pump was repositioned. the events were determined to possibly be related to the device and the procedure. the impella was noted to have been successfully explanted after five days of support and the patient recovered native heart function. | [5 empty] | The investigation into the positioning issues and the vf/vt/af/at has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue was most likely use related. as the necessary information was not provided, the root cause of the vt/vf was not determined. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28941 | 1220648-2024-08783 | 18985329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985329&pc=OZD | 03/11/2024 | 05/29/2024 |  | 03/27/2024 |  |  | 03/11/2024 | 03/27/2024 |  |  | 06/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375160 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during high risk percutaneous coronary intervention procedure, a 74-year-old female patient was on impella cp support. when the case was initiated, the device was noted to have gotten kinked in the outlet area. there were no signals and flow was around 1.3 liters. therefore, decision was made to explant the device. during removal, the physician was unable to explant the impella cp through the peel away. the physician took away half part of the device and left the other inside the peel away. the peel away and the other half were then removed. there was no reported patient harm and the procedure was successfully completed. | [5 empty] | The investigation has been completed. the returned complaint devices were visually inspected. the cannula was kinked and detached at the outflow cage. the kink was observed in the returned introducer. the cause of the product damaged (product damage (detached inflow/outflow - peripheral vessel) was sheath kink due to improper technique. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. | H.6 type of investigation code was revised. code 4121 was entered incorrectly in the previously submitted report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28942 | 1220648-2024-08782 | 18985330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985330&pc=OZD | 07/08/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 07/08/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204039 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 77-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support, patient had little hematoma at the groin site. therefore, the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28943 | 1220648-2024-08784 | 18985339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985339&pc=OZD | 07/09/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 07/09/2022 | 03/27/2024 |  |  | 10/31/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152010 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/09/2022 |  |  |  |  | The user facility reported a 61-year-old patient needing mechanical circulatory support. it was reported that the patient on impella cp had limb ischemia. lfa site showed distal perfusion down limb. no flow down the leg. given continued improvement in patient status, pump was weaned off and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28944 | 1220648-2024-08785 | 18985453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985453&pc=OZD | 03/11/2024 | 05/01/2024 |  | 03/27/2024 |  |  | 03/11/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453882 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a 64-year-old female patient had an attempted impella 5.5 implant, but it was unable to pass through the artery. lubrication was applied without success. the implant was aborted and an impella cp was placed. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery is determined to be patient condition because the impella 5.5 was unable to pass through innominate artery of the patient and later they were able to implant impella cp successfully. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. to conform to updated procedures, section d6 and h10 have been revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28945 | 1220648-2024-08786 | 18985482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985482&pc=OZD | 06/20/2022 | 05/20/2025 |  | 03/27/2024 |  |  | 06/20/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201027 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/20/2022 |  |  |  |  | The user facility reported a 67-year-old patient needing mechanical circulatory support. it was reported that when taking out and repositioning the sheath, the impella got pulled back and kinked. the impella was removed, alternate sheath was placed, and the impella was replaced. echo was done at bedside and placement appeared shallow, but it was reported that the patient's left ventricle was small and it would not be possible to advance without ectopy. during support, the patient experienced atrial fibriliation. the impella was repositioned and then the flow increased. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28946 | 1220648-2024-08787 | 18985535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985535&pc=OZD | 03/12/2024 | 04/19/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1185332 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | Additional information states that the console was in use during a clinical procedure, but there were no adverse effects related to this event. console was sent to field service for further investigation. | The user facility's biomedical engineer reported that the automated impella controller (aic) experienced battery failure. |  |  |  |  | The investigation for the console (aic) red alarm has been completed. the console was returned for evaluation. during functional testing, the reported battery failure was able to be reproduced. logs were reviewed which showed console having a battery failure alarm on the reported occurrence date. it also showed multiple instances of transport connection blown/missing alarms (event set #360). batttest results show no signs of battery cell imbalance. it is suspected that the battery failure was use related. it is suspected that the user booted up the console while it was connected to ac power and the battery switch was disengaged. as soon as the console finished booting up the battery failure (transport connection blown/missing) alarm was triggered. after that, user tried engaging the battery switch but since there is a delay with the pbm noticing that the batteries are engaged, the alarm would still persist which would lead the user to assume that the failure was still ongoing. instead of waiting long enough for the pbm to notice that the battery switch was engaged, user most likely disengaged it again which caused the alarm to continue. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28947 | 1220648-2024-08788 | 18985553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985553&pc=OZD | 03/07/2024 | 05/08/2024 |  | 03/27/2024 |  |  | 03/07/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025451086 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Device Dislodged or Dislocated | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 69-year-old male patient, the pump was unable to be repositioned successfully leading to pump having to be explanted. the pump had been dislocated during resuscitation attempt the previous night. there was no reported patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. the impella cp pump was not returned. the cause of the positioning issue was use related due to cpr being done during support. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
28948 | 1220648-2024-08789 | 18985558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985558&pc=OZD | 06/21/2022 | 05/13/2025 |  | 03/27/2024 |  |  | 06/21/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206478 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The user facility reported a 71-year-old male was implanted with an impella device for mechanical circulatory support. while they were loading the patient into the ambulance, the emergency medical technician without warning pushed the stretcher with some force into the ambulance and were not paying attention to the console and this power cable to the cp was pushed into the door of the ambulance which snapped it out the console, breaking the pins in the connector cable to aic. the pump stopped and immediately alarmed and the connector cable was broken. onsite, the physician withdrew impella and the patient was emergently taken to cath lab. a hematoma noted at access site of impella. decision made to proceed with exchanging of cp. a new impella cp was then advanced over the .018 wire and into place. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction : section b1 and h6 (medical device problem code) has been left out inadvertently in the initial report which is added to this report .
28949 | 1220648-2024-08790 | 18985588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985588&pc=OZD | 07/10/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 07/10/2022 | 03/27/2024 |  |  | 04/30/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, the device had numerous positioning alarms upon the patient's admission to the cardiac care unit (ccu) for ¿impella in ventricle¿. the device was repositioned at bedside. shortly after repositioning and during patient transport, the medical team observed additional position alarms. it was recommended to verify impella position with echocardiography and reposition as needed. no further adjustments were needed for impella and no further alarms were reported by the medical staff. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28950 | 1220648-2024-08791 | 18985589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985589&pc=OZD | 03/08/2024 | 04/18/2024 |  | 03/27/2024 |  |  | 03/08/2024 | 03/27/2024 |  |  | 11/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024433030 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Injury | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 69-year-old male of unknown race and ethnicity in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that when impella cp was removed the peel away sheath was removed over wire and perclose tightened but did not take. additional perclose deployed and angioseal was placed. no patient harm was reported. | [5 empty] | The investigation into the removal issues (vascular damage - peripheral vessel) has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage was most likely use related as there an issue with the perclose after impella was removed. there were no reported issues with impella pump. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28951 | 1220648-2024-08792 | 18985642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985642&pc=OZD | 03/12/2024 | 04/30/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  | 07/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443397 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient's right foot became mottled due to a lack of flow in the leg following impella cp implant. in response to the noted condition, the decision was made to remove the pump. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in order to conform to updated procedures, sections d6 and h10 have been revised with updated informatoin. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28952 | 1220648-2024-08793 | 18985657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985657&pc=OZD | 03/12/2024 | 03/03/2025 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1493604 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 64-year-old male in st-segment elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that automated impella controller (aic) placement signal was not reliable and displayed controller failure alarms. the aic was exchanged and there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) controller failure-red alarm has been completed. data logs revealed that that a controller error alarm (opm comm loss ¿ 1045) occurred after a placement signal not reliable alarm. following the set of errors, the optical bench was intentionally reset, which resolved the placement signal issue. real time (rt) logs confirmed that immediately prior to the controller error alarm, the optical bench reported an oscillating contrast value far outside the normal range of contrast. console was returned for evaluation. the issue was not able to be reproduced in the lab when running a known-good pump and purge system. the cause of the aic issue was a defective optical bench. corrections to the initial mfr report: d2-updated common device name. f6/f8 should have been left blank. g1 mdr reporting contact email.
28953 | 1220648-2024-08794 | 18985670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985670&pc=OZD | 06/26/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/26/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023195325 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding around the venoarterial extracorporeal membrane oxygenation (va ecmo) site and reported that a hematoma developed following the impella implant. the patient was transfused with blood as a result of the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28954 | 1220648-2024-08795 | 18985677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985677&pc=OZD | 03/12/2024 | 04/19/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  | 09/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403718 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support due to the patient coding. while on support, a large hematoma was noted, sutures were loose. dressing was removed, sutures replaced, angle matching dressing placed and the hematoma was washed out. the bleeding subsided. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue because it was noted that the sutures were loose ang hemostasis was achieved upon replacing the sutures. d4: updated model number. g1: updated mfr site name and address | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28955 | 1220648-2024-08799 | 18985822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985822&pc=OZD | 06/21/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/21/2022 | 03/27/2024 |  |  | 12/31/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177726 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2022 |  |  |  |  | The user facility reported a 1-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was oozing at the insertion site. a suture was placed at site. additionally, it was reported that the patient had hemolyzed labs and was septic. patient was brought to operating room for explant and exploratory surgery. device was weaned and explanted successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Access site bleeding: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. hemolysis: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. infection/sepsis: analysis summary: an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician. physical examination findings. results of special investigations : laboratory test results & imaging studies. microbiological culture and sensitivity results. response to previous treatments and therapies. any relevant epidemiological information or exposure history. concomitant devices in use. preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined.
28956 | 1220648-2024-08800 | 18985839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985839&pc=OZD | 03/13/2024 | 04/29/2024 |  | 03/27/2024 |  |  | 03/13/2024 | 03/27/2024 |  |  | 11/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431419 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed hematoma at the impella access site. heparin was held and the patient had a washout of the 5.5 pocket. additionally, a social media post stated that the patient was taken to the operation room for a procedure due to clots from the impella. the patient is stable. | [5 empty] | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of access site bleeding was determined to use issue because it was mentioned that the bleeding issue was resolved after angle matching. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 and h10 have been revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28957 | 1220648-2024-08801 | 18985844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985844&pc=OZD | 06/27/2022 | 04/11/2025 |  | 03/27/2024 |  |  | 06/27/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183477 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. the patient was transfused with red blood cells. a sterile dressing and a pressure bandage were applied at impella insertion site to resolve the bleed. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28958 | 1220648-2024-08802 | 18985857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985857&pc=OZD | 03/12/2024 | 04/19/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1427009 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 64-year-old male in st-segment elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the automated impella controller (aic) purge system failure alarm occurred. it was reported that inside the purge door where the cassette fits in which rotates for purge cassette function seemed to be locked up or frozen according to customer. the aic was exchanged and there was no patient harm reported. | [5 empty] | The investigation for the reported console (aic) purge issues. the impella was returned for evaluation. visual inspection of the purge assembly revealed excessive purge fluid residue, indicating leakage of purge fluid during clinical use. purge fluid residue was also found within the purge motor assembly, preventing the purge motor from spinning both by hand and via telnet commands to the console. logs were reviewed which confirmed the reported complaint. there was an instance of a purge system failure alarm due to a blocked piston. the root cause of the purge system failure was due to buildup of glucose on the purge stepper motor gasket. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28959 | 1220648-2024-08803 | 18985866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18985866&pc=OZD | 06/22/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/22/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a 44-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient was experienced oozing at the access site. manual pressure was applied, platelets administered, and dressings were changed. subsequently, the patient presented with good hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28960 | 1220648-2024-08804 | 18986017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986017&pc=OZD | 06/24/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/24/2022 | 03/27/2024 |  |  | 01/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185117 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 54-year-old male with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. while on support, a decision was made to remove the impella device after the echocardiogram showed a possible clot at the inlet. the impella was successfully explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus analysis summary the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
28961 | 1220648-2024-08805 | 18986026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986026&pc=OZD | 07/12/2022 | 05/12/2025 |  | 03/27/2024 |  |  | 07/12/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197759 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The complainant reported that the patient on impella cp was explanted due to extensive bleeding in the groin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28962 | 1220648-2024-08806 | 18986039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986039&pc=OZD | 06/27/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/27/2022 | 03/27/2024 |  |  | 12/23/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180180 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, a hematoma formed at the impella groin access site. the hematoma remained under observation with no intervention until the pump was explanted. after the pump was explanted, two perclose devices were used in attempt to resolve the hematoma, but the attempts were unsuccessful. a peripheral balloon was inserted and inflated over the impella site to achieve hemostasis. it was reported the groin dry and intact with no hematoma upon conclusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28963 | 1220648-2024-08807 | 18986056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986056&pc=OZD | 06/22/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/22/2022 | 03/27/2024 |  |  | 01/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187150 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a 87-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a hematoma at the access site. systemic heparin was held and the patient reportedly stabilized. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28964 | 1220648-2024-08808 | 18986086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986086&pc=OZD | 07/12/2022 | 04/16/2025 |  | 03/27/2024 |  |  | 07/12/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209625 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 56-year-old male was implanted with an impella cp. the complainant reported that the patient on impella cp had positioning issues. when patient was boosted in bed, positioned in aorta. md repositioned once at the bedside but the impella ¿came out again.¿ the decision was made to pull the impella into descending aorta and put on p-2 with plans to potentially explant. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28965 | 1220648-2024-08809 | 18986165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986165&pc=OZD | 03/12/2024 | 05/25/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438840 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pressure Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | Upon review by abiomed's medical safety officer, the patient is on comfort care only and the doctor will not do anything more. there is not enough information to associate the use of the device with the patient¿s outcome. | The user facility reported a 66-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that a high-pressure alarm occurred while on support. it was reported that while on support, the patient developed heparin-induced thrombocytopenia (hit). the physician elected no intervention as the patient was placed on comfort measures only. |  |  |  |  | The investigation into the high purge pressure and the thrombocytopenia issues have been completed since the original report was submitted. the pump was not returned for investigation the data logs were returned and investigated. there was a motor current spike followed by purge pressure high alarms ending with purge system blocked alarm. the data logs showed characteristics of biomaterial ingestion. per the data log review, the root cause of the high purge pressure issue was most likely biomaterial. the root cause of the thrombocytopenia issue was not determined since the product was not returned and insufficient clinical information was provided. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 and 3233 were inadvertently omitted from the previously submitted report. the device was not returned and the investigation has been completed since the previous report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28966 | 1220648-2024-08810 | 18986186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986186&pc=OZD | 06/23/2022 | 09/11/2024 |  | 03/27/2024 |  |  | 06/23/2022 | 03/27/2024 |  |  | 06/30/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022100778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Death | No | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 88-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was suction issues with the device leading to pericardial tamponade being performed. however, the suction reoccurred causing the pressure to drop. subsequently, the patient expired while on support. no allegations were made towards the impella device as the cause of death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08810 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28967 | 1220648-2024-08811 | 18986224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986224&pc=OZD | 03/14/2024 | 04/24/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445452 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. following implant of an impella 5.5 pump, the sterile sleeve was found to be torn and no longer secured. the sleeve was reattached and tegaderm was placed to secure the component. support continued as planned with the implanted pump. there was no patient harm. | [5 empty] | The investigation for the product damage has been completed. the 5.5 pump was not returned. clinical details were not sufficient to determine the root cause. the cause of the product damage (sterile sleeve) issue was not determined since product/images were not returned for investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28968 | 1220648-2024-08812 | 18986309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986309&pc=OZD | 06/23/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/23/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023195346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a 36-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma. the device had to be explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28969 | 1220648-2024-08813 | 18986421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986421&pc=OZD | 07/10/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 07/10/2022 | 03/27/2024 |  |  | 04/30/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213587 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/10/2022 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump experienced persistent positioning issues. it was initially noted that the pump was under the papillary muscle. the pump was removed, flushed, and re-inserted, but continuous suction waveforms/alarms were seen. it appeared that the aortic valve leaflets were pushing the pump towards the wall of the heart confirmed via fluoroscopy/echo. the hospital staff elected to run support at p6 to avoid suction to manage the poor positioning. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28970 | 1220648-2024-08814 | 18986436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986436&pc=OZD | 03/14/2024 | 05/15/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  | 07/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024435214 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 was unable to be placed due to the anatomy of the 75-year-old male patient. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the pump was retuned and no abnormalities were found. the root cause of the delivery issue was most likely due to patient condition since impella 5.5 couldn't be placed due to vascular anatomy of the 75-year-old male patient. added- d9 device return date in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
28971 | 1220648-2024-08815 | 18986450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986450&pc=OZD | 03/12/2024 | 05/24/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447942 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | Upon review by abiomed's medical safety officer, patient was not a candidate or able to survive extensive surgery due to the patient¿s severe mixed septic and cardiogenic shock. ards and multiple vasopressors were in use. comfort care measures were initiated. there is not have enough information to exclude the impella as an associated factor in the patient's outcome. | The user facility reported a 76-year-old female in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed some mottling of the right lower extremities. ultrasound scan (uss) showed extensive clot burden in right femoral artery (rfa). a point of care ultrasonography (pocus) was done at bedside by the physicians and the clinician noted a free floating clot in the left ventricle. the impella was dropped to p2, and the patient was taken to the catheterization lab for impella cp removal. the skin mottling continued after removal. vascular surgeons were consulted and a computed tomography was done. however, the patient was not a surgical candidate or able to survive extensive surgery due to the patient's severe mixed septic and cardiogenic shock. comfort care measures were initiated. |  |  |  |  | The investigation into the limb ischemia ¿ irreversible and the thrombus issues has been completed since the original report was submitted. the pump was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia and the thrombus was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury"
28972 | 1220648-2024-08816 | 18986502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986502&pc=OZD | 07/21/2022 | 09/12/2024 |  | 03/27/2024 |  |  | 07/21/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic |  | Death | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | Patient had very poor prognosis. family withdrew care and medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 44-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp pump experienced positioning issues during support. several persistent alarms were noted as a result of the issue. hospital staff decided not to reposition. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08816 was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28973 | 1220648-2024-08817 | 18986523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986523&pc=OZD | 03/14/2024 | 04/18/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  | 11/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432577 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | Unintended Movement |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 87-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having surgical mitral valve repair. while on support, the impella migrated into the left ventricle with patient movement. the surgeon stepped in to reposition the pump. the impella was pulled back across the aortic valve. the valve was initially very hard to cross and both the surgeon and cardiologist were unable to push it back across the valve. the impella was replaced. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not returned for evaluation. data logs were confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
28974 | 1220648-2024-08819 | 18986577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986577&pc=OZD | 06/24/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 06/24/2022 | 03/27/2024 |  |  | 01/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 39-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while attempting to deploy the first perclose, a malfunction with the perclose device occurred. while troubleshooting the perclose device, the wire access was lost and manual pressure held by the physician. protamine test dose and subsequent full dose protamine was administered. the wire was able to be readvanced through perclose and manual pressure was held. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records
28975 | 1220648-2024-08818 | 18986579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986579&pc=OZD | 03/14/2024 | 05/24/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  | 09/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445668 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | Upon review by abiomed's medical safety officer, the patient expired on support. there is not enough information to associate the use of the device with the patient¿s outcome. | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the doctor added an additional stitch and the oozing stopped. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was determined to be use issue because the bleeding was resolved by placing extra stitch at the oozing site. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28976 | 1220648-2024-08820 | 18986598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986598&pc=OZD | 06/27/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/27/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, the patient experienced bleeding from the impella access site. patient was transfused with one unit of packed red blood cells and heparin was discontinued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28977 | 1220648-2024-08821 | 18986619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986619&pc=OZD | 03/15/2024 | 04/29/2024 |  | 03/27/2024 |  |  | 03/15/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1591091 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a battery failure alarm. there was no patient involvement. the aic was returned for evaluation. | [5 empty] | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. the console was returned and tested. the failure mode was replicated on boot up in the lab. visual inspection revealed the lcd screen of battery 2 was blank. the root cause of the battery failure was a defective pbm not being to charge one of the batteries leading to depletion and battery lockout. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
28978 | 1220648-2024-08822 | 18986660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986660&pc=OZD | 03/15/2024 | 04/03/2024 |  | 03/27/2024 |  |  | 03/15/2024 | 03/27/2024 |  |  | 11/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024430240 |  | 0048-0014 EU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the purge line was noted to be leaking and fluid was coming out. the defect was between the purge cassette and the disc. the purge cassette was replaced. the automated impella controller (aic) displayed 'low purge pressure' alarm. there was no reported patient harm | [5 empty] | Manufacturer report number 1220648-2024-08822 was submitted in error. this is not a reportable event. this supplemental report is to retract the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
28979 | 1220648-2024-08823 | 18986718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986718&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/27/2024 |  |  | 03/15/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461228 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump being delivered for support of a patient via the axillary access. the patient was having the ecmo (extracorporeal membrane oxygenation) removed and 5.5 placed for the transfer patient with history of cabg (coronary artery bypass grafting). the 5.5 failed all deliveries and kinked with the attempts. the pump was removed and a 5.5 placement was to be done via direct access. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump was unable to pass through vasculature due to resistance at axillary artery. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. added- d4-updated model number. to conform to updated procedures, sections d6 & h10 were revised with updated information. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28980 | 1220648-2024-08824 | 18986787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986787&pc=OZD | 03/16/2024 | 05/29/2024 |  | 03/27/2024 |  |  | 03/16/2024 | 03/27/2024 |  |  | 01/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024296592 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | There is not enough information to associate use of device with outcome | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 had high purge pressure and purge system blocked alarms. tpa infusing was done through purge with some improvement, overnight flows up at 2.1. the plan is to repeat tpa protocol until alarms resolve. |  |  |  |  | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. according to clinical data provided, there was a gradual rise in purge pressure and the purge line blocked alarms were seen. there was also a motor current spike seen at the start on the rise corresponding to suction alarm. the root cause of the high purge pressure issue was most likely biomaterial. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
28981 | 1220648-2024-08825 | 18986792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986792&pc=OZD | 03/12/2024 | 04/26/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025442063 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support and had high purge pressure. tissue plasminogen activator was added to the purge, and this did not resolve the issue. the doctor was concerned with how long the impella could last with high purge pressure. the patient remained on support and there was no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella was returned for evaluation. the pump was returned cut near the butterfly area. the cannula portion was not returned and only the part connecting to the aic was returned. data logs were not returned as well making it impossible to investigate this complaint. the root cause of the high purge pressure issue was not determined since insufficient product was returned and data logs were not returned. the instructions for use for impella 5.5 with smartassist states in section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ g1-corrected spelling reporting contact last name. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28982 | 1220648-2024-08826 | 18986812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986812&pc=OZD | 06/28/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/28/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212179 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced a loss of pedal pulses that were not able to be found by doppler which resulting in limb ischemia. the physician attempted to perform an external bypass and the attempt was unsuccessful. the physician made decision to remove impella cp and place a balloon pump to resolve the limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.