2511 | 1220648-2025-49318 | 23888927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888927&pc=OZD | 09/27/2025 |  |  | 12/24/2025 |  |  | 09/27/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711267 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient lost pulsatility which was resolved with albumin administration. suction also occurred, which was resolved with packed red blood cells. hemolysis was also reported, with echocardiography confirming satisfactory position. staff eventually decided to explant the pump. the patient¿s outcome is unknown. | [5 empty] | Investigation summary: hypovolemia: the cause of the injury was not determined due to insufficient clinical details. pump suction: clinical reported suction that resolved with prbc's. the product was not returned for investigation. data log review confirms suction alarms triggering. at this time, placement signal dropped and started to trend down. placement signal then started to slowly increase again, which was likely due to albumin and prbcs being given. the cause of the suction was most likely related to the patients condition, as placement signal was trending down prior to volume administration, and clinical reported the patient was in poor conditions. hemolysis: clinical reported hemolysis occurring which led to decision to remove the pump. the pump was in proper positioning. the product was not returned for investigation. data log review shows a downward trend in placement signal and clinical stated the patient was in poor condition. there were no positioning issues or motor current spikes observed. the cause of the hemolysis was most likely related to the patients condition, as placement signal was trending down prior to volume administration, and clinical reported the patient was in poor conditions. device history lot: device lot number: 1954317 device history batch: subcomponent lot: n/a device history review: pump sn (b)(6) passed all post sterile inspection criteria.
2512 | 1220648-2025-49324 | 23888939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888939&pc=OZD | 10/09/2025 | 02/17/2026 |  | 12/24/2025 |  |  | 09/30/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724440 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support the device remained at p1 and was unable to flow higher without suction, the resolution for this issue is unknown. it was also reported that the patient experienced continous ventricular fibrillation, the suction alarm was disabled due to persistant suction at p2. the patient improved after suction was changed from p1 to p2. it was later reported that care was withdrawn and the patient died. | [5 empty] | Investigation summary: clinical symptom (hemodynamic instability): the cause of the injury was not determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. according to the clinical details, the patient was on ecpella support from the start of the case, and crrt was initiated the following day and continued until pump explant. on the first day of pump use, data log shows the pump was started at a higher p-level, but the patient was unable to tolerate it, with a gradual decreasing trend in pump flow and fluctuating placement signal. the p-level was reduced, resulting in some resolution of pump suction. later, improvement was noted in pump flow while running at a lower p-level, and the patient was able to tolerate higher p-levels without suction at the end of the case run when ecmo was explanted. no abnormalities were reported regarding motor current spikes or optical signal issues. the cause of the pump suction issue was most likely related to patient condition. device history lot: device batch: 1956262. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks. | Corrected info: the initial report stated in b5 that the p level was increased from p1 to p2. however, additional information provided states that the issue resolved by reducing p levels. d4 catalog number updated.
2513 | 1220648-2025-49323 | 23888940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888940&pc=OZD | 09/29/2025 |  |  | 12/24/2025 |  |  | 09/29/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738728 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. sepsis has been reported. after a high risk percutaneous coronary intervention (hrpci), the patient began to decompensate hemodynamically. the patient had developed a retroperitoneal bleed from the hrpci. the patient was administered multiple blood products and went to the operating room for repair. the patient stabilized. the pump was eventually successfully weaned and explanted, and the patient survived the event. | [5 empty] | Investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. sepsis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot : 1970179 device history batch: sub-component lot : n/a device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
2514 | 1220648-2025-49326 | 23888943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888943&pc=OZD | 10/01/2025 |  |  | 12/24/2025 |  |  | 10/01/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594240 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the nurse accidently sat the patient up in bed and this caused the impella to become dislodged. the device was replaced with another impella cp. it was reported that the patient was stable. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump became dislodged when the rn accidentally sat pt up in bed. device history lot: device lot: 1870336. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2515 | 1220648-2025-49327 | 23888944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888944&pc=OZD | 10/01/2025 |  |  | 12/24/2025 |  |  | 10/01/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627535 |  | 000000000001000361 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that device was malpositioned and it was decided to abort treatment. no further information was provided. | [5 empty] | Investigation summary ==> no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1895029 device history batch ==> subcomponent lot: . device history review ==> device (sn: (b)(6) passed all post sterile inspection checks.
2516 | 1220648-2025-49328 | 23888949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23888949&pc=OZD | 10/10/2025 | 11/12/2025 |  | 12/24/2025 |  |  | 10/10/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757665 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | Valvular Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient underwent a transcatheter aortic valve replacement (tvar) procedure on (b)(6) 2025, which was immediately complicated by a st-elevation myocardial infarction due to a left anterior descending (lad) artery occlusion. the patient received lad stenting and was cannulated for extracorporeal membrane oxygenation (ecmo) support, then successfully decannulated on (b)(6) 2025. on (b)(6) 2025, the patient experienced pulseless electrical activity arrest and was recannulated for ecmo and then brought to the cardiac catheterization laboratory for impella cp placement for left ventricular unloading. on day 2 of impella cp and emco support, the patient was given volume, a unit of blood, and albumin. it was noted that the patient had hematuria, however the patient¿s bladder had a known bleed and there was no concern from the physician for hemolysis. the following day it was noted that the patient continued to bleed and required blood products. on day 4 of impella cp and emco support left heart catheterization was performed and found a severe 70% mid-distal lad lesion, and two severe blockages in the right coronary artery: a 90% blockage and a 50% blockage in the proximal and mid sections. it was noted that the previously placed stents in the left main (lm) to lad and lm to left circumflex (lcx) arteries were patent. there was no immediate interventions planned for the patient. on day 5 of impella cp and emco support a transesophageal echocardiogram (tee) was to assess for thrombus on the tavr valve leaflets before considering an escalation to the impella 5.5. the tee revealed no thrombus on the aortic valve leaflets but indicated poor intrinsic heart function and restricted leaflet opening. a thrombus was noted in the left atrial appendage (laa). intermittent impella suction alarms occurred when the patient coughed. the decision was made to escalate the patient to an impella 5.5 for advanced support. the impella 5.5 was successfully placed; the impella cp was weaned and explanted. the patient remained on ecmo support with the impella 5.5. of note, 3 units of packed red blood cells were transfused to the patient during the implantation due to blood loss from the right axillary arterial line which was removed prior to the case. on support day 1 2 of impella 5.5 and ecmo, large left pleural effusion was identified with a chest x-ray. a pigtail catheter was placed for drainage. the patient required emergency reintubation this morning in the intensive care unit (icu) due to acute respiratory distress and hypoxia, occurring prior to a planned procedure in the operating room. there were reported impella suction alarms, which the medical team suspected to be related to hypovolemia. the impella performance was consequently reduced to p4 with flows of 1.6 l/min to because of the alarms. on support day 2 of impella 5.5 and ecmo, the patient¿s arterial blood gas results indicated the patient had hypoxia and hypercarbia upon admission from the operating room. ecmo flow rates were increased and the impella p level was increased to p3. overnight, the patient developed unequal pupils, prompting a head computed tomography (ct) and the temporary holding of anticoagulation therapy. the ct revealed no acute changes but noted findings consistent with "cortical laminar necrosis." this morning, the patient is stable and following commands in all extremities. the patient continued continuous renal replacement therapy (crrt). on support day 9 of impella 5.5 and ecmo, the patient experienced an acute failure of the ecmo system when the oxygenator clotted off. during the subsequent procedure to exchange the oxygenator and the entire circuit, the impella heart pump performance level was increased to p8. shortly after the increase, the impella experienced suction events and low flow. an echocardiogram performed at the bedside showed the left ventricle to be completely decompressed with no volume, and failed ramp tests suggested acute right ventricular failure. the patient was successfully placed back onto the ecmo circuit, achieving stable flow at 5 l/min, which stabilized the patient's condition. post-stabilization, the impella was stable at performance level p3. there were no immediate concerns or questions reported regarding the impella device function itself after the circuit exchange was completed. on support day 16 of impella 5.5 and ecmo, the patient¿s family decided to withdraw the care of the patient. the patient¿s care was withdrawn and the patient expired. this complaint involves two automated impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 1: impella cp, with serial number 628224, which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event and noted the following: the patient experienced pulseless electrical activity (pea) and was placed on extracorporeal membrane oxygenation (ecmo) support. impella cp was implanted for left ventricular unloading. during support, hematuria was observed, attributed to a pre-existing bladder bleed, and one unit of blood was administered. the patient was subsequently escalated to impella 5.5 with continued ecmo support. during impella 5.5 implantation, significant bleeding occurred at the axillary access site, requiring additional blood products. the patient later developed acute respiratory distress secondary to a left pleural effusion, which was drained. after 16 days of impella support, the patient¿s family elected to withdraw care, and the patient expired. adverse events impella cp: bleeding event (hematuria requiring transfusion); of note, a thrombus was noted in the left atrial appendage (laa). intermittent impella suction alarms occurred when the patient coughed. thrombosis is a high risk with ecmo use. impella 5.5: bleeding event (axillary site bleeding requiring transfusion) and hypoxia/respiratory failure (pleural effusion). the death is conservatively reported on impella 5.5; unlikely device-related, attributed to patient¿s underlying acuity. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, thrombosis, in order to make a cause determination, all of the clinical details would have to be provided. the cause of both the thrombosis and heart valve incompetence was not determined due to insufficient clinical details. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
2517 | 1220648-2025-49329 | 23889205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889205&pc=OZD | 11/28/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 11/28/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742309 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that they were having ¿placement signal issues¿. an imagine was sent and placement signal was gone but pump was functioning correctly. the placement signal alarms were disabled and instructed staff to monitor motor current closely. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product or data logs were returned. the clinical details state that there were placement signal issues, but the pump was functioning appropriately. due to a lack of product and data logs returned, the root cause of the optical sensor issue could not be established. device history review the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2518 | 1220648-2025-49331 | 23889499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889499&pc=OZD | 11/14/2025 |  |  | 12/24/2025 |  |  | 11/14/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768574 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Vascular Dissection | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the impella cp was prepped, inserted, and support was initiated without complications. following the hrpci, the impella was successfully weaned and removed in the procedural area. it was reported that the physician encountered complications with a perclose closure device following explant to close the access site. an additional perclose device and angioseal were deployed. hemostasis was achieved, however the patient lost distal flows after closure. the physician believed the perclose device may have attached to the anterior of the wall of the vessel and closed off the flow. vascular surgery was consulted to repair vessel. the patient's outcome at the time of explant was survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical symptoms (ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. peri-procedural adverse event (major bleed, vascular dissection): the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 code: a01 incorrectly submitted in initial.
2519 | 1220648-2025-49330 | 23889501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889501&pc=OZD | 11/13/2025 |  |  | 12/24/2025 |  |  | 11/13/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026771258 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Backflow | Restricted Flow rate | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Renal Impairment | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction (stemi)/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the hospital on (b)(6) 2025 following a witnessed cardiac arrest with 40 minutes of total resuscitation before achieving return of spontaneous circulation. the patient was intubated by emergency medical services and was found to be in ventricular tachycardia upon their arrival. upon admission, a 12-lead electrocardiogram revealed an inferior lateral stemi, and subsequent cardiac catheterization confirmed multivessel disease with a severely reduced ejection fraction (ef) of 15¿20%. although the cardiothoracic surgery team declined surgical intervention, the patient was stabilized with an intra-aortic balloon pump (iabp) at a 1:2 ratio and multiple vasopressors. neurologically, the patient had intact brainstem reflexes but remained non-responsive to commands, prompting the initiation of targeted temperature management (ttm), which reached its goal the following day. it was noted that the patient had an elevated lactate of 9.3, unsuccessful bedside swan-ganz catheter placement, and continued mechanical ventilation as the team prepared for rewarming the next day. the decision was made to implant an impella cp and the patient was transferred to the cardiac catheterization laboratory. an impella cp was prepped and inserted via the right axillary artery without complications. after support was initiated, the patient¿s iabp was placed on standby to assess if continued support was needed. the patient became increasingly less pulsatile with the impella cp at performance (p) level 5 with mean arterial pressures rapidly declining. the iabp was placed back at 1:1 and pressors were restarted. there were frequent suction alarms and retrograde flow alarms from the impella cp on the automated impella controller. the iabp was adjusted to 1:2 and the patient was transferred to the intensive care unit (icu) for continued management. after arriving to the icu, it was reported that there were continued suction alarms from the impella. the iabp was temporarily turned down to 1:4 and augmentation to 66% which resulted in decreased suction alarms and increased flows on the impella. oozing was noted at the impella insertion site and the physician was notified and discussed tightening the perclose sutures, however no intervention was performed at that time. on support day 2, the suction alarms were reported to have continued overnight and additionally, there were placement signal low alarms. an echocardiogram was performed to confirm impella placement. the patient¿s hemoglobin dropped to 7.4 g/dl overnight and the patient was transfused with 2 units of packed red blood cells. it was noted that when the patient¿s sedation was paused the patient was able to follow commands and the patient was being rewarmed following hospital protocol. the patient¿s fibrinogen was measured to be in the 130s md/dl and the patient was to be transfused with a unit of cryoprecipitate and a plan was made to restart heparin therapy. on support day 4, the patient was transfused with one unit of platelets, 1 unit of packed red blood cells, and 1.5 liters of albumin. on support day 5, it was noted that the impella access site was oozing. the physician tightened perclose sutures and applied pressure dressing. the patient was placed on continuous renal replacement therapy to help the patient¿s kidneys recover. on support day 7, the patient was weaned down to performance level 2. it was noted that the patient received an additional two units of packed red blood cells. the patient was brought to the cardiac catheterization laboratory for a high-risk percutaneous coronary intervention (pci). following a successful pci, the impella was weaned and explanted. following explant, the impella access site at the axillary artery continued to ooze. multiple attempts with balloon inflations did not achieve hemostasis. a covered stent was placed, and pressure dressing was applied to achieve hemostasis. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data log was not provided or could not be retrieved from the remote server. the cause of the injuries, low blood pressure/hypotension and acute kidney failure could not be determined due to insufficient clinical details. the root cause of pump suction, retrograde pump flow and optical sensor issues could not be determined without returned product investigation and sufficient clinical information, including data logs. for the access site hemorrhage/bleeding, clinical details indicate multiple instances of oozing and bleeding at the axillary artery access site throughout support. initial oozing was noted on (b)(6) 2025, with further episodes requiring transfusion of blood products and repeated clinical interventions. perclose sutures were tightened and pressure dressings applied as bleeding recurred. on (b)(6) 2025, persistent oozing during device removal required multiple balloon inflations and ultimately placement of a covered stent, after which the bleeding was controlled. while bleeding appeared to be managed with these interventions, it is not clear from the available details whether the bleeding was fully resolved or if there were any lasting complications. the cause of the access site adverse event issue was not determined without returned product investigation and sufficient clinical details.
2520 | 1220648-2025-49334 | 23889523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889523&pc=OZD | 08/08/2025 |  |  | 12/24/2025 |  |  | 11/18/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698585 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 95 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry noted a patient experienced a left femoral artery access site hematoma. it was stated that the left groin had a small hematoma but was stable, right groin no hematoma. the relationship was listed as possibly related to impella use. patient outcome was noted to be stable. additional information was not available in the complaint record accessible for mdr assessment at time of reporting. | [5 empty] | A5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> no product returned. asae/hematoma: loqi reported that left femoral artery access site hematoma. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot ==> device lot: 1943229 device history review ==> this pump (b)(6) passed all post sterile inspection checks
2521 | 1220648-2025-49335 | 23889525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889525&pc=OZD | 08/26/2025 |  |  | 12/24/2025 |  |  | 11/18/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026725503 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry noted a patient experienced compromised blood flow to the distal left leg. it was noted that the impella device was implanted in the left common femoral artery and as a result of the limb ischemia the device was removed. an intra-aortic balloon pump was placed in the patient¿s right common femoral artery for support after the impella was removed. the relationship between the impella and the lower limb ischemia was listed as probably. the patient's outcome at the time of explant was survived. additional information was not available in the complaint record accessible for mdr assessment at time of reporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2522 | 1220648-2025-49336 | 23889534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889534&pc=OZD | 11/18/2025 |  |  | 12/24/2025 |  |  | 11/18/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705978 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient, with no known prior medical history, experienced a witnessed out-of-hospital cardiac arrest at a 7-eleven establishment with a reported downtime of 15 minutes before emergency medical services arrived and initiated cardiopulmonary resuscitation (cpr). the patient presented to the emergency department with a lactate level of 8.0 mmol/l. a repeat electrocardiogram confirmed a stemi (st-elevation myocardial infarction). the patient was immediately transferred to the cardiac catheterization laboratory (cl) for emergent intervention. a subsequent left heart catheterization (lhc) revealed a chronic total occlusion of the right coronary artery (rca), a complex trifurcating lesion involving the left main (lm), left anterior descending (lad), and left circumflex (lcx) arteries, and an elevated left ventricular end-diastolic pressure (edp) of 26 mmhg. the decision was made to place an impella cp prior to a percutaneous coronary intervention (pci). an impella was prepped, inserted via left femoral artery and support was initiated without complications. following a successful pci, the physician decided to leave the patient on impella support overnight for rest and recovery. the patient was sent to the unit. later that day, it was reported that the patient¿s laboratory samples were hemolyzed due to poor position of the impella. an echocardiogram was performed, and the impella was repositioned and weaned to performance level 2. the physician stated that the goal was to wean and explant the pump. the device was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary. product and data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review. pump passed all post sterile inspection checks.
2523 | 1220648-2025-49337 | 23889539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889539&pc=OZD | 11/18/2025 |  |  | 12/24/2025 |  |  | 11/18/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651078 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female | 101 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that there was a patient listed as status 3 for transplant and planned for axillary impella 5.5 support. it was reported that the purge door stuck of the automated impella controller and was unable to open. swapped to back up console. there was no patient harm. | [5 empty] | After evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally. a regulatory report is no longer necessary. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2524 | 1220648-2025-49339 | 23889544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889544&pc=OZD | 11/19/2025 |  |  | 12/24/2025 |  |  | 11/19/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755665 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Cardiac Arrest | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp. while on support the it was reported that there was multiple rounds cardiopulmonary resuscitation, prolonged arrest periods. it was also reported later that the patient had atrial fibrillation with rapid ventricular response. later the patient was brought to cath lab in relatively stable condition. remains ventilated. device weaned explanted in standard fashion, remained hemodynamically stable. | Additional information was received that reported the patient experienced an atrial fibrillation (afib) and cardiac arrest prior to implantation of the impella. while on support, the patient had a reoccurrence of afib spontaneously. an echocardiogram verified the afib was unrelated to the positioning of the impella. the device was explanted and disposed of. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information added to b5, d9 and h6. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the cardiac arrest, paroxysmal and atrial fibrillation, could not be determined due to insufficient clinical details.
2525 | 1220648-2025-49340 | 23889547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889547&pc=OZD | 11/20/2025 | 12/19/2025 |  | 12/24/2025 |  |  | 11/20/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764033 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female | 50 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 85 year old female patient on impella cp for support during a high risk cardiac procedure. during support, the patient developed a gastrointestinal bleed in both upper and lower tracts. a decision was made to remove impella so that the patient could be discontinued from heparin. the patient's outcome is unknown. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2526 | 1220648-2025-49342 | 23889559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889559&pc=OZD | 11/21/2025 |  |  | 12/24/2025 |  |  | 11/21/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026774814 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 125 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp on a 75 year old male for support for cardiomyopathy. it was reported that during support, patient has frequent runs of premature ventricular contractions (pvc) after weaning off of levo. there was no additional information provided for treatment. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. there are three associated devices for the reported event. the first insertion kit is addressed under manufacturer report number xxxx. the second insertion kit is addressed under manufacturer report number xxxx. the pump (impella cp) is addressed under manufacturer report number xxxx. | Correction: e4, h11. did not include related mfr numbers in the initial report. the investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. there are three associated devices for the reported event. the first insertion kit is addressed under manufacturer report number 1220648-2025-49341. the second insertion kit is addressed under manufacturer report number 1220648-2025-49343. the pump (impella cp) is addressed under manufacturer report number 1220648-2025-49342. device history review: the device passed all post sterile inspection checks.
2527 | 1220648-2025-49344 | 23889561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889561&pc=OZD | 11/21/2025 |  |  | 12/24/2025 |  |  | 11/21/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026694305 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Ischemia | Hematuria | [39 empty] | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. one day after implant there was placement signal alarm with notable minimal dark red urine. continuous veno-venous hemofiltration initiated. overnight impella leg (rle) was warm and well perfused per team however the next day morning the leg is notably colder than lle and the r foot was turning blue. patient was in critical condition so family decided to withdrew care and patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | Complaint coding was updated to better reflect the reported event. consequently, section h6 health effect clinical and medical device problem codes were updated/corrected and repopulated accordingly. | Correction: b1. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood (hemolysis) / placement signal issue: the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data log review. device history review: the device passed all post sterile inspection checks.
2528 | 1220648-2025-49345 | 23889571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889571&pc=OZD | 12/03/2025 |  |  | 12/24/2025 |  |  | 12/04/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1133144 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller showed "failure system, please switch to back up controller". there was no patient involved. | [5 empty] | E1 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2529 | 1220648-2025-49348 | 23889577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889577&pc=OZD | 12/06/2025 |  |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701718 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | False Alarm | Inaccurate Flow Rate | Increase in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella 5.5 experienced saturated flow with high purge pressure. tissue plasminogen activator was used to resolve the issue. impella support continues. | Additional information was received indicating that the impella device was explanted. the patient is stable. |  |  |  |  | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no product was returned. no logs available. false alarm/saturated flows/high purge pressure: the cause of the failure events was unable to determined since no data logs or product was returned for review. device history review completed and passed all post sterile inspection checks. | H6: added code a160105 based on information in the complaint file. | Additional information received for d6 explant.
2530 | 1220648-2025-49349 | 23889581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889581&pc=OZD | 12/07/2025 |  |  | 12/24/2025 |  |  | 12/07/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026776001 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 57 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation. the patient was treated with amio and gtt to resolve the issue. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. tachycardia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2531 | 1220648-2025-49350 | 23889582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889582&pc=OZD | 12/07/2025 |  |  | 12/24/2025 |  |  | 12/07/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686409 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding and hematuria. the patient was transfused three units of packed red blood cells, and the p level of the pump was decreased in response. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. data log was not provided or was not found on the remote server. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood (hemolysis): the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data logs. clinical symptoms (hemorrhage/bleeding): the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. h6 code a24 was incorrectly reported.
2532 | 1220648-2025-49353 | 23889584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889584&pc=OZD | 12/08/2025 |  |  | 12/24/2025 |  |  | 12/08/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1658143 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was a controller error on the automated impella controller once staff turned the device on. there was no patient injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 and h6 updated. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the console was sent back for further investigation. the console ran a known-good pump and purge without reproducing the momentary loss of placement signal and controller error alarm. the root cause for the momentary loss of placement signal and controller error alarm was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic operated as designed.
2533 | 1220648-2025-49352 | 23889585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889585&pc=OZD | 12/08/2025 |  |  | 12/24/2025 |  |  | 12/08/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559450 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma and vascular damage after the sheath was exchanged. the impella was explanted. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. asae/hematoma: the cause of the access site adverse event was use related as the hematoma occurred after sheath exchange and the first perclose had failed. device/patient interaction: the cause of the device/patient interaction was use related as the vascular damage occurred during the sheath exchange and the first perclose had failed and made the vessel tight per clinical. device history review: the device passed all post sterile inspection checks.
2534 | 1220648-2025-49351 | 23889588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889588&pc=OZD | 12/07/2025 |  |  | 12/24/2025 |  |  | 12/07/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026775502 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia which was treated with an amio bolus. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. (tachycardia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2535 | 1220648-2025-49354 | 23889591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889591&pc=OZD | 12/09/2025 |  |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710389 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the pump was repositioned and the patient experienced hemolysis. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood (hemolysis): data log review was not provided or retrieved from the remote server. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details, including data log review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2536 | 1220648-2025-49355 | 23889592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889592&pc=OZD | 12/09/2025 | 01/13/2026 |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765731 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the groin site. a femstop was placed to contain the hematoma. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary no product returned asae/hematoma: hematoma noted at left groin site. with the femstop device hematoma quit expending. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot device lot- 1979787 device history batch subcomponent lot- n/a device history review the pump (b)(6) passed all the post sterile inspection checks.
2537 | 1220648-2025-49356 | 23889614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889614&pc=OZD | 12/09/2025 |  |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757671 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 110 KG |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a drift of the placement signal. the signal did not correlate to the arterial line. the sensor was manually recalibrated. no patient harm was reported. | [5 empty] | Additional information was provided therefore d6b was updated. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. no data logs. placement signal issue: the cause of the placement signal issue is not determined due to lack of product and data logs available. device history review: device passed all post sterile inspection checks.
2538 | 1220648-2025-49357 | 23889615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889615&pc=OZD | 12/09/2025 | 01/28/2026 |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771321 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella 5.5 experienced an infection from a pacemaker. the device was explanted and a temporary pacemaker was placed. the patient was treated for sepsis. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was discarded. | Additional information was received therefore d6b was updated | The investigation was completed. (sepsis ) : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2539 | 1220648-2025-49359 | 23889618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889618&pc=OZD | 12/09/2025 |  |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026779169 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and hemolysis. the impella was positioned shallow potentially the inlet was contacting the mitral valve. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: pump passed all post sterile inspection checks.
2540 | 1220648-2025-49360 | 23889620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889620&pc=OZD | 12/10/2025 | 12/29/2025 |  | 12/24/2025 |  |  | 12/11/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1733097 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when an automated impella controller was powered on, an error appeared. there was no patient involvement. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
2541 | 1220648-2025-49361 | 23889624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889624&pc=OZD | 12/09/2025 |  |  | 12/24/2025 |  |  | 12/09/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026770436 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient presented to the hospital on (b)(6) 2025, with intermittent chest pain and was diagnosed with an nstemi. following a left heart catheterization (lhc) that revealed multivessel disease and an echocardiogram which showed severe mitral regurgitation, the patient was scheduled for a triple coronary artery bypass graft and mitral valve replacement. the procedure was not well tolerated by the patient, and upon sedation the patient became hypotensive. additionally, the physician noted that the patient¿s left ventricle was hypokinetic and made the decision to place an impella 5.5 to provide necessary hemodynamic support. the impella 5.5 was placed successfully via the axillary artery, exact location was not provided, with no reported complications. following a successful procedure, the patient was transported to the surgical intensive care unit (sicu) for rest and recovery with impella support. upon arrival to the sicu, the automated impella controller waveforms were noted to be separating, and intermittent suction alarms began to occur. a bedside echocardiogram was performed and revealed that the impella was at an appropriate depth of 4.9 cm but was noted to be close to the septal wall. the physician pulled the device back approximately 0.5 cm. the waveforms improved, and the performance (p) level was able to be increased back to p7, and the pump flows were as expected at 4 l/min. on impella support day 2, there was noted concern that the patient needed to return to the operating room for excess chest tube output, however the output decreased prior to the decision to bring the patient to the operating room. the patient received a unit of packed red blood cells due the blood loss. of note, there were no impella concerns or problems. on impella support day 3, the patient was noted to be experiencing hematuria. it was noted that the patient¿s mean arterial pressure increased to 90-100 mmhg the previous day and there was an increased afterload prior to the hematuria. subsequently, the p level was increased to p7 and the patient was given a bolus of intravenous fluids. additionally, it was noted that there were intermittent suction events the previous day. the impellla position was reviewed with and echocardiogram, and the impella was determined to be in an appropriate position. as a result of the hematuria, the p level was decreased to p5. as a result of the increased afterload, the patient was given cardene. the urine was noted to be slightly pink tinged but clearing the following day. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: ppae (hematuria/major bleed): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2542 | 1220648-2025-49362 | 23889626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889626&pc=OZD | 12/11/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 12/11/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026787424 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64 year old male was implanted with a impella cp for support during a high risk cardiac procedure. a small hematoma was reported at site after peel away removed and repo advanced. resolution with fem stop. no blood products were required. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product returned. asae/hematoma: hematoma was noted after peel way was removed. using femstop it resolved. the cause of the access site adverse event was not determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks.
2543 | 1220648-2025-49365 | 23889627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889627&pc=OZD | 10/03/2025 | 02/12/2026 |  | 12/24/2025 |  |  | 12/11/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726411 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was being implanted with an impella cp for, mechanical circulatory support. during the insertion of the device, it was reported that the device would not advance into the left ventricle. it was reported that it was unclear if the device had an issue that contributed. it was noted that the procedure was aborted due to a massive aortic valve stenosis. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device is reportedly available for return; however, as of the time of this report the device has not yet been returned for evaluation. | [5 empty] | Section a patient information has been left blank, as patient demographics are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2544 | 1220648-2025-49364 | 23889629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889629&pc=OZD | 12/11/2025 | 01/09/2026 |  | 12/24/2025 |  |  | 12/11/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760961 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female | 68 KG | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp via the left femoral artery for mechanical circulatory support. on the first day of impella support, it was reported that the patient did not have flow distal to the patient¿s left foot. the patient was taken to the cardiac catheterization laboratory for a femoral-to-femoral bypass, which successfully restored flow to the patient¿s left foot. additionally, it was noted that the patient expired during impella support. it was noted that the cause of death was unknown, and the physician did not believe the death was impella related. the physician stated he believed the cause of death was due to other issues and that the family chose to withdraw care of the patient. additional clinical details, including perioperative course and contributing factors, were unavailable at the time of reporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: ppae (ischemia): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot- 1980833. device history batch subcomponent lot- n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2545 | 1220648-2025-49367 | 23889632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889632&pc=OZD | 12/11/2025 |  |  | 12/24/2025 |  |  | 12/11/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026765575 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 117 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp via the right femoral artery for mechanical circulatory support. it was reported that minutes after implanting the impella cp, there was an alarm that stated placement signal not reliable. it was reported that there were no issues with the pump flow during the surgical case. it was noted that the alarm self-resolved and there were no further issues with the placement waveform. the pump position was confirmed with echocardiography. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2546 | 1220648-2025-49366 | 23889636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889636&pc=OZD | 12/11/2025 |  |  | 12/24/2025 |  |  | 12/11/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678640 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a nurse mentioned that a patient had hematuria with red urine. the impella position was verified and impella p-level dropped from p8 to p6. waveforms were appropriate at this time. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
2547 | 1220648-2025-49368 | 23889637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889637&pc=OZD | 12/12/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026768692 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Care was withdrawn while the patient was till on venous-arterial extracorporeal membrane oxygenation. the patient was declared "brain-dead" due to not awakening after 2-3 days off of sedation/pain medication, no stimuli, and fixed and dilated pupils. | The user facility reported a 36 year old male was implanted with a impella cp for support during high risk cardiac procedure. it was reported that the patient came in cardiogenic shock. the patient had some ectopic beats where cath lab tech noted pulseless electrical activity where cardiopulmonary resuscitation was started for approximately 5-10 minutes. patient was given two units of blood. once patient arrived in intensive care unit, there was bleeding noted in the right femoral artery (place of insertion for impella cp) and in the left femoral artery and vein at the extracorporeal membrane oxygenation drainage and return cannulas insertions. |  |  |  |  | B5 updated. the hospital disposed of the pump. should any new information be received, a supplemental mdr will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product returned. arrhythmia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. bleeding noted at cp and extracorporeal membrane oxygenation (ecmo) femoral access sites during support. unknown how bleeding was resolved. the cause of the access site adverse event was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2548 | 1220648-2025-49370 | 23889640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889640&pc=OZD | 12/12/2025 | 02/12/2026 |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738834 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 70 KG | Hispanic |  | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the patient was treated with lasix which did not resolve the issue. the patient expired on impella support. | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. no product or data logs were available. the cause of the hemolysis was not determined due to lack of product, data logs, and sufficient clinical information.
2549 | 1220648-2025-49369 | 23889641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889641&pc=OZD | 12/12/2025 | 02/12/2026 |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026746219 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 34 year male was implanted with a impella 5.5 for a high risk cardiac procedure. it was reported that the patients purge pressure had increased over 1000 and purge flow less than 1. tpa was started, as purge fluid and they were notifying us that they were changing back to bicarbonate solution since purge pressure had normalized and purge flow back to 5ml and purge pressure 480. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b updated. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. the cause of the high purge pressure was not determined due to lack of product and data logs returned.
2550 | 1220648-2025-49371 | 23889645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889645&pc=OZD | 12/12/2025 | 12/30/2025 |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65 year old male was implanted with a impella cp for support for a high risk cardiac procedure. it was reported that the decision was made to leave pump in. a echocardiogram was performed at bedside and pump needed to be repositioned. while repositioning sleeve covering catheter ripped. the doctor wanted to just secure rip tuohy and t lock secured. device locked at 91cm and echo measurement 4.08 cm after repositioning. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Further details were provided stating that the sleeve was patched up with tegaderm.
2551 | 1220648-2025-49373 | 23889651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889651&pc=OZD | 10/01/2025 | 12/24/2025 |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619794 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient had a prior wound healing disorder with proven klebsiella oxcytoca. medical intervention as well as any medications given are unknown. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product return. infection / post-procedural adverse event: clinical details note a wound healing disorder with proven klebsiella oxcytoca. no product was returned and logs are not applicable. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot : 1893944. device history batch: subcomponent lot : n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2552 | 1220648-2025-49374 | 23889658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889658&pc=PYX | 12/12/2025 |  |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1699285 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 78 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella pump experienced an automated impella controller failure. the controller was replaced with another to continue patient support. no patient harm was reported. | This is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the impella rp flex. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 added information that was omitted from the initial report that was submitted. d2b updated procode as it was previously entered incorrectly on the initial report that was submitted. d9 the device was returned and the date was added. g4 revised pma/510(k) as it was entered incorrectly on the initial report that was submitted. h6 updated codes to reflect that the device was returned. h10 related report number was added as it was omitted from the initial report that was reported. h11 updated additional manufacturer narrative to reflect that the device was returned. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2553 | 1220648-2025-49375 | 23889662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889662&pc=OZD | 12/12/2025 | 02/26/2026 |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026781924 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 43 year old male was implanted with a impella 5.5 for support during a high risk cardiac procedure. it was reported that the patient had some suction and impella position unknown alarms this morning. the team did a point-of-care ultrasound which showed impella slightly deep. they pulled back 2 cm with new measurement at hub of 42cm and was previously 44cm. no atrioventricular to inflow measurement recorded yet. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Explant date has been updated. h6 health effect - clinical code was updated from e2321 to e2403.
2554 | 1220648-2025-49376 | 23889666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889666&pc=OZD | 12/12/2025 |  |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611365 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced heparin induced thrombocytopenia. argatroban was given instead of heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information: the customer reported that there was no other harm noted, patient was switched to argatroban and no bleeding events. product is not available for return. the patient was alive at explant. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was returned for evaluation. ppae (thrombocytopenia): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2555 | 1220648-2025-49377 | 23889667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23889667&pc=OZD | 12/12/2025 |  |  | 12/24/2025 |  |  | 12/12/2025 | 12/24/2025 |  |  |  |  | 12/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026789259 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 109 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced positioning alarms and hematuria. the pump was repositioned and eventually replaced. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 erroneously entered on the initial manufacturer device report number therefore updated to the correct code on this report. the investigation was completed. the pump was not returned for investigation. data log was not returned for investigation. device in wrong position: the pump was not returned for investigation. clinical details indicate that while removing the peel-away sheath, the device was advanced too far into the ventricle. the device was then pulled back to an angiographically appropriate position. later in the icu, echo revealed the impella at 6 cm. cardiology performed a 3 cm pullback, with appropriate waveforms and the touhy tightened. an "impella in aorta" alarm activated after a blood pressure spike; attempt to advance the device at the bedside with echo was unsuccessful. the patient was taken back to the cath lab and the impella was replaced. the cause of the device positioning issue was not determined without returned product investigation and sufficient clinical details. clinical symptom (hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
2556 | 1220648-2025-49230 | 23876269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23876269&pc=PYX | 11/24/2025 |  |  | 12/23/2025 |  |  | 11/24/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026713549 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a physician encountered two types of problems with an impella rp. immediately after insertion, negative curves on the positioning signal occurred which normalized after about 30-40 minutes and subsequently during support. they have repeatedly reset the flow to restore and reactivate the device, but without any particular resolution. they also noted a low purge fluid flow rate of 6-7 milliliter per hour which has always remained the same. they vented the system but the flow rate never increased. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data logs returned. placement signal (ps) issue: data logs confirm the narrative that there are negative ps values at the start of support for 59hours. the ps is widespread throughout the entirety of support. the cause of the placement signal issue was not determined due to lack clinical information and no product returned for investigation. device history review: the device passed all post sterile inspection checks.
2557 | 1220648-2025-49231 | 23876319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23876319&pc=OZD | 11/24/2025 |  |  | 12/23/2025 |  |  | 11/24/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762938 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the patient was in the intensive care unit while on impella support and had ventricular tachycardia (vt). it was further reported that the patient expired because of overall bad situation. the patients death was not related to the impella. additionally, it was noted that the patient had preexisting vt prior to impella use. there was no reported product malfunction with this reported event. | [5 empty] | Medical safety/clinical reviewed the event. an 84-year-old male with a history of coronary artery disease presented in cardiogenic shock (scai stage d) secondary to acute myocardial infarction. the patient was reported to have experienced ventricular tachycardia (vt) prior to impella implantation. in an effort to stabilize the patient, an impella cp was implanted for mechanical circulatory support. following implantation, the patient appeared to stabilize and was transferred to the unit. during the subsequent course of support, the patient experienced additional episode of ventricular tachycardia. after approximately 36 hours of impella cp support, care was withdrawn. the patient expired. given the limited information available, the post-implant ventricular tachycardia cannot be definitively dissociated from impella cp support, although vt was documented prior to device implantation. the impella cp will be reported as a death-type reportable event. however, the patient's underlying condition (acute myocardial infarction complicated by cardiogenic shock scai stage d and recurrent ventricular arrhythmias) appears to have contributed most to the demise outcome. note: h6. component code and investigation type/findings/conclusion coding remain unchanged as previously reported. | Investigation summary: the investigation has been completed. peri-procedural adverse event (ventricular tachycardia): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2558 | 1220648-2025-49232 | 23876422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23876422&pc=OZD | 11/12/2025 |  |  | 12/23/2025 |  |  | 11/24/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026750040 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was escalated from an impella cp to an impella 5.5 device for mechanical circulatory support. it was reported that the patient got hypotensive while switching from automated impella controller (aic) to aic on the impella 5.5. the patient was noted as stable and continued on support. the pump is not available for return. this is one of two related reports. this report addresses the impella 5.5 and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. | [5 empty] | The impella device is not available for return from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not available for evaluation. ppae: (hypotension): the cause of the injury was not determined due to insufficient clinical details. device history review: this pump passed all post sterile inspection checks.
2559 | 1220648-2025-49233 | 23877479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23877479&pc=OZD | 11/26/2025 | 02/27/2026 |  | 12/23/2025 |  |  | 11/26/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1839845 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Operating System Becomes Nonfunctional | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during therapy, the screen froze and than shut down with incomplete restart (hanging in bios). there was a successful quick automated impella controller change. the pump runs well on the new aic. no further information is possible. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2560 | 1220648-2025-49234 | 23877641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23877641&pc=OZD | 11/25/2025 | 12/23/2025 |  | 12/23/2025 |  |  | 11/27/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766400 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that upon implantation of an impella cp the pump was not recognized by the supporting automated impella controller. a new pump was used to continue patient support. the patient was in stable condition. | [5 empty] | A2 and a4 were unknown. the impella device was received from the customer for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
2561 | 1220648-2025-49235 | 23878021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878021&pc=OZD | 11/27/2025 |  |  | 12/23/2025 |  |  | 11/27/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744582 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. dr. (b)(6) reports a decreased pf from around 4ml/h to 2,5ml/h and an increased pp; recommendation given; local lysis discussed; smr forwarded; no further questions so far. csc highlighted for further support; pointed out to store the pump after explant. | [5 empty] | The impella device was received from the customer for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
2562 | 1220648-2025-49236 | 23878280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878280&pc=OZD | 11/27/2025 |  |  | 12/23/2025 |  |  | 11/26/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722095 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on support day 3 the impella was planned to be removed. however, prior to the explantation of the device, suction alarms occurred on the device approximately 15 minutes after heparin was discontinued. it was noted that the motor current consumption leveled off, and the pump flow rate fell to 0.0 l in both systolic and diastolic modes. therefore, the pump was immediately removed. there were no visible clots in the pump after removal. as it was the scheduled removal date, the pump was removed without any other mechanical support, and the patient was stable following the explantation of the impella cp. the medical team performed a computed tomography scan after removal as a precaution, and no significant blood clots were found. additionally, the medical team requested the automated impella controller to be investigated as a possible cause of the flow and suction issues. this complaint involves two impella devices: impella cp with serial number (b)(6). automated impella controller with serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: cp pump sn (b)(6) and a purge cassette returned. after 1 day on support, suction alarms occurred 15 minutes after heparin was discontinued. mc leveled off, and the pump flow rate fell to 0.0 l in both systolic and diastolic modes. act 181 with no visible clots seen on explanted pump. data log review showed suction alarms occur with 30 minutes of case remaining. at which time of the suction alarm triggering, mc spikes and drops to 200ma with pump flow dropping to zero. seemingly no impact to purge nor ps. the returned pump was found to have biomaterial wrapped around the impellor blade as well as coming out of the outflow cage. pump suction - the cause of the pump suction was biomaterial ingestion due to biomaterial found wrapped around the impellor blades and the suction alarm beginning the period of the case with a mc spike and pump flows dropping to zero. low or blocked pump flow - the cause of the low pump flow was biomaterial ingestion due to biomaterial found wrapped around the impellor blades. a suction alarm triggered with a mc spike that then reduced the pump flow to zero. device history lot: device lot: 1939765. device history batch: subcomponent lot: n/a. device history review: device sn (b)(6) passed all post sterile inspection checks. | A2, a3, a4, a6 are unknown. e. initial reporter information was not available in the complaint record accessible for mdr assessment at time of reporting. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2563 | 1220648-2025-49209 | 23878317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878317&pc=OZD | 11/21/2025 | 02/17/2026 |  | 12/23/2025 |  |  | 11/26/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026786633 |  | 1000851 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. long episodes of suction alarms during impella runtime reported, short term support, no hemolysis reported. expired | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction info: d4 catalog number. additional info: the investigation for the pump suction has been completed. the root cause of the pump suction was not determined due to suction occurring outside of positioning alarms as well and insufficient clinical details with unreturned product for further investigation.
2564 | 1220648-2025-49237 | 23878323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878323&pc=OZD | 01/18/2025 |  |  | 12/23/2025 |  |  | 11/28/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747694 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | False Alarm | Increase in Pressure | [5 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a patient implanted with an impella 5.5 for mechanical circulatory support experienced an adverse event of tachycardia. the patient was previously supported by an impella cp and was escalated to an impella 5.5. on day two of impella cp support, the patient experienced hematuria, and the pump was found to be inserted into the ventricle deep when reviewed with echocardiography. the pump was repositioned, and the urine began to clear. additionally, the decision was made to escalate the patient¿s care by cannulating for extracorporeal membrane oxygenation (ecmo). on day three of impella cp support, the patient experienced bleeding from the groin access site. the patient was transfused with one unit of packed red blood cells. on day five of impella cp support, the patient¿s platelet count and hematocrit dropped despite no active signs of bleeding. the patient was transfused with an additional unit of packed red blood cells. on impella cp support day eight, the decision was made to escalate the patient to an impella 5.5 as a bridge to advanced therapy options. on day three of impella 5.5 support, the patient was successfully weaned and decannulated from ecmo. it was reported that the impella 5.5 was repositioned during the ecmo decannulation, and the placement signal was disabled with appropriate waveforms noted. on day five of impella 5.5 support, it was reported that the patient developed tachycardia with a heart rate of 115-120 beats per minute and an electrocardiogram was performed. on support day eight, it was noted that pump flow was improved with nipride. an echocardiogram was performed and revealed that the impella was positioned too deep in the ventricle. a plan was made to reposition the device and optimize performance levels on the impella after repositioning. the following day, it was reported that there had been suction alarms and low placement alarms overnight. a bedside echocardiogram was performed and measured the impella at 5 centimeters form the inlet and was oriented towards the left ventricle wall. the physician opted not to reposition the device as it appeared to be free of interaction with the inlet. the performance level of the device was decreased from p5 to p4, which resolved alarms. at the time of writing this report, the patient continues to be supported by impella 5.5 while awaiting advanced therapy options. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this report specifically addresses pump 2: impella 5.5. a separate report will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. high purge pressure: clinical details report a drop in purge flow on 18/jan/2026 compared to the baseline at case start. there were no alarms or changes in mc. the patient's acts were in target range. no other details were provided. data logs were not available for evaluation and product wasn't returned. since neither data logs nor product were available for investigation, the cause of the high purge pressure could not be determined. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2565 | 1220648-2025-49238 | 23878390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878390&pc=OZD | 12/06/2025 |  |  | 12/23/2025 |  |  | 12/06/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720799 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female | 70 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced oozing from the access site during support. additionally it was reported that suction was experienced and a clot was suspected. in response, the pump was pulled back. during this time it was determined that the patient no longer needed support and decision made to explant the device. the impella was removed and exchanged for 8 french sheath. the patient received a dye shot and it was found that there was no flow down the leg. the sheath was exchanged for 9 french sheath and vascular was called for intervention. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2566 | 1220648-2025-49239 | 23878860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878860&pc=OZD | 11/30/2025 |  |  | 12/23/2025 |  |  | 11/30/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690389 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the impella was explanted because the arterial cannula of the extracorporeal membrane oxygenation (ecmo) exerted pressure on the lock of the impella. this pressure led to extremity ischemia. it was reported that the ischemia was decreasing after explant. neither the impella or the repo lock were guilty but the vascular condition and the ecmo cannula. the product was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
2567 | 1035166-2025-00067 | 23878907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23878907&pc=DYB | 12/04/2025 | 12/23/2025 |  | 12/23/2025 |  |  | 12/17/2025 | 12/23/2025 |  |  | 02/01/2026 |  | 12/23/2025 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3025 |  |  | INTEGER |  |  | US |  | DP-18438 |  | 0052-3025 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Distributor |  |  |  |  | 87 YR |  | Male |  |  |  | Injury | Yes | No | Difficult to Advance | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complaint summary: we have been notified of a complaint involving a oscor intro kit, 14fr x 13 and 25 cm (material #0052-3025, lot # dp-18438). "right groin access was successfully obtained for impella placement. left groin access attempted next by the interventional cardiologists using a 6 french sheath, however the sheath could not be advanced. the cardiologist removed the sheath from the left side and achieved hemostasis by manual compression. right groin used for single access pro pci. pro pci was successfully performed using single right groin access, the impella device provided left ventricular support throughout the procedure. at completion of the procedure the patient developed tachycardic and hypotension. clinical assessment revealed a left groin hematoma with active bleeding. hemodynamics stabilised. vascular surgeon called to assist with left groin bleed. hemostasis was successfully achieved. impella was explanted 18.00 once patient was stabilised. patient was taken to theatres for groin exploration and surgical repair of the left groin." the product return was requested but has not yet been received. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. integer is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by integer which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, integer, or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, integer, or its employees, that the report constitutes an admission that the device, integer, or its employees caused or contributed to the event described in this report.
2568 | 1220648-2025-49243 | 23880201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23880201&pc=OZD | 12/01/2025 |  |  | 12/23/2025 |  |  | 12/01/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768094 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Cardiogenic Shock | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp device via the right femoral artery for mechanical circulatory support. during support, the patient required cardiopulmonary resuscitation. ventricular fibrillation was subsequently reported. the patient expired while on impella cp support. it was reported that the physician did not link the patient¿s death with impella use, and that the patient suffered from fulminant cardiogenic shock and therapy was discontinued after maximum care was given. | [5 empty] | A4, a5, a6 are unknown. investigation summary: the investigation has been completed. ventricular fibrillation/peri-procedural adverse event: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2569 | 1220648-2025-49242 | 23880263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23880263&pc=OZD | 11/26/2025 |  |  | 12/23/2025 |  |  | 11/26/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669769 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was implanted via the right axillary artery and it was reported that on day 3 of impella support the position of the impella was found to be not optimal which was confirmed with fluoroscopy and echocardiography. the patient was taken back to the operating room to reposition the device. the device was removed, rewired, and reinserted. the patient subsequently returned to the operating room for further repositioning of the impella, and evacuation of a chest wall hematoma and drain insertion. the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product was returned. hematoma: the cause of the hematoma was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the device in wrong position was unable to be determined due to insufficient clinical details. device history review device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2570 | 1220648-2025-49245 | 23880376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23880376&pc=OZD | 11/25/2025 | 12/23/2025 |  | 12/23/2025 |  |  | 11/25/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766022 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that there was oozing from the insertion site. the decision was made to angle match of the impella device and to shut off heparin and place purse string. the patient was on bivalirudin and partial thromboplastin time was 61.2. in addition the bleeding from the right groin improved and the purse string was placed. the impella cp was removed post extracorporeal membrane oxygenation placement and the device was pulled out of the left ventricle. this is one of two related reports. this report addresses the second impella cp used and another report was submitted to address the first impella cp used. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary; no product returned: asae/ major bleed: bleeding at right groin cp site improved by purse string placed and holding bival. the cause of the access site adverse event was user related as the bleeding was resolved by placing a purse string suture and holding anticoagulation. device history review: the pump passed all the post sterile inspection checks.
2571 | 1220648-2025-49246 | 23881525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23881525&pc=OZD | 12/03/2025 |  |  | 12/23/2025 |  |  | 12/03/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1077403 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that prior to explanting an impella cp the flow was manually turned down to p0 but the automatic impella controller automatically turned the flow back up to p2. the rn turned the flow back down to p0 again and again the controller turned itself back up to p2. the rn pulled the white cable and the impella was removed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2572 | 1220648-2025-49248 | 23881767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23881767&pc=OZD | 12/09/2025 |  |  | 12/23/2025 |  |  | 12/09/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757678 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 was suspected of having a gastrointestinal bleed. heparin was withheld. impella support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was provided therefore d6b was updated. the investigation was completed (major bleed): the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks
2573 | 1220648-2025-49249 | 23881798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23881798&pc=OZD | 12/09/2025 |  |  | 12/23/2025 |  |  | 12/09/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763892 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. no patient harm was reported. impella support continues. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary data review: data logs showed pump ran without issue during support. there was flow drop with auto suction respond in early of the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history lot ==> device lot #: 1978908 device history batch subcomponent lot#: n/a device history review pump sn: (b)(6) passed all post sterile inspection checks.
2574 | 1220648-2025-49250 | 23881900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23881900&pc=OZD | 12/06/2025 |  |  | 12/23/2025 |  |  | 12/06/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026676321 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient developed rapid atrial fibrillation with rapid ventricular response (rvr) overnight; amiodarone was initiated at that time. impella device was removed today at noon, with native cardiac output measured at 6 l/min at the time of removal. the patient outcome at explant was reported as survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. ppae (tachycardia ) : in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The reporter clarified that the patient experienced atrial fibrillation prior to implantation of the impella.
2575 | 1220648-2025-49252 | 23882136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882136&pc=OZD | 12/08/2025 |  |  | 12/23/2025 |  |  | 12/08/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026776615 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon the removal of the device, they could no longer locate the patients posterior tibial pulse that was present before the removal. a vascular ultrasound was ordered and it was determined that the patient had a partial occlusion in the mid femoral artery and no flow below the right knee. the patient was taken to the operating room for an emergent vascular repair surgery. the patient outcome at explant was reported as survived. | Additional information was received that the patient is still sick in the intensive care unit. having multiple episodes of ventricular tachycardia. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received therefore b5 was updated. the investigation was completed. (arrhythmia/ ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
2576 | 1220648-2025-49253 | 23882144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882144&pc=OZD | 12/09/2025 |  |  | 12/23/2025 |  |  | 12/09/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026731290 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that there were multiple attempts that were made to insert the pump, due to the patient¿s anatomy as visualized under fluoroscopy, insertion was unsuccessful. the procedure was aborted because of patient anatomy, with no concerns related to the pump itself. patient went on to receive an axillary intra-aortic balloon pump (iabp). | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and tortuosity preventing successful delivery of impella. device history review: device (sn: (b)(6)) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2577 | 1220648-2025-49254 | 23882150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882150&pc=OZD | 12/11/2025 |  |  | 12/23/2025 |  |  | 12/11/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026748991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 71 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient had runs of ventricular tachycardia/ventricular fibrillation and required several shocks to obtain normal sinus rhythm. impella position good on echo and did not require repositioning. additional information was received that indicated that the patient care was withdrawn and the patient expired. | The complainant reported additional information upon request: "1. this pump is not available for return. 2. 3. i was not told cause of death. right ventricle (rv) failure was noted on transesophageal echocardiogram in the operating room and rp flex placement was attempted but was unsuccessful. patient placed on inotropic support for rv but function did not improve. despite maximum medical therapy, patient condition did not improve and family made decision to withdraw care. the physician did not make any comments about impella contributing to patient death." |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: ppae (tachycardia/ventricular fibrillation): in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1978837. device history batch: subcomponent lot: n/a. device history review: the pump sn: (b)(6) passed all the post sterile inspection checks. | Medical safety/clinical reviewed the event. this event involves a patient diagnosed with postcardiotomy cardiogenic shock (pccs) and low cardiac output syndrome (lcos), stage d, who underwent implantation of an impella 5.5 device. following direct placement of the impella 5.5, the patient experienced episodes of ventricular tachycardia and ventricular fibrillation (vt/vf) and required multiple defibrillation shocks to achieve normal sinus rhythm. an echocardiographic assessment was performed and documented appropriate impella 5.5 positioning; no device repositioning was required. subsequent attempts were made to provide biventricular mechanical circulatory support using an impella rp flex device. the attempts were unsuccessful due to the inability to advance the device into the pulmonary artery. the rp flex procedure was aborted, and no procedural complications were reported based on the information available. approximately 24 hours later, a decision was made to withdraw life-sustaining care, and the patient subsequently expired. note: h6. investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | B5 updated.
2578 | 1220648-2025-49240 | 23882191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882191&pc=OZD | 12/11/2025 | 01/15/2026 |  | 12/23/2025 |  |  | 12/11/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1545509 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Network Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complaint reported that the field team called stating that automated impella controller had both the airplane button and cloud led simultaneously blinking continuously while in use for a patient case. after the case was completed, the physician troubleshooted the controller. the team attempted to troubleshoot, however there were unable to perform factory reset. there was no patient injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. e1 (initial reporter) is unknown. | D9 updated; the product has been returned. the investigation is still ongoing.
2579 | 1220648-2025-49255 | 23882269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882269&pc=OZD | 12/10/2025 |  |  | 12/23/2025 |  |  | 12/10/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026776429 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted as the sheath was unable to cross. shockwave was utilized but was unsuccessful. the patient was in stable condition. | [5 empty] | Product complaint # (b)(4). investigation summary: no product returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had extensive pvd preventing successful delivery of impella. device history lot: device lot: 1995554. device history batch: subcomponent lot: n/a. device history review: device sn (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2580 | 1220648-2025-49256 | 23882285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882285&pc=OZD | 12/10/2025 | 01/08/2026 |  | 12/23/2025 |  |  | 12/10/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026778712 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 60 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during implantation of an impella cp the device kinked and could not obtain adequate flows. the impella was explanted and a new device was used for patient support. no patient harm was reported. | The complainant reported further information that the impella kinked where the catheter and cannula meet above the motor. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. h6 updated. the investigation is still ongoing. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
2581 | 1220648-2025-49257 | 23882713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882713&pc=OZD | 12/10/2025 |  |  | 12/23/2025 |  |  | 12/11/2025 | 12/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764602 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Damaged by Another Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the access site when a femoral sheath was placed for a swan. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added code a01 based on information in the complaint file.
2582 | 1220648-2025-49258 | 23882732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882732&pc=OZD | 12/10/2025 | 01/05/2026 |  | 12/23/2025 |  |  | 12/10/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026753172 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that implantation of an impella 5.5 was aborted as the tip of the pump could not pass a calcified lesion at the origin of the innominate artery. multiple wires and a shockwave balloon were attempted, and the patient was repositioned, all without success. no patient harm was reported. | The patient was placed on intra-aortic balloon therapy after the 5.5 placement was abandoned. the patient was then returned to the icu with no harm from the procedure. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
2583 | 1220648-2025-49259 | 23882736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882736&pc=OZD | 12/10/2025 | 02/10/2026 |  | 12/23/2025 |  |  | 12/10/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026764672 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 118 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a placement signal not reliable alarm. no kinks were identified, and proper pump placement was confirmed via echocardiogram. the alarm was disabled. no patient harm was reported. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. optical sensor issue: the pump was not returned for investigation. data log review was not provided or retrieved from the remote server. the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data log review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e1 added initial reporter phone as it was omitted from the initial report that was submitted. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The following sections have been updated: d6b explant date.
2584 | 1220648-2025-49260 | 23882763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882763&pc=OZD | 12/10/2025 |  |  | 12/23/2025 |  |  | 12/10/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644276 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal that did not correlate with the blood pressure cuff and arterial line. later, the patient expired. | [5 empty] | Section h6 has been updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned. cardiac arrest: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. optical sensor issue: the data logs show that upon starting the pump, the ps appeared aortic and there were impella position in aorta alarms. the ps was very erratic and snr and contrast dropped after the pump was turned on but were still within expected range. the ps can be seen to gradually decrease after and there was no clear or sinus pulsatility in the ps. the diastolic and systolic values were variable as well as mc. ps low alarms were accurately triggered throughout the entire pump run. due to a lack of clinical details about what was occurring with the patient or positioning of the pump, as well as no product returned, the root cause of the optical sensor issue cannot be established. device history review: impella (b)(6) passed all post sterile inspection checks.
2585 | 1220648-2025-49261 | 23882771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882771&pc=OZD | 12/10/2025 |  |  | 12/23/2025 |  |  | 12/10/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026737561 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 73 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the pump was confirmed to be in proper position. after the pump was explanted, the patient lost pulses in the right lower extremity. contralateral access was obtained and a blockage was observed at the arterial site. a balloon was used to open the artery and a stent was placed to achieve distal perfusion. the patient was stable. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the hematuria and ischemia could not be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2586 | 1220648-2025-49262 | 23882834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882834&pc=OZD | 02/19/2024 |  |  | 12/23/2025 |  |  | 02/19/2024 | 12/23/2025 |  |  | 11/30/2025 |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436278 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced an impella position in aorta alarm. it is unclear how the issue was resolved. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2587 | 1220648-2025-49263 | 23882931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23882931&pc=OZD | 12/11/2025 |  |  | 12/23/2025 |  |  | 12/11/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735926 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Ischemia | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | It was reported that during weaning from an impella cp the patient became hypotensive and developed ischemia of the leg. the pump was explanted and the patient was in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information: it was reported that the patient did have a return of pulses to the leg and has no other issues. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. ppae (ischemia): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. ppae (hypotension): the cause of the injury was not determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection checks.
2588 | 1220648-2025-49264 | 23883230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23883230&pc=OZD | 11/04/2025 |  |  | 12/23/2025 |  |  | 12/17/2025 | 12/23/2025 |  |  |  |  | 12/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811664 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller was experiencing performance issues requiring servicing. | [5 empty] | The reported event did not involve patient use. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. investigation summary: data review: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was tested. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
2589 | 23862904 | 23862904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23862904&pc=OZD | 11/27/2025 | 12/01/2025 | 12/01/2025 | 12/22/2025 | 12/01/2025 | 12/22/2025 |  | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE CASSETTE |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025100 |  |  |  | No answer provided |  |  | Not Available |  |  |  | Health Professional |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | While changing the purge cassette on the impella, a message displayed stating 'defective purge cassette. replace purge cassette and tubing'. bedside rn retrieved a new impella purge cassette and changed the tubing and fluid. the product did not reach the patient, and there was not harm.
2590 | 1220648-2025-49172 | 23862932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23862932&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026767275 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female | 80 KG |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. on echocardiogram, a thrombus was seen in the left ventricle with little to no flow on impella. a discussion with the physician about possible thrombus aspiration and potential pump stop was had and the, physician decided to leave the pump in place because they did not want to risk mobilizing the thrombus. the physician informed about possible retrograde flow. the survival outcome at explant was noted as the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Low or blocked pump flow: the pump was returned for investigation. no physical damage was noted on the returned product. the pump cannula and impeller were free of biomaterial. the cause of the low pump flow was most likely biomaterial ingestion based on data log review. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2591 | 1220648-2025-49173 | 23863062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863062&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653773 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a hemorrhagic stroke in the right frontal lobe resulting in weakness on the left side. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2592 | 1220648-2025-49174 | 23863095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863095&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026750040 |  | 000000000001000482 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the medical doctor called to report that a red alarm "¿purgepressure high¿ pf = <1ml/h and pp 1060mmhg, mc on (b)(4)." recommendation was given to replace the pump and discussed recombinant tissue plasminogen activator (rtpa) lysis. instructions were sent by email and the local team and medical team were informed. the next day the medical doctor on duty reported "successful rtpa - lysis; now pf 8ml/h and pp 500mmhg". no further information was possible. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
2593 | 1220648-2025-49176 | 23863212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863212&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674981 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 95 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support during a percutaneous coronary intervention. the patient arrived to the cardiac catheterization laboratory where diagnostic images found a highly calcified right coronary artery occluded proximally. the patient became hypotensive during the diagnostic imaging and required a levophed drip at 15 mcg/min, and the medical team decided to place an impella cp for support. the impella cp was inserted via the right femoral artery and support was initiated without any reported complications. following a successful pci, the patient was transferred back to the intensive care unit for recovery. on support day three, it was noted that the patient was placed on antibiotics following a positive blood culture. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Investigation summary: the investigation has been completed. infection: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2594 | 1220648-2025-49175 | 23863262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863262&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701886 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | Malposition of Device | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem |  |  |  | Heart Block | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced malpositioning of the pump and a complete heart block. the patient was resuscitated, and the pump was repositioned. | [5 empty] | H6 health effect - clinical code e2403 was removed. a health effect - impact code was added. medical device problem codes a24 and a150202 were removed and new codes were added. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2595 | 1220648-2025-49177 | 23863293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863293&pc=OZD | 08/25/2024 | 12/22/2025 |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | IMPELLA |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on extracorporeal membrane oxygenation (ecmo) was transported to a center. percutaneous coronary intervention and impella cp were implanted. ecmo was weaned and the patient needed to be upgraded to an impella 5.5. impella 5.5 was then implanted via a. lusoria and then later the patient was transitioned to a heartmate 3 from the impella and thereby dissection of a. lusoria. there was also apical thrombus formation noted and the impella cp was escalated to an impella 5.5. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the necessary materials were not returned for analysis so the serial/lot number could not be identified to complete a phr inspection. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2596 | 1220648-2025-49179 | 23863371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863371&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026771520 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, heparin was on pause as the patient had gastrointestinal (gi) bleeding. two units of red blood cells were given. it was reported that the patient had a history of gi surgery approximately 45 days prior. later, the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the access site bleed could not be determined due to insufficient clinical details.
2597 | 1220648-2025-49178 | 23863494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863494&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771210 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Inflammation | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ectopy. the pump was repositioned in response. the patient also experienced edema and heparin induced thrombocytopenia. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. inflammation and thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2598 | 1220648-2025-49182 | 23863620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863620&pc=OZD | 11/22/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763627 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 72 year old male was implanted with an impella cp for support during high-risk cardiac procedure. it was reported that when physician was placing impella across the aortic valve, he was pushed to hard to try and get across the valve which had severe stenosis. the physician removed pump and saw a kink in the catheter right behind the motor so he decided to place a new pump. the new impella was placed in good position. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had severe stenosis and resistance at axillary aortic valve preventing successful delivery of impella. product damage (cannula kink): the cause of cannula kink was most likely patient condition related since the patient had severe stenosis and resistance at axillary aortic valve. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2599 | 1220648-2025-49181 | 23863672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863672&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC |  |  | US |  | 2026774907 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Hypovolemia | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Clinical rationale: a 75-year-old male underwent high risk percutaneous coronary intervention (hrpci) and was supported with an impella device. his pre support clinical condition was consistent with scai shock stage d. overnight, the bedside nurse reported slight oozing and bleeding from the access site, which resolved with simple redressing. the care team noted hypovolemia with a cvp of -2, for which the patient received two units of packed red blood cells and normal saline. during the course of care, the patient sat up, causing the pump to become malpositioned; repositioning was performed and confirmed by imaging, after which the positioning issue was resolved. additional volume resuscitation included 2 units prbc and 2l of fluid, and dobutamine was weaned off. the patient later developed ventricular fibrillation arrest. he required 45 minutes of resuscitation, including 28 defibrillation attempts. after discussion with the medical team, the family elected to withdraw care, and the patient subsequently expired. the patient was receiving anticoagulation with heparin at 600 units/hr, and act at the time of bleeding was 170. no underlying patient conditions contributed to the bleeding aside from patient movement during support. the patient¿s subsequent ventricular fibrillation arrest and death were not attributed to device malfunction, as the device was functioning appropriately at the time, and no pump related complications were identified. the death is conservatively being reported to the impella cp but is unlikely a contributing factor and is most likely attributed to patients underlying critical condition as they presented in scai stage d shock on multiple inotropes and pressors and was in need of mitral valve surgery. | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the support, the patient experienced slight oozing/bleeding from the access site overnight reported by the bedside nurse. this was resolved by changing the dressing. the team also noted that the patient was hypovolemic with a central venous pressure measurement of -2. due to this, they have given two units of blood among normal saline to resolve his volume challenge. additionally, the pump came out of position due to the patient sitting up. two days later, the patient ventricular fibrillation arrested overnight. the patient coded for 45 minutes totaling 28 defibrillations to resuscitate him. the family withdrew care after they were notified and the patient expired. |  |  |  |  | B5 additional information was received. section d catalog number was corrected. h6 medical device problem code has been updated. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
2600 | 1220648-2025-49183 | 23863752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863752&pc=OZD | 11/24/2025 |  |  | 12/22/2025 |  |  | 11/25/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687899 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 112 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding at the chest site due to surgical complications. the patient expired. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: ppae (major bleed): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2601 | 1220648-2025-49168 | 23863905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863905&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026771231 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced thrombocytopenia. the patient was put on continuous renal therapy. the patient experienced abdominal pain and was treated with increased pressors. bleeding occurred around the liver and kidneys, and the patient developed ischemic bowel. care was withdrawn and the patient expired. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. ppae (thrombocytopenia/major bleed): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump sn (b)(4) passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
2602 | 1220648-2025-49184 | 23863981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23863981&pc=OZD | 11/23/2025 |  |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744963 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia and was treated with amio boluses. additionally, there were purge pressure high and purge pressure blocked alarms. no kinks were observed in the lines. the purge cassette was changed without resolution. tissue plasminogen activator was added to the purge to resolve the alarms. the patient was in stable condition. | [5 empty] | H6: updated codes based on the results of the investigation. investigation summary: the device was returned for evaluation. tachycardia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. purge pressure high: clinical reported purge pressure high alarms with tissue plasminogen activator (tpa) resolving the issue. the returned product was inspected and there was foreign decayed material in the cannula and by the impeller. there was also biomaterial beneath the impeller. the pump was run in the lab and high purge pressure did not reproduce. the cause of the issue was not established as no product defects were found and no logs were returned for analysis. device history review: the pump passed all post sterile inspection checks | D9 and h3 updated. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2603 | 1220648-2025-49169 | 23864060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864060&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753538 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. urine is now pink. repeat hemoglobin and hematocrit produced the same results/no change. heparin drop was turned down and the pink urine is now clear. the p level of the pump was reduced to p-5 to resolve the issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Device identification information, section d4, has been updated. the cause of the injury was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2604 | 1220648-2025-49185 | 23864165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864165&pc=OZD | 11/24/2025 |  |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719090 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 became septic and expired. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the sepsis was unable to be determined due to insufficient clinical details. | D6b, date explant, has been updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2605 | 1220648-2025-49186 | 23864284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864284&pc=OZD | 12/15/2025 |  |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1229940 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an automated impella controller failure. | [5 empty] | D6, date of implant and explant has been added. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2606 | 1220648-2025-49189 | 23864461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864461&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705297 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the nurse reported bleeding from the groin site. manual pressure has been held. the patient was taken for a computed tomography then was upgraded to an impella 5.5. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the access site bleed could not be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma¿. section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site¿. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
2607 | 1220648-2025-49187 | 23864601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864601&pc=OZD | 11/24/2025 |  |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026771587 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the access site. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2608 | 1220648-2025-49190 | 23864742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864742&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662810 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 78 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp for support during high risk cardiac procedure. the impella was placed via right femoral artery with ease and started on auto. the patient experienced increased ectopy so the pump placed on p7 and transferred back to the cardiothoracic intensive care unit with improved hemodynamics and rhythm maintenance. the next day, it was reported that there was bleeding from impella site resulting in drop in hemoglobin from 13 to 8. one unit of packed red blood cells were given, pressure was held and anticoagulation goals were adjusted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the root cause of access site adverse event was use related as it required readjusting anticoagulation goals.
2609 | 1220648-2025-49192 | 23864824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864824&pc=OZD | 11/28/2025 |  |  | 12/22/2025 |  |  | 12/01/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711659 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp device via the right femoral artery for mechanical circulatory support. it was reported that following impella insertion a hematoma formed at the impella access site. the bleeding temporarily stopped when the patient was transferred back to the intensive care unit, however there was noted to be pulsative bleeding that was observed later that day. as a result of the bleeding, the impella was explanted and the bleeding was controlled by surgical repair. it was the physician¿s opinion that the physician who implanted the device may have caused artery damage during the arterial punction or peel-away sheath removal as it was their first time performing the procedure. follow-up efforts were unsuccessful in obtaining additional information, and the final patient outcome is unknown. | [5 empty] | H6 added medical device problem code that was omitted from the initial report that was submitted. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma/asae: the cause of the access site adverse event was not determined due to insufficient clinical details. vessel perforation/patient device interaction: the cause of the vessel perforation was not determined due to insufficient clinical details. major bleed: armpit bleed reported. the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2610 | 1220648-2025-49191 | 23864864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23864864&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026768637 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient went into junctional rhythm overnight. the doctor switched him to a-pacing this morning. pulsatility and hemodynamics were improved but then an echocardiogram was done and the impella was repositioned. while repositioning the impella, ectopy occurred and pulse pressure dropped from ~20mmhg to 5mmhg. the doctor stated he could not visualize measurement under transthoracic echocardiogram. leaving the patient a-paced on p8 and will be getting daily echocardiograms to confirm ejection fraction and impella placement. later, the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. ppae (arrhythmia, fever): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2611 | 1220648-2025-49194 | 23865183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23865183&pc=OZD | 11/24/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768329 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | Sepsis | Hematuria | Heart Block | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with a impella cp for support during high-risk cardiac procedure. it was reported that there was a hematoma on the right femoral access site. manual pressure was held and a femstop was in place upon transport to the intensive care unit with 30mmhg of pressure. they were able to palpate distal pulses without an issue. the patient was taken to the cardiac catheterization laboratory on levophed/dopa/epinephrine for transvenous pacing placement for complete heart block. the impella initially increased to p-8 for support with an attempt to wean inotropes. heart failure at bedside was noted to reposition the impella which now is at 91 centimeters and measuring 3.7 centimeters past the atrioventricular. after continued hematuria on bumex 0.5mg/hr infusion per nephrology. with mixed sepsis/cardiogenic shock picture, vasopressin was added and dopamine was weaned to 2.5. epinephrine was off and the impella to p-4 with improved ejection fraction of 35%. urine improved on bumex infusion and after repositioning yesterday it is now clear yellow. dressing was clean, dry, and intact. the patient was on lower ventilator settings today with good oxygenation on 70% and five positive end-expiratory pressure, a febrile. the medical doctor was to review the echocardiogram and determined weaning. | [5 empty] | Investigation summary: no product was returned. hematuria: the cause of the injury was unable to be determined due to insufficient clinical details. sepsis, fever, heart block: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the cause of the hematoma was unable to be determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
2612 | 1220648-2025-49193 | 23865426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23865426&pc=OZD | 11/19/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1705772 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Operating System Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. customer switched the pump from an external aic to ic 10259. customer told that when switching the purge cassette later on the aic, the aic hung up. pump disconnected and restart of the aic. then no problem was observed. after this issue switch to a different aic to send the aic back to aachen. customer send a letter to aachen. patient got hypotensive while switching the pump, but no harm occurred. stable after restart the aic. stable on next aic. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2613 | 1220648-2025-49188 | 23865515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23865515&pc=OZD | 11/26/2025 |  |  | 12/22/2025 |  |  | 11/26/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026768664 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation. the pump was repositioned. later, the patient experienced ventricular tachycardia requiring defibrillation. the patient expired. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the tachycardia, paroxysmal and atrial fibrillation were unable to be determined due to insufficient clinical details. | Medical safety/clinical reviewed the event. a 61-year-old female with a known history of coronary artery disease presented in cardiogenic shock classified as society for cardiovascular angiography and interventions (scai) stage d. an impella cp was implanted for hemodynamic support. on post-implant day 0, the patient experienced bleeding at the groin access site. nursing staff applied manual pressure; however, due to continued bleeding, the impella cp was explanted, and the patient was escalated to impella 5.5 support. approximately four days after initiation of impella 5.5 support, the patient developed paroxysmal atrial fibrillation and required cardioversion. the device was noted to have migrated. point-of-care ultrasound (pocus) assessment demonstrated the inlet measuring approximately 5.8 cm and contacting the atrioventricular septum. imaging views were limited, and transesophageal echocardiography (tee) was recommended to assist with repositioning prior to extubation. lactate dehydrogenase (ldh) levels remained elevated. urine output improved following intravenous (ivp) lysix. seven days later, the patient experienced episodes of ventricular tachycardia overnight and again the following morning, requiring defibrillation on both occasions. nine days thereafter, care was withdrawn, and the patient expired. the access site bleeding associated with the impella cp is being reported as a serious injury type of reportable event, and the impella 5.5 is being reported as a death. the patient experienced device positioning complications and arrhythmic events. however, the patient¿s underlying conditions (advanced coronary artery disease, cardiogenic shock scai stage d) contributed most to the demise outcome. correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: remove peri-procedural adverse event/adverse event without identified device or use problem (a24) and insert malposition of device (a150202). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. this is one of two related reports. refer to section h10 for the related report number.
2614 | 1220648-2025-49195 | 23865797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23865797&pc=OZD | 11/23/2025 |  |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616772 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was implanted with a impella cp for support during high-risk cardiac procedure. it was reported that the impella measured at 3.2 centimeters in the left ventricle on echocardiogram. an impella position in aorta alarm occurred. under the echocardiogram, both the pigtail and cannula completely were out of the left ventricle past the aortic valve. they were unable to advance and the impella was pulled back into the descending aorta and ran at p1 overnight. in addition, the impella was explanted. the issue is that the cm marking did not change and the impella was locked. somehow, the impella migrated out of the left ventricle into the aorta. no position changes were made, no ectopy, and no change to drug regimen were noted. the impella migrated on its own. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2615 | 1220648-2025-49196 | 23865966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23865966&pc=OZD | 11/23/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026740994 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella 5.5 for support during a high-risk cardiac procedure. it was reported that the pan-computed tomography was due to fever. one unit of plasma was also given. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary ppae (infection/fever): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review ==> the pump sn 625545 passed all the post sterile inspection checks.
2616 | 1220648-2025-49197 | 23866398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23866398&pc=OZD | 10/19/2025 | 02/12/2026 |  | 12/22/2025 |  |  | 11/25/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753147 |  | 000000000001000851 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. no ao-placement signal after implant of cp smart assist. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary optical sensor issue: clinical details report the aops was absent after inserting the pump in the body, so the pump was explanted as a result. data logs were reviewed and there were no abnormalities noted. in imc logs, the pump was connected and the optical sensor was calibrated without issue. then rt logs show the pump being inserted in the body, based on typical pulsatile pressure measurements. during this entire period, optical signal 5s parameters remained stable at expected values, and there were no psnr alarms. the pump was then explanted and replaced with a new one. product was not returned for investigation. based on data log review, there was no problem found in relation to the reported optical sensor issue.
2617 | 1220648-2025-49198 | 23866415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23866415&pc=OZD | 11/22/2025 |  |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660281 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella cp for support during high-risk cardiac procedure. it was reported that lower limb ischemia was noted. the patient also had a right iliac fracture. there was no access and blood flow could not be secured with an antegrade sheath. after impella cp removal, an intra-aortic balloon pump was inserted and the nicotinamide adenine dinucleotide and dobutamine were increased, and the patient's condition stabilized. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2618 | 1220648-2025-49200 | 23866631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23866631&pc=OZD | 11/24/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519164 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that the groin access site was cleaned, dried and intact with no hematoma for about an hour after implant. after the patient was moved to a stretcher for medflight transport, the team noted a hard hematoma and bleeding from the access site. pressure was held with good effect. in addition, there was continued suction at any p levels higher than at p4 and above. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2619 | 1220648-2025-49199 | 23867001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23867001&pc=OZD | 11/24/2025 |  |  | 12/22/2025 |  |  | 11/25/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1388684 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Signal not reliable. new lamp & cable did not solve the problem. | [5 empty] | Device received, inv in progress: the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2620 | 1220648-2025-49201 | 23868101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23868101&pc=OZD | 11/23/2025 |  |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728373 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 120 KG |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Anemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. hematuria that is clearing after pump reposition was reported. the next day positioning issue and anemia were reported. an echocardiogram showed ejection fraction of 10-15% on extracorporeal membrane oxygenation/impella cp. the impella was at p-2 and the left ventricle was dilated with the impella cp measuring deep. urine was pink-tinged. the plan was to pull back the impella and increase the p-level to p-3. a head computed tomography was to be done to assess neurology function. one unit of packed red blood cells was given for hemoglobin less than seven. an additional unit to be transfused the next day. there was no evidence of bleeding from the impella site. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. anemia, hematuria: the cause of the injury was not determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot: device lot: 1967742 device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2621 | 1220648-2025-49202 | 23868668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23868668&pc=OZD | 11/24/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026758325 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that there was a sudden alarm indicating "purge system blocked"; no kinks, purge system already replaced without success. patient was in an overall worsening condition, and care was withdrawn. withdrew of care was not related to the device or the ci. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product with event logs were received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. the pump had a small kink on the catheter near the kink protector. no other physical damage was noted on the returned product. small, free-floating biomaterial was found on the impeller at the outflow cage, which was easily retrieved and saved. the pump was then primed and ran in a bucket of water with normal specified pump flow; purge parameters were around 9 ml/hr. and 500 mmhg. data log shows purge pressure and motor current spikes, with noted saturated pump flow and high spikes with high purge pressure alarm, which resolved after the second cassette change procedure. a few hours later in the morning, a sudden high purge pressure issue occurred again with a motor spike and pump stop. pump was successful restarted after multiple pump start followed by a gradual decrease in purge pressure. the pump was utilized and functioned fine without issue for the next three days. the purge system blocked alarm was seen along with high purge pressure alarm for both instances. the cause of the issue was most likely related to biomaterial, based on data log review and returned product investigation. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
2622 | 1220648-2025-49203 | 23869260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23869260&pc=OZD | 11/23/2025 |  |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 |  |  | 10/31/2025 |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408699 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Infection | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was presented to the emergency department after a ventricular tachycardia (vt) episode at home that self-converted. the patient was admitted to the cardiothoracic intensive care unit (cticu). increasing pressors were required and the team decided for impella 5.5 support. while on support, the patient went into vt and was cardioverted once and the pump was repositioned under transthoracic echocardiogram. the patient then went into ventricular fibrillation, got intubated and was cardioverted once. the impella was repositioned again. also, there was a concern for bacteremia and a hard twisting bend in the catheters at the red handle. there was a concern for placement signal drift. it was further reported that the patient had positive blood cultures with gram negative bacteria. the patient was started on antibiotics. the cticu team believes the likely source is from the peripherally inserted central catheter line that was in for about three weeks. the plan was to redraw blood cultures today. per the bedside nurse, the impella incision site did not look infected. later, the pump was explanted and the patient survived. | Additional information was received stating that the pump experienced placement signal drift. |  |  |  |  | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation investigation summary: the pump was not returned for investigation. data log was not provided or was not found on the remote server. placement signal issue: the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data logs. product damage (catheter kink): clinical details indicate concern for a hard twisting bend in the catheters at the red handle during support. the cause of the issue was not determined without returned product investigation and sufficient clinical details. clinical symptoms (tachycardia, unspecified infection): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical details: (b)(6) 2025: impella 5.5 s2 implanted via right axillary artery. (b)(6) 2025: patient went into ventricular tachycardia (vt) and was cardioverted once; the pump was repositioned under transthoracic echocardiography (tte) guidance. patient then went into ventricular fibrillation (vf), was intubated, and cardioverted once more. impella was repositioned again. there was concern for bacteremia and a hard twisting bend in the catheters at the red handle. concern for placement signal drift was also noted. (b)(6) 2025: patient had positive blood cultures with gram-negative bacteria and was started on antibiotics. the team believes the likely source is the peripherally inserted central catheter (picc) line, which has been in place for about 3 weeks. plan to redraw blood cultures today. per bedside nurse, the impella incision site does not look infected. (b)(6) 2025: pump explant day. device history review the pump passed all post sterile inspection checks. | The impella device is no longer available from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B5 additional information was received and d6 medical code updated. e1 added initial reporter facility address as it was omitted from the initial report that was submitted.
2623 | 1220648-2025-49205 | 23869946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23869946&pc=OZD | 11/23/2025 |  |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765660 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 77 KG |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was aborted insertion due to patient's anatomy. the patient survived. the product was not retained for evaluation. | [5 empty] | The impella device was not retrained from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | Correction: d3. the investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and resistance at femoral artery preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
2624 | 1220648-2025-49206 | 23870663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23870663&pc=OZD | 11/25/2025 |  |  | 12/22/2025 |  |  | 11/25/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654879 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient was having continuous suction with the device "needing to be moved urgently." they were able to reposition and improve the continuous suction alarm, but the patient continued to have suction related to device interaction with the intra-aortic balloon pump. | [5 empty] | H6 medical device problem code a0502 was inadvertently omitted on the initial manufacturer device report. | Medical safety/clinical reviewed the event. a 65-year-old female patient presented in cardiogenic shock (society for cardiovascular angiography and interventions [scai] stage c) secondary to advanced heart failure. an impella 5.5 was implanted for mechanical circulatory support. prior to device placement, the patient had a documented ruptured interventricular septum resulting in a ventricular septal defect (vsd). the septal rupture occurred before impella implantation and was not attributed to the device or the implant procedure. due to the septal defect, the impella was positioned deeper than typical to optimize support. on post-implant day 1, the patient experienced continuous suction alarms. given the presence of the vsd, there was concern for altered flow dynamics and potential right-sided blood entrainment. at that time, the patient remained hemodynamically stable. the impella was repositioned at the bedside, with the most stable position noted at 4.1 cm. an intra-aortic balloon pump (iabp) remained in operation concurrently. the patient continued to experience suction events, suspected to be related to device interaction and hemodynamic effects from iabp counterpulsation, including back pressure during balloon inflation. the patient was not considered a candidate for durable ventricular assist device therapy or cardiac transplantation per family wishes. ventricular function remained severely depressed. on the following day at approximately midnight, the patient expired. per the treating medical team, the patient was considered to be in advanced cardiogenic shock with progression beyond recoverable status. suction events are a known potential complication in patients receiving mechanical circulatory support, particularly in the presence of a ventricular septal defect and concurrent iabp therapy. the impella 5.5 was present at the time of death and will be reported as associated with the event. however, the patient¿s underlying clinical condition, including pre-existing ventricular septal rupture with vsd and advanced cardiogenic shock (scai stage c) was determined by the medical team to be the primary contributing factor to the patient's demise outcome. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2625 | 1220648-2025-49204 | 23871630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23871630&pc=OZD | 11/23/2025 | 12/22/2025 |  | 12/22/2025 |  |  | 11/24/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749771 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Crack | Fluid/Blood Leak | Decrease in Pressure | Increase in Pressure | Pumping Stopped |  |  | Cardiac Arrest | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the purge-arm was broken, between the filter and the purge-reservoir and the fluid drops out. the patient's pulmonary condition worsened, and the impella ran without problems. the doctoe reported that the impella was running well, presenting no issues. withdrawal of care was decided with the doctor and family. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. product damage (purge sidearm crack): the product was not returned and cause of the issue was not established. purge leak-pump/low purge pressure: the cause of the issue was not established as no product or images were returned. purge pressure high: the cause of the issue was not established as no product or logs were returned. device history review: the device passed all post sterile inspection checks. complaint coding was updated to better reflect the reported event. consequently, section h6 health effect clinical/impact codes and medical device problem codes were updated/corrected and repopulated accordingly. | This report is being submitted to provide medical safety/clinical review and to correct information provided in a previously submitted mdr corrected information: section b2 ¿ outcome previously reported as death. this has been removed. h1 ¿ type of reportable event previously reported as death. updated to serious injury. h6 ¿ health effect ¿ clinical code e2343 (hemodynamic instability) has been added. h6 ¿ health effect ¿ impact code f02 (death) was previously reported. this impact code has now been removed. 6. medical device problem code a0404 (crack) has been added. medical safety/clinical review: medical safety/clinical reviewed the event. patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support (mcs). the event describes a crack in the purge arm resulting in a leak. the pump ran on performance level p-9 with motor current in range (986/823/888mmhg). a different impella cp device was successfully inserted. patient's pulmonary condition worsened and withdrawal of care was decided with the physician and family. in this case, the patient's underlying condition most likely contributed to the demise outcome. regarding the impella cp pump 2, the physician reported that the impella cp pump 2 was running well, presenting no issues. however, conservatively, the death will be reported as the impella cp pump 2 was in use at the time of expiration. note: h6 investigation type/findings/conclusion remain unchanged. related medical device reports: this impella cp (pump 1) device report is one of two devices associated with the event. following internal review, a new complaint and mdr have been generated for pump 2 (impella cp). refer to the related manufacturer report number as noted in section h10 for the medical device report on the pump 2 impella cp. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2626 | 1220648-2025-49207 | 23872625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23872625&pc=OZD | 09/17/2025 |  |  | 12/22/2025 |  |  | 11/25/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1115977 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | After review of the complaint by the investigation team, the automated impella controller (aic) was noted as the potential cause of the pump not detected issue for the following initially reported event: the complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. after connecting to the aic, "defective impella catheter" alert was presented along with the instruction to "disconnect and replace impella catheter." medical team tried to disconnect and reconnect three times with the same results. the physician agreed to proceed with a new catheter. additional information was subsequently received and noted the following: [alarm of impella defective] the defective impella 5.5 device serial number (sn) (b)(6) was replaced with impella 5.5 device sn (b)(6). the physician met resistance in passing the impella 5.5 device sn (b)(6). this impella 5.5 device was removed, and the decision was made to perform contrast examination of the axillar / subclavian artery to determine why the impella 5.5 device sn (b)(6) was not advancing. significant stenosis of the distal subclavian artery was discovered and recognized as the reason for not being able to advance the impella sn (b)(6). a stent was subsequently placed and a second attempt to re-place the impella 5.5 sn (b)(6) failed as the pump was unable to pass through the stent. discussion was made regarding placing an impella cp from the prepared axilla; however, the team opted to not escalate care as notification was received intra-procedure that family wanted to go route of comfort care. the patient expired after care was withdrawn. | [5 empty] | Medical safety review. medical safety reviewed the event and noted the following: the patient presented with worsening cardiogenic shock and experienced cardiac arrest prior to the procedure. an impella 5.5 device was initially prepared, but three ¿defective impella catheter¿ alarms occurred during setup. a second impella 5.5 catheter was prepared without error; however, insertion was unsuccessful due to axillary artery stenosis confirmed by angiography. the physician placed a stent in the stenotic segment, but a subsequent attempt to insert the device also failed. the implant procedure was aborted, and care was withdrawn in accordance with the family¿s wishes. the death is conservatively reported under the second impella 5.5 as a failed device delivery. however, the patient¿s critical condition requiring cardiopulmonary resuscitation and poor prognosis were the most likely contributors to the patient's outcome. device status: the automated impella controller device was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this automated impella controller device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for first impella 5.5 device serial number (b)(6) and the medical device report on the second impella 5.5 serial number (b)(6), which is associated to the demise outcome.
2627 | 1220648-2025-49208 | 23872691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23872691&pc=OZD | 11/25/2025 |  |  | 12/22/2025 |  |  | 11/26/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675947 |  | 000000000001000482 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. customer called me the day before for initially a leakage of probably dextrose solution from the reservoir or the side arm without any problem for the pump, followed by a problem with the purge pressure (high pressure) on novembre 26th.they tried all procedures for troubleshoot the alarm (no kinks, change of purge cassette, deairing). | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2628 | 1220648-2025-49210 | 23872726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23872726&pc=OZD | 11/27/2025 |  |  | 12/22/2025 |  |  | 11/26/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1359283 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on support day 3 the impella was planned to be removed. however, prior to the explantation of the device, suction alarms occurred on the device approximately 15 minutes after heparin was discontinued. it was noted that the motor current consumption leveled off, and the pump flow rate fell to 0.0 l in both systolic and diastolic modes. therefore, the pump was immediately removed. there were no visible clots in the pump after removal. as it was the scheduled removal date, the pump was removed without any other mechanical support, and the patient was stable following the explantation of the impella cp. the medical team performed a computed tomography scan after removal as a precaution, and no significant blood clots were found. additionally, the medical team requested the automated impella controller to be investigated as a possible cause of the flow and suction issues. this complaint involves two impella devices: impella cp with serial number: (b)(6). automated impella controller with serial number: (b)(6). this mdr specifically addresses automated impella controller with serial number: (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A2, a3, a4, a6 are unknown. e. initial reporter information was not available in the complaint record accessible for mdr assessment at time of reporting. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2629 | 1220648-2025-49211 | 23872738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23872738&pc=OZD | 11/26/2025 |  |  | 12/22/2025 |  |  | 11/29/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026754134 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male | 90 KG |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. pt. is extubated for approx. 48 hours and shows signs of delirium. according to physician no correlation to impella therapy. no issues with the device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product or data logs returned. purge pressure high: the cause of purge pressure high was unable to be determined as limited clinical details were provided and no product was returned for review. device history the complaint pump passed all post sterile inspection checks. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. purge pressure high: the cause of purge pressure high was unable to be determined as limited clinical details were provided and no product was returned for review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2630 | 1220648-2025-49212 | 23872994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23872994&pc=OZD | 11/29/2025 |  |  | 12/22/2025 |  |  | 11/29/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750321 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Malposition of Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced optical issues. the pump was found to be in the aorta. the p level of the pump was reduced and the console was turned off. later, purge pressure issues occurred. the pump was explanted and replaced with a new pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated code a141204 based on information in the complaint file.
2631 | 1220648-2025-49213 | 23873000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873000&pc=OZD | 11/23/2025 |  |  | 12/22/2025 |  |  | 11/23/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764195 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 66 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [27 empty] | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information was received stating that the impella was explanted by the surgeon and exchanged for central extracorporeal membrane oxygenation. surgery was performed the following morning. per update, the patient did not do well postoperatively and expired at a later date; the subsequent clinical course is unknown. | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient had a small left ventricle, and subsequent to impella support a left ventricle thrombus and low ejection fraction of 5% were identified. the decision was made to explant the impella cp and centrally cannulate the patient for extracorporeal membrane oxygenation (ecmo). the following day, the patient underwent a surgical procedure. a few days later, the patient expired. it was reported that the patient experienced a poor outcome following surgery. the exact cause of death was not provided. additional clinical details, including perioperative course, contributing factors, and cause of death, were unavailable at the time of reporting. |  |  |  |  | A2 date of death is unknown. b5 additional information received. h6 type of reportable event was updated from serious injury to death and codes updated. | A5 is unknown the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. ppae (thrombosis): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | Medical safety/clinical reviewed the event. a 73-year-old male with a significant past medical history of coronary artery disease (cad), atrial fibrillation (afib), hypertension (htn), and hyperlipidemia (hld) presented to the emergency department (ed) and was found to be in ventricular tachycardia (vt). the patient underwent cardioversion by the electrophysiology (ep) service. due to respiratory distress, the patient was intubated. electrocardiogram (ecg) findings were consistent with st-segment elevation myocardial infarction (stemi), and the patient was transferred emergently to the cardiac catheterization laboratory. coronary angiography revealed 100 percent occlusion of the left anterior descending (lad) artery and chronic total occlusion of the right coronary artery (rca), complicated by cardiogenic shock (cgs). given the severity of shock, an impella cp was implanted for mechanical circulatory support. cardiothoracic surgery was consulted, and possible escalation to an impella 5.5 was discussed. after approximately four hours of impella cp support, and in the setting of reduced left ventricular ejection fraction (approximately 5 percent) with concern for possible left ventricular thrombus, the decision was made to transition the patient to central extracorporeal membrane oxygenation (ecmo). upon arrival to the operating room, the patient was receiving dobutamine (dba) at 5 mcg/kg/min. impella cp flows and waveforms were reported as appropriate at performance level p-2 with flows of approximately 1.7 liters per minute. the surgeon elected to explant the impella cp prior to initiation of central ecmo support. additional follow-up information indicated that the impella cp was discontinued per surgeon preference before surgery. the patient subsequently underwent cardiovascular surgery the following day and was maintained on central ecmo support. the patient survived the impella cp explant procedure. postoperatively, the patient¿s clinical condition deteriorated while on ecmo support. the patient expired several days later (exact timeline unknown) while not on impella support. events occurring during impella cp support are being reported as serious injury due to the patient¿s critical condition at the time of use. the death will be conservatively reported as well although it is important to note the patient was on ecmo as primary support device prior to surgery and expired post surgery. note: h6. investigation type/findings/conclusion remain unchanged.
2632 | 1220648-2025-49215 | 23873066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873066&pc=OZD | 12/01/2025 |  |  | 12/22/2025 |  |  | 12/01/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661377 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that following successful implantation of an impella cp an aortic dissection was observed. the pump was explanted and the patient is in stable condition. | [5 empty] | Patient information was not provided. section a was left blank. the cause of the injury was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2633 | 1220648-2025-49214 | 23873068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873068&pc=OZD | 09/24/2025 | 10/24/2025 |  | 12/22/2025 |  |  | 09/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026716253 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 85 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient had impella 5.5 smartassist implanted and small right axillary incision with small hematoma. jackson-pratt drain (jp) placed, anti-coagulation on hold. jp tube had intermittent drainage in various amounts. wash out of right axillary impella access site completed 5 days post implant. reported to have epistaxis and oral bleeding being managed by ear, nose, and throat (ent). | [5 empty] | No product was returned for investigation. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of the bleeding is determined to be patient condition because of the bleeding from non impella locations. the cause of the epistaxis was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2634 | 1220648-2025-49216 | 23873071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873071&pc=OZD | 12/01/2025 |  |  | 12/22/2025 |  |  | 12/01/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026752835 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Increase in Pressure | [6 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge flow blockage. the patient expired from multi organ failure while on impella support. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: clinical symptom (organ failure): the cause of the injury was not determined due to insufficient clinical details. purge system blocked: the pump was not returned for investigation. console serial number was unknown. data log was not provided or retrieved from the remote server. the cause of the purge system block issue was not determined without returned product investigation and sufficient clinical information, including data log. device history review: the complaint pump passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2635 | 1220648-2025-49217 | 23873107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873107&pc=OZD | 09/21/2025 |  |  | 12/22/2025 |  |  | 09/22/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025581166 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced malpositioning of the pump when the patient coughed. the pump was explanted and replaced with another impella 5.5. no patient harm was reported. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the cause of positioning issue was most likely use issue related since pump came loose from the aortic valve during backing of the ventilator. | A2 and a4 are unknown. no product was returned for investigation. the cause of positioning issue was most likely use issue related since pump came loose from the aortic valve during backing of the ventilator. the device passed all post sterile inspection checks.
2636 | 1220648-2025-49218 | 23873112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873112&pc=OZD | 12/08/2025 | 01/13/2026 |  | 12/22/2025 |  |  | 12/08/2025 | 12/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026771944 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Tachycardia | [40 empty] | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with and impella cp experienced intermittent tachycardia and ventricular fibrillation requiring defibrillation. | [5 empty] | Investigation summary ppae (ventricular tachycardia): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot: 1995216 device history batch subcomponent lot: n/a device history review the pump sn 640879 passed all the post sterile inspection checks. | It was reported the patient also experienced a fever. health effect code e230101 has been added. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2637 | 1220648-2025-49219 | 23873134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873134&pc=OZD | 12/09/2025 |  |  | 12/22/2025 |  |  | 12/09/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769802 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced renal failure. the patient was transfused one unit of packed red blood cells. | [5 empty] | Additional information was received indicating that the initial report was submitted in error due to a misunderstanding of the patient¿s clinical condition. at the originating facility, it was initially believed that the patient had renal failure and may require dialysis. following transfer to a subsequent facility and completion of further clinical evaluation, it was determined that dialysis was not required and no clinical issue was present. as a result, the previously submitted complaint has been redacted. the regulatory report submitted on 22-dec-2025 should be considered void and disregarded. no further information will be provided, as there was no adverse event or malfunction associated with the subject device. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2638 | 1220648-2025-49220 | 23873138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873138&pc=OZD | 09/23/2025 | 10/23/2025 |  | 12/22/2025 |  |  | 09/23/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615118 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 65 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient progressed into hemodynamic collapse after a stent placed in left anterior descending artery. decision for impella cp was made and successfully placed. during cp support the patient developed ventricular tachycardia with loss of pulsatility and flows on cp. cardiopulmonary resuscitation initiated with return of spontaneous circulation achieved. drips remain the same. decision made to escalate patient to impella 5.5 with cp removal from groin. | [5 empty] | The pump was not returned for investigation. the cause of the low pump flow issue was most likely biomaterial ingestion during the case run, based on data log review. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2639 | 1220648-2025-48838 | 23873143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873143&pc=OZD | 11/08/2025 |  |  | 12/22/2025 |  |  | 11/08/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717820 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia. no patient harm was reported. | [5 empty] | The cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2640 | 1220648-2025-49221 | 23873163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873163&pc=OZD | 09/17/2025 |  |  | 12/22/2025 |  |  | 09/17/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683986 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | Device Damaged by Another Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the guidewire was difficult to remove after implanting an impella cp. the lumen of the hemostatic valve was narrowed due to the repositioning sheath, leading to increased shear forces. the pump was explanted and replaced with a new one. no patient harm was reported. | [5 empty] | G1 added manufacturing site name and address as they were omitted from the initial report that was submitted. h6 revised the medical device problem code since the initial report was submitted. | A4 is unknown. the cause of the pump interaction by another device was most likely unintentional use-related, specifically the interaction between the repositioning sheath and the peel-away sheath, which led to narrowing of the valve lumen and increased shear forces, based on the provided clinical details. the guide wire should be removed before starting the impella, as per the steps mentioned in the ifu. the device passed all post sterile inspection checks.
2641 | 1220648-2025-49223 | 23873219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873219&pc=OZD | 09/22/2025 |  |  | 12/22/2025 |  |  | 09/22/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703497 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | Tachycardia | Ventricular Fibrillation | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a hematoma and bleeding at the access site which was treated with surgical foam, surgicel, and nu-knit. the patient experienced hypotension that quickly resolved. some clots were removed. the patient went into ventricular tachycardia and ventricular fibrillation requiring defibrillation, versed, amio, and gtt. the patient was transfused one unit of packed red blood cells due to anemia. a thombus on the left ventricle was observed on a ct scan. the patient was in stable condition. | [5 empty] | A4 is unknown. the impella was returned for investigation. the cause of the thrombosis, ventricular tachycardia, and ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of the hypotension was not determined due to insufficient clinical details. the cause of the hematoma could not be determined due to insufficient clinical details and no abnormalities observed with the product. the device passed all post sterile inspection criteria.
2642 | 1220648-2025-49224 | 23873222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873222&pc=OZD | 08/11/2025 |  |  | 12/22/2025 |  |  | 09/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697883 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male | 110 KG | Hispanic |  | Injury | Yes | No | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a thrombus in the left ventricle and positional issues. the patient was treated with dabigatran and the pump was explanted. the patient survived. | [5 empty] | No product was returned for investigation. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2643 | 1220648-2025-49225 | 23873227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873227&pc=OZD | 09/24/2025 |  |  | 12/22/2025 |  |  | 09/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659944 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced back bleeding through the graft and the small pectoral muscle. the patient was transfused volume and blood products. the graft was resolved with graft ties. the muscle bleed was stabilized surgically. the patient also experienced ectopy which resolved by repositioning the pump. | [5 empty] | The cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2644 | 1220648-2025-49226 | 23873234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873234&pc=OZD | 09/24/2025 |  |  | 12/22/2025 |  |  | 09/24/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695458 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella cp was aborted due to patient anatomy. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy along with tortuous calcified abdominal aorta and pseudoaneurysm in the right iliac preventing successful delivery of impella. the device passed all post sterile inspection checks.
2645 | 1220648-2025-49227 | 23873259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873259&pc=OZD | 09/15/2025 |  |  | 12/22/2025 |  |  | 09/08/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619150 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Fever | Hemorrhage/Blood Loss/Bleeding | Inflammation | Tachycardia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring defibrillation. the patient also had a fever and hypotension. vasos were given. oozing from the graft anastomosis caused tamponade within the chest cavity requiring cvor and sutures to be placed. later, the patient developed a gastrointestinal bleed with ischemic colitis and tarry stools. systemic heparin was turned off and the patient was transfused blood products. | [5 empty] | The root cause of the tachycardia and atrial fibrillation was unable to be determined due to insufficient clinical details. the cause of the fever and inflammation was not determined due to insufficient clinical details. the cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks.
2646 | 1220648-2025-49228 | 23873276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23873276&pc=OZD | 12/09/2025 |  |  | 12/22/2025 |  |  | 12/09/2025 | 12/22/2025 |  |  |  |  | 12/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1647627 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller generated a "change aic" yellow alarm. the procedure was completed without patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2647 | 1220648-2025-49147 | 23848986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23848986&pc=OZD | 11/20/2025 | 01/13/2026 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758857 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 81 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that 66-year-old male patient from a different facility was on impella cp support for cardiomyopathy and required escalation to 5.5. it was reported that after the impella cp was removed, after 5.5 insertion, and perclose tightened, distal pulses were absent on doppler post closure. a decision was made to cut down and repair with a patch. the cutdown was successful and revascularization was achieved with dacron patch. the surgeon stated that perclose placed at outside facility was deployed in the posterior wall of the common femoral artery. the patient survived the procedure. | [5 empty] | Investigation summary: ppae ( vascular injury): the cause of the injury was not determined due to insufficient clinical details. device history lot; device lot- 1978889. device history batch; subcomponent lot- n/a. device history review the pump sn (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2648 | 1220648-2025-49148 | 23849101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849101&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724742 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during support on impella cp the patient was positive for hematuria since pump placement and back to unit. position was confirmed with echocardiogram per nurse. mean arterial pressures in the high 90's-low 100's. stated no issues with foley placement. discussed troubleshooting options and medical doctor opted to leave as is. later the patient was successfully weaned and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ppae ( hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2649 | 1220648-2025-48590 | 23849135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849135&pc=OZD | 10/25/2025 | 11/25/2025 |  | 12/19/2025 |  |  | 10/25/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744011 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | Pericardial Effusion | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was placed in a 60 year old male patient for non-st-elevation myocardial infarction support. an echocardiogram showed the impella 5.3 centimeters from the aortic valve with a small pericardial effusion. the patient had a couple of runs with ventricular tachycardia (vt). the patient had six to ten beats of vt. lidocaine was being used. the next day the patient remained intubated and went into atrial fibrillation. a sternotomy was performed and left internal mammary artery taken down. hemashield platinum graft anastomosed to the aorta and tunneled out to the left supraclavicular space. an axillary sheath was placed and the graft was de-aired. the impella 5.5 was advanced under transesophageal echocardiogram guidance. once proper positioning was obtained, gradually flows increased. the doctor proceeded with off-pump coronary artery bypass grafting x 3. at the conclusion of the case, the graft was clamped, the peel-away sheath was removed, the graft was trimmed and secured to blue suture hub and secured blue suture hub to skin. the chest was closed and final measurement was 5.2 centimeters. the patient was sent to intensive care unit. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2650 | 1220648-2025-49149 | 23849417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849417&pc=OZD | 11/20/2025 | 12/18/2025 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757950 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 106 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 83-year-old male on impella 5.5 for support for acute myocardial infarction. during support, the patient had bleeding leading to the heparin being paused to get nose and throat bleeding under control. additionally, it was noted that the patient's white blood cell count was elevated requiring blood cultures. it was later reported that the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined due to insufficient clinical details. infection: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2651 | 1220648-2025-49153 | 23849493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849493&pc=OZD | 11/24/2025 | 12/19/2025 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743990 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 48-year-old male on impella 5.5 for cardiomyopathy and bi-ventricular failure. it was reported that during support, the console had a spike in motor current and a drop in flow. patient was on p6 flowing 2.2l with a motor current in the 250s. once the spike happened, impella flow dropped to 1.5l and motor current was still in the low to mid 300s. previously under insertion of rp, there was a clot visualized but then it went away. clot has not been visualized since. ptts have been in the 50s. additionally, the patient experienced atrial fibrillation and electrophysiologist was consulted. the patient survived the procedure. | Additional information states that low flow was incorrectly linked to this patient. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two associated devices for the reported event. the controller (aic) is addressed under manufacturer report number xxxx. the pump (impella 5.5) is addressed under manufacturer report number xxxx. | B5 additional information was received and h6 code a140508 was corrected to a24. | B1 product problem is no longer selected due to information that was received on the previously submitted report. b2 selection was revised as it was incorrect on the initial report that was submitted. b3 revised date due to event that happened on the originally reported date did not occur on this patient. e1 added zip code extension as it was omitted from the initial report that was submitted. g3 date on initial report was incorrect and should of reflected 24-nov-2025 due to event that happened on the originally reported date did not occur on this patient.. h6 code e0514 was removed as it did not occur on this patient. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. h10 related report number was added as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned. arrythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post-sterile inspection checks.
2652 | 1220648-2025-49150 | 23849513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849513&pc=PYX | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502640 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 53 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while on impella rp flex support the patient was having hemolysis when running impella at p9 as evidenced by hemolysis resolving when dropped to p7. device unable to be repositioned successfully. the patient required more flow than impella could deliver on p7, decision made to explant and switch patient to right ventricular assist device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was returned. hemolysis/mechanical interaction with blood: the pump was returned and evaluated. there was biomaterial found wrapped around the underside of the impeller and a cannula kink observed near the motor housing. no other abnormalities were found. no logs were returned. the cause of the hemolysis was not determined because the relation between the product findings and hemolysis was inconclusive without data logs. device in wrong position: the pump was returned and evaluated. there was a cannula kink observed near the motor housing. no logs were returned. the cause of the positioning issues was not determined since insufficient clinical details were provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2653 | 1220648-2025-49154 | 23849515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849515&pc=OZD | 11/20/2025 | 12/19/2025 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707479 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male on impella 5.5 for support for cardiomyopathy. it was reported that during support, the impella was unable to be repositioned. cvicu team confirmed they pressed and held the button and the catheter was fully stationery. the impella is deep requiring repositioning. however, the patient bridged to transplant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2654 | 1220648-2025-49155 | 23849532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849532&pc=OZD | 11/20/2025 | 12/19/2025 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728977 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 61 YR |  | Male | 107 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported on a 61 year old male on impella 5.5 for cardiovascular support. it was reported that the impella was not able to be delivered into the lv. the physician could only get the tip of the impella a few cm pass the anastomosis and as result, was unable to pass because of an issue at the anastomosis. physician tried multiple times but was unsuccessful. the procedure was aborted. the patient's outcome was unknown. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2655 | 1220648-2025-49151 | 23849545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849545&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664241 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. while on support it was reported that a hematoma was noted to right groin at insertion. the medical doctor aware, appears to not be getting worse. activated clotting time >400 and actively bleeding from chest tubes. multiple blood products given. also went on and off bypass multiple times making clotting factors affected. there was also no flow past the impella repositioning sheath. got left femoral arterial access to assess grafts. a 4fr catheter was used up and over, past the sheath. connected that to the extracorporeal membrane oxygenation circuit for distal perfusion. later the patient was successfully weaned off impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia : no flow past the impella repositioning sheath. unable to get reprofusion catheter in vessel distal to impella sheath. the root cause of the injury was unable to be determined due to insufficient clinical details. asae/hematoma : hematoma noted to right groin at insertion. act >400 and actively bleeding from chest tubes. multiple blood products given. the cause of the access site adverse event was use related as the acts were >400 and the hematoma was improved by reducing acts. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2656 | 1220648-2025-49152 | 23849555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849555&pc=OZD | 11/20/2025 | 12/19/2025 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251204 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported on a 48-year-old male patient on impella 5.5 for support due to cardiomyopathy. it was reported that the patient moved from cath lab table to transport gurney and a placement signal malfunctioned intermittently accompanied by intermittent console error alarm both occurred. both alarms resolved but intermittently returned during transport. impella flows and motor current remained consistent. the console was not swapped. there was no patient harm reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two associated devices for the reported event. the controller (aic) is addressed under manufacturer report number xxxx. the pump (impella 5.5) is addressed under manufacturer report number xxxx. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
2657 | 1220648-2025-49156 | 23849574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849574&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766999 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Decision to leave impella device in place. peel-away backed out and manual pressure held after device removed from body. repositioning sheath inserted into arteriotomy. despite angle matching and forward tension sutures placed, site still oozing. attempted to tighten pre-close perclose sutures around site without success. manual pressure held at site and fem stop placed over insertion site. activated clotting time still >250 but the doctor did not want protamine. hemostasis achieved with current measures. the doctor says this happens often when impella is placed into existing arteriotomy (device inserted into re-wired iabp site). the impella cp was explanted and the patient was placed on an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary; the device was returned for evaluation. returned complaint cp pump visually inspected, no abnormality found. major bleed: peel-away backed out and manual pressure held after device removed from body. repositioning sheath inserted into arteriotomy. manual pressure held at site and fem stop placed over insertion site. the cause of the access site adverse event was user related since activated clotting time (act) was >250 and hemostasis achieved when manual pressure held at site and fem stop placed over insertion site. device history review; this pump passed all post sterile inspection checks.
2658 | 1220648-2025-49157 | 23849578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849578&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614222 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was on impella 5.5 support when some bleeding from groin and right chest site was noted. heparin was stopped and one packed red blood cells were given to the patient. it was further reported that the patient was transferred to another facility a couple of weeks ago where the impella was exchanged for a durable ventricular assist device. | [5 empty] | E1 added zip code extension as it was omitted from the initial report that was submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. access site adverse event/major bleed: the cause of the bleed was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | Additional information received d6b: explantation day, month and year. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2659 | 1220648-2025-49159 | 23849586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849586&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. upon returning from cath lab a moderate size hematoma was felt by palpitation at medial side of insertion site. dressing was removed and no bleeding noted at site and no discoloration at time of dressing change. the doctor ordered fem-stop to be applied. once applied pulses still present. no issues. the patient was later weaned and survived. | [5 empty] | B2 revised selection as it was previously entered incorrectly on the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. access site adverse event/hematoma: a moderate size hematoma was observed at the insertion site. the cause of the hematoma cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2660 | 1220648-2025-49158 | 23849597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849597&pc=OZD | 11/20/2025 | 01/03/2026 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263049 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during prep, biomed stated that they are getting a system self check failed alarm when powering ic 4700. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
2661 | 1220648-2025-49161 | 23849609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849609&pc=OZD | 11/20/2025 | 12/19/2025 |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753188 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 58 KG | Non Hispanic |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported on a 58 year old male on impella cp support for cardiogenic shock. it was reported that the device was pulled across aortic valve upon echo assessment. the physician was unable to readvance across aortic valve and as a result, the device was removed. the patient's outcome is unknown. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted.
2662 | 1220648-2025-49160 | 23849621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849621&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695438 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 104 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while on impella cp support the patient's urine became bloody after extracorporeal membrane oxygenation removed. the following day the pump was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. ppae ( hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2663 | 1220648-2025-49162 | 23849935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23849935&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688386 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 55 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Hematuria | Thrombosis/Thrombus | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old female on an impella 5.5 for support during a high-risk cardiac procedure. the patient was also reported to have arhythmias and was cardioverted. it was reported the plasma free hemoglobin (pfhg) and hematuria in urine. patient is having bleeding into the foley with formed clots, resulting in a drop in hemoglobin. a mobile mass was noted on the catheter. additional information was provided stating that the patient had expired. | [5 empty] | Correction : e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis, arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed, hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2664 | 1220648-2025-49163 | 23850194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23850194&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702707 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported there was leaking from purge side arm of the impella cp. no further information provided. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. purge leak - pump: clinical details noted that leaking from the purge sidearm was noted. returned product was visually inspected and no damage to purge sidearm was observed. pump was purged with colored fluid for approximately 24 hours with no purge leak able to be reproduced. the cause of the purge leak - pump could not be determined as the issue was unable to be reproduced and no damage to the purge sidearm was observed on returned product. device history: pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2665 | 1220648-2025-49164 | 23850261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23850261&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026756855 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 85 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Difficult to Advance |  |  |  |  | Cardiac Arrest | Thrombosis/Thrombus | Heart Block | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient presented in shock after stent to left anterior descending artery at outside hospital. intra-aortic balloon pump placed at outside hospital and patient cannulated for extracorporeal membrane oxygenation (ecmo) after going into ventricular fibrillation. team decision for impella 5.5. thick heart muscle and small left ventricle cavity was noted. while on support more episodes of heart block was noted. team decision to place pacemaker bedside. it was further reported the patient coded and flows of both the ecmo and impella 5.5 were drastically reduced (impella to 0.1 lpm). team in the room manipulated the 5.5 and it came back across the aortic valve. flows of the ecmo returned to 4+ lpm. family was consulted and the decision was made to withdraw care. the patient expired. impella device was saved and will be returned. while debriefing case, surgeon noted that team reviewed the transesophageal echocardiogram after the case and identified a clot in patient's left ventricle. team suspects pump ingested the clot while patient was on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. device in wrong position: based on the clinical details provided that the pump was inadvertently pulled back into the aorta by the user, the cause of the device in wrong position was determined to be a use issue. difficult to advance: clinical details report that it was difficult to advance the pump through the anastomosis. it was noted that the patient had a small left ventricle. manual manipulation, a papaverine and 0.025" guidewire were used to resolve the issue, but no further information was provided. as noted above, there were no abnormalities noted on the returned catheter/cannula. since limited clinical details were provided and no defects were found on the returned pump, the cause of the difficulty to advance could not be determined. low pump flow: clinical details report that the patient coded around 07:00 on 20/nov and both ecmo and impella flows drastically reduced. the team attempted to reposition the pump and administer volume, but neither resolved the complication. the pump was inadvertently pulled back across the aortic valve and ecmo flows improved. post-explant, it was discussed that tee identified a clot in the patient's lv. data logs were reviewed and the report of low flows were confirmed. these signatures indicate that biomaterial was ingested by the pump, which resulted in a purge gap blockage and a reduction of hemodynamic support and deteriorating condition. returned product was investigated, and there were no defects noted on the product. there was, however, significant biomaterial in the outflow cage, near the impeller, and throughout the cannula. based on the signatures noted in data logs and low pump flows reproducing on returned product with biomaterial, the cause of the low pump flows was determined to be biomaterial ingestion. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest/heart block: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2666 | 1220648-2025-49165 | 23851470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23851470&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770684 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp and while on support the patient experienced limb ischemia on the impella access side. external fem performed and the patient was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2667 | 1220648-2025-49167 | 23853841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23853841&pc=OZD | 11/22/2025 |  |  | 12/19/2025 |  |  | 11/22/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700910 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp the pump was pulled into the aorta. multiple attempts to cross the valve were unsuccessful so a new pump was implanted. the patient experienced hypotension and a hematoma. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hematoma: the cause of the hematoma was not determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position into aorta when the doctor was trying to deploy a second pre close and maintain wire access. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2668 | 1220648-2025-49166 | 23853928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23853928&pc=OZD | 11/20/2025 |  |  | 12/19/2025 |  |  | 11/20/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692887 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 101 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella cp device and two days later experienced an arrhythmic event. the patient was thereafter escalated to an impella 5.5 device and during percutaneous coronary intervention (pci) with the impella support, the patient decompensated and expired. the patient was initially placed on an intra-aortic balloon pump (iabp) and was then transferred to another facility for the impella placement. iabp site, right femoral artery, was used for impella cp placement with no report of complications. day 2 update noted, per the nurse, the patient required two shocks for atrial flutter. shortly thereafter the patient was escalated to an impella 5.5 and planned for pci (as the patient was denied coronary artery bypass graft surgery). the following day, the patient's status was noted to be "significantly" improving. the next day, the patient was on schedule for the pci and status noted as still improving. however, during the pci with impella 5.5 support, the patient decompensated going into pulseless electrical activity (pea) and then scores of ventricular rhythms before code was called. additional information noted the following: the cause of death is noted as asystole after balloon angioplasty due to severe heart failure. the physician does not allege the impella caused the event. the physician was worried about the patient being a candidate for pci and had even upgraded the patient to a higher flow impella 5.5. both impella devices were in for approximately a week and there were no rhythm disturbances nor any asystolic events. the asystolic event coincided with the angioplasty of the left anterior descending artery. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the arrhythmia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | This report is being submitted to provide medical safety/clinical review no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review: medical safety/clinical reviewed the event. the patient presented in cardiogenic shock (scai stage e) secondary to acute myocardial infarction. an intra-aortic balloon pump (iabp) was placed, and the patient was transferred to another facility for impella support. the right femoral artery iabp access site was utilized for placement of an impella cp device. no complications related to placement were reported. on day two, the nurse noted the patient required two electrical cardioversions for atrial flutter. subsequently, the patient was escalated to an impella 5.5 device, and percutaneous coronary intervention (pci) was planned; coronary artery bypass grafting (cabg) was not pursued. on (b)(6), the patient¿s clinical status was documented as significantly improved. on (b)(6), the patient remained scheduled for percutaneous coronary intervention (pci) with continued improvement noted. during pci performed with impella 5.5 support, the patient acutely decompensated, developing pulseless electrical activity followed by multiple ventricular arrhythmias. a code was initiated. the patient expired, with the cause of death documented as asystole following balloon angioplasty in the setting of severe heart failure. the treating physician did not attribute the event to the impella devices. the death will be conservatively reported for the impella 5.5 device as it was in use at the time of expiration. note: the impella cp is a serious injury type of reportable event. the event story involves 2 product complaints: impella cp - mdr 1220648-2025-49146: serious injury. impella 5.5 - mdr 1220648-2025-49166: death.
2669 | 1220648-2025-49170 | 23855488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23855488&pc=OZD | 07/26/2025 |  |  | 12/19/2025 |  |  | 07/26/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698641 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that access site bleeding occurred while obtaining access to the peel away sheath for impella cp implantation. the needle and wire were then removed. proglide sutures were tightened but the bleeding did not resolve. it was suspected that the bleeding was coming from a crack in the peel away sheath. the peel away sheath was removed, a repo sheath was advanced, and the sutures were tightened to obtain hemostasis. the patient expired. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2670 | 1220648-2025-49171 | 23855496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23855496&pc=OZD | 08/07/2025 |  |  | 12/19/2025 |  |  | 08/07/2025 | 12/19/2025 |  |  |  |  | 12/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694971 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male | 32 KG | Non Hispanic | White | Injury | Yes | No | Break | Restricted Flow rate | Patient-Device Incompatibility | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  | Anemia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Inflammation | Renal Failure | Sepsis | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced hemolysis and anemia. the patient was transfused one unit of packed red blood cells. due to suction alarms and decreased output the pump was repositioned and the p level of the pump was increased. suction events occurred and the pump was repositioned again. the suction events did not resolve. during repositioning of the device, the sheath broke. the patient then exhibited an increased white blood cell count and positive blood culture and drainage results. antibiotics were administered. a skin defect was observed at the access site. a wound vac was placed in addition to irrigation and partial graft removal. the patient also developed endocardidtis and vegetation. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2671 | 1220648-2025-49131 | 23837113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23837113&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1455741 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while using an impella 5.5, the position waveform and numerical values suddenly stopped being displayed and a "controller error (yellow)" alarm occurred. while preparing to replace the controller, the display returned to normal, but the controller was replaced just to be safe. there have been no problems since then with operation. there was no reported patient harm. | [5 empty] | Additional information was received and note the impella pump that was connected to the automated impella controller (aic) at the time of the aic complaint was eventually explanted approximately 36 days later as a bridge to another impella with a survived patient outcome. note: h6 investigation coding previously reported remains unchanged. patient status update only with no impact to the investigation outcome. | Investigation summary: the investigation has been completed. data review: console logs were consistent with what was reported in the complaint. the logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. real time logs confirmed that all signals received from optical bench (placement, signal to noise ration, contrast, lamp, gain) are represented as -16384 values due to the crc error. device analysis: the console was sent back for further investigation. the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller software operated as designed. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of console.
2672 | 1220648-2025-49132 | 23837239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23837239&pc=OZD | 08/25/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026715965 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Obstruction/Occlusion | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received via long-term outcome and quality indicator (loqi) registry that a patient had acute femoral artery access site occlusion with relationship to an impella pump. left femoral artery access site bleeding was noted. loqi states: "after removal of impella device and deployment of 2x perclose and 1x 6 fr angioseal, angiography of the l cfa showed acute vessel closure. initially, an antegrade approach was taken via the existing l radial access, with unsuccessful crossing of the stenosed lesion. a l sfa retrograde arteriotomy was performed to cross the lesion and initial balloon dilatation, then an absolute pro 8.0x60mm self-expanding stent was placed across the lesion in cfa. ivus was utilized to assess lesion, and to optimize stent length. balloon tamponade of the l sfa arteriotomy site was attempted for a total duration of 20 minutes. however, bleeding persisted. thus, a lifestream 6.0x26mm covered stent was deployed across the arteriotomy site, followed by an absolute pro 8.0x60mm self-expanding stent inside the covered stent. patient received 3 u prbc. hemoglobin baseline 15.1 nadir 6.5." | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. arterial occlusion: the cause of the injury was not determined due to insufficient clinical details. asae / major bleed: left femoral artery access site bleeding was noted after impella removal and closure with 2 perclose devices and a 6 fr angioseal. angiography revealed acute cfa occlusion, requiring both antegrade and retrograde attempts to cross the lesion. a self-expanding stent was placed, but bleeding from the sfa arteriotomy persisted despite 20 minutes of balloon tamponade. a covered stent followed by an additional self-expanding stent was deployed to control the bleeding. the patient received 3 units of prbc, with hemoglobin decreasing from 15.1 to 6.5. the cause of the access site adverse event was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2673 | 1220648-2025-49133 | 23837534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23837534&pc=OZD | 11/21/2025 | 12/18/2025 |  | 12/18/2025 |  |  | 11/21/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026715438 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male | 80 KG |  | White | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that during the procedure it was noticed that the pump was deep into the left ventricle but no alarm occurred. the procedure went well and the patient was weaned after protected percutaneous coronary intervention. | [5 empty] | The device and data logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding device in wrong position due to lack of clinical information, it cannot be established how the pump became deep in the lv. optical sensor issue due to lack of data logs and product returned, the root cause of the optical sensor issue cannot be established as to why there was no placement alarm that triggered. e1 added initial reporter title as it was omitted from the initial report that was submitted. | D6.a. and d6.b.: the device was implanted and explanted in (b)(6) 2025. a follow up is being conducted to confirm the exact dates. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2674 | 1220648-2025-49135 | 23837763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23837763&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719403 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced bleeding at the impella access site in the right femoral artery after the peel away sheath was removed. the physician decided to remove the pump. the patient was then implanted with an additional impella cp via the left femoral artery. it was noted that the second impella was removed due to the patient experiencing ventricular fibrillation. it was noted that the patient¿s outcome at the time of explant was survived. no further information was provided or obtained. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). 14 french introducer, used with pump 1, with lot number s9557231. pump 2: impella cp, with serial number (b)(6). this report specifically addresses pump 1: impella cp, with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Investigation summary access site adverse event (major bleed): the product was not returned for investigation. data log review was not required for this case. clinical details indicate that the patient experienced bleeding issues after the peel-away sheath was removed, leading the physician to remove the pump on the same day. the cause of the access site bleeding issue was not determined without returned product investigation and sufficient clinical details. device history review the pump (b)(6) passed all post sterile inspection checks.
2675 | 1220648-2025-49136 | 23837850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23837850&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716286 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was admitted in cardiogenic shock. the patient presented to the operating room for exchange to an impella 5.5 for additional hemodynamic support. while on support, heparin was stopped due to oozing at the access site around the sheath. later, the patient was successfully weaned and bridged to a durable left ventricular assist device in the operating room. the patient's outcome at explant was noted as survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the asae/ major bleed: heparin stopped d/t oozing. the cause of the access site adverse event was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2676 | 1220648-2025-49138 | 23838159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23838159&pc=OZD | 11/19/2025 |  |  | 12/18/2025 |  |  | 11/19/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. it was reported that there was oozing at groin site and heparin was put on hold. the patient remained on support for 3 more days until successful explanation. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. additional information was received: the customer does not believe that the device was returned and does not require a return kit. patient had a biliary stent that was occluded leading to liver issues. device was successfully explanted and customer felt impella adequately achieved intended use. investigation summary the device was not returned for evaluation. the cause of the access site adverse event was patient related as patient has coagulopathy. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2677 | 1220648-2025-49139 | 23838359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23838359&pc=OZD | 10/13/2025 |  |  | 12/18/2025 |  |  | 11/20/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026715018 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female | 90 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with impella devices (indication unknown) for mechanical circulatory support (mcs). the patient underwent a mitral clip procedure and subsequently required mechanical circulatory support. an impella cp device was placed post-procedure. due to worsening right heart failure, an impella rp device was attempted 24 hours later. the initial rp insertion via the left femoral vein was unsuccessful due to tortuous venous anatomy; the device was removed. a second rp device was successfully placed in the left femoral vein. at the conclusion of the procedure, the physician noted a split at the tip of the peel-away sheath; no groin-related adverse event occurred. the following day, the impella rp device generated a ¿placement signal not reliable¿ alarm thought to be related to insertion technique. there was no impact on the pump function. however, the next day, the impella rp device was explanted (for reason unknown), and the patient remained on impella cp mcs for approximately 1.5 weeks. the patient¿s condition deteriorated resulting in systemic inflammatory response syndrome (sirs) and subsequent demise. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. medical safety review. medical safety reviewed the event and noted the following: a patient underwent a mitral clip procedure and subsequently required mechanical circulatory support. an impella cp device was placed post-procedure. due to worsening right heart failure, an impella rp device was attempted 24 hours later. the initial rp insertion via the left femoral vein was unsuccessful due to tortuous venous anatomy; the device was removed. a second rp device was successfully placed in the left femoral vein. at the conclusion of the procedure, the physician noted a split at the tip of the peel-away sheath; no groin-related adverse event occurred. on the following day, the rp device generated a ¿placement signal not reliable¿ alarm, suspected to be related to insertion technique. pump function was not affected. the rp device was explanted the next day, and the patient remained on impella cp support for approximately 1.5 weeks. despite ongoing support, the patient¿s clinical condition deteriorated, resulting in systemic inflammatory response syndrome (sirs) and eventual death. the initial unsuccessful rp insertion and the split at the introducer tip are considered device malfunctions. the ¿placement signal not reliable¿ alarm on the second rp device is also considered a device malfunction, despite no impact on pump function. the patient¿s death is conservatively attributed to the impella cp device, as it was in place at the time of death; however, the patient¿s worsening clinical condition likely contributed significantly to the outcome. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device report is one of four devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp pump 2, impella rp pump 3 and impella introducer.
2678 | 1220648-2025-49140 | 23838636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23838636&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756537 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was on impella cp and after about 20 minutes after getting to back to intensive care unit, the patient went laminar, mean arterial pressure dropped, hypotension and bradycardia. atropine was given and possible vagal episode. the patient recovered. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. bradycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2679 | 1220648-2025-49141 | 23839680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23839680&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769544 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 73 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was presented with chest pain, cough, shortness of breath, dyspnea on exertion, and vomiting diagnosed as non-st-elevation myocardial infarction and treated with antibiotics for possible pneumonia. left heart catheterization showed severe 90% distal left main coronary artery extending to left anterior descending artery. ejection fraction of 35-40%. the patient was hypotensive during left heart catheterization and required levophed and was sent to the intensive care unit. overnight the patient required bipap for respiratory support and was brought down for primary percutaneous coronary intervention of left main coronary artery and left circumflex artery with impella cp to support the patient. it was reported that there was a hematoma formed at access site, manual pressure was held and required peripheral stent which was successful. blood pressure was soft and noted a hematoma on right anterior chest not relieved by in vitro fertilization bolus and manual pressure. right radial access was obtained and the patient required 10 millimeters covered stent to stop extravasation which the physician assessed to be successful based on post-intervention angiogram. end-diastolic pressure measured at this time was three. blood pressure recovered to systolic blood pressure greater than 90. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. hypotension: noted hematoma on right anterior chest not relieved by intravenous fluid (ivf) bolus and manual pressure. the cause of the injury was not determined due to insufficient clinical details. hematoma: hematoma formed at access site, manual pressure and required peripheral stent which was successful. the cause of the access site adverse event was not determined due to insufficient clinical details. device history this pump passed all post sterile inspection checks | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2680 | 1220648-2025-49137 | 23839774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23839774&pc=OZD | 11/19/2025 |  |  | 12/18/2025 |  |  | 11/19/2025 | 12/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753383 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Crack | Decrease in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was in cardiogenic shock and was implanted with an impella cp device for mechanical circulatory support. during support it was reported by the medical doctor there was a hairline crack between the purge filter and the pressure relief valve, causing the pure pressure to fall off. they recommended replacing the pump. later, it was reported there was a hematoma at the groin site. no information about the further therapy. additional information was received. the impella cp showed a defect in the plastic of the purge side arm which required s temporary purge bypass to be set up. the impella was explanted. regarding the hematoma, this was noted to be a user error. the impella had been sutured flat against the thigh. the dressing was renewed and the sheath was correctly repositioned which resolved the bleeding. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Complaint coding was updated to better reflect the reported event. consequently, section h6 medical device problem codes were updated/corrected and repopulated accordingly.
2681 | 1220648-2025-49142 | 23840650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23840650&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667316 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. the patient presented to the emergency department with palpitations, nausea, vomiting, and progressive exertional dyspnea. upon arrival, the patient was found to be in severe hyperglycemia with a of glucose less than500 mg/dl and exhibited tachycardia, tachypnea, and hypoxia requiring emergency intubation. laboratory and diagnostic testing revealed significant a troponin result of greater than 60,000 and an electrocardiogram showing st-elevation. a subsequent left heart catheterization (lhc) identified severe multi-vessel disease, including severe calcified distal left main (lm) bifurcation involving the ostial circumflex with 99% stenosis and ostial left anterior descending with 75% stenosis, with an elevated left ventricular end-diastolic pressure (lvedp) of 30 mmhg. an intra-aortic balloon pump (iabp) was placed for hemodynamic support, and the patient was transferred to the complainant hospital for further treatment. although briefly extubated and alert, the patient experienced decompensation on the morning of the day of impella implantation, following new-onset arrhythmias and chest pain, requiring re-intubation and emergency stabilization. the medical team decided to implant an impella 5.5 to perform hrpci. the impella 5.5 was prepped, inserted via the right axillary artery, and support was initiated without complications. the hrpci was successfully performed, and the iabp was removed. the impella position was imaged and found to be too deep, and the physician successfully repositioned the pump. the patient was noted to be in stable condition and was transferred to the cardiac care unit for rest and recovery. on support day two, the patient required an extra suture at the impella access site. the patient was transfused with one unit of packed red blood cells and 1.5 liters of albumin. additionally, heparin was withheld until the patient¿s hemoglobin stabilized. the medical team stated that there was previous concern for bleeding that was unrelated to the impella 5.5 placement, and a gastrointestinal consultation was pending. on support day three, the impella site continued to ooze and required redressing. on support day four, the medical team attempted to wean the patient to performance (p) level 4 from p5, however the patient became hypotensive. the patient received two additional units of blood the previous night and was started on antibiotics for an elevated temperature of 38.5 degrees celsius. the patient had a unit of platelets ordered, and it was noted that a heparin induced thrombocytopenia test was ordered and resulted as negative. on support day five, the patient was noted to have bruising and hematoma at the impella access site. additionally, it was noted that the patient was extubated. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2682 | 1220648-2025-49145 | 23842732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23842732&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722743 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 64 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. the patient did not have pre-existing ep issues per the associate clinical consultant's knowledge. | It was reported that a patient was scheduled for coronary artery bypass graft but was canceled due to porcelain aorta. primary percutaneous coronary intervention was elected with a plan to fix the left anterior descending (lad) artery, right coronary artery, and chronic total occlusion. the patient was prepped in sterile fashion. best practices were followed to access the right femoral artery and insert an impella into the left ventricle maintaining p-3 until dampening occurred. p-level was increased to p-7 while shockwaving the left circumflex artery (lcx) and lad artery. diabetic ketoacidosis crush to the lad and lcx, ballooning to the lymphatic malformation, and the patient went into ventricular fibrillation (vf) on three occasions requiring defibrillation. the patient maintained airway and awareness after being shocked but the impella supported the patient during events of the procedure. lactic was 1.4 after the procedure and vital signs were stable without recurrent arrhythmias so the doctor removed the impella per best practices. the doctor reported predisposition to vf because of coronary disease. |  |  |  |  | Investigation summary: the investigation has been completed. peri-procedural adverse event (ventricular fibrillation): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2683 | 1220648-2025-49144 | 23842960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23842960&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660971 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 86 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient presented to an outside hospital with chest pain and angina. the patient was found to have non-st-elevation myocardial infarction and ejection fraction less than 15%. left heart catheterization found mitral valve disease with left ventricular end-diastolic pressure of 34. the patient was transferred to another hospital for advanced care. heart failure consult and the decision was made to place an impella 5.5 as a bridge to a possible transplant. while on support, intermittent suction was reported, but pressures were stable. it was also reported that the patient had 10 minutes of ventricular tachycardia. the patient was reported to have bloody stools which resulted in blood transfusion and heparin being paused. the patient was later successfully explanted and survived. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> the pump was not returned for investigation. data log was not provided and was not found on the remote server. pump suction: the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data log review. clinical symptoms (ventricular tachycardia): the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (major bleed): the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device batch: 1919490. device history review ==> the pump sn(b)(6) passed all post sterile inspection checks. | H6 medical device problem code: a24 was erroneously entered on the initial manufacturer device report number. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2684 | 1220648-2025-49146 | 23845551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23845551&pc=OZD | 11/18/2025 |  |  | 12/18/2025 |  |  | 11/18/2025 | 12/18/2025 |  |  |  |  | 12/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724597 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella cp device and two days later experienced an arrhythmic event. the patient was thereafter escalated to an impella 5.5 device and during percutaneous coronary intervention (pci) with the impella support, the patient decompensated and expired. the patient was initially placed on an intra-aortic balloon pump (iabp) and was then transferred to another facility for the impella placement. iabp site, right femoral artery, was used for impella cp placement with no report of complications. day 2 update noted, per the nurse, the patient required two shocks for atrial flutter. shortly thereafter the patient was escalated to an impella 5.5 and planned for pci (as the patient was denied coronary artery bypass graft surgery). the following day, the patient's status was noted to be "significantly" improving. the next day, the patient was on schedule for the pci and status noted as still improving. however, during the pci with impella 5.5 support, the patient decompensated going into pulseless electrical activity (pea) and then scores of ventricular rhythms before code was called. no further information surrounding the event is known. | [5 empty] | This report is being submitted to provide medical safety/clinical review no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review. medical safety/clinical reviewed the event. the patient presented in cardiogenic shock (scai stage e) secondary to acute myocardial infarction. an intra-aortic balloon pump (iabp) was placed, and the patient was transferred to another facility for impella support. the right femoral artery iabp access site was utilized for placement of an impella cp device. no complications related to placement were reported. on day two, the nurse noted the patient required two electrical cardioversions for atrial flutter. subsequently, the patient was escalated to an impella 5.5 device, and percutaneous coronary intervention (pci) was planned; coronary artery bypass grafting (cabg) was not pursued. on 19-nov, the patient¿s clinical status was documented as significantly improved. on 20-nov, the patient remained scheduled for percutaneous coronary intervention (pci) with continued improvement noted. during pci performed with impella 5.5 support, the patient acutely decompensated, developing pulseless electrical activity followed by multiple ventricular arrhythmias. a code was initiated. the patient expired, with the cause of death documented as asystole following balloon angioplasty in the setting of severe heart failure. the treating physician did not attribute the event to the impella devices. the death will be conservatively reported for the impella 5.5 device as it was in use at the time of expiration. note: the impella cp is a serious injury type of reportable event. the event story involves 2 product complaints. impella cp - mdr 1220648-2025-49146: serious injury. impella 5.5 - mdr 1220648-2025-49166: death. | Additional information received noted the following: the cause of death is noted as asystole after balloon angioplasty due to severe heart failure. the physician does not allege the impella caused the event. the physician was worried about the patient being a candidate for pci and had even upgraded the patient to a higher flow impella 5.5. both impella devices were in for approximately a week and there were no rhythm disturbances nor any asystolic events. the asystolic event coincided with the angioplasty of the left anterior descending artery. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the arrhythmia, the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome. | Patient-specific information a4. weight is unknown at the time of this mdr writing. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device report is one of two devices associated with the event. the manufacturer report number for the related impella 5.5 device report, which is associated with the demise outcome, will be shared in a supplemental report.
2685 | 1220648-2025-49081 | 23826472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23826472&pc=OZD | 11/29/2025 | 01/12/2026 |  | 12/17/2025 |  |  | 11/29/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1493604 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Unknown | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in was implanted with an impella 5.5 device for mechanical circulatory support as an escalation from an impella cp and to bridge to a coronary artery bypass graft. the patient presented to the hospital with a two-week history of shortness of breath and an associated episode of chest pain. diagnostic testing upon admission revealed elevated troponin and b-type natriuretic peptide levels and a subsequent cardiac catheterization was performed and right heart catheterization identified severe distal left main (lm) coronary artery disease. due to the severity of the findings and the patient's clinical instability, a "shock call" was initiated and the decision was made to place impella cp support as a bridge to transport to another hospital for a planned escalation to impella 5.5 and coronary artery bypass graft. the impella cp was prepped, inserted via the right femoral artery, and support was initiated with flows at 3.5 liters per minute without complications. the patient was stabilized and prepped for transfer to another hospital for advanced treatment. the patient was successfully transferred to the outside hospital on support day two, and the patient was noted to be in stable condition with some shortness of breath. there was noted to be oozing at the impella access site and as a result of the oozing, heparin was discontinued. the patient continued to decompensate on the impella cp and was taken later that morning to the operating room for impella cp removal and an escalation to impella 5.5. the impella 5.5 was prepped, inserted via the right axillary artery, and support was initiated as the impella cp was pulled across the aortic valve, support discontinued and removed. during the implant of the impella 5.5, it was reported there was a controller error on the automated impella controller (aic). support from the new impella 5.5 was not compromised due to the error, however the decision was made to replace the console. after the back up unit was obtained, it was noted that the pump flows and purge systems were running, and the controller alarm resolved, however there was a loss of placement signal and motor current. the aic was replaced with the back up unit, and the issues were resolved. the impella cp access site was surgically repaired for closure. the patient was transported to the intensive care unit for rest and recovery. on impella 5.5 support day three, it was noted that the patient was believed to be in septic shock and cardiogenic shock. the patient had a fever of 38 celsius and was placed on a cooling blanket. blood cultures and a sputum sample were sent for bacterial growth. the patient was noted to be critically ill. on support day four, the patient remained febrile. the patient¿s care plan was to stabilize the patient until a percutaneous coronary intervention (pci) could be performed. on support day six, the patient¿s urine was noted to be pink tinged. the patient was noted to be in stable condition. on support day seven, the medical team continued to have concern for the patient having sepsis. the medical team did not believe that the impella access site was infected. a plan was made to start the patient on antibiotics and to pull more cultures to identify the pathogen. there was additionally a concern for an acute kidney injury as the patient¿s creatinine levels were increasing. on support day eight, it was noted that the patient received multiple fluid boluses for low central venous pressure and suction alarms. on support day nine, it was noted that the patient had some episodes of ventricular tachycardia when the patient coughed or moved overnight. later that day, the patient had a successful pci. following the pci the patient remained on impella 5.5 support for rest and recovery. on support day 10, the patient remained febrile with unclear source of infection. on support day 11, the concern for infection continued. the patient received a whole body computed tomography scan to attempt to find a source of infection. the patient remained on antibiotics, and of note the impella site was without infection concern. on support day 13, it was noted that the patient was not responding to neurological stimuli. it was reported that the aic had an alarm for ¿unexpected controller shut down¿ which self-resolved. a back up aic was placed on standby, and the aic was monitored. on support day 17, the patient had bloody drainage from a nasogastric tube and had slight hematuria. the medical team was advised to verify placement of the impella 5.5 with ultrasound. the patient continued to not follow commands but was noted to move spontaneously. on support day 18, a bedside ultrasound revealed the impella had migrated, requiring the physician to reposition the device by pulling it back 2cm into the correct position. the patient¿s hematuria was noted to be clearing; while the cause is unconfirmed, it may be related to foley catheter trauma, as plasma free hemoglobin levels were within normal limits. additionally, a large clot was identified during an upper endoscopy and removed. the heparin in the impella purge solution was discontinued, and the patient was extubated to a high flow nasal cannula. on support day 19, the patient experienced flash pulmonary edema and required emergency intubation. the planned wean of the impella 5.5 was stopped, and the performance level was increased to p5 from p2. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves four impella devices: pump 1: impella cp, with serial number (b)(6), pump 2: impella 5.5, with serial number (b)(6), automated impella controller 1 with serial number (b)(6), automated impella controller 2 with serial number (b)(6). this report specifically addresses automated impella controller 2 with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h10 for the related report numbers. | This complaint involves four impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). automated impella controller 1 with serial number (b)(6). automated impella controller 2 with serial number (b)(6). this report specifically addresses automated impella controller 2 with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h10 for the related report numbers. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | B5 corrected to identify the proper subject device. the investigation is still ongoing.
2686 | 1220648-2025-49087 | 23826569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23826569&pc=OZD | 11/19/2025 |  |  | 12/17/2025 |  |  | 11/19/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026764678 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | Sepsis | Tachycardia | Hematuria | Thrombosis/Thrombus | Renal Impairment | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in was implanted with an impella 5.5 device for mechanical circulatory support as an escalation from an impella cp and to bridge to a coronary artery bypass graft. the patient presented to the hospital with a two-week history of shortness of breath and an associated episode of chest pain. diagnostic testing upon admission revealed elevated troponin and b-type natriuretic peptide levels and a subsequent cardiac catheterization was performed and right heart catheterization identified severe distal left main (lm) coronary artery disease. due to the severity of the findings and the patient's clinical instability, a "shock call" was initiated and the decision was made to place impella cp support as a bridge to transport to another hospital for a planned escalation to impella 5.5 and coronary artery bypass graft. the impella cp was prepped, inserted via the right femoral artery, and support was initiated with flows at 3.5 liters per minute without complications. the patient was stabilized and prepped for transfer to another hospital for advanced treatment. the patient was successfully transferred to the outside hospital on support day two, and the patient was noted to be in stable condition with some shortness of breath. there was noted to be oozing at the impella access site and as a result of the oozing, heparin was discontinued. the patient continued to decompensate on the impella cp and was taken later that morning to the operating room for impella cp removal and an escalation to impella 5.5. the impella 5.5 was prepped, inserted via the right axillary artery, and support was initiated as the impella cp was pulled across the aortic valve, support discontinued and removed. during the implant of the impella 5.5, it was reported there was a controller error on the automated impella controller (aic). support from the new impella 5.5 was not compromised due to the error, however the decision was made to replace the console. after the back up unit was obtained, it was noted that the pump flows and purge systems were running, and the controller alarm resolved, however there was a loss of placement signal and motor current. the aic was replaced with the back up unit, and the issues were resolved. the impella cp access site was surgically repaired for closure. the patient was transported to the intensive care unit for rest and recovery. on impella 5.5 support day three, it was noted that the patient was believed to be in septic shock and cardiogenic shock. the patient had a fever of 38 celsius and was placed on a cooling blanket. blood cultures and a sputum sample were sent for bacterial growth. the patient was noted to be critically ill. on support day four, the patient remained febrile. the patient¿s care plan was to stabilize the patient until a percutaneous coronary intervention (pci) could be performed. on support day six, the patient¿s urine was noted to be pink tinged. the patient was noted to be in stable condition. on support day seven, the medical team continued to have concern for the patient having sepsis. the medical team did not believe that the impella access site was infected. a plan was made to start the patient on antibiotics and to pull more cultures to identify the pathogen. there was additionally a concern for an acute kidney injury as the patient¿s creatinine levels were increasing. on support day eight, it was noted that the patient received multiple fluid boluses for low central venous pressure and suction alarms. on support day nine, it was noted that the patient had some episodes of ventricular tachycardia when the patient coughed or moved overnight. later that day, the patient had a successful pci. following the pci the patient remained on impella 5.5 support for rest and recovery. on support day 10, the patient remained febrile with unclear source of infection. on support day 11, the concern for infection continued. the patient received a whole body computed tomography scan to attempt to find a source of infection. the patient remained on antibiotics, and of note the impella site was without infection concern. on support day 13, it was noted that the patient was not responding to neurological stimuli. it was reported that the aic had an alarm for ¿unexpected controller shut down¿ which self-resolved. a back up aic was placed on standby, and the aic was monitored. on support day 17, the patient had bloody drainage from a nasogastric tube and had slight hematuria. the medical team was advised to verify placement of the impella 5.5 with ultrasound. the patient continued to not follow commands but was noted to move spontaneously. on support day 18, a bedside ultrasound revealed the impella had migrated, requiring the physician to reposition the device by pulling it back two centimeters into the correct position. the patient¿s hematuria was noted to be clearing; while the cause is unconfirmed, it may be related to foley catheter trauma, as plasma free hemoglobin levels were within normal limits. additionally, a large clot was identified during an upper endoscopy and removed. the heparin in the impella purge solution was discontinued, and the patient was extubated to a high flow nasal cannula. on support day 19, the patient experienced flash pulmonary edema and required emergency intubation. the planned wean of the impella 5.5 was stopped, and the performance level was increased to p5 from p2. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves four impella devices: pump 1: impella cp, with serial number (b)(6) pump 2: impella 5.5, with serial number (b)(6) automated impella controller 1 with serial number (b)(6) automated impella controller 2 with serial number (b)(6) this report specifically addresses pump 2: impella 5.5, with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h10 for the related report numbers. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings, ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
2687 | 1220648-2025-49088 | 23826628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23826628&pc=OZD | 11/17/2025 |  |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322586 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in was implanted with an impella 5.5 device for mechanical circulatory support as an escalation from an impella cp and, to bridge to a coronary artery bypass graft. the patient presented to the hospital with a two-week history of shortness of breath and an associated episode of chest pain. diagnostic testing upon admission revealed elevated troponin and b-type natriuretic peptide levels and a subsequent cardiac catheterization was performed and right heart catheterization identified severe distal left main (lm) coronary artery disease. due to the severity of the findings and the patient's clinical instability, a "shock call" was initiated and the decision was made to place impella cp, support as a bridge to transport to another hospital for a planned escalation to impella 5.5 and coronary artery bypass graft. the impella cp was prepped, inserted via the right femoral artery, and support was initiated with flows at 3.5 liters per minute without complications. the patient was stabilized and prepped for transfer to another hospital for advanced treatment. the patient was successfully transferred to the outside hospital on support day two, and the patient was noted to be in stable condition with some shortness of breath. there was noted to be oozing at the impella access site and as a result of the oozing, heparin was discontinued. the patient continued to decompensate on the impella cp and was taken later that morning to the operating room for impella cp removal and an escalation to impella 5.5. the impella 5.5 was prepped, inserted via the right axillary artery, and support was initiated as the impella cp was pulled across the aortic valve, support discontinued and removed. during the implant of the impella 5.5, it was reported there was a controller error on the automated impella controller (aic). support from the new impella 5.5 was not compromised due to the error. however, the decision was made to replace the console. after the back up unit was obtained, it was noted that the pump flows and purge systems were running, and the controller alarm resolved. however, there was a loss of placement signal and motor current. the aic was replaced with the back up unit, and the issues were resolved. the impella cp access site was surgically repaired for closure. the patient was transported to the intensive care unit for rest and recovery. on impella 5.5 support day three, it was noted that the patient was believed to be in septic shock and cardiogenic shock. the patient had a fever of 38 celsius and was placed on a cooling blanket. blood cultures and a sputum sample were sent for bacterial growth. the patient was noted to be critically ill. on support day four, the patient remained febrile. the patient¿s care plan was to stabilize the patient until a percutaneous coronary intervention (pci) could be performed. on support day six, the patient¿s urine was noted to be pink tinged. the patient was noted to be in stable condition. on support day seven, the medical team continued to have concern for the patient having sepsis. the medical team did not believe that the impella access site was infected. a plan was made to start the patient on antibiotics and to pull more cultures to identify the pathogen. there was additionally a concern for an acute kidney injury as the patient¿s creatinine levels were increasing. on support day eight, it was noted that the patient received multiple fluid boluses for low central venous pressure and suction alarms. on support day nine, it was noted that the patient had some episodes of ventricular tachycardia when the patient coughed or moved overnight. later that day, the patient had a successful pci. following the pci the patient remained on impella 5.5 support for rest and recovery. on support day 10, the patient remained febrile with unclear source of infection. on support day 11, the concern for infection continued. the patient received a whole body computed tomography scan to attempt to find a source of infection. the patient remained on antibiotics, and of note the impella site was without infection concern. on support day 13, it was noted that the patient was not responding to neurological stimuli. it was reported that the aic had an alarm for ¿unexpected controller shut down¿ which self-resolved. a back up aic was placed on standby, and the aic was monitored. on support day 17, the patient had bloody drainage from a nasogastric tube and had slight hematuria. the medical team was advised to verify placement of the impella 5.5 with ultrasound. the patient continued to not follow commands but was noted to move spontaneously. on support day 18, a bedside ultrasound revealed the impella had migrated, requiring the physician to reposition the device by pulling it back two centimeters into the correct position. the patient¿s hematuria was noted to be clearing; while the cause is unconfirmed, it may be related to foley catheter trauma, as plasma free hemoglobin levels were within normal limits. additionally, a large clot was identified during an upper endoscopy and removed. the heparin in the impella purge solution was discontinued, and the patient was extubated to a high flow nasal cannula. on support day 19, the patient experienced flash pulmonary edema and required emergency intubation. the planned wean of the impella 5.5 was stopped, and the performance level was increased to p5 from p2. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves four impella devices: pump 1: impella cp, with (b)(6) pump 2: impella 5.5, with serial number (b)(6) automated impella controller 1 with serial number (b)(6) automated impella controller 2 with serial number (b)(6) this report specifically addresses automated impella controller 1 with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h10 for the related report numbers. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2688 | 1220648-2025-49089 | 23826707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23826707&pc=OZD | 11/17/2025 |  |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767844 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in was implanted with an impella cp device for mechanical circulatory support as a bridge to an impella 5.5 and coronary artery bypass graft. the patient presented to the hospital with a two-week history of shortness of breath and an associated episode of chest pain. diagnostic testing upon admission revealed elevated troponin and bnp levels and a subsequent cardiac catheterization was performed and right heart catheterization identified severe distal left main (lm) coronary artery disease. due to the severity of the findings and the patient's clinical instability, a "shock call" was initiated and the decision was made to place impella cp support as a bridge to transport to another hospital for a planned escalation to impella 5.5 and coronary artery bypass graft. the impella cp was prepped, inserted via the right femoral artery, and support was initiated with flows at 3.5 liters per minute without complications. the patient was stabilized and prepped for transfer to another hospital for advanced treatment. the patient was successfully transferred to the outside hospital on support day two, and the patient was noted to be in stable condition with some shortness of breath. there was noted to be oozing at the impella access site and as a result of the oozing, heparin was discontinued. the patient continued to decompensate on the impella cp and was taken later that morning to the operating room for impella cp removal and an escalation to impella 5.5. the impella 5.5 was prepped, inserted via the right axillary artery, and support was initiated as the impella cp was pulled across the aortic valve, support discontinued and removed. during the implant of the impella 5.5, it was reported there was a controller error on the automated impella controller (aic). support from the new impella 5.5 was not compromised due to the error, however the decision was made to replace the console. after the back up unit was obtained, it was noted that the pump flows and purge systems were running, and the controller alarm resolved, however there was a loss of placement signal and motor current. the aic was replaced with the back up unit, and the issues were resolved. the impella cp access site was surgically repaired for closure. the patient was transported to the intensive care unit for rest and recovery. on impella 5.5 support day three, it was noted that the patient was believed to be in septic shock and cardiogenic shock. the patient had a fever of 38 celsius and was placed on a cooling blanket. blood cultures and a sputum sample were sent for bacterial growth. the patient was noted to be critically ill. on support day four, the patient remained febrile. the patient¿s care plan was to stabilize the patient until a percutaneous coronary intervention (pci) could be performed. on support day six, the patient¿s urine was noted to be pink tinged. the patient was noted to be in stable condition. on support day seven, the medical team continued to have concern for the patient having sepsis. the medical team did not believe that the impella access site was infected. a plan was made to start the patient on antibiotics and to pull more cultures to identify the pathogen. there was additionally a concern for an acute kidney injury as the patient¿s creatinine levels were increasing. on support day eight, it was noted that the patient received multiple fluid boluses for low central venous pressure and suction alarms. on support day nine, it was noted that the patient had some episodes of ventricular tachycardia when the patient coughed or moved overnight. later that day, the patient had a successful pci. following the pci the patient remained on impella 5.5 support for rest and recovery. on support day 10, the patient remained febrile with unclear source of infection. on support day 11, the concern for infection continued. the patient received a whole body computed tomography scan to attempt to find a source of infection. the patient remained on antibiotics, and of note the impella site was without infection concern. on support day 13, it was noted that the patient was not responding to neurological stimuli. it was reported that the aic had an alarm for ¿unexpected controller shut down¿ which self-resolved. a back up aic was placed on standby, and the aic was monitored. on support day 17, the patient had bloody drainage from a nasogastric tube and had slight hematuria. the medical team was advised to verify placement of the impella 5.5 with ultrasound. the patient continued to not follow commands but was noted to move spontaneously. on support day 18, a bedside ultrasound revealed the impella had migrated, requiring the physician to reposition the device by pulling it back two centimeters into the correct position. the patient¿s hematuria was noted to be clearing; while the cause is unconfirmed, it may be related to foley catheter trauma, as plasma free hemoglobin levels were within normal limits. additionally, a large clot was identified during an upper endoscopy and removed. the heparin in the impella purge solution was discontinued, and the patient was extubated to a high flow nasal cannula. on support day 19, the patient experienced flash pulmonary edema and required emergency intubation. the planned wean of the impella 5.5 was stopped, and the performance level was increased to p5 from p2. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves four impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). automated impella controller 1 with serial number (b)(6). automated impella controller 2 with serial number (b)(6). this report specifically addresses pump 1: impella cp, with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h10 for the related report numbers. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the asae/ major bleed: minimal oozing at cp site, heparin off but unknown resolution, the cause of the access site adverse event was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
2689 | 1220648-2025-49091 | 23826821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23826821&pc=OZD | 11/17/2025 | 12/24/2025 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768852 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 100 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp for support during a high-risk cardiac procedure. it was reported that the patient was having runs of ventricular tachycardia and would resolve with repositioning of the impella. the doctor had to reposition a couple times, and the decision was made to leave the impella shallow due to sensitivity in the left ventricle. flows decreased to 2-2.5liters per minute, and the doctor was okay with positioning. | [5 empty] | The investigation for the reported tachycardiahas been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the tachycardia could not be determined. | The impella device is not available for return from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed.
2690 | 1220648-2025-49093 | 23826931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23826931&pc=OZD | 02/28/2025 | 12/16/2025 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611781 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported in the long-term outcome and quality indicator (loqi) that a patient was implanted with an impella cp for support during a high-risk cardiac procedure. the patient had chronic kidney disease with worsening renal function requiring continuous veno-venous hemodialysis (hd), after placement of the impella cp. normal filling pressures and cardiac output making the acute kidney injury unlikely due to heart failure/cardiogenic shock. no plasma free hemoglobin drawn, so unclear if this was secondary to hemolysis. dialysis catheter removed on (b)(6) 2025 with gradual improvement, followed by re-decline of renal function. ongoing concerns for minimal urine output, resulting in tunneled hd catheter placed on (b)(6) 2025 in the right internal jugular vein. the subject was discharged with tunneled hd catheter and ongoing outpatient dialysis. | [5 empty] | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. renal failure/ cardiac failure: the cause of the injury was not determined due to insufficient clinical details. hemolysis: complaint device not returned for evaluation. data log reviewed, no suction, no mc spike observed, no position issue or placement signal trend observed. the cause of hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it¿. impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine¿. section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction, suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure¿.
2691 | 1220648-2025-49094 | 23827117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23827117&pc=OZD | 11/17/2025 | 01/05/2026 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740371 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that the patient had decreased platelets function. no signs of hemolysis. alanine aminotransferase was improving. the patient was bleeding from all peripheral sites including the right femoral artery (rfa) impella. the patient had known decreased platelet function. they gave a unit of platelets while on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. clinical details states that the patient has decreased platelets function and was bleeding from all peripheral sites. the cause of the access site adverse event was patient related to the decreased platelets function and bleeding from all peripheral sites.
2692 | 1220648-2025-49095 | 23827332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23827332&pc=OZD | 11/17/2025 | 12/16/2025 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026760981 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | Device Sensing Problem |  |  |  |  | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was implanted with a impella 5.5 for support during a high-risk cardiac procedure. it was reported that the patient had a run of ventricular tachycardia and was shocked. the placement signal on the automated impella controller went away. two days later, the pump stopped and motor current high alarm were noted. the pump had critical purge rates less than one cubic centimeter per hour and purge pressures greater than 1100. tissue plasminogen activator was independently ordered by hospital providers and started. the nurse notified local team that impella placement signal had gone out. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Investigation summary: 5.5 pump sn (b)(6) and a purge cassette returned. optical sensor issue: clinical details report that the placement signal went away on (b)(6) 2025. no further details were provided. data logs were not available for investigation. returned product was investigated and there were no abnormalities noted. the pump's 5s parameters were within the expected range. the pump passed laser continuity testing. when connected to a console for several hours, all 5s parameters remained stable and the ps readings were typical/remained stable. since insufficient clinical details were provided, data logs weren't available, and there were no defects on returned product, the cause of the optical sensor issue could not be determined. clinical details report a high mc pump stop alarm on 19/nov. earlier that day, a high purge pressure alarm was reported. tpa was reportedly used to troubleshoot. data logs were not available for investigation. returned product was investigated and the pump passed electrical resistance testing. there was noted to be biomaterial in the purge gap and around the impeller. the pump was connected to a console and high purge pressure reproduced. the pump was unable to be restarted, reproducing high mc pump stop alarms. several windows were cut in the catheter to visualize the purge line, and it was found that blood had ingressed into the purge line up to the 75 cm mark (near the red plug), but not any further. there were no kinks noted in the patient cable or catheter. after several hours of attempting to purge and soaking the sidearm overnight, purge fluid had not gotten further than the connection between the sidearm filter and the purge lumen, indicating some sort of blockage. the red plug was opened and there was no ingress there, however, where the purge line transitioned from the larger to smaller tube, it was noted that there was a region of brown cloudiness in the tubing. the purge line was cut at this location, purge fluid was injected via a syringe, and then purge fluid was able to be injected from sidearm and expel at the location of the cut, indicating that this location was a potential cause of the blockage at the time of investigation. based on clinical details provided, however, the high purge pressure event did not occur until the 6th day of support, suggesting this issue was not present from case start, as would be expected with a blockage in the red plug. there are several potential reasons that blood ingressed as high as it did in the purge line, though none could be confirmed with returned product. purge pressure high: limited clinical details were provided. additionally, the complication did not occur until 6 days into support, so the findings of blood ingress and a blockage in the purge lumen within the red plug does not explain the cause. finally, data logs were not provided. for these reasons, the cause of the high purge pressure could not be determined. pump stop: based on the findings on returned product and clinical details provided, the cause of the pump stop was determined to be due in part to the blood ingress into the motor. the proximal cause of the blood ingress could not be established. ventricular fibrillation: clinical details report that the patient had a run of ventricular tachycardia on (b)(6). in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot. device lot number: 1986735. device history batch. subcomponent lot: n/a. device history review the pump sn (b)(6) passed all post sterile inspection checks.
2693 | 1220648-2025-49096 | 23827850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23827850&pc=OZD | 11/17/2025 | 12/24/2025 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760544 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 90 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella cp for support during a high risk cardiac procedure. it was reported that the patient had a foley but had the staff remove it due to discomfort and urine was reported to be wine colored. however, no alarms were present, and activated clotting time sub-therapeutic and heparin were titrated up. the pump was eventually removed independently and hemostasis was achieved. | [5 empty] | The impella device is not available from the customer. therefore, investigation of the device is not possible. should any new information be received, a supplemental report will be filed. | The investigation for the reported hematuria has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematuria could not be determined.
2694 | 1220648-2025-49097 | 23828201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23828201&pc=OZD | 02/27/2025 | 12/16/2025 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533151 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported in the long-term outcome and quality indicator (loqi) that a patient was implanted with impella cp for support during a high risk cardiac procedure. it was reported that the patient was not able to lift the right upper extremity, with right hemiplegia (resolved) and hemisensory loss (resolved). a head computed tomography indicated multiple small infarcts in the left hemisphere in a watershed distribution, bilateral cerebellar hemispheres and occipital lobes suggesting central embolic etiology. no acute surgical intervention. this resolved prior to discharge without sequelae. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system cardiogenic shock. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
2695 | 1220648-2025-49098 | 23828452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23828452&pc=OZD | 11/17/2025 |  |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026765825 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon arrival to the emergency department, a patient was in and out of ventricular tachycardia storm. the patient was shocked 12 times total. the plan was for an impella 5.5 as a bridge to ablation. post procedure, multiple false impella in aorta alarms were noted. placement was verified by transesophageal echocardiogram. the placement monitoring alarm was disabled by the perfusionist. upon arriving to coronary care unit, it was noted that the left ventricle (lv) and aorta (ao) were missing. flows were appropriate for p-level as well as appropriate motor current. placement signal (ps) waveform spontaneously returned, immediately followed by a controller error alarm. automated impella controller (aic) to aic transfer was completed by perfusion. ao ps artifact was still present but not alarming. about ten minutes after aic to aic transfer the ao/lv ps returned and correlated to the patients a-line. care was withdrawn and the patient expired. this report is one of two related reports. this report for the impella 5.5 and an additional report will be submitted for the aic (serial number (b)(6)). refer to section h10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. false alarm: data logs were reviewed and confirmed the alarms. the cause of the false alarm could not be determined. optical sensor issue: data logs were reviewed and the report was confirmed. upon review of data logs, it was determined that the console was the cause of the optical sensor issue. device history review: the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | The impella device was not saved by the customer post explant and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
2696 | 1220648-2025-49099 | 23828557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23828557&pc=OZD | 11/17/2025 | 01/30/2026 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102529 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 115 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon arrival to the emergency department, a patient was in and out of ventricular tachycardia storm. the patient was shocked 12 times total. the plan was for an impella 5.5 as a bridge to ablation. post procedure, multiple false impella in aorta alarms were noted. placement was verified by transesophageal echocardiogram. placement monitoring alarm was disabled by the perfusionist. upon arriving to coronary care unit, it was noted that the left ventricle (lv) and aorta (ao) were missing. flows were appropriate for p-level as well as appropriate motor current. placement signal (ps) waveform spontaneously returned, immediately followed by a controller error alarm. automated impella controller (aic) to aic transfer was completed by perfusion. ao ps artifact was still present but not alarming. about 10 minutes after aic to aic transfer, the ao/lv ps returned and correlated to the patients a-line. care was withdrawn and the patient expired. this report is one of two related reports. the is report is for the aic and an additional report will be sent for the impella 5.5 (serial number 634688). refer to section h10 for the related report number. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation is still ongoing. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2697 | 1220648-2025-49102 | 23828874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23828874&pc=OZD | 08/26/2025 |  |  | 12/17/2025 |  |  | 08/26/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026708854 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the patient experienced a drop in hemoglobin and hematocrit raising concern for a gastrointestinal bleed. a continuous furosemide infusion at 5 milligrams per hour was ongoing, and two units of packed red blood cells were initiated in response to the suspected bleeding. it was reported that the patient was not accepted for transfer, and the impella device was being weaned with possible removal anticipated. the patient continued to be hemodynamically supported with continued monitoring of respiratory status and laboratory values. | [5 empty] | No product was returned. optical sensor issue: clinical details note psl alarms during the case. impella was repositioned. it is unknown if the repositioning resolved the alarms. data log review confirmed psl alarms on (b)(6) 2025. snr/contrast had a downward trend leading up to the psl alarms. during this time, there was variation in the ps waveforms with some periods of more ventricular-looking waveforms compared to the expected aortic waveforms. however, with a lack of clinical details, a positioning issue cannot be confirmed. the cause of the optical sensor issue was not determined since insufficient clinical details were provided and product was not returned. major bleed: clinical details state that the patient had a drop in h/h and there was concern for a gi bled. 2 prbcs were given. the cause of the major bleed was not determined since insufficient clinical details were provided. the pump passed all post-sterile inspection checks.
2698 | 1220648-2025-49103 | 23828882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23828882&pc=OZD | 11/17/2025 |  |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026783584 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that groin access was obtained. the sixth french sheath was inserted into the right femoral artery. the 14 french peel away was inserted and impella was primed and prepped. the impella cp was inserted through the peel away sheath but were unable to advance due to vessel size. the impella cp was removed, the peel away was removed, and the patient was taken to the intensive care unit. | [5 empty] | H6: type of investigation: omit code b17 and added code b18. additional information: the device was discarded. the serial number for the pump was (b)(6) and lot number 2026783584. patient demographics are not available. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | D4 catalog and primary udi number were updated as they were reported incorrectly on the previously submitted report. h4 revised device manufacture date as it was reported incorrectly on the previously submitted report. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had small vessel size preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
2699 | 1220648-2025-49105 | 23829102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829102&pc=OZD | 09/16/2025 |  |  | 12/17/2025 |  |  | 09/10/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706397 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 110 KG | Non Hispanic | White | Death | Yes | No | Difficult or Delayed Positioning | [6 empty] | Arrhythmia | Atrial Fibrillation | Unspecified Infection | Sepsis | Asystole | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed atrial fibrillation with rapid ventricular response, which resolved after restarting furosemide. the impella became malpositioned, flipping toward the mitral valve and causing ectopy, and multiple attempts to reposition the device were unsuccessful. the patient subsequently developed pneumonia and underwent tracheostomy. the patient later experienced a cardiac arrest with return of spontaneous circulation; however, the patient ultimately expired while on mechanical circulatory support after the family elected not to pursue further resuscitative efforts. | [5 empty] | No product was returned. arrhythmia, paroxysmal atrial fibrillation, infection, sepsis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. asystole: the root cause of the injury was unable to be determined due to insufficient clinical details. difficult to place or position (positioning issue): the root cause of the positioning issue was unable to be determined due to insufficient clinical details. device passed all post sterile inspection checks.
2700 | 1220648-2025-49104 | 23829205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829205&pc=OZD | 11/17/2025 |  |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733776 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was presented to the emergency department with dizziness and high blood pressure. a chest x-ray revealed pulmonary edema. electrocardiogram revealed a new right bundle branch block and non-st-elevation myocardial infarction. the patient was sent to the catheterization lab where a right heart catheterization revealed biventricular failure and cardiogenic shock. lactate was noted to be 2.2. the shock team discussed and unfortunately, there was concerns regarding femoral access and the patient was placed on an intra-aortic balloon pump and was transferred to another hospital for evaluation. the patient was turned down for any surgical revascularization but was evaluated for placement of the impella 5.5. the patient was taken to the operating room and the impella 5.5 was inserted into the graft and was clearly visualized on fluoroscopy. multiple attempts were made to advance the catheter but were unable to traverse the angle necessary for insertion. after multiple attempts the wire was exchanged for platinum plus and despite best efforts the impella 5.5 still was unable to traverse the patients anatomy. the decision was made to abort the impella 5.5 insertion and attempt an impella cp. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and tortuosity preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
2701 | 1220648-2025-49107 | 23829220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829220&pc=OZD | 09/11/2025 |  |  | 12/17/2025 |  |  | 09/11/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026724353 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp following cardiac arrest. the patient remained critically ill on multiple vasopressors, receiving blood products and volume resuscitation, with a hypercoagulable state and renal failure. a hematoma developed at the impella access site, and the patient experienced severe bleeding from the pericardial drain, requiring eleven units of packed red blood cells and manual pressure at the groin. the device was explanted, and the procedure outcome is unknown. | [5 empty] | Hematoma/ major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. renal failure : the cause of the injury was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
2702 | 1220648-2025-49109 | 23829263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829263&pc=OZD | 09/11/2025 |  |  | 12/17/2025 |  |  | 09/11/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691913 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella cp repeatedly alternated between ¿suction alarm¿ and ¿impella position in the aorta¿ alarms. echocardiography showed that the pump outlet was positioned over the atrioventricular junction, and attempts to reposition the device were unsuccessful. despite volume administration, suction alarms persisted, and even when the pump was pulled back into the aorta, flow could not be increased beyond 0.5¿l/min. the device was removed and successfully replaced, and no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The product with event logs were received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. data review: data logs confirmed low flow when pump started with multiple suction alarms & remained to the rest of the case. logs also revealed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. the returned pump was analyzed. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. pump suction: the cause of pump suction was unable to be determined as limited clinical details were provided and no issue was found on the pump. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no issue was found on the pump. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
2703 | 1220648-2025-49110 | 23829577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829577&pc=OZD | 11/17/2025 | 01/02/2026 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484233 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was implanted with an impella 5.5 for support during a high risk cardiac procedure. it was reported that the purge pressure alarmed was too high and alteplase was administered through the purge. the patient was switched back to bicarbonate and purge pressure was back to normal. a week later, bivalirudin was still not on due to low platelets. three days later, the patient received one unit of packed red blood cells for low hemoglobin, which is attributed to the chest tube. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The device or logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: clinical details noted that patient received 1 unit prbcs for low hemoglobin due to patient's chest tubes. the cause of the injury was not determined due to insufficient clinical details. high purge pressure: clinical details noted that patient had high purge pressure alarms overnight that were resolved with tpa. no additional details were provided. the cause of the high purge pressure could not be determined as limited clinical details were provided and no product or logs were returned for evaluation. h1type of reportable event correctly as it was entered incorrectly on the initial report that was submitted. | Pump wasn't returned after explanation and it is unknown if any blood products were given.
2704 | 1220648-2025-49112 | 23829730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829730&pc=OZD | 09/18/2025 | 12/17/2025 |  | 12/17/2025 |  |  | 11/19/2025 | 12/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322914 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump implanted to assist a patient with cardiomyopathy who was being assessed and monitored for a potential heart transplant. it was reported that during support there was placement signal unreliable alarm with no waveform or numerics. intensive care unit team checked for cable kinks and for proper securement in automated impella controller. provider ordered alarm audio disabled. there was no patient harm. the impella 5.5 was already submitted in a precious report. the investigation team requested on november 19th 2025 that the aic be investigated also. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2705 | 1220648-2025-49114 | 23829902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23829902&pc=OZD | 11/18/2025 | 01/22/2026 |  | 12/17/2025 |  |  | 11/18/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762824 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Intraoperative Cardiac Injury | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that there was mention of a possible reason for the patient's transfer to another hospital after the removal of the impella. the mitral valve might have been injured by the pigtail as the patient worsened after impella weaning and was transferred for further therapy. the patient was completely weaned and the pump was discarded by the customer. no further information was available. | [5 empty] | Investigation summary no product returned. heart injury/patient-device interaction problem: clinical reported the mitral valve might have been injured by the pigtail, as the patient worsened after impella weaning. the cause of the issue was not established as limited clinical information was provided and no product was returned. device history lot device lot: 1990645. device history batch subcomponent lot: n/a. device history review device sn (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6 codes were corrected from a24 to a01.
2706 | MW5180714 | 23830204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23830204&pc=OZD | 12/04/2025 | 12/11/2025 |  | 12/17/2025 |  |  |  | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA PURGE TUBING |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 83 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | Cardiogenic Shock | [12 empty] | Other | [28 empty] | Initial submission | [9 empty] | Purge tubing for an impella device connected to pump, for emergent impella placement on an unstable patient in cardiogenic shock. upon connecting the purge tubing to the pump, the tubing would not prime and error message was displayed on console. tried several more times, unable to prime the line- repeated error messages. moved the purge tubing to a different impella console, and the same error message occurred. during this time (while troubleshooting), the patient grew increasingly unstable and required epinephrine injections to stabilize. pulled new purge tubing from a new sealed impella cp box and changed the tubing that was connected to the impella- the new tubing immediately primed. once the pump was primed, inserted it and the patient began to stabilize. the faulty tubing resulted in a critical delay in care.
2707 | 1220648-2025-49116 | 23830371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23830371&pc=OZD | 11/18/2025 | 01/05/2026 |  | 12/17/2025 |  |  | 11/18/2025 | 12/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729361 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 for support during a high-risk cardiac procedure. it was reported that the patient had a hematoma at the axillary site. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device is not available from the customer. therefore, investigation of the device is not possible. should any new information be received, a supplemental report will be filed.
2708 | 1220648-2025-49115 | 23830443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23830443&pc=OZD | 11/18/2025 | 12/17/2025 |  | 12/17/2025 |  |  | 11/18/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770717 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Low Blood Pressure/ Hypotension | Unspecified Infection | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 68 year old male was implanted with a impella cp for support during high risk cardiac procedure. it was reported that the patient became hypotensive with arrhythmias requiring increase in pressors. ramp down was also attempted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: omit code f1904, this code was reported in the initial report in error. h11: updated with additional information from the customer. additional information: the device was discarded after explant. the patient had extensive past medical history of congestive heart failure (chf) and atrial fibrillation (afib). ppae (infection, arrhythmia): investigation summary ppae (infection, arrhythmia): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. (fever, hypotension): the cause of the injury was not determined due to insufficient clinical details. device history review the pump passed all the post sterile inspection check.
2709 | 1220648-2025-49117 | 23830548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23830548&pc=OZD | 11/19/2025 |  |  | 12/17/2025 |  |  | 11/21/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026756474 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male | 95 KG |  | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was implanted with a impella 5.5. it was reported that according to the doctor, there was a problem with the purge pressure. the purge pressure was too high (1100 mmhg) and the purge flow low ( less than one milliliter per hour). they started tissue plasminogen activator lysis after consulting. the purge pressure and the flow improved to normal levels. after some time, the increase of the purge pressure started again, when the ward doctor improved the p-level of the pump. the surgeons made the decision to change the pump. the patient was not harmed. | [5 empty] | H6 medical device problem code: a140508 was erroneously entered on the initial manufacturer device report number. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | Investigation summary: purge pressure high: clinical reported that the purge pressure was too high (1100mmhg) and the purge flow low (<1 ml/h). tissue plasminogen activator (tpa) was administered which resolved the issue, but high purge pressure reoccurred. the data logs confirm purge pressure high alarms. lt log analysis shows no motor current spikes outside p-level changes. logs show a gradual rise in purge pressure with a corresponding decrease in purge flow. first gradual purge pressure rise was for about 2 hours before alarms triggered. high purge pressure was sustained for over 2 hours before decreasing towards normal values aligning with clinical mention of tpa administration. about 7 hours after first high purge pressure event there was another gradual rise in purge pressure (over about 25 minutes) before high purge pressure alarms triggered. issue resolved after about 44 minutes. the cause of the issue was determined to be biomaterial build up as evidenced by log analysis showing gradual purge pressure rise with no motor current spikes device history review: the pump passed all post sterile inspection checks.
2710 | 1220648-2025-49120 | 23830759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23830759&pc=OZD | 11/21/2025 | 12/11/2025 |  | 12/17/2025 |  |  | 11/21/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764427 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male | 109 KG |  |  | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a doctor could not place the impella cp due to the patients ventricle. the doctor tried to use a buddy wire to improve the kinking, but impella placement was not achieved. the patient vessel was treated without impella. the procedure was completed without the impella and the patient remained stable the entire time. at the end of the procedure, the vessel was closed with a double proglide. throughout the entire examination, no additional categories were necessary and no harm occurred to the patient. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had tortuosity and resistance at aorta preventing successful delivery of impella. product damage (catheter kink): the cause of the product damage was unable to be determined due to insufficient clinical details. device history lot: device lot: 1990693. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2711 | 1220648-2025-49113 | 23830969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23830969&pc=OZD | 11/18/2025 | 12/17/2025 |  | 12/17/2025 |  |  | 11/18/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703336 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 69 year old male was implanted with a impella cp for support during high risk cardiac procedure. it was reported that the patient was having multiple runs of ventricular fibrillation which led to the decision to remove pump. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2712 | 1220648-2025-49121 | 23831049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23831049&pc=PYX | 11/20/2025 | 01/09/2026 |  | 12/17/2025 |  |  | 11/20/2025 | 12/17/2025 |  |  | 02/28/2026 |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478759 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Bradycardia | Cardiac Arrest | Low Blood Pressure/ Hypotension | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella rp flex for support during a high risk cardiac procedure. it was reported that after an impella rp flex was inserted, the patient became progressively more hypotensive and bradycardic. the patient arrested requiring cardiopulmonary resuscitation. the device position was checked with transesophageal echocardiogram and fluoroscopy and showed to be in good position. however device flows now are down to 2.5 liters per minute. the dialysis line was removed during the same time. the patient arrived to catheterization lab for impella rp flex placement with blood infusing for low hemoglobin and hematocrit. the team felt once blood was administered, device flows would improve. throughout the afternoon flows progressively got worse. it was felt that the pump may have ingested a thrombus during the code. the team/family did not wish to replace the device at this time. the patient prognosis was poor and the family was moving towards comfort measures. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary no product returned. no data logs available. clinical states after rp flex was inserted, patient became progressively more hypotensive and bradycardic. flows continued to worsen while on support despite blood given and pump repositioning. they suspected thrombus ingestion. no cannula kinks noted. pt had end organ failure prior to case start. patient prognosis poor and family moving towards comfort measures and patient ultimately expires. bradycardia / thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension / cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow: no product or data logs available. the cause of the low flow was not determined due to lack of product and data logs returned for investigation. device history review: device sn (b)(6) passed all post sterile inspection checks.
2713 | 1220648-2025-49122 | 23831244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23831244&pc=OZD | 11/24/2025 |  |  | 12/17/2025 |  |  | 11/25/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729965 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female | 35 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced anuria and anemia. a computed tomography scan revealed bleeding in the right pleural cavity. the patient was transfused 12 units of packed red blood cells. the patient expired due to multi organ failure. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed, organ failure: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2714 | 1220648-2025-49118 | 23831293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23831293&pc=OZD | 09/13/2025 | 01/30/2026 |  | 12/17/2025 |  |  | 11/19/2025 | 12/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322187 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Decrease in Pressure | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the purge system was successfully changed and the purge pressures were within the normal range. they want to began weaning the pump. the defective purge system can be picked up from the cardiotechnics department in the operating room. | [5 empty] | Investigation summary: data analysis: console logs from the reported day of event are consistent with complaint and show several purge-related alarms purge pressure high (#408) and purge system blocked (#409) which seem to self-resolve. ltlog and rtlog data shows values of purge pressure increasing up to 1200mmhg or drop to zero, with both purge cassettes that were used during this case. replacing the purge cassette did not resolve the issue, reduced its severity and frequency of alarms. device analysis: console was tested at fs and its purge system was found to be operating within specification - both, the purge motor and purge pressure sensor were confirmed to be working as expected. separate investigation ((b)(4)) into purge cassettes 61648, 190694 and pump 611116 did not identify root cause, as all devices were working as expected during testing. root cause: the root cause of purge pressure fluctuations could not be determined because it could not be reproduced. device history lot n/a. device history batch n/a. device history review: q1) service history review - are there an qns associated with the complaint device¿s most recent service report? no. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a. q3) complaint history review - have there been any other complaints against this console? no. there were no issues related to complaint discovered when reviewing the product history of console (b)(6). | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
2715 | 1220648-2025-49123 | 23831422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23831422&pc=OZD | 10/01/2025 |  |  | 12/17/2025 |  |  | 10/01/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703340 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the access site. pressure was held and heparin was withheld. | [5 empty] | No product was returned for investigation. the cause of the bleeding is determined to be use issue because the bleeding was resolved by matching the angle and holding the pressure at the site. the device passed all post sterile inspection checks.
2716 | 1220648-2025-49124 | 23831441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23831441&pc=OZD | 10/01/2025 |  |  | 12/17/2025 |  |  | 10/01/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722786 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Malposition of Device | Difficult to Open or Close |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure alarms. lysis protocol was performed and the purge flow stabilized. later, suction alarms occurred and the pump was found to be positioned deep in the left ventricle. the pump could not be repositioned. the patient was weaned and explantation was planned. | [5 empty] | Section d6b, date of explant has been added. no product or logs were returned for evaluation. the cause of the difficulty to open/close catheter anchor could not be determined as no product was returned for evaluation. the cause of the high purge pressure could not be determined as no product or log were returned for evaluation and limited clinical details were provided. the cause of the pump suction could not be determined as limited clinical details were provided and no product or logs were returned for evaluation. the cause of the device positioning issue could not be determined as limited clinical details were provided as to how the pump came out of position. the pump passed all post inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2717 | 1220648-2025-49111 | 23831577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23831577&pc=OZD | 11/17/2025 | 01/03/2026 |  | 12/17/2025 |  |  | 11/17/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026746193 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 120 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella 5.5 for support during a high risk cardiac procedure. the patient was having non-sustained premature ventricular contraction, and the surgeon repositioned under point-of-care ultrasound after admission to the intensive care unit and the ectopy resolved. the device was eventually explanted at beside and the patients outcome was stable. the pump is not available for return. | [5 empty] | The impella device is not available for return from the customer. therefore, investigation of the device is not possible. should any new information be received, a supplemental report will be filed. | The investigation for the reported ppae (arrhythmia) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ppae (arrhythmia) could not be determined.
2718 | 1220648-2025-49125 | 23832303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23832303&pc=OZD | 10/01/2025 |  |  | 12/17/2025 |  |  | 10/02/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026696137 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the access site. the pump was explanted and the patient was in stable condition. | [5 empty] | A2 and a4 are unknown. the cause of the bleeding was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks.
2719 | 1220648-2025-49126 | 23834793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23834793&pc=PYX | 10/01/2025 |  |  | 12/17/2025 |  |  | 10/02/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025578143 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with impella devices (indication unknown) for mechanical circulatory support (mcs). the patient underwent a mitral clip procedure and subsequently required mechanical circulatory support. an impella cp device was placed post-procedure. due to worsening right heart failure, an impella rp device was attempted 24 hours later. the initial rp insertion via the left femoral vein was unsuccessful due to tortuous venous anatomy; the device was removed. a second rp device was successfully placed in the left femoral vein. at the conclusion of the procedure, the physician noted a split at the tip of the peel-away sheath; no groin-related adverse event occurred. the following day, the impella rp device generated a ¿placement signal not reliable¿ alarm thought to be related to insertion technique. there was no impact on the pump function. however, the next day, the impella rp device was explanted (for reason unknown), and the patient remained on impella cp mcs for approximately 1.5 weeks. the patient¿s condition deteriorated resulting in systemic inflammatory response syndrome (sirs) and subsequent demise. | [5 empty] | Medical safety reviewed the event and noted the following: a patient underwent a mitral clip procedure and subsequently required mechanical circulatory support. an impella cp device was placed post-procedure. due to worsening right heart failure, an impella rp device was attempted 24 hours later. the initial rp insertion via the left femoral vein was unsuccessful due to tortuous venous anatomy; the device was removed. a second rp device was successfully placed in the left femoral vein. at the conclusion of the procedure, the physician noted a split at the tip of the peel-away sheath; no groin-related adverse event occurred. on the following day, the rp device generated a ¿placement signal not reliable¿ alarm, suspected to be related to insertion technique. pump function was not affected. the rp device was explanted the next day, and the patient remained on impella cp support for approximately 1.5 weeks. despite ongoing support, the patient¿s clinical condition deteriorated, resulting in systemic inflammatory response syndrome (sirs) and eventual death. the initial unsuccessful rp insertion and the split at the introducer tip are considered device malfunctions. the ¿placement signal not reliable¿ alarm on the second rp device is also considered a device malfunction, despite no impact on pump function. the patient¿s death is conservatively attributed to the impella cp device, as it was in place at the time of death; however, the patient¿s worsening clinical condition likely contributed significantly to the outcome. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, unable to deliver or advance, data log review was not required for this case run. the cause of the delivery issue was most likely related to the patient¿s tortuous vessel condition, based on the given clinical details. related medical device reports: this impella rp report is one of four devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp pump 3, impella introducer and impella cp pump 1, which is associated with the demise outcome.
2720 | 1220648-2025-49129 | 23834799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23834799&pc=OZD | 10/05/2025 |  |  | 12/17/2025 |  |  | 10/05/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720044 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Tachycardia | Thrombosis/Thrombus | Hemodynamic instability | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced labile blood pressure. the patient was transfused fluid and blood products. the groin site was surgically closed and a jp drain was placed in the oozing axillary site. the patient continued to be hemodynamically tenuous. the patient was in sinus tachycardia and an echocardiogram showed the pump was positioned deep and close to the anterior wall. a sandbag was placed at the insertion site due to bleeding. the patient was also treated with dobutamine and levo. the pump flow was reduced in response to suction, and the pump was repositioned. due to decreased hemoglobin the patient was transfused one unit of packed red blood cells. the patient's blood culture was positive so the swan was replaced. heparin was withheld due to a nosebleed and hematuria. upon explantation clots were observed on the pump inlet. | [5 empty] | No product was returned for investigation. data logs were returned for analysis. the root cause of the pump suction was most likely use related due to improper positioning based on provided clinical details and data log analysis. the root cause of the access site bleeding was not determined due to inconclusive evidence from the provided clinical details and unreturned product for further evaluation. the root cause of the thrombosis, tachycardia, and infection was unable to be determined due to insufficient clinical details. the cause of the hemodynamic instability was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks.
2721 | 1220648-2025-49128 | 23834802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23834802&pc=OZD | 10/04/2025 |  |  | 12/17/2025 |  |  | 10/04/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703251 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a leaking purge line near by the purge filter. the patient was successfully weaned from the pump and no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2722 | 1220648-2025-49130 | 23834818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23834818&pc=OZD | 10/06/2025 |  |  | 12/17/2025 |  |  | 10/06/2025 | 12/17/2025 |  |  |  |  | 12/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703235 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced cardiac arrest and suction alarms. later, the patient expired. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The returned product was investigated. biomaterial was found in the purge gap and around the impeller blades. the pump initially produced suction alarms and motor current high pump stop, after which the biomaterial was flushed out. after the biomaterial was removed, the pump ran without any issues and performed per specification. data logs were not returned for analysis, so suction alarms could not be confirmed. the root cause of the pump suction was not determined due to unreturned data logs and inconclusive evidence based on returned product evaluation and provided clinical details. the cause of the cardiac arrest was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | Medical safety/clinical reviewed the event. the patient acute myocardial infarction/cardiogenic shock (scai stage c). the patient supported by impella 5.5 since (b)(6) 025 and was in stable condition. approximately 9 days after start of support, the patient went into cardiac arrest while cardiac rhythm was paced by an internal defibrillator, and suction alarms occurred while on impella support. the patient expired three days later (based on event and explant dates) on the same device. it does not appear the impella device was exchanged. per the reported information no reason for the patient's unexpected expiration on support was discovered by the medical team. initially, the team suspected events like a massive pulmonary embolism, stroke or coronary thrombosis. per the report the medical team does not think the impella is involved with the "poor" outcomes. based on the clinical information and physician's comment there is not enough information at hand to determine direct association between the device and the patient's demise. note: h6 adverse event/investigation coding remains unchanged.
2723 | 1220648-2025-49048 | 23813097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813097&pc=OZD | 11/16/2025 | 01/05/2026 |  | 12/16/2025 |  |  | 11/16/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728817 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 62 yr old female was implanted with a impella cp for support during a high risk cardiac procedure. it was reported patient received 7 left normal saline plus albumin, when bolus¿s stopped patient central venous pressure dropped and would not raise above 10. patient gets hypotensive and suction alarms occurred. it was noted that there was continued oozing from groin site, appears to be impella related vs premature atrial contraction. patient¿s hemoglobin dropped from 10.2 to 6.7. however 2 units of packed red blood cells were given. the impella was successfully weaned and explanted . | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the hypotension, access site bleed, arrhythmia, and anemia has been completed. no product was returned. without the device nor sufficient clinical details, the root cause of anemia, hypotension, arrhythmia, and access site bleed could not be determined.
2724 | 1220648-2025-49047 | 23813098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813098&pc=OZD | 11/14/2025 | 01/12/2026 |  | 12/16/2025 |  |  | 11/16/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754019 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 90 KG | Hispanic |  | Death | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 39 yr old male was implanted with a impella cp for support during high risk cardiac procedure. it was reported that purge system blocked and the patient had oozing at sites. there was a discussion of switching back to bicarb purge solution. patient eventually explanted independently, and the team withdrew care and also removed extracorporeal membrane oxygenation. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product or logs were returned for evaluation. major bleed/minor bleed: clinical details noted that patient had oozing from access sites and was taken to ct to check for retroperitoneal bleed, secondary to ecmo cannulas. the cause of the injury was not determined due to insufficient clinical details. access site adverse event: clinical details noted that patient was oozing from access sites. the cause of the access site adverse event could not be determined as limited clinical details were provided. purge system blocked: clinical details noted that purge system blocked alarms were noted and tpa was added to purge. purge values were able to resolve. no additional details were provided. the cause of the purge system blocked could not be determined as no product or logs were returned and limited clinical details were provided. device history lot: device lot: 1973970. device history batch: subcomponent lot: n/a. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
2725 | 1220648-2025-49049 | 23813110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813110&pc=OZD | 11/16/2025 | 01/21/2026 |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026721545 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57 yr old female that was implanted with a impella cp for support during a high-risk cardiac procedure. it was reported that the patient's urine was light pink. in addition, the pump was measuring 4.6 on echocardiogram and was adjusted in ventricle and measured 3.6cm. the pink urine continued. | [5 empty] | Investigation summary: no product was returned. hematuria: the cause of the hematuria was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot device lot: 1956242. device history batch subcomponent lot: n/a. device history review device (sn: (b)(6)) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2726 | 1220648-2025-49050 | 23813111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813111&pc=OZD | 11/17/2025 | 01/03/2026 |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150571 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the biomed cherry foffer has informed of controller failure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. data analysis: the logs confirmed the reported failure mode given there was a controller failure alarm triggered. device analysis: upon bootup, it is confirmed that the console displayed a controller failure with a red alarm message. an attempt to extract and cross verify the code, but due to the software corruption, the computer was unable to establish a connection to the impellatronic msc1210. conclusion: the root cause of the controller failure was due to the epm software corruption. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(4). | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
2727 | 1220648-2025-49052 | 23813113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813113&pc=OZD | 11/17/2025 | 01/05/2026 |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523044 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 48 yr old female was implanted with a impella cp for support during a high risk cardiac procedure. it was reported that the patient had small hematoma several hours after insertion, resolved within 12 hours. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
2728 | 1220648-2025-49044 | 23813158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813158&pc=OZD | 11/16/2025 | 12/15/2025 |  | 12/16/2025 |  |  | 11/15/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739614 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 57 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Inaccurate Flow Rate | Increase in Pressure | Device Sensing Problem |  |  |  | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 48 yr old male was implanted with a impella 5.5 for support during high risk cardiac procedure. it was reported that a transthoracic echocardiography was performed to check impella placement as aorta signal appeared slightly damp and impella showed left ventricle placement alarm in white. a hematoma remained at the 5.5 axillary insertion site and was monitored every hour. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. optical sensor issue: since the pump was not returned for investigation and limited clinical details were provided, the cause of the dampened waveform could not be determined. false alarm: since the pump was not returned for investigation and clinical details provided were unclear as to what specific false positioning alarm was being reported, the cause of the false alarm could not be determined. high purge pressure: clinical details report a hpp event on 14/dec. at the time, the aic stand was sitting on the impella catheter for 5 minutes. the tpa protocol was initiated to troubleshoot. data logs were reviewed and the high purge pressure event was confirmed on 14/dec. the purge pressure suddenly spiked to 1000 mmhg over a 1.5 minute period without any changes in mc. within the first few minutes of the purge pressure spike, there were several quick drops to zero. additionally, mc started to trend up, resulting in saturated flows. roughly 1 hour after the initial pp spike, there was a mc spike with a ms deviation. purge pressure returned to baseline in a stepwise fashion between 15-16/dec. based on the signature of the rise in purge pressure seen in data logs along with clinical information provided that the aic stand was on the impella catheter, the cause of the high purge pressure was determined to be a kinked purge line. saturated flow: clinical details report that the day after a high purge pressure event (14/dec/2025), the csc noted that pump flows were saturated. tpa had been administered the day before to troubleshoot hpp, however, it had not yet resolved the issue. the patient remained on support through 3/jan/2026. data logs were reviewed and confirmed that the high purge pressure event was due to a kinked purge line. following the initial purge pressure spike, mc began to trend up, resulting in saturated flows. roughly 1 hour after the initial purge pressure spike, there was a mc spike with ms deviation, indicating biomaterial ingestion. product was not returned for investigation. based on the signatures noted in data logs, the cause of the saturated flow was determined to be due to biomaterial buildup. asae/hematoma: clinical details report that there had been a hematoma at the impella 5.5 access site since the date of insertion. no further details were provided, and product wasn't returned for investigation. since insufficient clinical details were provided and the pump wasn't returned for investigation, the cause of the access site adverse event/hematoma could not be determined. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2729 | 1220648-2025-49043 | 23813193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813193&pc=OZD | 11/15/2025 | 12/15/2025 |  | 12/16/2025 |  |  | 11/15/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643672 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 55 yr old male was implanted with a impella cp for support during high-risk cardiac procedure. it was reported that patient had pink-tinged urine and hemolysis was suspected before dropping p-level. there was a hrombus in the left ventricle that was noted on echocardiogram, and the decision was made to leave impella at p2 to unload due to lack of native left ventricle. however, impella was escalated to 5.5 then withdrew care. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. (hematuria): the cause of the injury was not determined due to insufficient clinical details. ppae (thrombosis): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all post sterile inspection checks.
2730 | 1220648-2025-49051 | 23813211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813211&pc=OZD | 11/17/2025 | 12/16/2025 |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750170 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 51 yr old male was implanted with impella cp for support during a high risk cardiac procedure. it was reported that the patient had ischemia that didnt have any flow distal to impella. the impella was explanted due to no flow past impella sheath. patient expired in cath lab. | The complainant reported that a patient was being implanted with an impella cp for mechanical circulatory support and experienced limb ischemia that led to explantation of the device. the patient presented to the hospital with complaints of indigestion that had been occurring for the past week, and pain that started radiating down his back and arm that morning. the patient was then transferred to the complainant hospital and was found to be in st elevation myocardial infarction. an impella cp was placed via the right femoral artery to support the patient during a percutaneous coronary intervention (pci). following the pci, the impella cp was explanted due to lack of distal flow past the impella access site. additionally, it was noted that the patient had a pseudoaneurysm. the patient was then transferred back to the intensive care unit, where the patient decompensated during the afternoon. the decision was made to bring the patient to the operating room and attempt to implant an impella 5.5. during the implant of the impella 5.5 via the right axillary artery, the device would not pass through the patient¿s anatomy. the medical team decided to abort the impella 5.5 implant and implant an impella cp instead. the impella cp was implanted via the right axillary artery, and support was initiated without complications. after impella cp support was initiated, the patient was evaluated for an impella rp flex. the decision was made to implant an impella rp flex. the impella rp flex was implanted via the right jugular vein. on support day two with impella cp and impella rp flex, it was noted that the patient had a hematoma under the right armpit area of the impella cp access site. the site was dry and intact, and the hematoma was monitored. the patient's outcome at the time of explant of the impella cp and impella rp flex was expired. this complaint involves four impella devices: pump 1: impella cp, with serial number (b)(6), pump 2: impella 5.5, with serial number (b)(6), pump 3: impella cp, with serial number (b)(6), pump 4: no malfunctions or adverse events were reported or observed for the impella rp flex, with serial number (b)(6). this report specifically addresses pump 1: impella cp, with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint with reportable events. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 is being corrected to include all relevant information that was not included in the initial report. the revised b5 provides a complete event narrative. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. related medical device reports: this impella cp device report is one of three devices associated with the event. refer to the related manufacturer report numbers as noted in section h10 for the medical device report on the other devices. investigation summary: ppae (ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1977159. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2731 | 1220648-2025-49042 | 23813213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813213&pc=OZD | 11/15/2025 | 12/15/2025 |  | 12/16/2025 |  |  | 11/15/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687250 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 61 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an 84 yr old female was implanted with a impella cp for support during a high-risk cardiac procedure. the patient¿s groin site had become bruised and her hemoglobin dropped from 12 to 7. patient recieved blood and was taken to cardiac computed tomography (ct) scan for a possible retro-peritoneal bleed. ct scan was negative, but she did have some bleeding internally around the femoral artery/impella sheath. the decision was made to explant the impella cp after receiving blood and albumin (and improved blood pressure). patient¿s slightly slurred speech did not improve throughout the afternoon, and she even developed a slight facial ¿droop¿. patient to be taken to magnetic resonance imaging to scan for a possible stroke. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: the cause of the internal bleed was not determined due to insufficient clinical details. stroke: the cause of the injury was unable to be determined due to insufficient clinical details. vessel perforation (peripheral): the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2732 | 1220648-2025-49053 | 23815623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23815623&pc=OZD | 11/16/2025 |  |  | 12/16/2025 |  |  | 11/16/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741017 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella 5.5 as escalation of therapy from an impella cp. the patient had presented to the hospital with worsening chest pain and hypotension. the patient was found to be in st elevation myocardial infarction and diabetic ketoacidosis. the patient was brought to the cardiac catheterization laboratory and was implanted with an impella cp via the right femoral artery in the setting of acute myocardial infarction and cardiogenic shock. impella cp support was initiated with no complications, and a percutaneous coronary intervention was performed with impella cp support. the patient was then transferred to another hospital for additional therapy. after transfer, the patient developed pericardial effusion and cardiac tamponade and was placed on venoarterial extracorporeal membrane oxygenation (va ecmo). the decision was then made to bring the patient to the cardiovascular operating room for a pericardial window and an impella 5.5 placement. following the procedure, the impella 5.5 was prepped and inserted via the right axillary artery without complications. the impella cp was explanted and a 12 french cook sheath was placed at the impella cp access site. impella 5.5 support was initiated, and the patient was transferred to the intensive care unit for rest and recovery. on impella 5.5 support day four, the patient developed a gastrointestinal (gi) bleed and received a blood transfusion. on support day six, the patient was transfused with one unit of packed red blood cells (prbcs). on support day seven, the patient was transfused with two units of prbcs, and it was noted there was no active signs of bleeding. on support day eight, the patient received another two units of prbcs. on support day 13, the patient was noted to continue to be transfused with two units of prbcs per shift. computed tomography was performed and no gi bleed was revealed. there were no signs of hemolysis, and no signs of another source of blood loss. on support day 14, the patient received one unit of prbcs. additionally, the patient experienced blood stools. on support day 18, the patient was transfused with one unit of prbcs. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: pump 1: impella cp, with serial number: (b)(6). pump 2: impella 5.5, with serial number: (b)(6). this report specifically addresses pump 2: impella 5.5. a separate report was submitted for the other device associated with this complaint. refer to section h10 for the related report number. | It was additionally noted that a clip was placed on support day 5 during an esophagogastroduodenoscopy as result of the ongoing gastrointestinal bleeding. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Additional narrative added to b5 to capture clip used to treat gastrointestinal bleed. h6 health effect - impact code added surgical intervention for additional information not previously reported. | Correction: h6 health effect - impact code 19 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2733 | 1220648-2025-49054 | 23815827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23815827&pc=OZD | 11/19/2025 |  |  | 12/16/2025 |  |  | 11/19/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728188 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge system blocked alarm. the perfusionist wants to use the lysis protocol. the pump was replaced and a thrombus was observed on the cannula of the removed pump. the patient is in stable condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Impella was returned. thrombosis: clinical details noted that thrombus was present on pump at explant. in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. purge system blocked: clinical details noted that a "purge system blocked" alarm was noted on day 2 of support. tpa was added to purge but decision was made to replace the pump instead. thrombus was observed on pump after explant. upon visual inspection of returned product, biomaterial was present around impeller blade and in purge gap. pump was connected to console for approximately 8 minutes before ¿purge system blocked¿ alarm occurred. pump was turned onto p-9 and biomaterial was kicked out of pump and purge pressure returned to normal range. data logs were reviewed and lt log showed and increase in purge pressure over approximately 3 hours before "high purge pressure" and "purge system blocked" alarm was triggered. no motor current spike was observed at time of purge pressure rise. the cause of the purge system blocked was due to biomaterial build-up based on data log review and returned product analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2734 | 1220648-2025-49055 | 23815919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23815919&pc=OZD | 11/19/2025 |  |  | 12/16/2025 |  |  | 11/19/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658523 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It is reported that a patient implanted with an impella cp experienced major bleeding from the suture site and chest tubes. the patient was treated surgically with a wash out of the chest. the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary- no product returned. asae/major bleed: patient was bleeding from his suture sites and his chest tubes. the cause of the access site adverse event was patient related since patient has coagulopathy device history lot- device lot: 1885297. device history review- this pump sn (b)(6) passed all post sterile inspection checks
2735 | 1220648-2025-49056 | 23816090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23816090&pc=OZD | 11/18/2025 | 01/05/2026 |  | 12/16/2025 |  |  | 11/19/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722125 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal not reliable alarm. the pump position was confirmed to be correct. the pump was removed, flushed, and reinserted but the alarm did not resolve. the p level of the pump was reduced to resolve the alarm. when the alarm repeated the next day, the pump was explanted. no patient harm was reported. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 revised initial reporter first and last name as they were entered incorrectly on the initial report that was submitted. h6 added codes that were omitted from the first follow up report that was submitted.
2736 | 1220648-2025-49057 | 23816745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23816745&pc=OZD | 11/18/2025 | 12/16/2025 |  | 12/16/2025 |  |  | 11/19/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687850 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Fracture | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a impella 5.5 was implanted for support during high risk cardiac procedure. the doctor reports that the catheter was severed at the red plug (cables are visible) and the impella stopped working. pump was explanted. patient survived. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
2737 | 1220648-2025-49059 | 23816934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23816934&pc=OZD | 11/13/2025 |  |  | 12/16/2025 |  |  | 11/19/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735713 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced pericardial effusion, and drainage was performed. the patient was diagnosed with a left ventricular free wall rupture. the circulatory failure led to the patient's death. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: ppae (pericardial effusion): the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1960353. device history batch: subcomponent lot: n/a. device history review: this pump sn: (b)(6) passed all post sterile inspection checks.
2738 | 1220648-2025-49060 | 23817247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23817247&pc=OZD | 11/17/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026738804 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 82 KG |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was inserted via the right axillary artery. the red impella plug started leaking blood at the end of the case, soon after the right axillary pocket was surgically closed. there was a concern for possible nick of impella catheter during closure. impella flows remained stable. the patient also started bleeding a large amount from the right axillary jp drain. the patient became hypotensive. the right axillary chest was re-opened and washed out. blood stopped leaking from the red plug after the washout. the medical doctor decided not to replace the impella 5.5 at this time. the patient remains on support. | [5 empty] | D6b, date of explant, has been added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2739 | 1220648-2025-49061 | 23817682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23817682&pc=OZD | 11/20/2025 | 02/17/2026 |  | 12/16/2025 |  |  | 11/20/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025538035 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | False Alarm | Difficult or Delayed Positioning | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced hemolysis. proper pump placement was confirmed. later, the patient experienced a placement alarm and cardiac output alarm so the pump was repositioned. impella support continues. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Complaint coding was updated to better reflect the reported event. consequently, section h6 medical device problem codes were corrected and repopulated accordingly.
2740 | 1220648-2025-49062 | 23817744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23817744&pc=OZD | 11/17/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084303 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient was recently discharged after an non-st-elevation myocardial infarction. the patient had a left heart catheterization that showed extensive multivessel disease. the patient was told his vessels were too small and were unable to place stents. later, the decision was made for percutaneous coronary intervention with impella support. it was reported that the impella aorta (ao) and left ventricle (lv) placement signal was not reliable at this time but no alarms were present. ao and lv signals then returned shortly but then went away again for several minutes, during this time a controller error alarm appeared. waveforms appeared again and controller error alarm stopped. throughout the case, the waveform was present until near the end. waveforms were lost again with no alarms. they were lost for minutes until controller error alarm appeared and automated impella controller (aic) to aic transfer was completed. throughout the rest of the case, there was no issues with the waveforms. it appeared to be an aic issue. no patient harm was noted. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2741 | 1220648-2025-49063 | 23817957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23817957&pc=OZD | 11/20/2025 |  |  | 12/16/2025 |  |  | 11/20/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688075 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to patient anatomy. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had heavily calcified iliacal and femoral tracjectory with a lot of tortuosity preventing successful delivery of impella. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e1 added initial reporter address and city as they were omitted from the initial report that was submitted.
2742 | 1220648-2025-49058 | 23818059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818059&pc=OZD | 11/17/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026768593 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was brought to the cardiovascular operating room for impella 5.5 placement. multiple attempts were made to pass the impella 5.5 through the right axillary artery. percutaneous transluminal coronary angioplasty and buddy wire were used. all were unsuccessful. the impella 5.5 was aborted. there was no patient harm reported. additional information was received. the product was not available for return. a second impella 5.5 was utilized and placed successfully. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella.
2743 | 1220648-2025-49065 | 23818372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818372&pc=OZD | 09/25/2025 | 01/07/2026 |  | 12/16/2025 |  |  | 10/01/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1920677 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during patient support the automated impella controller failed to recognize the purge cartridge. the controller was exchanged for a new controller to continue patient support. the patient is in stable condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2744 | 1220648-2025-49066 | 23818390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818390&pc=PYX | 10/01/2025 |  |  | 12/16/2025 |  |  | 10/01/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025578131 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female | 90 KG | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Multiple Organ Dysfunction Syndrome | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | IMPELLA CP SMART ASSIST SET, EU | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella rp experienced loss of the placement signal. impella support continued and no patient harm was reported. | The complainant reported a patient was implanted with impella devices (indication unknown) for mechanical circulatory support (mcs). the patient underwent a mitral clip procedure and subsequently required mechanical circulatory support. an impella cp device was placed post-procedure. due to worsening right heart failure, an impella rp device was attempted 24 hours later. the initial rp insertion via the left femoral vein was unsuccessful due to tortuous venous anatomy; the device was removed. a second rp device was successfully placed in the left femoral vein. at the conclusion of the procedure, the physician noted a split at the tip of the peel-away sheath; no groin-related adverse event occurred. the following day, the impella rp device generated a ¿placement signal not reliable¿ alarm thought to be related to insertion technique. there was no impact on the pump function. however, the next day, the impella rp device was explanted (for reason unknown), and the patient remained on impella cp mcs for approximately 1.5 weeks. the patient¿s condition deteriorated resulting in systemic inflammatory response syndrome (sirs) and subsequent demise. |  |  |  |  | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event and noted the following: a patient underwent a mitral clip procedure and subsequently required mechanical circulatory support. an impella cp device was placed post-procedure. due to worsening right heart failure, an impella rp device was attempted 24 hours later. the initial rp insertion via the left femoral vein was unsuccessful due to tortuous venous anatomy; the device was removed. a second rp device was successfully placed in the left femoral vein. at the conclusion of the procedure, the physician noted a split at the tip of the peel-away sheath; no groin-related adverse event occurred. on the following day, the rp device generated a ¿placement signal not reliable¿ alarm, suspected to be related to insertion technique. pump function was not affected. the rp device was explanted the next day, and the patient remained on impella cp support for approximately 1.5 weeks. despite ongoing support, the patient¿s clinical condition deteriorated, resulting in systemic inflammatory response syndrome (sirs) and eventual death. the initial unsuccessful rp insertion and the split at the introducer tip are considered device malfunctions. the ¿placement signal not reliable¿ alarm on the second rp device is also considered a device malfunction, despite no impact on pump function. the patient¿s death is conservatively attributed to the impella cp device, as it was in place at the time of death; however, the patient¿s worsening clinical condition likely contributed significantly to the outcome. type of reportable event update. after review of the case by medical safety, it was determined the impella rp event is a malfunction type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. note: h6: device problem/component/investigation coding initially reported remains unchanged. correction: the following fields were corrected to provide information known: a4 weight, b5 event description (to provide complete event story with associated devices), b7 medical history and d10. concomitant device. related medical device reports: this impella rp report is one of four devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp pump 2, impella introducer and impella cp pump 1, which is associated with the demise outcome. | The investigation of the reported failure mode has been completed. no product was received for evaluation. multi organ failure: the cause of the injury was not determined due to insufficient clinical details. placement signal (ps) issue: the cause of the placement signal issue is not determined due to no product returned, lack of clinical information, and inconclusive data logs review. device history review: the device passed all post sterile inspection checks.
2745 | 1220648-2025-49067 | 23818409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818409&pc=OZD | 10/01/2025 |  |  | 12/16/2025 |  |  | 10/01/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681531 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a leak from the purge filter. the pump was replaced to continue patient support. the patient was in stable condition. | [5 empty] | No product or logs were returned for evaluation. the cause of the purge leak - pump, could not be determined as no product or logs were returned to confirm the location of the leak. the device passed all post-sterile inspection checks. | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2746 | 1220648-2025-49069 | 23818416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818416&pc=OZD | 10/01/2025 |  |  | 12/16/2025 |  |  | 10/02/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp arrived at the customer site with a broken seal. the shipping box was also broken. no patient use was reported. | [5 empty] | G: added additional manufacturer information. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary packaging damage (sterility compromised): the impella cp device was not returned, and no images were returned to inspect the packaging damage from original packaging box. it was reported that the impella cp pump box seal was broken prior to device removal from original packaging. the cause of packaging issue was not determined due to device/images not being returned for investigation. clinical review: 1-oct-2025: the device was send to the distributor and it was observed that seal of the impella cp was broken. when the customer opened the box the seal of one of the boxes was broken. the device was not used. the device serial number and lot info is unknown. device history review the necessary materials were not returned for analysis so the serial/lot number could not be identified to complete an inspection. | The reported event did not occur with patient use. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2747 | 1220648-2025-49068 | 23818419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818419&pc=OZD | 10/01/2025 |  |  | 12/16/2025 |  |  | 10/01/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711451 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Priming Problem | [5 empty] | Cardiac Arrest | Fever | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | Embolism/Embolus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced suction alarms. the patient was transfused fluid and one unit of packed red blood cells. systemic anticoagulation was withheld for nasal bleeding post nasal-gastric tube placement. air in the purge was resolved with de-airing. suction alarms continued but self-resolved. the patient also had delirium and coded. the patient was treated with epinephrine and cardiopulmonary resuscitation (cpr). additionally, the patient experienced acute respiratory distress syndrome and will possibly have a bronchoscopy today. the patient continued with treatment of bival and levophed. the patient was sedated and intubated due to bleeding issues around the endotracheal tube. next, bivalirudin was withheld as a bronchoscopy was perfomed and clots were removed. treatment with levophed continued. another bronchoscopy was performed and blood was removed. the patient was transfused with one unit of packed red blood cells, bival was withheld, and the patient was scheduled for a tracheostomy and a percutaneous endoscopic gastrostomy (peg). treatment with levosimendan and another unit of packed red blood cells was given. the tracheostomy and gastrostomy were declined by the patient's family. due to another suction alarm the pump was repositioned via echocardiogram. after the patient was rolled he lost pulsatility. levophed was restarted and pulsatility returned within 15 minutes. a peg was placed and resulted in bleeding so another unit of packed red blood cells was transfused. sporatic loss of pulsatility continued. a tracheostomy was placed, and suction overnight. the patient was transfused fluids. the patient then experienced a fever and elevated white blood cell count. the patient was treated with dobutamine. later, the impella was successfully weaned and the patient survived. | [5 empty] | A4 is unknown. no product was returned for investigation. the cause of the suction was patient condition related as the alarms self-resolved or were treated with volume administration. the cause of the air in purge system was not determined since product was not returned and insufficient clinical details were provided. the cause of the epistaxis, respiratory failure, major bleed, embolism, fever, and cardiac arrest cannot be determined as insufficient clinical details were provided. the device passed all post-sterile inspection checks.
2748 | 1220648-2025-49070 | 23818422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818422&pc=OZD | 10/02/2025 |  |  | 12/16/2025 |  |  | 10/02/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661715 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced hypotension and agitation. the patient was treated with an albumin bolus. upone explant, the patient exhibited an increase in white blood cell count. the patient was treated with antibiotics. additionally, the position of the pump was deep. the impella was pulled back under echocardiogram and continues to be in appropriate positioning. epinephrine was swapped for levophed, and heparin was withheld. | [5 empty] | The cause of the infection was unable to be determined due to insufficient clinical details. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2749 | 1220648-2025-49071 | 23818429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818429&pc=OZD | 10/03/2025 |  |  | 12/16/2025 |  |  | 10/03/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700930 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and oozing at the access site. heparin was withheld. the patient is in stable condition. | [5 empty] | The pump was not returned for investigation. the cause of the hemolysis was not determined without returned product investigation and sufficient clinical details. the cause of the access site bleeding issue could not be determined without returned product investigation and sufficient case details. the device passed all post sterile inspection checks.
2750 | 1220648-2025-49072 | 23818443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818443&pc=OZD | 10/03/2025 |  |  | 12/16/2025 |  |  | 10/03/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694364 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the device was replaced with an impella 5.5 to continue patient support. | [5 empty] | A4 is unknown. no product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2751 | 1220648-2025-49073 | 23818674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818674&pc=OZD | 11/17/2025 | 01/12/2026 |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747046 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 112 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Malposition of Device | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was transferred to a different hospital for further care. the patient was present for a planned impella 5.5 with a potential extracorporeal membrane oxygenation (ecmo) reconfiguration/groin repair of ecmo site. the medical doctor (md) gained left ventricle access with the impella, removed the wire, and started the impella at p2. the md removed the peel away and sutured the impella at the blue suture hub in place. the doctor then tried to manipulate the positioning and the impella fell back into the aorta. this required another axillary insertion kit and 0.018 wire was to be used. nine days later, it was reported that the impella purge became blocked overnight. purge flow was less than 1.0 milliliter. the md team switched to tissue plasminogen activator (tpa) for purge on their own. purge flow remains less than 1.0 now. the device was removed due to a suspected clot and purge flow blocked. the purge flow had normalized after tpa was used by the staff on their own accord. however, they still decided to remove the pump. the pump had a large clot on motor housing in a circumferential manner above the outlet portion. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Investigation summary: no product or logs were returned for evaluation. thrombosis: clinical details noted that a clot was present around the motor housing on explant. in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: clinical details noted that after pump was inserted in lv, physician was unable to visualize the tip of the impella and pulled back the pump position. the pump was pulled back into the aorta and had to be removed. the pump was rinsed with sterile saline, then reinserted with a new axillary insertion kit. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled into the aorta when physician was pulling back pump. purge system blocked: clinical details noted that a "purge system blocked" alarm occurred overnight. tpa was added to purge and was able to resolve issue. the cause of the purge system blocked could not be determined as no product or logs were returned for evaluation. device history lot: device lot: 1980876. device history batch: subcomponent lot; n/a. device history review: pump sn 630261 passed all post-sterile inspection checks.
2752 | 1220648-2025-49074 | 23818949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23818949&pc=PYX | 11/17/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1824431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A biomedical engineer reported that on a automated impella controller they pulled the power plug off the end of the power cord. there was no patient involvement. | [5 empty] | The device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the console logs were not relevant and showed no errors. the console ic11628 was returned to fs for further investigation. during the inspection, it was noted that all three wires were disconnected from their terminals, and the ac cord plug had detached from the cord. additionally, the outer sheath was improperly cut, in accordance with the vendor wiring instructions. also, the light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose. the reported issue of the impella controller power plug disconnected from the end of the cord was determined to be caused by defective ac power cord set.h6 investigation: h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2753 | 1220648-2025-49075 | 23819119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23819119&pc=OZD | 11/17/2025 | 01/21/2026 |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026761547 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 95 KG | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was being implanted with an impella 5.5 for mechanical circulatory support following decompensating after having an impella cp explanted. the patient presented to the hospital with complaints of indigestion that had been occurring for the past week, and pain that started radiating down his back and arm that morning. the patient was then transferred to the complainant hospital and was found to be in st elevation myocardial infarction. an impella cp was placed via the right femoral artery to support the patient during a percutaneous coronary intervention (pci). following the pci, the impella cp was explanted due to lack of distal flow past the impella access site. additionally, it was noted that the patient had a pseudoaneurysm. the patient was then transferred back to the intensive care unit, where the patient decompensated during the afternoon. the decision was made to bring the patient to the operating room and attempt to implant an impella 5.5. during the implant of the impella 5.5 via the right axillary artery, the device would not pass through the patient¿s anatomy. the medical team decided to abort the impella 5.5 implant and implant an impella cp instead. the impella cp was implanted via the right axillary artery, and support was initiated without complications. after impella cp support was initiated, the patient was evaluated for an impella rp flex. the decision was made to implant an impella rp flex. the impella rp flex was implanted via the right jugular vein. on support day two with impella cp and impella rp flex, it was noted that the patient had a hematoma under the right armpit area of the impella cp access site. the site was dry and intact, and the hematoma was monitored. the patient's outcome at the time of explant of the impella cp and impella rp flex was expired. this complaint involves four impella devices: pump 1: impella cp, with serial number (b)(6), pump 2: impella 5.5, with serial number (b)(6), pump 3: impella cp, with serial number (b)(6). pump 4: no malfunctions or adverse events were reported or observed for the impella rp flex, with serial number (b)(6). this report specifically addresses pump 2: impella 5.5. separate reports will be submitted for the other devices associated with this complaint with reportable events. | [5 empty] | D9 updated; the product has been returned. the investigation is still ongoing. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2754 | 1220648-2025-49076 | 23819962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23819962&pc=OZD | 11/16/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A circulating tech reported that the console would not recognize purge disc after three attempts. consoles were swapped and a new console recognized the purge disc. there was no patient harm. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. data review: console logs were consistent with what was reported in the complaint. purge disc not detected alarms [event set #402] were observed in the imc logs. device analysis: console was tested after arriving in field service. the purge disc not detected issue was able to be reproduced. further inspection of the console revealed that the purge disc flag was broken (missing at blue retainer). conclusion: the root cause of the reported purge disc recognition issue was determined to be a broken purge disc flag. | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2755 | 1220648-2025-49077 | 23820600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23820600&pc=OZD | 11/17/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694371 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was brought in with st-elevation myocardial infarction and taken to the catheterization lab for an emergent angiogram. as soon as the patient arrived they coded multiple times (ventricular tachycardia/ventricular fibrillation). post defibrillation, the patient was hypotensive requiring vasopressors and the decision was made to place an impella cp for acute myocardial infarction/cardiogenic shock. while on support, the patient was very restless and the groin started oozing upon arrival to the intensive care unit. the doctor came to the bedside and tightened perclose and placed a fem stop without pressure. oozing has stopped and the groin looked soft. no hematoma was noted per the nurse. the pump was later explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. peri-procedural adverse event: the cause of the bleed was most likely use issue related since the bleed started occurring after arrival to the intensive care unit (icu) and hemostasis was achieved by tightening the perclose. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2756 | 1220648-2025-49078 | 23821117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23821117&pc=OZD | 11/17/2025 |  |  | 12/16/2025 |  |  | 11/17/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753405 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient decompensated and was sent to the catheterization lab for impella placement. while on support, hemolysis was reported due to low volume. this resolved with fluid bolts. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely patient condition related since clinical team confirmed hemolysis due to low volume & resolved with fluid compensation. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
2757 | 1220648-2025-49079 | 23821271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23821271&pc=OZD | 11/20/2025 |  |  | 12/16/2025 |  |  | 11/20/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026759940 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge pressure high alarm. tissue plasminogen activator was added to the purge to resolve the issue. no patient harm was reported. the pump was exchanged due to high purge pressure despite lysis therapy. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
2758 | 1220648-2025-49080 | 23821439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23821439&pc=OZD | 11/19/2025 |  |  | 12/16/2025 |  |  | 11/20/2025 | 12/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760589 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted was implanted with an impella cp. after a successful implantation of the pump, the wire was removed and the pump was started. the pump stopped at 0.8 liters per hour and the flow did not increase. the pump was repositioned. the flow was still at 0.8 liters per hour. the pump was exchanged for a new one. the patient was in stable condition. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. low or blocked pump flow: returned complaint cp pump visually inspected. no cannula kink or broken blade observed. some webbing/sticky material around impeller observed which dissolved after soaking. pump run for over 10 minutes at p9 in the investigation lab and flows were within specified limit. data logs not available to review. the cause of the low flow issue cannot be determined due to insufficient clinical details, issue not reproduced on returned product, and lack of field data logs available. | The product was received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. low or blocked pump flow: the returned impella cp pump was visually inspected, and no cannula kink or broken blade was observed. some webbing/sticky material around impeller observed which dissolved after soaking. pump run for over 10 minutes at p-9 in the investigation lab and flows were within specified limit. data logs were not available to review. the cause of the low flow issue cannot be determined due to insufficient clinical details, issue not reproduced on returned product, and lack of field data logs available. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A4, a5, and a6 are unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2759 | 1220648-2025-49082 | 23824274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23824274&pc=OZD | 10/06/2025 |  |  | 12/16/2025 |  |  | 10/06/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719719 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Pulmonary Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported unreliable values and placement signal issues on the automated impella controller (aic). staff confirmed proper positioning with echocardiography. on day five of impella support the patient developed lung edema and hemolysis values in lab. the patient continued with support until the patient was bridged to left ventricular assist device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, pulmonary edema, the clinical details state that the patient developed lung edema. the cause was not determined due to insufficient clinical details. for the hemolysis, the clinical details state that there are hemolysis values in the labs on october 11th. there were no other reported incidents of hemolysis or issues with the pump noted. the data logs do not show any motor current spikes, but there are position unknown alarms that are triggering from the lower placement signal. the clinical additionally confirmed pump positioning throughout support. due to a lack of clinical details and no product returned, the root cause of the hemolysis cannot be determined. regarding the optical sensor issue, the clinical details state that there was no placement signal throughout the case with confirmed pump positioning. the data logs show that there were no placement signal alarms. however, the placement signal was low throughout support and the start of the pump plug in - but the waveform was still present. the 5s parameters are stable and there are no issues seen with them throughout support. the logs show that the g0 was recalibrated to be 196 lower than the original g0 value during calibration but it cannot be confirmed there were any errors that would contribute to a lower placement signal. due to a lack of product returned and insufficient information, the root cause of the optical sensor issue cannot be determined. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation.
2760 | 1220648-2025-49083 | 23824277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23824277&pc=OZD | 10/06/2025 |  |  | 12/16/2025 |  |  | 10/06/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719708 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. therapy ended due to sepsis, endocarditis and multiple organ failure. additional information noted that the patient expired. | [5 empty] | The investigation into the reported issue has been completed no product was returned for investigation. sepsis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. organ failure: the cause of the injury was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2761 | 1220648-2025-49084 | 23824278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23824278&pc=OZD | 10/07/2025 | 12/31/2025 |  | 12/16/2025 |  |  | 10/07/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757911 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with heart failure and cardiogenic shock (cgs) underwent implantation of an impella cp device for mechanical circulatory support. during support, a hematoma was noted at the impella access site. manual pressure was applied for 20 minutes, an additional suture was placed, and the patient received 2 units of packed red blood cells (prbcs) and 1 unit of fresh frozen plasma, achieving hemostasis. oozing from the access site subsequently decreased significantly. the patient was a bridge-to-transplant candidate and later received a heart transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed/ hematoma has been completed. the root cause of the injury is most likely use related since hematoma noted at impella site, additional suture placed, and blood given hemostasis achieved.
2762 | 1220648-2025-49085 | 23824284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23824284&pc=OZD | 10/08/2025 |  |  | 12/16/2025 |  |  | 10/07/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 0048-0024-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient underwent percutaneous cardiopulmonary support and percutaneous coronary intervention under intra-aortic balloon pumping for acute myocardial infarction, including left main coronary artery disease. due to worsening heart failure, support was escalated to an impella device. the patient subsequently stabilized, and the percutaneous cardiopulmonary support device was weaned. during impella weaning, concomitant mitral regurgitation was identified. hemodynamic assessment was performed using a pulmonary artery catheter while adjusting the impella flow rate. a decrease in impella flow resulted in worsening mitral regurgitation and increased pulmonary artery wedge pressure. based on these findings, a percutaneous mitral valve clip was implanted, allowing successful impella weaning without hemodynamic deterioration. | [5 empty] | The investigation into the reported issue has been completed. no product was returned for investigation. mitral valve incompetence: the cause of the injury was not determined due to insufficient clinical details. the necessary materials were not returned for analysis so the serial number could not be identified to complete a device history record inspection.
2763 | 1220648-2025-49086 | 23824293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23824293&pc=OZD | 10/07/2025 |  |  | 12/16/2025 |  |  | 10/07/2025 | 12/16/2025 |  |  |  |  | 12/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731524 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the physician could not implant the pump because of a calcified subclavian artery. the pump could not be advanced. no patient harm has been reported. | [5 empty] | No product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had calcification of a. subclavia into lv preventing successful delivery of impella. the device passed all post sterile inspection checks.
2764 | 1220648-2025-48988 | 23801558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23801558&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716981 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated to a second impella cp device for mechanical circulatory support. while on support, hemolysis was reported. no additional information was provided. the pump was later successfully weaned and explanted. the patients outcome at explanted was noted as survived. this is one of two related reports.. this report address the second impella cp. another report will be submitted to address the first impella cp. refer to section h10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
2765 | 1220648-2025-48987 | 23801619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23801619&pc=OZD | 11/15/2025 | 01/29/2026 |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695387 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that implantation of an impella cp was aborted because the device kinked during insertion. the patient was in stable condition and a new impella cp was implanted. this is one of two related reports. this report addresses the first impella cp. another report was submitted to address the second impella cp. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. product damage (cannula kink): the cause of the cannula kink was most likely related to challenging patient anatomy. unable to deliver or advance: the cause of the delivery issue was most likely related to cannula kink, which occurred due to challenging patient anatomy. | D9 updated; the product has been returned for investigation. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the unable to deliver or advance: the cause of the delivery issue was most likely related to challenging patient anatomy. product damage (cannula kink): the cause of the cannula kink was most likely related to challenging patient anatomy; however, the exact product damage could not be verified without returned product investigation. the pump and introducer were not returned for investigation. data log review was not required for this case. clinical details indicate that the impella cp c9+ was advanced through a peel-away sheath into the right axillary artery, but final placement was unsuccessful due to kinking of the device and vessel tortuosity. the device was explanted after unsuccessful attempts at placement. the patient was reported to have obesity and significant vessel tortuosity, which likely contributed to the difficulty in advancing the device. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2766 | 1220648-2025-48989 | 23801695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23801695&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688261 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in non-st elevation myocardial infarction (nstemi)/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient presented on (b)(6) 2025 with worsening chest pain and nstemi. a left heart catheterization on (b)(6) 2025 revealed multivessel coronary artery disease with a 99% left main artery occlusion and an ejection fraction (ef) of 55-60%. cardiothoracic surgery was consulted, and the patient was accepted for emergent quadruple coronary artery bypass grafting (cabg). coming off cardiopulmonary bypass, the patient experienced severe mitral regurgitation and elevated pulmonary artery pressures, necessitating the placement of an intra-aortic balloon pump (iabp) and four high-dose vasoactive medications. on (b)(6) 2025, the iabp was successfully removed, though a transesophageal echocardiogram showed new inferior wall hypokinesis and a reduced ef of 40-45% off primary blood pressure support medications. by (b)(6) 2025, heart function continued to decline with worsening global hypokinesis and a further decreased ef of 30%. a follow-up angiogram to check graft patency suggested possible occlusion in both vein grafts, prompting the initiation of heparin. the patient¿s clinical status continued to decline over the subsequent days, becoming severely acidotic and hypotensive. the decision was made to bring the patient to the operating room (or) for a possible repeat coronary artery bypass graft and to place an impella 5.5. it was determined in the or that the cabg did not need to be redone, and an impella 5.5 was placed via direct aortic approach and support was initiated with no complications. on support day three, the patient was extubated and weaned off all pressor support. it was noted that the patient was having short runs of ventricular tachycardia and the patient was treated with lidocaine and volume. on support day nine, the patient¿s hemoglobin was noted to be seven g/dl, and the patient was transfused with blood. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. ppae (tachycardia): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. ppae (anemia): the cause of the injury was not determined due to insufficient clinical details. device history review ==> the pump sn 600973 passed all the post sterile inspection checks.
2767 | 1220648-2025-48990 | 23801893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23801893&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755891 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 79 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced cardiac arrest at home and cardiopulmonary resuscitation (cpr) was initiated and emergency medical services (ems) were called. return of spontaneous circulation (rosc) was achieved after five minutes of cpr. ems arrived and transported the patient to the hospital and during transport the patient experienced another episode of cardiac arrest and cpr was initiated. upon arrival to the emergency department, the patient was intubated, and rosc was achieved after 10 minutes of cpr and intubation. an electrocardiogram revealed the patient was in st elevation myocardial infarction and was emergently brought to the cardiac catheterization laboratory for impella cp placement and percutaneous coronary intervention. the impella cp was prepped, inserted, and support was initiated without complications. additionally, a temporary heart pacer was placed, and the decision was made to transport the patient to the cardiovascular intensive care unit (cvicu) for rest and to plan for the pci later that week. upon arrival to the cvicu, it was noted that the patient had a large hematoma at the impella access site. manual pressure was held for 15 minutes and hemostasis was achieved. the access site was redressed with angle matching and a femstop was placed with mild pressure overnight. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the sae/hematoma: rn noticed large hematoma at the impella access site. two rns held manual pressure for 15 min achieving hemostasis. the cause of the access site adverse event was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2768 | 1220648-2025-48991 | 23801959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23801959&pc=OZD | 11/15/2025 | 01/06/2026 |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721357 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was transferred for coronary artery bypass graft evaluation. on arrival st-elevation myocardial infarction elevation and hematemesis were noted. the patient was suspicious for a gastrointestinal (gi) bleed. the decision was made to take the patient emergently for a catheterization. left main, left anterior descending artery and left circumflex artery lesions were noted. the lesion was unable to wire. the doctor decided to place an impella 5.5 to stabilize. while on support, a gi scope was completed. three clips were performed for bleeding. multiple units of blood pressure bagged in prior to oncoming nurse shift. the day nurse had given cryo time two, frozen fresh plasma times two and platelets times two. lactic acid was trending up suspected due to hypovolemic shock. later, the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The cause of the injury was not determined due to insufficient clinical details. the pump sn 615968 passed all the post sterile inspection checks.
2769 | 1220648-2025-48992 | 23802137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802137&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700010 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented to the emergency department complaining of chest pain that started three hours prior to arrival. st-elevation myocardial infarction were found and the patient was brought to the catheterization lab. access for percutaneous coronary intervention was gained and the patient was found to be in cardiogenic shock with index of 1.8. the patient was decompensating and required stat intubation and impella placement. while on support, it was reported that the patient's urinary output was greater than 30 cubic centimeters per hour and pink red. in addition, the patient was noted to have absent distal right lower extremity signals (impella side) with rising lactate. drips and hemodynamics were stable. the decision was made to remove the impella cp in the catheterization lab. left leg artery (femoral artery) intra-aortic balloon pump was placed. the patient's hemodynamics remained stable. per the nurse, palpable popliteal pulse right side. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: (ischemia). in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. ppae (hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2770 | 1220648-2025-48993 | 23802155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802155&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667268 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced suction. the pump was explanted. the patient is in stable condition. | [5 empty] | A4, a5, and a6 are unknown. the impella device was received from the customer for evaluation. upon completiong of the investigation, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was returned for evaluation. pump suction: the pump was returned for investigation. significant biomaterial was found on the impeller. no other physical damage was noted. the pump was then started in a bucket of water, with reproducible low flow and suction alarms. biomaterial was flushed out during the test run, and after that, the pump was able to run per specification. data log was not provided for the reported suction incident. clinical details indicate that after repositioning, pump suction was noted at all p-levels. the left ventricle was filled, no lack of volume was reported, and right ventricular function was normal. purge flow and purge pressure were normal and unchanged, and motor current was damped but not elevated, remaining within range. the patient remained stable with low-dose norepinephrine and was well recovered prior to explant. the cause of the suction issue was most likely related to biomaterial, based on returned product investigation. device history review the pump passed all post sterile inspection checks.
2771 | 1220648-2025-48994 | 23802212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802212&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660382 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support after becoming hypoxic and hypotensive during a right heart catheterization. the patient arrived at the hospital via emergency medical services with severe chest pain. the patient had a stent placed in the left anterior descending (lad) coronary artery on (b)(6) 2025 prior to hospitalization. the patient was brought to the cardiac catheterization laboratory for right heart catheterization, during which the patient became hypoxic and hypotensive. the patient was intubated and an impella cp was placed. the impella cp was prepped, inserted via the right femoral artery, and support was initiated without complications. a percutaneous coronary intervention was performed to the patient¿s lad after discovering an in-stent thrombosis. following the procedure, the patient was transferred to the intensive care unit. during support day one, the patient experienced runs of ventricular tachycardia. the patient received two grams of calcium and two ampules of bicarbonate. a point of care ultrasound was performed at the bedside and revealed the impella to be close to the mitral valve. blood was drawn for laboratory testing and the medical team planned on escalating the patient to an impella 5.5 based on the results. the medical team decided to escalate the patient to an impella 5.5. the patient was brought to the operating room where the impella cp was explanted, and the impella 5.5 was successfully implanted without complications. the patient remained on impella 5.5 support as a bridge to heart transplant or durable left ventricular assist device. additional information regarding impella 5.5, any complications, device explant, and patient outcome following impella 5.5 support is unknown as outcome information was not available in the complaint record accessible for report assessment at time of reporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the tachycardia the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2772 | 1220648-2025-48996 | 23802274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802274&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770788 |  | 000000000001000080 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in non-st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on support day two, the patient experienced ventricular tachycardia that required two shocks of 200 joules to resolve. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: added type of investigation codes b13 and b24. h11: added the results of the investigation. investigation summary: (arrhythmia). in order to make a cause determination, all clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2773 | 1220648-2025-48995 | 23802298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802298&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768684 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma at the access site. a fem stop was placed to improve the patient condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks.
2774 | 1220648-2025-48997 | 23802406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802406&pc=OZD | 11/16/2025 |  |  | 12/15/2025 |  |  | 11/16/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026761516 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male | 62 KG |  |  | Injury | Yes | Yes | Crack | Fluid/Blood Leak | [5 empty] | Stroke/CVA | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient first presented in the emergency department at an outside hospital with chest pain and syncope episodes. the patient went into atrial fibrillation with rapid ventricular response and was cardioverted twice and transferred. the plan was to place an impella 5.5 to support while working up for left ventricular assist device. while on impella 5.5 support, the nurse reported the patient was having an episode of ventricular tachycardia (vt), overnight, not sustained, and lytes were replaced. no further episodes were noted. it was also reported that the patient had a stroke that required thrombectomy. furthermore, it was reported that the impella position was deep in the ventricle. the patient had vt. echocardiogram measured 6.2 centimeters. the impella was repositioned under echocardiogram to 5.5 centimeters and ectopy was noted to be gone. cm marking remained the same from the previous day when the echocardiogram measured 5.1 centimeters compared to when the echocardiogram measured 6.2 centimeters. the care team had no explanation for this. it was also reported that the nursing staff was changing the purge cassette and tubing when accidentally broke the purge side arm on the catheter which contributed to a purge fluid leak and purge pressure volume to be reduced to 30-110. the staff called the company for help with troubleshooting the issue but was advised that the pump needed to be exchanged immediately due to the fact that the motor was no longer being protected and also the patient was now at a risk for infection due to the crack on the purge sidearm. the patient was taken to the operating room for impella exchange. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the cause of the purge leak was damage to the purge sidearm; however, the cause of the damage is not known due to no product nor images returned and insufficient clinical details. the root cause of the stroke and tachycardia were unable to be determined due to insufficient clinical details.
2775 | 1220648-2025-48998 | 23802476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802476&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768545 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 106 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a change in the pump position after the patient was placed on ecmo. despite multiple attempts, the pump could not be repositioned. the pump was explanted and replaced with another impella cp. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position post ecmo insertion and cardiologist was unable to readvance. the device passed all post sterile inspection checks.
2776 | 1220648-2025-48999 | 23802498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802498&pc=OZD | 11/15/2025 | 01/22/2026 |  | 12/15/2025 |  |  | 11/16/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767566 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis which were identified by alanine aminotransferase and lactate dehydrogenase labs. the pump was replaced with another impella cp. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Investigation summary: product & data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history lot. device lot #: 1979714. device history batch. subcomponent lot#: n/a. device history review. pump sn: (b)(6) passed all post sterile inspection checks.
2777 | 1220648-2025-49000 | 23802582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802582&pc=OZD | 11/14/2025 | 01/05/2026 |  | 12/15/2025 |  |  | 11/16/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744993 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella 5.5 for support while working up for a heart transplant. it was reported that a small hematoma was noted around the surgical cut down site of the impella 5.5. the hematoma was resolved with sandbag pressure times two hours per the nursing staff. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
2778 | 1220648-2025-49001 | 23802803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802803&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766064 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 85 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Respiratory Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp became hypoxic and experienced tachyarrhythmia. later, the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the respiratory failure was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks
2779 | 1220648-2025-49002 | 23802819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802819&pc=OZD | 11/16/2025 |  |  | 12/15/2025 |  |  | 11/16/2025 | 12/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754208 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support during a surgical workup. the impella cp was prepped and inserted per instructions for use via the right femoral artery. it was reported that there was severe bleeding at the insertion site and the bleeding was not able to be controlled with manual pressure. the decision was made to remove the impella cp, insert a 14 french cook sheath, and send the patient for vascular surgery repair. the patient's outcome at the time of explant was survived and was noted to be in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary: asae/major bleed: severe bleeding at insertion site. unable to stop with manual pressure. the cause of the access site adverse event was not determined due to insufficient clinical details. device history lot: device lot: 1973983. device history batch: subcomponent lot: n/a. device history review: this pump sn (b)(6) passed all post sterile inspection checks.
2780 | 1220648-2025-49003 | 23802896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23802896&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731134 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a pump stop. the pump could not be restarted. the pump was explanted, and the patient expired. | Additional information received reporting the patient was off systemic anticoagulation for multiple days prior to the pump stop. the medical team was aware of the risks associated with being off systemic anticoagulation while on support. the pump is not available for return. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. pumo stop: the data logs were reviewed and on (b)(6) at 16:05 there were 2 motor current spikes with speed deviations and effect on pump flows right before alarm #119 pump stop. the logs on the day prior (b)(6) show a small motor current spike before alarm #119 pump stop. after pump was restarted, there was a motor current spike with speed deviation four minutes later with a sudden drop in pump flows. the cause of the issue was not established as no product was returned and limited clinical information was provided device history review: the device passed all post sterile inspection checks. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary 5.5 pump (b)(6) and purge cassette returned. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. pump stop: clinical reported a pump stop on (b)(6) after which restart attempts were unsuccessful which ultimately led to pump explant. the data logs were reviewed and on (b)(6) at 16:05 there were 2 motor current spikes with speed deviations and effect on pump flows right before alarm #119 pump stop. the logs on the day prior (b)(6) show a small motor current spike before alarm #119 pump stop. after pump was restarted, there was a motor current spike with speed deviation four minutes later with a sudden drop in pump flows. the returned product was inspected and a major kink in the catheter was found right beside the kink protector as well as biomaterial by the impeller. the electrical readings were obtained and an open line between phase 1 and 2 was observed. when the location of kink was flexed, electrical readings were obtained confirming location as source of electrical open. the root cause of the pump stop was determined to be due to an electrical open issue as evidenced by returned product analysis confirming varying electrical readings at catheter kink location and log analysis with electrical related alarm #119. device history review: the pump sn (b)(6) passed all post sterile inspection checks. | B5: additional event information received.
2781 | 1220648-2025-49005 | 23803024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803024&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714110 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bigeminy and arrhythmia. during repositioning of the pump, the patient went into ventricular fibrillation requiring defibrillation. the pump was explanted but the arrythmia and need for defibrillation continued. amio and gtt were initiated. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2782 | 23803063 | 23803063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803063&pc=OZD | 10/23/2025 | 10/27/2025 | 10/27/2025 | 12/15/2025 | 10/27/2025 | 12/15/2025 |  | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0000-US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Nurse | [6 empty] | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella did not alarm at all during pt trip to ct. minutes after returning to pt's room, impella device began alarming "placement signal not reliable". the impella screen showed no placement signals but still shows motor currents, flow, and purge system numbers. pocus [point-of-care ultrasound] showing current position of the impella was appropriate. dr. [redacted] repositioned the impella, which was later checked with a chest xray. 1-800 number listed on device called and the impella rep informed us that the sensor for the device placement likely just stopped working or had gone bad.
2783 | 1220648-2025-49006 | 23803220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803220&pc=OZD | 11/15/2025 | 12/15/2025 |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742164 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 79 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 71 yr old male was implanted with a impella cp to support for bridge to transplant. it was reported that bilateral lower extremity ischemia on impella with known illiac disease on non impella leg. patient decompensated overnight despite fem-fem bypass. however patient subsequently expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. (ischemia): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2784 | 1220648-2025-49004 | 23803226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803226&pc=PYX | 11/18/2025 |  |  | 12/15/2025 |  |  | 11/18/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026622141 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp. the patient's urine turned a bit red the following day. the pump was explanted the same day. the patient is in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. ppae/hematuria: after starting rp support, the patient's urine turned red. the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2785 | 1220648-2025-49007 | 23803275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803275&pc=OZD | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527820 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 118 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the access site. the patient was transfused two units of blood. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the access site bleed was most likely user related to high activated clotting time. the device passed all post sterile inspection checks.
2786 | 1220648-2025-49011 | 23803332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803332&pc=PYX | 11/15/2025 |  |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  | 01/31/2026 |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025471866 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced hematuria. an x-ray showed the pump was positioned deep. the pump was repositioned to resolve the issue. the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: data review: data logs showed pump ran without issue for about 44 minutes, then mc suddenly spiked to 1600 ma with ps shift and dropped in flows that triggered auto suction response and suction alarms. after that, pump ran without issue to the rest of the case. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely biomaterial ingestion since the data meets the pattern of biomaterial ingestion. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: pump passed all post sterile inspection checks.
2787 | 1220648-2025-49008 | 23803376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803376&pc=OZD | 11/15/2025 | 12/12/2025 |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763799 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 64 yr male was implanted with a impella 5.5 for support during high-risk cardiac procedure. it was reported that the patient has started oozing around the impella site. a hematoma has also formed at the site. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed and hematomahas been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed and hematoma could not be determined.
2788 | 1220648-2025-49012 | 23803401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803401&pc=OZD | 11/17/2025 |  |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689279 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ectopy due to the angle of insertion. the pump was removed, rinsed in d5w, and successfully reinserted. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product was returned. arrhythmia: the cause of the positioning issue was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was determined most likely to be an unintentional use error as the pump was out of position due to poor angle of impella insertion. device history review: device passed all post sterile inspection checks.
2789 | 1220648-2025-49013 | 23803411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803411&pc=OZD | 11/17/2025 | 02/27/2026 |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650646 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a sensor failure but pump function was not affected. the patient remains on impella support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | Corrections: d4-udi number updated e1-removed doctor name h6 component code changed from g04105 to g07001 h6 type of investigation-removed b13 the investigation for the optical sensor issue has been completed. since data logs were not available beyond calibration on the first console or on the second console and limited clinical details were provided, the cause of the optical sensor issue could not be determined. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2790 | 1220648-2025-49016 | 23803455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803455&pc=OZD | 11/17/2025 | 12/11/2025 |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026759866 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 75 yr old male was implanted with impella cp for support during high-risk cardiac procedure. it was reported that after extracorporeal cardiopulmonary resuscitation and clinical and professional skills assessment. insertion was attempted, but due to a concomitant dissecting aortic aneurysm, the false lumen expanded due to extracorporeal membrane oxygenation blood flow, preventing the wire from entering the true lumen. insertion was abandoned. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. vessel perforation: the cause of the injury was not determined due to insufficient clinical details. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and aortic aneurysm preventing successful delivery of impella. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted.
2791 | 1220648-2025-49014 | 23803492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803492&pc=OZD | 11/17/2025 |  |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766935 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient went into cardiac arrest and was shocked with automated external defibrillator x1 and received bystander cardiopulmonary resuscitation (cpr). the patient was intubated in field and taken to cath lab emergently. local team called for impella cp placement. while on support the patient had episode of ventricular fibrillation (vf) requiring cpr and defibrillation x1 and amio bolus and drip. echo was done post-code and impella measure slightly shallow so the doctor repositioned at bedside. arrhythmia not thought to be related to impella per the doctor. the patient had another episode of vf requiring defibrillation x1 and lidocaine drip started. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: ppae (ventricular fibrillation): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1982197. device history batch: subcomponent: n/a. device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2792 | 1220648-2025-49015 | 23803517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803517&pc=OZD | 11/16/2025 | 01/27/2026 |  | 12/15/2025 |  |  | 11/19/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026746547 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 suffered a cerebrovascular hemorrhage which was confirmed via head computed tomography scan. despite the insertion of the impella 5.5, the patient was unable to be weaned off extracorporeal membrane oxygenation (ecmo) and maintained at an activated clotting time of 200 or higher. before switching from an impella cp to an impella 5.5, maylon purge was being used as the purging fluid. the doctor said he did not believe the impella was the cause. the patient died due to loss of brainstem function following brainstem hemorrhage. ecmo could not be weaned. heparin administration was discontinued upon confirmation of cerebral hemorrhage. purging with sodium bicarbonate was also continued. the exact cause of death was unknown, but the doctor commented that the impella was not considered to be the cause of death. | The complainant reported additional information: there was no impella access site complication. |  |  |  |  | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B5 updated. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6. investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
2793 | 1220648-2025-49017 | 23803669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803669&pc=OZD | 11/18/2025 | 02/27/2026 |  | 12/15/2025 |  |  | 11/18/2025 | 12/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627820 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) failed to detect the purge disc while preparing it with the purge solution. a second purge cassette was tried but it was also not recognized. the aic was exchanged for a new controller to continue patient support. no patient harm was reported. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Corrections: b1-updated to add adverse event. b2-selected required intervention. h6 clinical code added e2343. h6 component code changed from g04035 to g07001. h6 type of investigation removed b13. the investigation for the purge disc not detected alarm has been completed. the root cause of the reported issue could not be determined.
2794 | 1220648-2025-49021 | 23803749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803749&pc=OZD | 09/18/2024 |  |  | 12/15/2025 |  |  | 09/18/2024 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025533271 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Death | [28 empty] | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. it was reported that the patient was escalated to the impella 5.5 from an impella cp after there was no hemodynamic improvement. there was small oozing noted from the tunnel site. the patient remained in critical condition and had a difficult night with frequent suction alarms. fluids could not be pulled on a continuous renal replacement therapy due to hypotension. the patient received epinephrine, levophed, vasopressin, and dobutamine. the patient was made do not resuscitate (dnr) and expired while on support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, data logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the data log analysis confirmed suction alarms occurring throughout the case. impella position unknown alarms were also seen throughout the case. no motor current spikes were seen outside of p-level changes. placement signal remained stable with no trends or spikes. no other abnormalities were seen. clinical details mention patient being escalated to impella 5.5 support due to no hemodynamic improvement on impella cp. there was small oozing noted from tunnel site of the 5.5 pump. the patient was in critical condition, and the pump had frequent suction alarms. patient was also on continuous renal replacement therapy (crrt) support which was unable to pull fluid due to hypotension. the patient expired while on support. the root cause of the pump suction was not determined due to inconclusive evidence from the provided clinical details and data log analysis. access site adverse event, the root cause of the access site bleeding could not be determined due to inconclusive evidence from the clinical details. and hypotension, the cause was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Medical safety/clinical reviewed the event. a 61-year-old male with a medical history significant for nonischemic cardiomyopathy (nicm) secondary to prior exposure to adriamycin and radiation therapy for hodgkin lymphoma, mitral regurgitation (mr) status post mitraclip implantation ×2, type 2 diabetes mellitus (dm2), and chronic kidney disease stage 3 (ckd3) presented to the emergency department on (b)(6) 2025, with decompensated heart failure. echocardiogram demonstrated moderate-to-severe mr with an ejection fraction (ef) of 25%. the patient was evaluated by the interventional cardiology team and deemed not a candidate for a third mitraclip procedure. an intra-aortic balloon pump (iabp) was placed on (b)(6) 2025, for hemodynamic support. on (b)(6) 2025, the patient experienced further clinical deterioration, including decreased urine output (<30 ml/hour), central venous pressure (cvp) of 15 mmhg, and mixed venous oxygen saturation (svo2) of 31%. aggressive diuresis was initiated overnight; however, urine output remained minimal. serum creatinine increased to 2.7 mg/dl and lactate to 2.3 mmol/l. blood-tinged and concentrated urine was observed prior to impella placement. the patient required escalating vasopressor support with norepinephrine (levophed) and epinephrine. a multidisciplinary shock team consultation was initiated on september 18, 2025. placement of an impella 5.5 was initially recommended but declined by cardiothoracic surgery. the treating physician elected to implant an impella cp with smartassist system in the catheterization laboratory. subsequently, abiomed¿s long-term outcome and quality indicator (loqi) impella registry reported that the patient developed acute kidney injury (aki) requiring dialysis, assessed as possibly related to the impella device. during hospitalization, the patient became anuric and required renal replacement therapy. automated impella controller (aic) waveforms were reported as appropriate. in the intensive care unit (icu), a bicarbonate-based purge solution and systemic heparin were administered per protocol. urine output remained minimal and bright red in color. hemolysis was considered; however, the team awaited laboratory results and echocardiographic evaluation prior to repositioning the device. there was no hemodynamic improvement four hours following impella cp implantation. on the following day, the impella 5.5 with smartassist system was implanted, and the impella cp was explanted. hemostasis was achieved at the explant site. a hematoma was noted at the impella cp insertion site post-explant; it was assessed as soft and without active drainage. on the same day as impella 5.5 implantation, minor oozing was reported from the device tunnel site. the pocket site was assessed as soft and without drainage. the patient remained critically ill and experienced recurrent suction alarms overnight. fluid removal was unsuccessful despite continuous renal replacement therapy (crrt) due to persistent hypotension. the patient¿s code status was changed to do not resuscitate (dnr), and supportive measures were continued. despite ongoing mechanical circulatory support, the patient could not be stabilized and subsequently expired. the event associated with the impella cp is being reported as a serious injury due to hematoma, hemolysis and acute kidney injury requiring dialysis. the impella 5.5 is associated with the outcome of death. however, the patient¿s underlying condition (society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock with multiorgan failure) contributed most to the demise outcome. note: h6. investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The complaint coding was updated to better reflect the reported event. therefore, h6 health effect - clinical / impact and medical device problem coding were updated/repopulated accordingly. note: h6 investigation: type, findings and conclusion codes remain unchanged as initially reported.
2795 | 1220648-2025-49022 | 23803752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803752&pc=OZD | 09/18/2024 |  |  | 12/15/2025 |  |  | 09/19/2024 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523306 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing high-risk percutaneous coronary intervention. while on support, the patient experienced acute kidney injury and became anuric requiring dialysis. in the intensive care unit (icu), bicarbonate purge solution and systemic heparin were ordered, awaiting pharmacy. urine output remained minimal with bright red color. in the intensive care unit (icu), bicarbonate purge solution and systemic heparin ordered, awaiting pharmacy. hemolysis was discussed but the team planned to await labs and an echocardiogram, prior to repositioning the pump again. there was no hemodynamic improvement 4 hours post impella cp implant; therefore, the impella cp was removed and an impella 5.5 was implanted. hemostasis was achieved. despite efforts, the patient declined, made do not resuscitate. | [5 empty] | A4 is unknown. no product was returned for investigation. data logs were returned for analysis. the root cause for the hemolysis was not determined due to lack of sufficient clinical details and inconclusive evidence from the data log analysis. the root cause of the access site hematoma was not determined due to lack of sufficient clinical details and unreturned product for further evaluation. the cause of the acute kidney failure was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | Medical safety/clinical reviewed the event. a 61-year-old male with a medical history significant for nonischemic cardiomyopathy (nicm) secondary to prior exposure to adriamycin and radiation therapy for hodgkin lymphoma, mitral regurgitation (mr) status post mitraclip implantation ×2, type 2 diabetes mellitus (dm2), and chronic kidney disease stage 3 (ckd3) presented to the emergency department on (b)(6) 2025, with decompensated heart failure. echocardiogram demonstrated moderate-to-severe mr with an ejection fraction (ef) of 25%. the patient was evaluated by the interventional cardiology team and deemed not a candidate for a third mitraclip procedure. an intra-aortic balloon pump (iabp) was placed on (b)(6) 2025, for hemodynamic support. on (b)(6) 2025, the patient experienced further clinical deterioration, including decreased urine output (<30 ml/hour), central venous pressure (cvp) of 15 mmhg, and mixed venous oxygen saturation (svo2) of 31%. aggressive diuresis was initiated overnight; however, urine output remained minimal. serum creatinine increased to 2.7 mg/dl and lactate to 2.3 mmol/l. blood-tinged and concentrated urine was observed prior to impella placement. the patient required escalating vasopressor support with norepinephrine (levophed) and epinephrine. a multidisciplinary shock team consultation was initiated on (b)(6) 2025. placement of an impella 5.5 was initially recommended but declined by cardiothoracic surgery. the treating physician elected to implant an impella cp with smartassist system in the catheterization laboratory. subsequently, abiomed¿s long-term outcome and quality indicator (loqi) impella registry reported that the patient developed acute kidney injury (aki) requiring dialysis, assessed as possibly related to the impella device. during hospitalization, the patient became anuric and required renal replacement therapy. automated impella controller (aic) waveforms were reported as appropriate. in the intensive care unit (icu), a bicarbonate-based purge solution and systemic heparin were administered per protocol. urine output remained minimal and bright red in color. hemolysis was considered; however, the team awaited laboratory results and echocardiographic evaluation prior to repositioning the device. there was no hemodynamic improvement four hours following impella cp implantation. on the following day, the impella 5.5 with smartassist system was implanted, and the impella cp was explanted. hemostasis was achieved at the explant site. a hematoma was noted at the impella cp insertion site post-explant; it was assessed as soft and without active drainage. on the same day as impella 5.5 implantation, minor oozing was reported from the device tunnel site. the pocket site was assessed as soft and without drainage. the patient remained critically ill and experienced recurrent suction alarms overnight. fluid removal was unsuccessful despite continuous renal replacement therapy (crrt) due to persistent hypotension. the patient¿s code status was changed to do not resuscitate (dnr), and supportive measures were continued. despite ongoing mechanical circulatory support, the patient could not be stabilized and subsequently expired. the event associated with the impella cp is being reported as a serious injury due to hematoma, hemolysis and acute kidney injury requiring dialysis. the impella 5.5 is associated with the outcome of death. however, the patient¿s underlying condition (society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock with multiorgan failure) contributed most to the demise outcome. note: h6. investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2796 | 1220648-2025-49023 | 23803759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803759&pc=OZD | 06/04/2025 |  |  | 12/15/2025 |  |  | 06/12/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675321 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced non-sustained ventricular tachycardia. the patient was treated with amio but expired on impella support. | [5 empty] | A4 is unknown. no product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2797 | 1220648-2025-49018 | 23803768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803768&pc=OZD | 11/18/2025 | 02/23/2026 |  | 12/15/2025 |  |  | 11/18/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714207 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Pumping Stopped | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was implanted with a impella 5.5 for support during high-risk cardiac procedure. it was reported that the impella plug showed defective cable on the red risks. however, it was discussed that the pump would not be replaced. it was reported that approximately two weeks later, pump stop motor current was increasing. re-start was tried without success the automated impella controller was replaced with no success. the patient was stable. | [5 empty] | D9 and h6 updated. the investigation is still ongoing. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Additional information updated: d.6b - explantation day, month and year.
2798 | 1220648-2025-49025 | 23803769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803769&pc=OZD | 08/11/2025 |  |  | 12/15/2025 |  |  | 09/24/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697144 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp for mechanical circulatory support. there was an attempt to use the impella after a minor acute myocardial infarction (ami), but due to the patient's large size, the medikit 16f short sheath could not be inserted. the impella implant was abandoned and the patient was supported by an intra-aortic balloon pump (iabp). in addition, the doctor commented that the problem was that the impella was opened too early. the patient recovered with iabp, so there was no impact on the outcome. | [5 empty] | No product was returned. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and large size vessel preventing successful delivery of impella. the device passed all post sterile inspection checks.
2799 | 1220648-2025-49026 | 23803914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23803914&pc=OZD | 11/18/2025 | 01/12/2026 |  | 12/15/2025 |  |  | 11/18/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026770117 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. the patient was implanted with an impella cp instead. | [5 empty] | Investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and resistance at aorta preventing successful delivery of impella. device history lot device lot: 1974008. device history batch subcomponent lot: n/a. device history review device (sn: (b)(6)) passed all post sterile inspection checks. | A2, a4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2800 | 1220648-2025-49028 | 23804264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23804264&pc=OZD | 11/28/2025 |  |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714639 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | Pumping Stopped | Malposition of Device |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced an unexpected pump stop. the pump was restarted with saturated flow and high motor current. the pump was explanted and the patient was put on ecmo. later, the patient was diagnosed with heparin induced thrombocytopenia. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. data logs returned for review. saturated flow/pump stop: lt logs from 11/17 to 11/27 showed no saturated flows, pump stop issue. data logs returned are insufficient for lt logs in relevance to pump stop timeline. the cause of saturated flow/pump stop issue was not determined due to inconclusive log evidence and since no product was returned. device in wrong position: as per clinical, heart surgeon had pulled out the impella out of ventricle prior to the explant. it was correctly positioned in ventricle prior to that. the cause of device in wrong position is use related since pump was pulled out of ventricle. thrombocytopenia/major bleed: the cause of the injury was not determined due to insufficient clinical details. thrombosis: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the complaint pump passed all post sterile inspection checks.
2801 | 1220648-2025-49009 | 23804477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23804477&pc=OZD | 11/15/2025 | 12/31/2025 |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768470 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 76 KG | Hispanic |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella cp for support during high-risk cardiac procedure. it was reported that the patient came in as a st-elevation myocardial infarction. an impella cp was placed in the cardiac catheterization laboratory (ccl). the sheath was exchanged, peel away to repo sheath at end of procedure. a small hematoma was noted. in recovery, the nurse noted that hematoma was significantly getting bigger. manual pressure was held and the doctor was called. the decision was made to take patient back to the ccl and do an angiogram of the right femoral artery. the doctor got access in the right pedal artery. the catheter was advanced and an angiogram was performed. a perforation was noted. a peripheral balloon was inserted to attempt to tamponade the area. they were unsuccessful after multiple inflations. the decision was then made to remove impella and deploy percloses times two. post angiogram revealed no perforation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported vessel perforation and hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vessel perforation and hematoma could not be determined.
2802 | 1220648-2025-49029 | 23804608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23804608&pc=OZD | 11/17/2025 |  |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308462 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) displays the error "controller failure¿ when starten up. the aic was replaced and there was no patient harm. | [5 empty] | The impella device was received from the customer, therefore, investigation of the device is underway. once its completed a supplemental mdr will be filed.
2803 | 1220648-2025-49030 | 23805027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23805027&pc=OZD | 11/16/2025 | 01/12/2026 |  | 12/15/2025 |  |  | 11/16/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026745769 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on support day 15, it was reported that there was an alarm for placement signal not reliable. the placement signal was noted to no longer be functioning. there was no effect on patient support and no patient harm was reported. at the time of writing this report, the patient continues to be supported by impella 5.5 while awaiting a donor heart for transplantation. | [5 empty] | D6b updated. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: the cause of the issue was not established as no product or data logs were returned for analysis. device history review: the device passed all post sterile inspection checks.
2804 | 1220648-2025-49031 | 23805153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23805153&pc=OZD | 08/27/2025 |  |  | 12/15/2025 |  |  | 08/25/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302709 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal not reliable alarm. the pump position was verified via echo. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2805 | 1220648-2025-49032 | 23805159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23805159&pc=OZD | 09/03/2025 |  |  | 12/15/2025 |  |  | 09/03/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Failure to Analyze Signal | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal low alarm. the device was repositioned to resolve the issue. | [5 empty] | No product was returned for investigation. the cause of the optical signal issue was improper positioning of the pump that resolved with pump repositing. the cause of the hemolysis was improper positing of the device. the pump passed all post sterile inspection checks.
2806 | 1220648-2025-49036 | 23805835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23805835&pc=OZD | 03/01/2024 |  |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  | 10/31/2025 |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420160 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team was notified of an adverse event of access site hematoma from st-elevation myocardial infarction door-to-unload determined relationship to impella device to be definite. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product returned. hematoma: the cause of the access site adverse event was not determined due to insufficient information. device history review device passed all post sterile inspection checks.
2807 | 1220648-2025-49035 | 23806586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23806586&pc=OZD | 11/16/2025 | 01/12/2026 |  | 12/15/2025 |  |  | 11/16/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728952 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support following a triple coronary bypass graft (cabg) procedure. the patient developed a thrombus in her left anterior descending coronary artery which led to an acute decreased ejection fraction from 40% to 20% at the beginning of the case and the physician decided to place the impella 5.5 for the patient to come off bypass. after a successful cabg, the impella 5.5 was inserted via direct aortic approach, and support was initiated without complications. the patient was transferred to the intensive care unit for rest and recovery with a plan to wean and explant the following monday. on support day 4, the patient experienced an episode of atrial fibrillation with rapid ventricular response. the impella 5.5 position was verified with an echocardiogram. no additional information was provided on treatment given to the patient for the arrhythmia. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: ppae (paroxysmal atrial fibrillation): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1962573. device history batch: subcomponent: n/a. device history review the pump sn (b)(6) passed all the post sterile inspection checks.
2808 | 1220648-2025-49027 | 23806741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23806741&pc=OZD | 11/17/2025 |  |  | 12/15/2025 |  |  | 11/18/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757296 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male | 95 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | Insufficient Information | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a thrombus in the descending aorta, attached to the impella catheter. venoarterial extracorporeal membrane oxygenation was in place. neurological status of the patient after long cardiopulmonary resuscitation (cpr) and mitral valve replacement was still uncertain. great amounts of platelets and plasma were given after surgery. three days later it was reported the a cardiac computed tomography scan revealed hypoxic brain damage. an embolic stroke was noted. therapy was to be limited/terminated. the pump was explanted and the patient expired shortly after the explant. the patient expired due to hypoxic brain damage, most likely as a result of prolonged cpr. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death". section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers¿. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. additional information: the following was information was received: in this case the prolonged cardiopulmonary resuscitation (cpr) took place before the impella was used. the hypoxic brain damage was according to the physician- a result of the low-flow time during the cpr. the impella 5.5 was removed due to the failure mode (thrombus at impella catheter) and the missing prognosis of the patient. investigation summary: the device and a purge cassette returned. returned complaint cp pump visually inspected. dried dextrose around outflow cage and small thrombus wrapped around pigtail. after soaking in water dextrose dissolved and thrombus remained. thrombosis/ stroke: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2809 | 1220648-2025-49038 | 23808269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23808269&pc=OZD | 11/17/2025 | 02/24/2026 |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026756361 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The medical doctor reported that the purge flow temporarily dropped below 1 ml/h and the purge pressure is at 1040 mmhg. no kinks are visible in the tubing. the doctor requests a lysis protocol. the lysis protocol was sent, and the doctor switched to tissue plasminogen activator. the purge pressure was now in range. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The pump and purge cassette was returned. high purge pressure: clinical details noted that a temporary drop in purge flow with increase in purge pressure. no visible kinks in purge line. tpa was administered and later the same day, purge values returned to normal range. pump ran without issue for an additional 6 days. returned product was visually inspected and biomaterial was present on impeller blade with but purge gap was clear. pump was connected to console and purged for approximately 10 minutes without "high purge pressure" alarm. pump was attempted to be started with immediate pump stop occurring. pump was continued to purged without high purge pressure alarms for 5 more minutes. no data logs were returned for analysis. the cause of the high purge pressure could not be determined as issue was unable to be reproduced with returned product and no logs were returned for analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2810 | 1220648-2025-49039 | 23808931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23808931&pc=OZD | 11/17/2025 |  |  | 12/15/2025 |  |  | 11/17/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766620 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female | 59 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient went to the emergency room with chest pain, dyspnea on exertion, increase in trop, rapid drop in ejection fraction. 45% dropped to 25%. massive increase in pressor and inotropic support. the patient was taken to the catheterization lab for impella cp placement. while on support, the medical team described the presence of left ventricle (lv) thrombus observed at the apex away from the impella inflow or pigtail. lv thrombus was believed to be a result of patients severe lv dysfunction, or other cause. the pump was later explanted and the patient survived. | Additional information was received. the patient was transported to a facility. it is unknown if the impella cp was kept for investigation. data logs will try to be sent for investigation. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | D1 revised brand name as it was entered incorrectly on the initial report that was submitted. d4 revised model as it was entered incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. peri-procedural adverse event/thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2811 | 1220648-2025-49040 | 23810192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23810192&pc=OZD | 06/28/2025 |  |  | 12/15/2025 |  |  | 06/29/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1209412 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella 5.5 was complicated by difficulty crossing the aortic valve. upon implantation, there was an optical signal issue with low pump flow and a placement signal. the pump was positioned deep so it was retracted and the p value of the pump was increased to resolve the issues. additionally, the patient developed a fever and a clot was suspected. after the pump was explanted the patient developed drainage from the graft site. a culture of the fluid tested positive and the patient was treated with antibiotics. the patient also underwent a washout in the operating room. the patient was in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2812 | 1220648-2025-49041 | 23813091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23813091&pc=OZD | 11/15/2025 | 12/15/2025 |  | 12/15/2025 |  |  | 11/15/2025 | 12/15/2025 |  |  |  |  | 12/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717557 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 60 yr old male was implanted with a impella 5.5 for support during high risk cardiac procedure. it was reported that fiber optic cable appears to have malfunctioned. the nurse called into the cardiac surgical center with an active placement signal not reliable alarm. the nurse was walked through on how to disable the audio to this alarm. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. no further information was provided regarding the placement signal alarm. due to lack of sufficient clinical information, and no product or data logs returned, the root cause could not be established.
2813 | 1220648-2025-31951 | 23787911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23787911&pc=OZD | 03/03/2025 | 12/12/2025 |  | 12/12/2025 |  |  | 03/03/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569311 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support there were placement signal alarms. there were no flow, motor current, or purge issues and support continued. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the device was not returned; however, based on data log review, the root cause of the optical signal issue was determined to most likely be a damaged optical fiber.
2814 | 1220648-2025-48952 | 23787956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23787956&pc=OZD | 11/13/2025 |  |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186355 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with sepsis in cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while on support it was reported that the left ventricle signal/display on the automated impella controller (aic) was not displaying at p2. no other information was provided. this report is for the aic and there will be an additional report sent for the impella 5.5 (serial number 581846). | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6a and d6b: added implant and explant dates. d9: added device return information. investigation summary: the device was returned for evaluation. no issues were found on console (b)(4). the reported symptom ¿ ¿left ventricle signal/display on the automated impella controller (aic) was not displaying at p2¿ ¿ was consistent with the design requirement that lv signals are displayed only when the p-level is above p3.
2815 | 1220648-2025-48953 | 23787996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23787996&pc=OZD | 05/25/2025 |  |  | 12/12/2025 |  |  | 11/12/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded publication entitled: "platelet derived prothrombotic microparticles and adverse events in patients supported with impella 5.5". case study patient #2 was in cardiogenic shock and was put on impella 5.5 support. while on support the patient had ventricular fibrillation and was cardioverted by one shock and medical iv therapy was escalated. after 13 days on support the 5.5 was explanted. four days after explant, they observed a neurological event and a minor stroke was diagnosed. causal relationship to pump is not known as the fibrillation occurred during the support. the patient survived. this complaint is related to another report for the same publication but for a different patient. this report is for case patient #2 and case patient #1 is reported in another report. refer to section h10 for the related report number. | [5 empty] | The device serial number, lot number and udi are unknown. therefore, the device expiration date and device manufacture device are unknown. implant and explant dates are also unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. d¿ettore, n., petruzza, a., cardinale, a., bensi, t., maj, g., ciriello, m. m., & pappalardo, f. (2025). platelet derived prothrombotic microparticles and adverse events in patients supported with impella 5.5. the international journal of artificial organs, 48(10), 794¿799. https://doi.org/10.1177/03913988251357199. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. ventricular fibrillation, stroke: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details.
2816 | 1220648-2025-48954 | 23788043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788043&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585409 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 52 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was taken to the operating room for possible tissue aortic valve replacement, tissue mitral valve regurgitation, left atrial appendage ligation, with impella 5.5 support. prior to closing, the patient became hypotensive and coagulopathic. the patient received cryo 10, two units of platelet packs, two units of fresh frozen plasma, and prothrombin complex concentrate. the chest was then closed and the impella was secured to the skin utilizing the provided three point fixation lock. the patient remains on support. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hypotension. the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2817 | 1220648-2025-48945 | 23788094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788094&pc=OZD | 10/08/2025 | 01/12/2026 |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711625 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for high risk cardiac procedure. percutaneous coronary intervention was performed after impella cp insertion for acute myocardial infarction, and for the first time, perclose was used for removal. in addition to control bleeding, perclose was performed while the balloon was inflated more proximally than the insertion site via the right atrium. although bleeding was successfully stopped, bleeding occurred at the balloon inflation site, requiring stent placement. vascular dissection was caused by the balloon used during removal. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | Investigation summary: no product was returned. vascular dissection: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the bleed was not determined due to insufficient clinical details. device history lot. device lot: 1932133. device history batch subcomponent lot: n/a device history review device (sn: 601947) passed all post sterile inspection checks.
2818 | 1220648-2025-48955 | 23788283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788283&pc=OZD | 11/12/2025 |  |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726704 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs). during the case, the physician elected to retain the peel-away sheath. associated risks were discussed with the physician and team. the patient was transferred to the unit on impella mcs which was maintained for a total of five days. there was no report or indication of device related complaints or performance issues. on day five of support, the patient's clinical symptoms worsened, and the patient subsequently expired on support. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, hemodynamic instability, the cause was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Patient-specific information a4. weight is unknown. medical safety review. medical safety reviewed the event and noted: a patient with cardiogenic shock underwent placement of an impella cp via femoral artery. during the procedure, the physician elected to retain the peel-away sheath despite the device instructions recommending removal. the associated risks were discussed with the physician and clinical staff. the patient remained on impella support for five days without device-related complaints or performance issues. on day five, the patient experienced worsening clinical symptoms and subsequently expired while still on support. no device malfunction or performance issue was reported. the event is conservatively reportable because the patient expired while on impella cp support however, the patient death was most likely due to worsening clinical symptoms and unlikely unrelated to device performance. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2819 | 1220648-2025-48956 | 23788465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788465&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686538 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support following a mitral valve replacement (mvr) procedure. the patient was originally admitted to the hospital on (b)(6) 2025 was found to be in cardiogenic shock. an impella cp was placed, and a percutaneous coronary intervention was performed to the right coronary artery. additionally, during this time the patient was septic with endocarditis which was possibly due to a recent foot infection that required a toe amputation. the patient has remained in the hospital since the impella cp explant, and the date of explant was not available in the complaint record at the time of writing this report. the patient was then taken for an mvr with a planned impella 5.5 placement following the surgery. following a successful mvr procedure, the impella 5.5 was prepped and placed via the axillary artery, specific side was not reported, without complications. a transesophageal echocardiogram (tee) was performed and revealed that the pumps inlet was facing towards the mitral valve, and the physician attempted to reposition the device but had difficulty. another physician came to assist and was able to successfully position the device out of the mitral valve area. following the impella 5.5 implant, the patient was taken to the intensive care unit for rest and recovery. on support day two, it was reported that the patient had an episode of self-converting ventricular tachycardia. on support day four, tee was performed and revealed the impella was deep in the ventricle. the physician pulled the pump back one centimeter, with a final distance of 5.5 centimeters from the aortic valve area. on support day five, it was reported that the patient had a moderate amount of serous fluid output from the chest tube. the patient received two units of packed red blood cells. it was noted that the impella had to be pulled back one centimeter during the morning when tee revealed the pump was too deep at seven centimeters. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: ppae (tachycardia). the root cause of the injury was unable to be determined due to insufficient clinical details. ppae (major bleed). the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1930001. device history review: the pump sn (b)(6) passed all the post sterile inspection checks. | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2820 | 1220648-2025-48957 | 23788492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788492&pc=OZD | 11/13/2025 |  |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728649 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced loss of pump flow and a suction alarm after patient movement. an echocardiogram confirmed the pump was in the aorta. the pump was repositioned but flows did not return. the p level of the pump was increased to resolve the issue. later, the patient experienced atrial fibrillation requiring cardioversion. the patient is in stable condition. | [5 empty] | H6 code was added as it was not on the initial report that was submitted. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2821 | 1220648-2025-48958 | 23788568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788568&pc=OZD | 11/12/2025 | 01/06/2026 |  | 12/12/2025 |  |  | 11/12/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688954 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 110 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was also administered methylene blue. | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support and experienced complications. the patient was transferred from an outside hospital on (b)(6) 2025 to the complainant hospital due to progressively worsening cardiogenic shock and the patient required veno-arterial extracorporeal membrane oxygenation (va ecmo) cannulation. on (b)(6) 2025 the patient had an ablation which caused worsening shock and bleeding and required the va ecmo to be reconfigured to veno-arterio-venous extracorporeal membrane oxygenation (vav ecmo). on (b)(6) 2024, the vav ecmo flows were reduced, and the decision was made to decannulate the arterial cannula and insert an impella 5.5 for continued cardiac support, with a plan to de-escalate veno-venous extracorporeal membrane oxygenation (vv ecmo) support with the impella 5.5. the impella 5.5 was prepped, inserted, and support was initiated without complications. on support day two, it was noted that the patient remained on vv ecmo with impella 5.5 support. the patient¿s distal extremities were noted to be cool and dusky. the patient¿s liver function tests were elevated, and a liver ultrasound was ordered. on support day four, it was noted that the patient had an elevated plasma free hemoglobin (pfhgb), and it was noted that the results had been decreasing. the patient was administered with methylene blue and flows on the vv ecmo were turned down. the following day, it was noted that the patient did not have signs of active hemolysis. on support day six, the patient was being weaned from vv ecmo and maintained on impella 5.5 support. the patient was noted to be hemolyzing with a pfhgb greater than 40 mg/dl. the medical team suspected the hemolysis was possibly due to the impella. on support day nine, the patient was on continuous renal replacement therapy, and it was noted that the urine output was red tinged. additionally, it was noted that the patient received six blood products that morning. the patient and family decided to withdraw care and the patient was placed on comfort measures only. the patient was noted to have expired at impella 5.5 explant. |  |  |  |  | B5 updated. the impella device has not been received from the customer; therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B2 added selection that was omitted from previously submitted reports. investigation summary: the investigation has been completed. product and data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2822 | 1220648-2025-48959 | 23788573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788573&pc=OZD | 11/13/2025 |  |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1388684 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller had a self check failure upon start up. the controller was removed from service. there was no patient involvement. | [5 empty] | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. data analysis: the console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the automated impella controller console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. during firmware execution, the sau_powermod_1 hardware revision displayed as ¿n/a¿ and the firmware version also showed ¿n/a.¿ this occurred due to a communication issue between the device and the system. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. product history review summary: there were no issues related to complaint discovered when reviewing the product history of the console. | The reported event did not involve patient use. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2823 | 1220648-2025-48961 | 23788759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788759&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770545 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with st-segment elevation myocardial infarction- severe left main into chronically diseased left anterior descending artery with acutely diseased circumflex. the patient was placed on impella cp for support to assisted percutaneous coronary intervention (pci). pci was performed on the left main to circumflex and left anterior descending - culotte stenting. the impella was weaned down and removed successfully. it was reported that femstop was applied due to hematoma. the patient survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the access site adverse event: the cause of the bleed was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2824 | 1220648-2025-48962 | 23788950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23788950&pc=OZD | 11/14/2025 | 12/11/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765638 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The decision made to pull back on impella to attempt to fix the kink that was observed on fluoroscopy. while attempting to reposition the device, the impella cp was pulled back across aortic valve. it was at this point that the impella cp was removed. | The user facility reported that a patient with implanted with an impella cp for support during high-risk cardiac procedure. it was reported that the impella cp was placed in standard fashion and a kink was noticed on fluoroscopy in the impella cannula below the outlet window. impaired flows caused the physician to remove the pump and place a second impella cp. |  |  |  |  | The investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the pump was returned for investigation. a cannula kink was noted on the returned product. no other damage or abnormalities were found on the returned product. the pump was then tested in a bucket of water without any issue. data log review was not provided or retrieved from the remote server. low or blocked pump flow / product damage (cannula kink): clinical details indicate that the cp was inserted during compressions. after implantation, flows were low at 1.5¿2 l/min. fluoroscopy revealed a kink in the impella cannula below the outlet window. despite standard troubleshooting and support measures (including repositioning, vasopressors, and inotropes), pump flows did not improve. the low flow was most likely related to the kink observed in the cannula, as confirmed by imaging and also observed on the returned product. the cause of the cannula kink was most likely related to unintended use, since the pump was implanted during compressions. device in wrong position: no defect was found on the returned pump related to the positioning issue. per clinical details, during attempts to fix the kink and improve flow, the impella cp was inadvertently pulled back across the aortic valve, resulting in malposition. the device was subsequently explanted and replaced following balloon aortic valvuloplasty. the positioning issue was most likely use-related, associated with procedural manipulation during the case. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | B5 updated. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2825 | 1220648-2025-48960 | 23789034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23789034&pc=OZD | 10/23/2025 |  |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026718852 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. an interview was conducted with a cardiovascular surgeon. we asked about the cerebral hemorrhage and whether it was related to impella, but he said the connection is unknown. he commented that it cannot be said that there is no connection at all, but that the detailed cause is unknown. death, cause of death: ami/vsp. indirect cause of death: subarachnoid hemorrhage. relation to impella: unknown based on doctor's comments. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
2826 | 1220648-2025-48963 | 23789133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23789133&pc=OZD | 11/14/2025 | 12/24/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733595 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ischemia | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for high risk cardiac procedure. there was slight ooze from the impella site, that was alleviated by manual pressure and purse string suture. after the impella was in auto, a seven french companion sheath was placed for percutaneous coronary intervention and percutaneous transluminal coronary angioplasty to orca and the patient began to have bradycardia. 1/2 amp of atropine was given. the patient was successfully weaned at bedside, remained hemodynamically stable during and post explant, but after explant it was noted that the right foot lost pulses and became pale but remained warm. the patient went to the operating room for an endarterectomy. the patient was now doing well and extubated. | [5 empty] | The investigation for the reported ischemia, bradycardia, hypotension, and major bleed has been completed. no product was returned for review. the root cause of the ischemia, bradycardia, and hypotension was unable to be determined due to insufficient clinical details. the root cause of the major bleed was user related because bleeding resolved with purse string suture. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
2827 | 1220648-2025-48965 | 23789225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23789225&pc=OZD | 11/12/2025 | 12/12/2025 |  | 12/12/2025 |  |  | 11/12/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026716045 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella cp for support during a high risk cardiac procedure. it was reported that the patient had synchronized cardioversion for initial heart rhythm 200 beats per minute (bpm) supraventricular tachycardia versus ventricular tachycardia. the patient was then converted to paced rhythm set at 100 bpm. heparin was paused due to low platelet count. the right distal perfusion catheter appeared to be low or no flow. the impella cp was discontinued and the patient was implanted with an impella 5.5. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary: no product was returned for investigation. tachycardia/ ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1958334. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
2828 | 1220648-2025-48966 | 23789318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23789318&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766537 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient had some short runs of ventricular tachycardia and the team repositioned the pump. later the pump was explanted and the patient survived. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: tachycardia. in order to make a cause determination, the clinical would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: this pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2829 | 1220648-2025-48964 | 23789341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23789341&pc=OZD | 11/12/2025 | 01/13/2026 |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735966 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient was on dialysis and had severe calcification (failure to advance) . the pig and gw crossed the a valve but not the impella. the complainant reported a patient was implanted with an impella device for mechanical circulatory support. a 12mm bav was attempted, but the condition did not improve, so insertion was abandoned. pci was performed with ecmo and iabp. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had calcification and resistance at aortic arch preventing successful delivery of impella. device history lot device lot: 1962843 device history batch subcomponent lot: n/a device history review device (sn: 618869) passed all post sterile inspection checks. | E1 corrected initial reporter name as it was in japanese in the initial report that was submitted.
2830 | 1220648-2025-48967 | 23791342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23791342&pc=OZD | 11/13/2025 | 02/12/2026 |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490417 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Inaccurate Flow Rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that a 67 yr old male was implanted with a impella cp for suport during high risk cardiac procedure. it was reported that the patient presented to cath lab for complex high-risk percutaneous coronary intervention with impella support. the right femoral artery prepped and accessed following best practices. impella inserted under best practice recommendations. upon start of device after a couple of minutes in auto mode, waveforms revealing inverted. ao and lv waveforms with trending high mc and saturated flows of 4-4.3 lpm. after discussion with the doctor. the decision was made to wean and remove cp device and exchange to new cp device. device weaned and removed successfully. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: cp pump sn(b)(6), usb and purge cassette . high or saturated pump flow: clinical details note that upon start of device after a couple of minutes in auto mode, waveforms revealing inverted. ao and lv waveforms with trending high mc and saturated flows of 4-4.3 lpm. data log review confirmed saturated flow upon pump startup. there were no mc spikes outside of the expected spike during pump startup. purge flow increased from ~15 ml/hr to almost 30 ml/hr after pump startup with purge pressure between 300 and 400 mmhg. there were no alarms during the saturated flow. product was returned and evaluated. there was no sign of biomaterial upon initial inspection and no blood ingress into the purge line just past the motor housing. after soaking in warm water, the pump would not start in a bucket of water when attempting to start it up. the pump was then soaked in tergazyme solution and then ran initially with saturated flow, which then resolved into normal pump flow after a few minutes. the pump motor housing was torn down and there was no clear evidence of biomaterial on the components. there was one speck of what appeared to be dried blood on the magnet. given the evidence that the pump initially reproduced saturated flow that then resolved into normal pump flow, the cause of the saturated flow was most likely biomaterial present within the motor housing. optical sensor issue: clinical details note that upon start of device after a couple of minutes in auto mode, waveforms revealing inverted. md suspected it was optical sensor issue, previous echo showing no lv thrombus. the pump was returned and evaluated. there were no abnormalities found with the optical parameters or sensor. data log review confirmed the uncoupling of the ao and lv signals with the lv trace being higher. however, there were no abnormalities with the optical parameters or the ao placement signal. since the lv signal is calculated from placement signal and motor current, a high motor current resulting from biomaterial interaction would result in a higher lv signal, likely decoupled from the ao signal as seen in this case. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. high or saturated pump flow: clinical details note that upon start of device after a couple of minutes in auto mode, waveforms revealing inverted. aortic valve (ao) and left ventricle (lv) waveforms with trending high mean corpuscular volume (mc) and saturated flows of 4-4.3 lpm. data log review confirmed saturated flow upon pump startup. there were no mc spikes outside of the expected spike during pump startup. purge flow increased from ~15 ml/hr to almost 30 ml/hr after pump startup with purge pressure between 300 and 400 mmhg. there were no alarms during the saturated flow. the cause of the saturated flow was not determined since product was not returned. optical sensor issue: clinical details note that upon start of device after a couple of minutes in auto mode, waveforms revealing inverted. the physician suspected it was optical sensor issue, previous echo showing no lv thrombus. data log review confirmed the uncoupling of the ao and lv signals with the lv trace being higher. however, there were no abnormalities with the optical parameters. the cause of the optical sensor issue was not determined since product was not returned. device history review ==> the pump passed all post-sterile inspection checks. | Correction: e4. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2831 | 1220648-2025-48968 | 23792469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23792469&pc=OZD | 11/12/2025 | 12/22/2025 |  | 12/12/2025 |  |  | 11/12/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769648 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp developed a hematoma and bleeding at the groin site. the patient was in stable condition. | [5 empty] | Date of explant, d6b, has been updated. | The cause of the hematoma and major bleed is most likely use related since the patient arrived to cicu with hematoma at groin site. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2832 | 1220648-2025-48969 | 23793462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23793462&pc=OZD | 11/12/2025 |  |  | 12/12/2025 |  |  | 11/13/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534791 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Pericardial Effusion | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp showed a possible pericardial effusion, thrombosis, and early tamponade via echocardiogram. the pump was explanted and the patient survived. | [5 empty] | No product was returned for investigation. the cause of the cardiac tamponade and pericardial effusion was not determined due to insufficient clinical details. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2833 | 1220648-2025-48971 | 23794095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794095&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735923 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the patients urine is pink and plasma free and hgb is elevated. there was concern for hemolysis. the pump was repositioned as the impella as it was deep in the left ventricle. it was also reported that bilateral lower extremities were cold and unable to doppler pulses. narrow pulse pressure (5-10mmhg). the pump was explanted and patient survived. | [5 empty] | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the pump was not returned for investigation. data log review was not provided or could not be retrieved from the remote server. the root cause of the hemolysis issue was not determined without returned product review and sufficient clinical details, including data logs. the root cause of clinical symptoms, ischemia, low blood pressure/hypotension and hematuria could not be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2834 | 1220648-2025-48972 | 23794164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794164&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026760968 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding from the chest tube. the patient was transfused five units of packed red blood cells and taken to surgery. during the surgery, the patient's left ventricle was perforated. the patient was placed on cardiopulmonary bybass while the left ventricle was surgically repaired. the patient was in stable condition. | [5 empty] | The device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. no physical damage was observed on the returned product. biomaterial was found on the outflow cage, which was removed before testing in a bucket of water. the pump ran as expected in the bucket of water. data log review was not required for this case. clinical symptoms (major bleed): the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with tissue (cardiac perforation): no pump damage was observed on the returned device. clinical details indicate that left ventricular perforation occurred during manipulation of the lv while exploring for bleeding sources in the or. the cause of the cardiac perforation was most likely related to unintentional use, specifically related to heart manipulation. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2835 | 1220648-2025-48970 | 23794165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794165&pc=OZD | 11/12/2025 |  |  | 12/12/2025 |  |  | 11/12/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Degraded | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the cover of the white plug on the automated impella controller broke upon insertion of the impella catheter. a new controller was used to support the patient. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2836 | 1220648-2025-48974 | 23794173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794173&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762933 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a placement signal not reliable alarm. the patient is in stable condition. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. optical sensor issue: clinical details report a placement signal not reliable alarm on the date of implant. shockwave was being used in the riva/lad at the time the issue was noted, though the distance between the device and the optical sensor was not clear. data logs were reviewed and the alarm was confirmed. roughly three hours into the case, there was a sudden drop of signal to noise ratio to 0, lamp increased, contrast dropped, gain dropped, and light remained stable. these are indications of a sensor break. product was not returned for investigation. based on the clinical details provided in conjunction with data log signatures, the cause of the optical signal issue was determined to be a device interaction that resulted in sensor damage. since the distance between the devices at the time of damage was unknown, the more proximal cause could not be determined.
2837 | 1220648-2025-48973 | 23794176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794176&pc=OZD | 11/14/2025 | 12/12/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 68 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 79 yr old male was implanted with a impella cp for support during high risk cardiac procedure. it was reported that the heparin was held due to pericardial bleeding. impella increase to p-8. impella was successfully explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: ppae: major bleed. the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1930021. device history review: this pump sn (b)(6) passed all post sterile inspection checks.
2838 | 1220648-2025-48977 | 23794182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794182&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763735 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced low flow and suction alarms. the position of the pump was confirmed to be correct. an embolus was seen in the inlet of the pump. the pump was exchanged for a new pump to resolve the issue. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. the impella cp was returned. low pump flow - the cause of the low pump flow was not determined due to low flows not recreating on returned product, no data logs recovered, and insufficient clinical details. patient codes - thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. | Correction : section h6 :removed pec high or saturated flows, and added pec low or blocked pump flows as the voc states the pump max flow at p-3 was 1 l/min and the md comment of 'pump only had 1l/min flow.' the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2839 | 1220648-2025-48980 | 23794223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794223&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707307 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 38 KG |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a stroke. and expired. | [5 empty] | The root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2840 | 1220648-2025-48978 | 23794224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794224&pc=OZD | 11/14/2025 | 12/10/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735081 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 50 yr old male with an impella 5.5 for support to wean off ecmo and bridge to recovery. it was reported that 5.5 suddenly stopped after about a week of use. ecmo had to be inserted. there was no visible blood clot. this is a problem with the 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2841 | 1220648-2025-48981 | 23794241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794241&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756435 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 81 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a fever and expired. | [5 empty] | The cause of the fever was not determined due to insufficient clinical details the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2842 | 1220648-2025-48976 | 23794254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794254&pc=OZD | 11/14/2025 | 12/11/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697655 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63 yr old male with a impella cp for support during high risk cardiac procedure.it was reported that patient went into cardiac arrest post impella placement. patient subsequently placed on venoarterial extracorporeal membrane oxygenation. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. (cardiac arrest). in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2843 | 1220648-2025-48982 | 23794262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794262&pc=OZD | 11/14/2025 | 12/10/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728761 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 59 yr old male was implanted with an impella cp for high risk cardiac procedure. impella placed for cardiogenic shock left femoral, patient on vaso, levo, and dobutamine, lactate 11. after completion of procedure and exchanging to repositioning sheath a femoral angiogram was done via the wire reaches port. it was noted that very little flow was making it beyond access site. decision made to place antegrade perfusion sheath on the left side and connect to a retrograde sheath on the right side creating an external fem-fem bypass. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2844 | 1220648-2025-48983 | 23794275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794275&pc=OZD | 11/14/2025 | 01/27/2026 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747051 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient had worsening heart failure symptoms and is now in stage 3 heart failure with an ejection fraction of 20%. due to most recent admission, he was scheduled for an impella 5.5 placement and will be listed for transplant the following day. during impella support it was reported that there were slight concerns for hemolysis as lactate dehydrogenase is ~300s, previous result hemolyzed and plasma free hemoglobin spiked to 100 and is now down to 29. heart failure team plans to do echo today to verify placement. the patient remains on support. | The complainant reported the patient had high fever and white blood cell count. suction alarms and placement signal low alarms/impella position unknown also occurred. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B2 revised selection as the selection on the initial report that was submitted was incorrect. e1 added initial reporter phone and zip code extension as they were omitted from the initial report that was submitted. h6 revised medical device problem code since the initial report has been submitted. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. product and data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the pump passed all post sterile inspection checks. | (b)(4). | D6b: added explant date.
2845 | 1220648-2025-48979 | 23794277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794277&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705975 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient came in to hospital for elect heart cath and cardioversion. elective coronary artery bypass graft surgery with impella 5.5 support. while on support the patient went into atrial fibrillation with rapid ventricular response. amiodarone drip initiated. the device was later explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Ppae (arrhythmia): in order to make a cause determination, all of the clinical details outlined would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. the pump sn 608064 passed all the post sterile inspection checks.
2846 | 1220648-2025-48984 | 23794284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794284&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713137 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was presented to support through high-risk percutaneous coronary intervention with impella cp. later the impella was successfully weaned and explanted. it was reported that the perclose failed to obtain hemostasis so femstop was placed achieving hemostasis. the patients outcome at explanted was noted as survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the peri-procedural adverse event: the cause of the bleed was use issue related since the perclose failed to obtain hemostasis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2847 | 1220648-2025-48985 | 23794344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794344&pc=OZD | 09/15/2025 | 12/11/2025 |  | 12/12/2025 |  |  | 11/18/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1424180 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 36 yr old male with an automated impella controller that was placed for support during high-risk cardiac procedure. it was reported that placement signal and left ventricle signal not displayed, beside this pump worked properly. | [5 empty] | In section g (manufacturing site name), the abiomed europe gmbh site address and required site information were inadvertently omitted from the initial report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2848 | 1220648-2025-48975 | 23794372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794372&pc=OZD | 11/14/2025 |  |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747608 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support it was reported that there was a plan to switch anticoagulant to bival, as there was a concern for heparin-induced thrombocytopenia and thrombosis. in addition amio bolus, lidocaine drip initiated after torsades event. episodes of torsades, required cardioversion. later the patient experienced ventricular tachycardia requiring debilitation x1. the patient remains on support. | Additional information was received, the patient was explanted. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. arrhythmia, tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. additional information received for d6 explant.
2849 | 1220648-2025-48986 | 23794580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23794580&pc=OZD | 11/15/2025 |  |  | 12/12/2025 |  |  | 11/15/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693633 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced paroxysmal atrial fibrillation with rapid ventricular response and ventricular tachycardia. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the tachycardia and atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2850 | 1220648-2025-31952 | 23795110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23795110&pc=OZD | 08/09/2023 | 12/12/2025 |  | 12/12/2025 |  |  | 11/14/2025 | 12/12/2025 |  |  |  |  | 12/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123074 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a 5.5 support ongoing for a patient awaiting heart transplant. the pump lost the optical placement signal but remained on for support for 26 more days. in total the pump supported for 38 days until it was explanted. no patient harm has been reported. | [5 empty] | Data analysis: the console (b)(6) imc log displayed the following error on the complaint date (09/08/23 00:05:37), indicating the onset of "oscillating contrast". this issue resolved automatically after approximately 12 days and 14 hours ( 21/08/23 14:28:17). processsampleforopminvalidsignalerror: calculator placement signal 1036 real time (rt) logs on 09/08/23 confirmed that contrast signal received from optical bench is represented between -3000 and 3000 values. see attachments section for ic7562_rtlog_459783_onset_detail_01.png. device analysis: this console was tested after arriving in fs. issue was not reproduced; however pattern of unrealistic contrast values is consistent with loss of light within the optical system, which could be caused by either defective lamp or optical bench electronics. inspection of the console also revealed corrupt microsd card on the rlm (b)(4), and minor damage to the optical pump connector - both unrelated to the complaint. ¿placement signal not reliable¿ alarms and intermittent optical sensor readings were caused by malfunction of the optical bench.
2851 | 1220648-2025-48922 | 23777010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777010&pc=OZD | 11/11/2025 |  |  | 12/11/2025 |  |  | 11/11/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759978 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the pump could not be brought into the correct position despite several attempts. the 0.018 wire was always in the apex. with every attempt the pump stood up, pulled the wire and twisted in the direction of the right ventricle. also, with a buddy wire it was impossible to introduce the pump. due to the patients condition, the physicians decided to implant a new pump, which could be introduced successfully. the patient's outcome was noted as stable. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. unable to deliver or advance: returned complaint device visually inspected, no abnormality observed. the root cause of the delivery issue most likely was due to patient condition related. device history review: the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2852 | 1220648-2025-48923 | 23777192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777192&pc=OZD | 11/12/2025 |  |  | 12/11/2025 |  |  | 11/12/2025 | 12/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709730 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated to another impella 5.5 due to delivery issues. it was reported against the 2nd imeplla 5.5 that the catheter bent at the level of the entrance to the airlock. pump not repositionable. implantation of new 5.5 is not possible due to anatomical reasons. the patient was noted to be in stable condition. this report is for the second imella 5.5 and the first impella 5.5 was reported on this report number 1220648-2025-48909. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding device in wrong position and product damage (catheter kink), the cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2853 | 1220648-2025-48924 | 23777222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777222&pc=OZD | 11/10/2025 |  |  | 12/11/2025 |  |  | 11/11/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709043 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient was under impella cp support for cardiogenic shock secondary to myocardial infarction complicated by ventricular septal defect. after six days of uncomplicated care with the impella cp, the patient underwent surgery to resolve the mechanical complication resulting from the heart attack. extreme vasoplegia developed postoperatively. due to high doses of norepinephrine, limb ischemia developed, and it was decided to explant the impella. the patient's condition deteriorated, and he died hours later. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". | Investigation summary: ppae (ischemia): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot. device lot: 1955206. device history batch. subcomponent lot: n/a . device history review: the pump sn (b)(6) passed all the post sterile inspection checks.
2854 | 1220648-2025-48925 | 23777343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777343&pc=OZD | 11/11/2025 |  |  | 12/11/2025 |  |  | 11/11/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the device could not be rewired once moved back across the aortic valve. successful insertion was noted and upon removal of intra-aortic balloon pump the impella came out and the physicians were unable to cross the aortic valve. the impella was removed and rewiring was not possible so another impella was opened. the serial number for this impella pump possibly could be (b)(4). the senior clinical consultant believes this serial number to be correct going by inventory. this is one of two related reports. this report represents the first impella cp and another report was submitted for the second impella cp. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, device in wrong position, the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
2855 | 1220648-2025-48927 | 23777571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777571&pc=OZD | 11/14/2025 |  |  | 12/11/2025 |  |  | 11/15/2025 | 12/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728746 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced difficulty in positioning the pump. the doctor attempted to advance the impella slightly but was unable to land the device near 3.5 centimeters (cm). the pump position was maintained at 2.8 cm in to the left ventricle via bedside transthoracic echocardiogram. the pump was explanted and the patient survived. this is one of two related reports. this report addresses the second impella cp. another report will be submitted to address the first impella cp. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2856 | 1220648-2025-48926 | 23777654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777654&pc=OZD | 11/14/2025 |  |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591253 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 97 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in with unstable angina and complex coronary artery disease was implanted with an impella cp device for mechanical circulatory support. the patient presented to the hospital and was urgently taken for an angiogram. the patient had a history of previous stenting in the main left anterior descending (lad) artery and diagonal branches which had developed recurrent in-stent. the angiogram revealed significant blockages with an 80% blockage in the lad stent and 70% in the left circumflex artery (lcx). ballon angioplasty was performed to the lad stent, and subsequently the patient became severely hypotensive, and blood flow became sluggish in both the lad and lcx arteries. the patient developed ventricular tachycardia (vt), requiring defibrillation and then subsequently developed ventricular fibrillation. the patient received cardiopulmonary resuscitation for seven minutes and three shocks before achieving return of spontaneous circulation. the patient was brought to the cardiac catheterization laboratory and urgently intubated, and the decision was made to implant an impella cp. an impella cp was prepped and inserted via the right common femoral artery without complications. upon initiating support, the pump would not flow above 1 liter per minute (l/min). the medical team ruled out patient volume status, pump position, thrombus, and right ventricular function as cause to the low flow. the decision was made to remove the impella cp and rewire and reinsert the same pump. upon insertion, the pump was started, and flows did not flow above 1 l/min. the impella cp was removed and a new impella cp was placed. it was noted that the second pump was sightly shallow at 2.8 centimeters (cm). the physician attempted to advance the impella but was unable to reposition the device. the device was maintained at 2.8 cm and a plan was made to explant the pump the following day. the second pump was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella cp, with serial number (b)(6). this report specifically addresses pump 1. a separate report was submitted for the other device associated with this complaint. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical details indicate that after insertion, the impella cp failed to achieve flows greater than 1 l/min despite troubleshooting volume status, pump positioning, rv function, and thrombus. the same issue persisted after removing, rewiring, and reinserting the same pump. without the returned product and sufficient clinical information, including data log, the cause of the low or blocked pump flow could not be determined.
2857 | 1220648-2025-48928 | 23777759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777759&pc=OZD | 11/13/2025 | 12/10/2025 |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770772 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79 yr old male that was implanted with a impella cp for high-risk cardiac procedure. there was hemolysis over night, and echocardiogram was obtained, with repositioning pulling back 1cm. the inlet was against a papillary now the plasma free hemoglobin were trending down and labs are not hemolyzing. patient was weaned from all drops and extubate prior to surgery. patient was then taken to surgery for impella removal. | [5 empty] | E4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in proper position. the cause of positioning issue was unknown. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2858 | 1220648-2025-48930 | 23777962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23777962&pc=OZD | 11/14/2025 | 12/10/2025 |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769627 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for high cardiac risk procedure. while shockwaving, the ostial left main coronary artery placement signals spiked then disappeared. no change in motor current nor impella flows. there was no reported change in patient treatment plan. at the end of the case, the impella was weaned, removed, and access was closed. | [5 empty] | Investigation summary: the clinical details state that while using shockwave that was proximal to the impella optical sensor, the ps spiked and then disappeared. the data logs show that after about 30 minutes the ps spikes to 3000, snr dropped to 0, contrast dropped, lamp maxed out, and gain and light were stable. psnr alarms were triggered and the ps remained out for the rest of the case. based on the clinical details and data log analysis, the root cause of optical sensor issue was due to a material crack due to shockwave. per capa-013581, ifu cp: 0048-9007 was updated to recommend physician users should ensure the shortest distance from the shockwave coronary ivl device to the impella optical sensor is =20 mm. device history lot: device lot: 1986690. device history batch: subcomponent: n/a. device history review: impella 633578 passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2859 | 1220648-2025-48929 | 23778001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778001&pc=OZD | 11/13/2025 |  |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667947 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 80 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Difficult to Insert | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that the device bent during insertion of the impella cp. suction occurred and flows were marginal upon startup. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 has been updated. | Added h6 codes due to product return. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. a cannula kink was noted near the outlet cage on the returned pump. no other damage was noted. the pump was tested in a bucket of water and ran without issue. data log was not provided or available on the remote server. low or blocked pump flow / pump suction / product damage (cannula kink): the cause of the low pump flow and suction issue was most likely related to the cannula kink based on returned product and provided clinical details. the cannula kink issue was most likely use-related, as it occurred during placement. device history review: the device passed all post sterile inspection checks.
2860 | 1220648-2025-48932 | 23778183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778183&pc=OZD | 11/14/2025 | 12/10/2025 |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768840 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Priming Problem | [5 empty] | Air Embolism | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with a impella cp for heart failure support. the impella cp was inserted via axillary artery utilizing a 10 mm x 20 cm vascutek gelweave. after insertion, large amounts of air were introduced into the left ventricle. the care team attempted to secure any potential areas of embolization externally. echocardiogram was used to check for air in the left atrium as well as other areas. the cardiac surgical center (csc) was called due to intermittent air boluses. the csc determined a potential cause could be cavitation. it was recommended dropping to p6. air significantly was reduced and completely went away at p5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 medical device problem code was changed since the initial report was submitted. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. priming problems- during insertion procedure, a large amount of air was introduced into the left ventricle. troubleshooting efforts and customer support center input determined a potential cause could be cavitation. dropping to p6 resolved the issue. data logs showed no purge related issues during the first few hours of support. the cause of the priming problems was not determined due to no product returned and insufficient clinical details. patient code - air embolism: the cause of the air embolism was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
2861 | 1220648-2025-48931 | 23778221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778221&pc=OZD | 11/14/2025 | 12/10/2025 |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744154 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 74 KG | Non Hispanic | White | Death | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 for support to wean off extracorporeal membrane oxygenation and bridge to recovery. the impella 5.5 perforated the lateral wall of the left ventricle and had to be removed. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the clinical details state that upon placing the 5.5, it perforated the lv and had to be removed. the heart was repaired and another 5.5 was placed. the doctor did not blame the 5.5 as he said echo imaging was not ideal. based on the clinical details, the root cause of the mechanical interaction with tissue is most likely due to use as the imaging was not ideal during placement of the impella. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2862 | 1220648-2025-48934 | 23778442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778442&pc=OZD | 11/14/2025 | 01/05/2026 |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758996 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced low pump flow. a second impella cp was placed following bicuspid aortic valve. the impella cp was placed in standard fashion through 14 french peel away sheath which was already in place. the wire was removed and the impella cp was started on auto. flows were 1.2 liters per minute. on fluoroscopy, a kink was found below the outlet window in the same area of previously placed impella cp. the catheter was torqued and remove kink from cannula to resolve the kink. flows were immediately up to 3.7 liters per minute. the impella cp was sutured in place and dressed accordingly. swan was placed in the right groin. following the case, the patient was transported to the intensive care unit. however, the patient expired on impella support. | [5 empty] | The investigation for the product kink and the low pump flow has been completed. the subject device was not returned for investigation. data logs were not provided or retrieved from the remote server. clinical details indicate that after placement of the second impella cp, flows were initially low at 1.2 l/min. on fluoroscopy, a kink was observed in the cannula below the outlet window, which likely contributed to the low pump flow. after the kink was corrected by torquing the catheter, flows immediately improved to 3.7 l/min. the cause of the low flow was not determined without returned product investigation and sufficient clinical details, including data log. clinical details further indicate that a kink was observed in the cannula of the second replaced impella cp, seen on fluoroscopy below the outlet window. the kink was resolved by torquing the catheter, resulting in improved pump performance. the cause of the cannula kink was not determined without returned product investigation and sufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2863 | 1220648-2025-48936 | 23778627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778627&pc=OZD | 11/14/2025 |  |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026770804 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Thrombosis/Thrombus | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to an outside hospital in st elevation myocardial infarction and cardiogenic shock following a percutaneous coronary intervention eight days prior and was found to have in-stent thrombosis. an impella cp was prepped, inserted via the left femoral artery, and delivered to the left ventricle without complications. following initiation of impella support, it was noted that the patient had poor pulse pressure and multiple episodes of ventricular tachycardia and the decision was made to cannulate the patient with extracorporeal membrane oxygenation (ecmo). a percutaneous coronary intervention was performed to the left anterior descending and right coronary arteries, and a right chest tube was placed with loss of flows. the patient was then transferred to another hospital for further treatment supported by ecmo and impella cp. the patient arrived to the hospital critically unstable and was found to have a large hemothorax in the right lung with five liters of blood from the chest tube. the patient was emergently taken to the operating room for exploratory surgery to determine the bleeding source and mass transfusion protocol was initiated. it was noted that the source of bleeding was unknown, and it was unknown if it was related to the impella placement or position. the patient¿s pericardial space was normal, which indicated the bleeding was not coming from a source in the heart but more likely from the lungs. the following day, the patient¿s chest bleeding had significantly decreased and was more stable. on support day seven, the medical team decided to decannulate ecmo and remove impella cp and escalate the patient to an impella 5.5. during the removal of the impella cp and ecmo, a large amount of thrombus was noted in the descending aorta. the clot was evacuated by the physician using a penumbra device. it was noted that other contributing factors of the thrombus was that the patient was admitted to the hospital for in-stent stenosis and the patient received cardiopulmonary resuscitation and massive blood transfusions after admission. additional information regarding the impella 5.5, and the final patient outcome information was not available in the complaint record accessible for report assessment at time of reporting. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary (tachycardia) : in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension, thrombosis, hemorrhage : the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
2864 | 1220648-2025-48937 | 23778666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778666&pc=OZD | 11/13/2025 |  |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718281 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. while on support the patient a few ventricular tachycardia storms with hemodynamical crash. defibrillation and iv drugs was how the complication was managed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported event, arrhythmia and hemodynamic instability issue, has been completed. the product was not returned. the root cause of the injuries was unable to be determined due to insufficient clinical details.
2865 | 1220648-2025-48938 | 23778673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778673&pc=OZD | 11/13/2025 |  |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758880 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the p level of the pump was lowered to resolved the issue. the patient is in stable condition. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hematuria: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2866 | 1220648-2025-48939 | 23778699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778699&pc=OZD | 11/14/2025 |  |  | 12/11/2025 |  |  | 11/14/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766934 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient undergoing a percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. during the pci, the patient¿s ostial right coronary artery underwent shockwave for treatment. during the shockwave treatment, the optical sensor on the impella was affected and stopped working. there was no effect to the pump flows or patient support. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: since neither data logs nor product were returned for investigation, the cause of the optical sensor issue could not be determined. device history review: the device passed all post sterile inspection checks. | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2867 | 1220648-2025-48940 | 23778717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778717&pc=PYX | 11/03/2025 |  |  | 12/11/2025 |  |  | 11/04/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026741909 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Renal Failure | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced elevated white blood cell count, renal failure, and hematuria. additionally, the patient had a gastrointestinal bleed. an esophagogastroduodenoscopy was performed and the patient was treated with protonix. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. infection : the cause of the injury was unable to be determined due to insufficient clinical details. hematuria/ renal failure : the cause of the injury was not determined due to insufficient clinical details. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2868 | 1220648-2025-48935 | 23778952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23778952&pc=PYX | 11/13/2025 | 01/05/2026 |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  | 02/28/2026 |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529187 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female | 61 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced bleeding at the right femoral vein so a femstop was placed. bleeding continued from the arterial line and groin site. the patient was transfused two units of packed red blood cells and systemic heparin was withheld. ultimately, the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
2869 | 1220648-2025-48941 | 23779421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23779421&pc=OZD | 11/13/2025 |  |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026754184 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was transported to the complainant hospital for a second opinion and advanced therapy workup for acute decompensated heart failure with cardiogenic shock after being denied at the previous hospital. the medical team ultimately decided to place the patient on impella 5.5 for mechanical circulatory support for advanced therapies workup. the patient was taken to the operating room and an impella 5.5 was prepped, inserted, and support was initiated without complications. on support day 2, the impella alarmed for suction events and the patient experienced ectopy. the suction alarms self-resolved. the patient¿s ectopy continued overnight, and the patient received milrinone. the impella placement was reviewed with transthoracic echocardiogram. the pump¿s depth was as expected; however, the inlet was pointed towards the intraventricular septum. the physician declined repositioning the device. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. no data logs returned for analysis. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the pump suction: clinical details mention that there were a couple of suction alarms observed on 13 nov 2025 which self resolved with ectopy and movement. tte showed that the pump was in good position and depth. the root cause of the pump suction was not determined due to lack of sufficient clinical details and unreturned product and data logs for further evaluation. arrythmia: clinical details mention that the patient suffered from ectopic heartbeats during impella support. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2870 | 1220648-2025-48942 | 23780006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23780006&pc=OZD | 11/13/2025 | 12/09/2025 |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026732595 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 100 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 51 yr old female was implanted with an impella 5.5 for support during a high-risk cardiac procedure. it was reported that the patient had a stroke and was confirmed to be hit+. no information was provided on the treatment. however, the patient subsequently expired. | [5 empty] | The impella device has not been returned for evaluation. if the device or new information is received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injuries, thrombocytopenia and stroke, could not be determined due to insufficient clinical details.
2871 | 1220648-2025-48943 | 23780430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23780430&pc=OZD | 11/13/2025 |  |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1144656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while on impella support there was automated impella controller failure. the console was switched out and there was no patient harm. | [5 empty] | D9 updated to device received date. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2872 | 1220648-2025-48946 | 23780717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23780717&pc=OZD | 05/25/2025 |  |  | 12/11/2025 |  |  | 11/12/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded publication entitled: "platelet derived prothrombotic microparticles and adverse events in patients supported with impella 5.5". case study patient #1 was on the impella 5.5 for support during percutaneous coronary intervention (pci). after 34 days on support the patient was weaned for explant, however during the weaning there was aortic valve heart block of 3rd degree, atrial fibrillation, and a pacemaker was placed. the patient survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. arrhythmia/heart block: the root cause of the injury was unable to be determined due to insufficient clinical details. | The device serial number, lot number and udi is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d¿ettore, n., petruzza, a., cardinale, a., bensi, t., maj, g., ciriello, m. m., & pappalardo, f. (2025). platelet derived prothrombotic microparticles and adverse events in patients supported with impella 5.5. the international journal of artificial organs, 48(10), 794¿799. https://doi.org/10.1177/03913988251357199.
2873 | 1220648-2025-48948 | 23780755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23780755&pc=OZD | 11/12/2025 | 01/06/2026 |  | 12/11/2025 |  |  | 11/12/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760472 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 63 KG | Hispanic |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was admitted as st-segment elevation myocardial infarction. the patient was brought to cath lab and referred to impella cp support due to mitral valve disease. while on support it was reported that the patient¿s urine turned dark red. hemolysis verified by haptoglobin test results from 70 mg/dl to 30 mg/dl. the pump was repositioned and the patient was noted in stable condition. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation investigation summary data review: insufficient data logs showed pump ran without issue during support. there were no alarms triggered during the case. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review pump passed all post sterile inspection checks. | The device was discarded. should any new information be received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2874 | 1220648-2025-48944 | 23781345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23781345&pc=OZD | 11/13/2025 | 01/22/2026 |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739644 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to heart transplant. the patient was admitted to the hospital with ventricular tachycardia storm, and the medical team decided to implant an impella 5.5 to bridge the patient to a heart transplant or durable left ventricular assist device. the impella 5.5 was prepped, inserted via the right axillary artery, and support was initiated without complications. on support day 6, it was reported that the patient was transfused with two units of packed red blood cells and hematoma was observed at the impella access site. a nerve block was administered at the site. imaging was performed to investigate the hematoma and revealed the hematoma to be stable and not actively bleeding. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary no product returned. asae/hematoma: hematoma has developed at rt. axillary artery cutdown site. pt was transfused with prbc x2. the cause of the access adverse event was not determined since product was not returned and insufficient clinical details provided. device history lot device lot: 1973996 device history batch sub-component lot: n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
2875 | 1220648-2025-48949 | 23782739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23782739&pc=OZD | 11/13/2025 | 12/10/2025 |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749867 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user faility reported that a 7i yr old female was implamted with impella cp. it was reported that the impella advanced into proper position and started in auto. quickly ramped up to >4.0 left and left ventricle waveform very elevated with elevated motor current. position appropriate on fluoro, doctor not satisfied with waveforms and concerned for pump malfunction or thrombus (although recent magnetic resonance imaging showed no left ventricle thrombus). pump removed and exchanged for new impella cp. old impella with no evidence of thrombus. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the high or saturated pump flow impella advanced into proper position and started in auto. quickly ramped up to >4.0 l and lv waveform very elevated with elevated motor current. the data logs confirm higher than expected flows for set p-level. pump was flowing at a mean of 4.3l/min. there were no motor current spikes. the logs show a slight increase in purge pressure with a decrease in purge flow. clinical mention position was appropriate. the cause of the issue was not established as limited clinical information was provided and no product was returned. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2876 | 1220648-2025-48951 | 23783258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23783258&pc=OZD | 11/12/2025 |  |  | 12/11/2025 |  |  | 11/13/2025 | 12/11/2025 |  |  |  |  | 12/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758957 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  |  |  | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was admitted for chest pain. was later taken to cath lab for intervention to left anterior descending artery. the patient was sent to intensive care unite and returned to cath lab due to worsening chest pain, found stents to be thrombosed. intervention completed to remove thrombus. the patient then coded and impella cp placed. while one support the impella moved into aorta and unable to reposition leading to pump removal. per icu nurse there was significant bleeding at impella insertion site that was not controllable that required a vascular cutdown. the patient outcome is unknown. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary no product was returned. device in wrong position: the cause of the device in wrong position was unable to be determined due to insufficient clinical details. access site adverse event: the cause of the bleed was unable to be determined due to insufficient clinical details. device history lot device lot: 1978858 device history batch subcomponent lot: n/a device history review device sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2877 | 1220648-2025-48878 | 23766814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23766814&pc=OZD | 11/10/2025 |  |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026763365 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 16 YR |  | Male | 71 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. after implantation of the impella 5.5, the patient was transferred to the intensive care unit. during this, a purge system blocked alarm occurred. the error could not be resolved and a new impella 5.5 was placed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the impella 5.5 was visually inspected and no biomaterial was found around impeller blade. some bending/waving of catheter around the 55 to 60cm mark near cath anchor. pump was connected to console and purged for around 5 minutes and high purge pressure values and alarm were reproduced. pump was run at p-9 and high purge pressure resolved when running. clinical details noted that a "purge system blocked" alarm was triggered during icu transfer. the cause of the purge system blocked could not be determined as no data logs were returned to confirm pump metrics at time of high purge pressure alarms, no cause of purge system blocked were observed on returned product, and limited clinical details at time of increased purge pressure instance.
2878 | 1220648-2025-48871 | 23766854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23766854&pc=OZD | 10/17/2025 | 12/09/2025 |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683207 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported by the impact team that a patient was implanted with impella 5.5 for support for atrial fibrillation. afterwards the patient was switched to oral amiodarone as prophylaxis against further atrial fibrillation. the pump was in the weaning process. the patient was awake and neurological was adequate. the pump runs without any problems on p2. the doctors want to explant as soon as possible. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2879 | 1220648-2025-48879 | 23766941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23766941&pc=OZD | 11/10/2025 |  |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751671 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male | 70 KG |  |  | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered pump suction and positioning issues during impella cp support. after brief cardiopulmonary resuscitation, positioning problems occurred. no red alarm occurred, only suction alarm. the pump was viewed under fluoroscopy, and it was not correctly positioned. the pump was too far in the left ventricle. repositioning initially was not possible. the pump could not be pulled back, because it was stuck in a structure. after consultation with the senior physician, an attempt was made to reposition the pump, as the impella could not currently provide support. during this attempt, the pump slipped in the aorta and it was not possible to reposition it. due to the patient's hemodynamics instability, the decision was made to reinsert a new pump. a new pump was inserted, however, shortly afterwards the patient expired. | [5 empty] | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device report. medical safety/clinical review. medical safety/clinical reviewed the event. an impella cp device was implanted for mechanical circulatory support prior to high-risk percutaneous coronary intervention (pci) involving the left main and left anterior descending artery, with treatment including coronary angioplasty, stent placement, and shockwave therapy. following cardiopulmonary resuscitation (cpr), device positioning issues were identified. no device alarms were noted; however, fluoroscopic assessment demonstrated that the pump was positioned too deeply within the left ventricle. initial attempts to reposition the impella cp were unsuccessful, as the device appeared to be stuck within a cardiac structure and could not be withdrawn. after consultation with a senior physician, further attempts were made to reposition the device due to inadequate hemodynamic support. during these attempts, the impella cp migrated into the aorta and could not be repositioned for support. given the patient¿s ongoing hemodynamic instability, the decision was made to remove and reinsert a new impella cp device. approximately 23 hours after implant of the new impella cp, the patient expired. no additional clinical or procedural details regarding the event were available. usmdr, there is insufficient information to exclude impella cp pump 1 as an associated factor in the patient's outcome. impella cp pump 2 was in use at the time of the patient¿s death; however, based on the available information, impella cp pump 1 is considered the device potentially associated with the demise outcome. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
2880 | 1220648-2025-48880 | 23766985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23766985&pc=OZD | 11/08/2025 |  |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750621 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered placement signal issues during impella cp support. it was noted that there was no placement signal as the left ventricle placement signal was not displayed on the screen. support continued and the patient remained stable. additional information was received. the problem resolved by itself. the pump runs without problems. the impella pump was removed three days later. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. no data logs returned. the product was returned and there was no biological material noted. the 5s parameters were within specification and the pump passed the laser continuity test. additionally, the pump ran with no issue. there were no abnormalities found with the pump regarding the optical sensor issue. due to lack of data logs returned, the root cause of the optical sensor issue cannot be determined. device history review: the device passed all post sterile inspection checks.
2881 | 1220648-2025-48882 | 23767063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767063&pc=OZD | 01/01/2025 | 12/10/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562075 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis. the p level of the pump was reduced to improve the hemolysis. the patient is in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data logs showed pump ran without issue during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks.
2882 | 1220648-2025-48884 | 23767127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767127&pc=OZD | 11/07/2025 | 12/08/2025 |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026727975 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was implanted with a impella cp to support percutaneous coronary intervention. from the time of insertion, there was a pressure difference of approximately 30-50 mmhg between the a-line blood pressure and the aortic pressure in the automated impella controller. for management purposes, a-line pressure was managed using a vital sign monitor, and this did not pose a problem, but as blood pressure and diastolic pressure dropped, "placement signal minimum and low value alarms" began to appear frequently. after removal, the impella cp was retrieved, and a request was made for analysis of the pump catheter. upon checking the insertion technique, it was determined that the catheter had initially been exposed to the surgical field, and that it was possible that it had come into contact with the white plug when it was connected. approximately a week later, the catheter was removed because the heart had recovered. the pressure difference was not the cause. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2883 | 1220648-2025-48881 | 23767194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767194&pc=OZD | 11/10/2025 | 12/10/2025 |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696065 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with impella cp for support for heart failure. after dc, the impella cp was found to have dislodged into the aorta. this is one of two related reports. this report addresses the impella cp and another report will be submitted to address that impella 5.5 refer to section h10 for the related report number. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for investigation. device in wrong position: the pump was not returned for investigation. clinical details indicate that the impella cp c8 was found to have dislodged into the aorta after discontinuation of support. additional information suggests that the device may have been placed a little too shallow initially. although the fixation rings were reported as properly secured. the cause of the device in wrong position issue was not determined without returned product and sufficient clinical information. device history review the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2884 | 1220648-2025-48883 | 23767228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767228&pc=OZD | 11/14/2025 |  |  | 12/10/2025 |  |  | 11/14/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703601 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during implantation of an impella 5.5 the patient suffered a perforation of the left ventricle. cardiopulmonary bypass was established with impella insertion to perform ventricular septal perforation (vsp) patch closure, but the left ventricular perforation occurred during impella insertion. cardiopulmonary support was started, the impella 5.5 was removed, and vsp patch closure and left ventriculoplasty were performed. the impella was facing the posterior wall, making it difficult to properly position it from the time of insertion. the height of the aortic clamp was checked and slightly deepened (approximately five centimeters from the aortic valve to the aspiration site, but since it was originally facing the posterior wall, it may have been close to or interfering with the left ventricular wall). this is one of two related reports. this report represents the impella 5.5 and another report was submitted to address the impella cp. refer to section h10 for the related report number. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical details indicate that left ventricular perforation occurred during impella 5.5 s2 insertion for ventricular septal perforation (vsp) patch closure. the perforation happened while the device was facing the posterior wall and during a vsp rotation maneuver on the posterior wall of the left ventricular septum. cardiopulmonary support was started, the impella was removed, and vsp patch closure plus left ventriculoplasty was performed. the cause of the cardiac perforation was most likely use-related, associated with device positioning and procedural technique during the vsp rotation maneuver. clinical details indicate that from the time of insertion, the impella was facing the posterior wall, making it difficult to properly position. the device may have been close to or interfering with the left ventricular wall, contributing to the perforation. the cause of the device in wrong position was not determined without returned product investigation and sufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | Complaint coding was updated to better reflect the reported event. consequently, section h6 medical device problem code was updated/corrected and repopulated accordingly.
2885 | 1220648-2025-48885 | 23767315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767315&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026753156 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 82 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient recently underwent coronary artery bypass grafting and valve with ejection fraction of 15-20% and elective on impella 5.5 placement. while on the impella 5.5 support it was reported that there was oozing from the graft anastomosed site. one unit of blood was given and hemostasis was achieved during surgery. the pump was later explanted and the patient was stable. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | The impella device is not available from the customer. therefore, investigation of the device was not possible. should any new information be received, a supplemental report will be filed.
2886 | 1220648-2025-48886 | 23767419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767419&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576523 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | False Alarm | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented to the emergency department (ed) with a st-elevation myocardial infarction. an echocardiogram in the ed demonstrated reduced ejection fraction versus prior. the patient was brought to catheterization lab and an impella cp was used for support. it was reported that during support, there was "impella in aorta" alarms and abnormal waveforms, despite a confirmed stable position. these are likely linked to the patient's underlying arrhythmia affecting the left ventricular function and hemodynamics. in addition, red urine and hemolysis on labs was reported. the patient expired the following day and no other information was provided. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. hemolysis: clinical details report that the patient had red urine and hemolyzed labs the morning of (b)(6) 2025. pump position had been confirmed to be adequate that morning, but overnight it had come out of position during transport. the team did not want to adjust p-level or administer volume. later that day, the patient passed away on support. neither data logs nor product were available for investigation. since neither data logs nor product were available for investigation, the cause of the hemolysis could not be determined. false alarm: clinical details report that "position in aorta" alarms were triggering falsely on the first day of support. at this time, the patient was experiencing arrhythmias, so the waveform looked abnormal. neither data logs nor product were available for investigation. since neither data logs nor product were available for investigation, the cause of the false alarm could not be determined. placement signal issue: around case start, it was also noted that the ps map was 20 mmhg lower than the invasive ao map. neither data logs nor product were available for investigation. since neither data logs nor product were available for investigation, the cause of the placement signal issue could not be determined. arrhythmia: clinical details report that the patient was having rapid atrial flutter prior to and after impella insertion. in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review pump sn (b)(6) passed all post-sterile inspection checks. | The impella device is not available for return from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
2887 | 1220648-2025-48887 | 23767506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23767506&pc=OZD | 11/10/2025 |  |  | 12/10/2025 |  |  | 11/10/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765474 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented with non-st-elevation myocardial infarction and was transferred for surgical evaluation. the doctor requested an impella cp be placed the morning of coronary artery bypass grafting due to tight left main coronary artery and low ejection fraction of 25%. the doctor felt anesthesia induction was risky and wanted support to get the patient in the operation. the patient arrived back in unit from two vessel bypass this afternoon and the case went well but after weaning from bypass, the patient was unable to maintain impella flows or blood pressure. multiple blood products were given over the next two hours trying to catch up. hemoglobin and hematocrit continued to drop and ultimately the doctor found the heel of the anastomosis of the left internal mammary artery to left anterior descending artery graft was bleeding. the decision was made to go back on bypass and arrest the heart to repair the suture line. while on the pump the second time, the patient was felt to have become cardioplegic and methylene blue was administered by the doctor. it was later reported that the patient took a turn overnight, continued to deteriorate and the expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the injuries, hypotension and major bleed were not determined due to insufficient clinical details. the pump was not returned for investigation. data log review was not provided or could not be retrieved from the remote server. the cause of the low pump flow issue was not determined without returned product investigation and sufficient clinical details.
2888 | 1220648-2025-48888 | 23768141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23768141&pc=OZD | 11/11/2025 | 12/10/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700773 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a was implanted with an impella cp device for mechanical circulatory support. the patient had presented to the hospital for left heart catheterization and percutaneous coronary intervention (pci). following the pci where a stent was placed in the left anterior descending coronary artery, the patient was moved from the table to a transport bed, and the patient became pulseless. advanced cardiac life support protocol was initiated, and the patient received 13 defibrillations. the medical team decided to implant an impella cp via the left common femoral artery for mechanical circulatory support. impella was prepped, inserted, and support was initiated with flows as expected for 3.3 to 3.4 l/min. a hematoma and bleeding were noted around the groin at the impella access site. it was noted that bleeding occurred during the placement of the introducer sheath. manual pressure was held for 30 minutes. the patient required blood products and increased vasoactive medications. despite increased chemical support the patient continued to decompensate. the bleed at the femoral artery was determined to require vascular surgical repair, and due to the bleeding and the patient¿s decompensation the decision was made to escalate the patient to an impella 5.5. the impella 5.5 was implanted via the right axillary artery without complications. the impella cp was removed, and the femoral artery was surgically repaired and closed. due to continuous bleeding and inability to control bleeding family agreed to stop further treatment and intervention efforts. the patient outcome at impella 5.5 explant was expired. this complaint involves two impella devices: 14f sheath with lot number s9554810 used with pump 1, impella cp with serial number (b)(6) pump 2: impella 5.5 with serial number (b)(6) this mdr specifically addresses pump 2: impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. the impella device was discarded by the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety/clinical reviewed the event. an 81-year-old male patient presented for a left heart catheterization (lhc). a stent was placed in the left anterior descending artery (lad) via right common femoral artery (rcfa) access. following completion of the procedure, while the patient was being transferred from the procedure table to the bed, he became pulseless. advanced cardiovascular life support (acls) protocol was initiated. due to the presence of an angio-seal vascular closure device at the right common femoral artery access site, vascular access was obtained via the left common femoral artery (lcfa). the patient received 13 defibrillation. an impella cp was implanted via the left femoral artery and mechanical circulatory support (mcs) was initiated. on post-implant day 0, the patient developed bleeding at the access site with associated hematoma formation. the patient was noted to have decompensation requiring escalating doses of vasoactive medications and transfusion of blood products. due to worsening hemodynamic instability, the patient was escalated to an impella 5.5, which was surgically implanted via the right axillary artery. support was initiated at performance level p-2 and gradually increased as tolerated. surgical closure was performed at both the right axillary artery and left common femoral artery access sites. despite these interventions, the patient experienced ongoing bleeding that could not be controlled. the patient had been anticoagulated for the cardiac catheterization procedure, which may have contributed to the severity and persistence of bleeding. due to continuous bleeding and inability to achieve hemostasis despite surgical intervention, the patient¿s family elected to withdraw further life-sustaining measures. the patient subsequently expired. the impella cp will be a serious injury event, while the impella 5.5 will be classified as the death type of reportable event. correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: surgical intervention (f19). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. note: h6. investigation type, findings, conclusion, and component coding remain unchanged as previously reported (update was made only to the impact code and did not affect the health effect adverse event and/or product experience coding). related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. access site adverse event: the cause of the bleed was most likely patient condition related since patient had friable vessels. device history lot: device lot: 1955446. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2889 | 1220648-2025-30744 | 23768454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23768454&pc=OZD | 06/30/2025 | 07/28/2025 |  | 12/10/2025 |  |  | 06/30/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683333 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that patient had a cp pump placed to support st-segment elevation myocardial infarction, drug and alcohol abuse, and with multivessel coronary artery disease. the pump was placed to allow for the percutaneous coronary intervention but, there was ischemia. the pump was explanted after 20 hours of use and successful weaning. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
2890 | 1220648-2025-30745 | 23768456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23768456&pc=OZD | 06/30/2025 | 07/28/2025 |  | 12/10/2025 |  |  | 06/30/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694585 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that roughly two days into impella cp support, the patient developed a gastrointestinal bleed. blood products were given to manage the bleed. two days later, decision was made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
2891 | 1220648-2025-30746 | 23768467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23768467&pc=OZD | 07/02/2025 | 07/28/2025 |  | 12/10/2025 |  |  | 07/02/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684250 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient had a cp pump placed to support ongoing acute myocardial infarction complicated by cardiogenic shock. the cp allowed for the patient to have the percutaneous coronary intervention for the coronary artery disease. the pump was weaned and explanted after 39 hours and groin became an issue. the pulses were lost. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
2892 | 1220648-2025-30740 | 23768468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23768468&pc=OZD | 07/06/2025 | 07/28/2025 |  | 12/10/2025 |  |  | 07/14/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2026697644 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Sepsis | Aortic Valve Insufficiency/ Regurgitation | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient became hypotensive requiring vasoactives. additionally, the patient had elevated white blood cells and sepsis was suspected. antibiotics were started and support continued. furthermore, the pump was confirmed to be deep. no known reposition was performed. moreover, the patient had a gastrointestinal bleed. heparin was stopped and one unit of blood products was given to the patient. eventually, a pump stop occurred. the device was attempted to be restarted but it did not work. it was then pulled back into the ascending aorta to stop aortic regurgitation. the patient was brought back to the lab for removal. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ impella cp with smartassist section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
2893 | 1220648-2025-48890 | 23768874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23768874&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718762 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding and hemolysis. the impella was explanted and the patient is in stable condition. | [5 empty] | A4, a5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. medical device problem code a24 was corrected to a01.
2894 | 1220648-2025-48892 | 23769042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769042&pc=OZD | 11/11/2025 | 12/08/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  | 12/31/2025 |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025440057 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Anemia | Fever | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 72yr old male was implanted with a impella 5.5 for support to wean off ecmo and bridge to recovery. hemolysis was reported that required a blood transfusion with 1 unit of packed red blood cells given. the physician noted to have small left ventricle and causing occasional " position in aorta alarms". in addition, no systemic heparin running due to melena stool. heparin stopped, and then continues to run in purge. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding mechanical interaction with blood, the was likely due to improper position issues caused by patient condition as confirmed in clinical that patient had small left ventricle affecting position and was later resolved. for the fever, in order to make a cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. additionally, the bleed (gi bleed)/anemia cause was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
2895 | 1220648-2025-48894 | 23769081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769081&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759740 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced loss of the placement signal. the pump was explanted and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | The investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the clinical details state that after shockwave was initiated the ps was lost on the pump. the data logs show that about 21 hours into support, psnr alarms occur while snr drops to 0, gain increase, light remains stable, contrast decreases and lamp was maxed out. the returned product did not reveal any kinks in the catheter. laser continuity showed light exiting the sensor face. 5s parameters were indicative of a damaged sensor face. the root cause of the optical sensor issue is due to a material crack due to a device interaction with shockwave. "per capa-013581, ifu cp: 0048-9007was updated to recommend physician users should ensure the shortest distance from the shockwave coronary ivl device to the impella optical sensor is =20 mm. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2896 | 1220648-2025-48891 | 23769104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769104&pc=OZD | 11/11/2025 | 12/30/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026766045 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 70 yr old male patient was implanted with a impella cp for cardiac arrest support. the decision was simultaneously made to place impella cp. patient was mildly febrile and was started on vancomycin. peel away removed, repo sheath inserted, slack removed, and sheath secured. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported fever has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the fever could not be determined.
2897 | 1220648-2025-48897 | 23769132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769132&pc=OZD | 08/13/2025 | 12/10/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Cardiac Tamponade | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported in the long-term outcome and quality indicator (loqi) that a 28 yr old female was implanted with impella cp for support during high risk cardiac procedure. impella was inserted as left ventricle vent. patient was coagulopathic and for this reason, patient was placed on wound vac. patient also had ongoing mediatinal hemorrhage and right hemothorax. exploration and replacement of vac dressing. baseline h/h unknown. in the operating room, patient received 3 units of packed red blood cells 3 units of fresh frozen plasma, 2 units of cryoablation and 1platelet count. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. cardiac tamponade: the cause of the injury was not determined due to insufficient clinical details. infection: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2898 | 1220648-2025-48899 | 23769169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769169&pc=OZD | 11/11/2025 | 12/08/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during pump priming of impella 5.5, the purge disc would not register as being inserted. it was attempted to restart the aic and begin pump priming process once more and yet again not registering purge disc. another aic was brought in to prime the device successfully. | [5 empty] | H6: added codes based on the results of the investigation. investigation summary data review: console logs were consistent with what was reported in the complaint. device analysis: console was tested after arriving in field service. the purge disc detection issue was able to be reproduced. a broken purge disc flag was observed. conclusion: the root cause of the purge disc not detected alarms was a broken purge disc flag (missing at blue retainer). | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2899 | 1220648-2025-48900 | 23769256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769256&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627610 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 100 KG | Hispanic | White | Death | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Bradycardia | Cardiac Arrest | Hematoma | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the access site. manual pressure was held and a fem stop was placed. the patient also experienced positioning issues, bradycardia, and cardiac arrest. the device was repositioned and a pacer was placed. the patient expired on impella support. | [5 empty] | No product was returned for investigation. the cause of the cardiac arrest was not determined due to insufficient clinical details. the cause of the bradycardia was unable to be determined due to insufficient clinical details. the cause of the hematoma was not determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2900 | 1220648-2025-48901 | 23769263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769263&pc=PYX | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026742024 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced a purge pressure low alarm. the pump was exchanged for a new pump. no patient harm was reported. | [5 empty] | Product complaint # (b)(4). investigation summary: rpf pump (b)(6) returned. no data logs were returned for analysis. purge pressure low: the returned pump was investigated. the red lumen was still in the cannula, which was removed and the pump was primed and tried to be run in the lab. no other visual abnormalities were seen. the low purge pressure alarms were reproduced in the lab, with purge flow at 30 ml/hr. the blue handle was opened to check for leaks in the purge line. no leaks were observed. it was seen in the lot traveler for the pump that the manufacturing purge flow was 26.6 ml/hr, which was close to the 28 ml/hr threshold for acceptable purge flow rate for the pump. the root cause of the purge pressure low alarms was not determined due to inconclusive evidence from the product investigation and clinical details. the failure was related to the device based on the product evaluation but could not be further traced more specifically. pump #615903 passed all post-sterile inspection checks. | Medical safety/clinical reviewed the event. a 68-year-old female patient underwent orthotopic heart transplantation (oht) on (b)(6) 2025. earlier on the day of the reported event, the patient was decannulated from extracorporeal membrane oxygenation (ecmo). following decannulation, the patient experienced acute hemodynamic decompensation requiring right ventricular mechanical circulatory support with an impella rp flex device. the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock. during device preparation and priming, a ¿purge pressure low¿ critical alarm was triggered. the team reported that normal setup procedures were followed in accordance with the device instructions for use (ifu). despite this, the purge pressure alarm persisted. as a result, the initial impella rp flex device was replaced with a second impella rp flex device. the replacement device was prepared and implanted without reported issues. support was initiated and maintained at performance level p-8, achieving flows of 3.8 l/min. the patient remained on mechanical circulatory support. approximately four days later, the patient¿s family elected to withdraw care and the patient subsequently expired. the reported device-related issue involved a ¿purge pressure low¿ alarm during preparation and priming of the first impella rp flex device and the device was replaced, and no further purge pressure alarms or other issues/complications were reported. the patient's underlying condition appears to have contributed to the demise outcome (cardiac condition that included recent orthotopic heart transplantation (oht) rejection, right ventricular failure, and cardiogenic shock stage e). however, the impella rp flex pump 1 will be conservatively reported as death. correction. after review of the case by medical safety/clinical, it was determined the impella is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. note: h6: medical device problem, component, and investigation type/finding/conclusion coding remains unchanged. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2901 | 1220648-2025-48896 | 23769328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769328&pc=PYX | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026679976 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Other | [28 empty] | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella rp flex device for mechanical circulatory support. on support day 6, the patient experienced pink tinged urine. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived and noted to be hemodynamically stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: added type of investigation codes b13 and b24. h11: added the results of the investigation. investigation summary. hematuria the cause of the injury was not determined due to insufficient clinical details. device history review the pump sn (b)(6) passed all post sterile inspection checks.
2902 | 1220648-2025-48903 | 23769922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23769922&pc=OZD | 11/11/2025 | 12/30/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741744 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 128 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 42 yr old male was implanted with a impella 5.5 for high risk cardiac procedure. it was reported that patient arrested two times with ventricular tachycardia and ventricular fibrillation. patient was defibrillated successfully. prolonged qt was present. patient defibrillated a few more times with runs of ventricular tachycardia heparin held due to heparin induced thrombocytopenia. impella 5.5 was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombocytopenia and ventricular tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia and ventricular tachycardia could not be determined.
2903 | 1220648-2025-48905 | 23770032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770032&pc=PYX | 11/11/2025 | 01/22/2026 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625192 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella rp flex experienced a placement signal not reliable alarm. the pump was explanted. no patient harm was reported. | [5 empty] | Investigation summary: rp flex pump sn (b)(6) returned. no field data logs available. placement signal issue: psnr alarm reported post implant with intermittent ps on the console. returned product passed laser continuity test and 5s acceptable (lamp at 100%). electrical and voltage/signal dp electrical passed. psnr alarm was not reproduced when plugged into the console. the pump was run and psnr did not reproduce. lab run showed raw dp not drifting and it was appropriately reacting to p-level changes. no damages/abnormalities on dp sensor face. no biomaterial around the impeller. the cause of the placement signal issue was not determined since the issue was not reproduced on returned product and no field data logs returned. device history lot device lot: 1887745. device history batch subcomponent lot: n/a. device history review device sn (b)(6) passed all post sterile inspection checks. | No product returned. no data logs available. the cause of the placement signal issue was not determined due to no product or data logs returned. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2904 | 1220648-2025-48904 | 23770070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770070&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743380 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on support day 19, there was a report of placement signal low on the impella 5.5. an echocardiogram was performed and revealed that the impella was slightly deep at 6cm. the physician declined to reposition the device unless there were additional alarms. two days later, it was observed that there was a sensor drift with the optical sensor readings. the device was noted to be in adequate position. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. there were ps low alarms and then it was noted two days later the sensor was drifting. due to lack of clinical information, and no product or data logs, the root cause cannot be determined. | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: optical sensor issue. there were ps low alarms and then it was noted two days later the sensor was drifting. no data logs were provided. the returned product did not have any damages noted. 5s parameters were within specification with the cavity length being 16386, and the pump also passed laser continuity. the pump was run in a bucket for approximately 15 minutes and the calculated ps was negative but did not reach values to trigger psnr alarms, but met criteria to trigger and reproduce placement signal low alarms. snr was stable along with the other 5s parameters. the ps can be seen to trend down during this run. the root cause of the optical sensor issue is most likely due to material wear at 20 days that occurred within design life. device history review: impella passed all post sterile inspection checks.
2905 | 1220648-2025-48906 | 23770080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770080&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/12/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152776 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during patient support an automated impella controller (aic) alarm occurred. the aic was exchanged for a new controller to continue patient support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2906 | 1220648-2025-48907 | 23770087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770087&pc=OZD | 11/12/2025 | 01/22/2026 |  | 12/10/2025 |  |  | 11/12/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739573 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 50 yr old male was implanted with an impella 5.5 for cardiac arrest to support percutaneous coronary intervention. it was reported that patient had ventricular tachycardia which attributed to impella hitting ventricular wall due to smaller cavity and hyper-dynamic left ventricle. right axillary hematoma was also noted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2907 | 1220648-2025-48909 | 23770136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770136&pc=OZD | 11/12/2025 |  |  | 12/10/2025 |  |  | 11/12/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709623 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Failure to Advance | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. the patient remains in stable condition. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the peripheral artery dissection: the cause of the injury was not determined due to insufficient clinical details. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had intima flap over the aortic arch preventing successful delivery of impella. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2908 | 1220648-2025-48911 | 23770164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770164&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026747727 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella cp failed due to patient anatomy. the patient's aorta was tortuous and caused the guidewire and pump to bend. a new guidewire and pump were implanted to support the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2909 | 1220648-2025-48912 | 23770177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770177&pc=OZD | 11/08/2025 |  |  | 12/10/2025 |  |  | 11/12/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766414 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Cardiac Tamponade | Pericardial Effusion | [6 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced atrial fibrillation requiring cardioversion. the patient also had oozing at the access site. a stitch was placed to obtain hemostasis. the patient is in stable condition. | Clinical rationale: a 55-year-old male with acute myocardial infarction and cardiogenic shock (scai stage e) underwent impella cp support in conjunction with peripheral va ecmo for rv support following post-operative decompensation after mitral valve replacement. on (b)(6) 2025, nursing staff reported the patient developed atrial fibrillation/flutter with rapid ventricular response requiring one episode of successful cardioversion to normal sinus rhythm. on (b)(6) 2025, the patient developed increasing hypotension with escalating vasopressor requirement, prompting a stat echocardiogram that revealed pericardial effusion with cardiac tamponade. the first pump (b)(4) was implanted and explanted on (b)(6) 2025, with observed red handle leak, major bleeding, and device revision/replacement. the second pump (b)(4) was implanted on (b)(6) and explanted on (b)(6) 2025 following the patient¿s clinical decline. documented adverse events included hypotension, arrhythmia, anemia, paroxysmal atrial fibrillation, major bleeding, pericardial effusion, cardiac tamponade, and required interventions such as medication, transfusion, resuscitation, hemostasis, surgical intervention, and ultimately the patient¿s death. based on the information provided, the impella cp demonstrated stable positioning with no suction events during the arrhythmia episode. bleeding at the access site required surgical management, and a red handle leak prompted device exchange of the first pump. the subsequent deterioration into pericardial effusion and tamponade appears related to the patient¿s overall critical condition and recent cardiac surgery rather than to device malfunction; however, a causal relationship to the impella cannot be definitively excluded. a causal relationship to the impella cannot be definitively excluded. returned product evaluation is pending and will be required to determine whether a device related failure contributed to the reported events. the patient withdrew care and expired. |  |  |  |  | D6a and d6b should of been left blank on the initial report that was submitted. e1 added initial reporter address line 2 and initial reporter phone number as they were omitted from the initial report that was submitted. h6 revised investigation conclusions code as it was entered incorrectly on the initial report that was submitted. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 clinical rationale was added. d6a and d6b dates were reported correctly on the initial report that was submitted and the note about them being incorrect on the previous report was incorrect. h10 adde related report number. this is one of two related reports. this report represents the second impella cp and another pump was submitted for the first impella cp. investigation summary: the investigation was completed. no product was returned. hypotension, anemia, pericardial effusion, cardiac tamponade: the cause of the injury was unable to be determined due to insufficient clinical details. arrhythmia, paroxysmal atrial fibrillation: the cause of the injury was unable to be determined due to insufficient clinical details. access site adverse event/major bleed: the cause of the bleed was unable to be determined due to insufficient clinical details.
2910 | 1220648-2025-48913 | 23770186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770186&pc=OZD | 11/12/2025 |  |  | 12/10/2025 |  |  | 11/12/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755524 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a pump stop. the pump could not be restarted. the pump was explanted. upon explant, a wire was observed coming from the motor and through the outlet cage. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was returned for evaluation. pump stop: during a percutaneous coronary intervention (pci), there was a pump stop. after multiple attempts to restart, it failed to start back up. after removing the pump, a wire was observed coming out of the motor and through the outlet cage. product was returned and evaluated. the wire reported to be in the outlet cage was returned in a separate bag. the edges of the impeller blades appeared rough at the surface. the pump was able to run in the lab and the pump stop did not reproduce. the returned wire had an intact end with the of the coiled wire unraveled. the diameter of the wire tip was approximately 0.3 mm. a data stick was also returned but there was no data found. logs were not provided. the cause of the pump stop was interaction with a non-abiomed guidewire based on clinical details and returned product. device history review the pump passed all post-sterile inspection checks. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. pump stop: during the procedure, there was a pump stop. after multiple attempts to restart, it failed to start back up. after removing the pump, a wire was observed coming out of the motor and through the outlet cage. no product/logs were returned. the cause of the pump stop was not determined since product/logs were not returned. device history review: the pump passed all post-sterile inspection checks.
2911 | 1220648-2025-48908 | 23770209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770209&pc=OZD | 11/11/2025 | 12/08/2025 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276150 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that when console was turned on, there was a notice that came up that said ¿system self check failed. controller was replaced immediately. it was turned off and from underneath and would not respond. then it suddenly shut itself off. | [5 empty] | Correction e4the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. there were no issues related to complaints discovered when reviewing the product history of console (b)(4). the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the reported failure mode was reproduced during testing at the field service. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination (pbm_contamination.jpg) which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2912 | 1220648-2025-48915 | 23770404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770404&pc=OZD | 11/11/2025 | 01/12/2026 |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205610 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that when plugging in new cp ¿controller failure¿ alarm showed. new aic used. | [5 empty] | Investigation summary: data analysis: console logs from the reported day of event, shows an controller error alarm (opm comm crc error ¿ event set # 1045) and controller error (optical bench ambient pressure out of range - event set # 1040) triggered during the clinical procedure. no controller failure related alarms were seen. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. device analysis: the console ic3922 was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the optical bench was evaluated for 5s parameters and the analog signals were checked for distortion, but no issues were found. additionally, the aic was tested with know good pump and purge cassette and no issues were found. root cause: the root cause of the ¿controller error alarm (opm comm crc error ¿ event set # 1045) and controller error (optical bench ambient pressure out of range - event set # 1040)" was due to a firmware issue (optical bench fw anomaly). device history lot: n/a. device history batch. subcomponent name: opm fw v2.20. material number: 0042-96254. lot number: n/a. serial number: n/a. device history review: q1) service history review - are there any qns associated with the complaint device¿s most recent service report? there were no nonconformances observed in most recent fsr. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? no. q3) complaint history review - have there been any other complaints against this console? 23-14909: aic - impella connect - usd card corruption not related to the reported complaint. phr summary: there were one complaint was discovered when reviewing the product history of console ic3922, which is not related to the reported complaint. | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
2913 | 1220648-2025-48916 | 23770661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770661&pc=PYX | 11/12/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATIC IMPELLA CONTROLLER |  |  | PYX |  |  | OPTICAL AIC W/IMPELLA CONNECT, PKGD, JP |  |  | ABIOMED, INC. |  |  | US |  | 1566737 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a controller error alarm occurred during the nighttime hours. the patient's hemodynamics were normal, and the alert was resolved within approximately two minutes. when the alert occurred, the position waveform was not displayed, but the motor waveform was. the alarm subsequently resolved naturally, and since it was nighttime and there were few staff members, the person in charge was contacted in the morning. aic replacement will be carried out today. *dr. comment below added by abj-gvp controller error alarm occurred. additional information (2025, nov, 12), problematic aic serial number: (b)(6), replaced aic serial number: (b)(6). | [5 empty] | Device received, inv in progress: the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The root cause of the controller error and the loss of placement signal is due to an optical bench communication protocol anomaly, resulting in misalignment of data communication packets.
2914 | 1220648-2025-48917 | 23770889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23770889&pc=OZD | 11/10/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728160 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Difficult to Insert | Detachment of Device or Device Component | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. impella was inserted through the left femoral artery but encountered significant difficulty advancing through tortuosity within the upper segment of the left femoral artery device and impella wire removed the cardiologist exchanged the short sheath for the long sheath and re inserted the impella following the usual steps again failed to advance the impella. fluoroscopy a kink within the peal away sheath and the outlet of the impella could not advance through. on removal of the impella device from the sheath it became apparent the device had fractured and the inlet and pigtail section of the device had near -totally detached from the cannula (still connected by the cannula wiring and impella insertion wire. the nitinol reinforced wire had unraveled the impella 0.18 wire and the almost detached inlet were pulled back into the long peel away sheath carefully. the sheath was pulled halfway out of the femoral artery with partially detached impella inside. the lfa access was regained by puncturing the sheath below the affected impella that was held with in the sheath using a kimmel needle and advancing a j wire through the wall of the peal away. having control of the artery the 14 ff impella sheath was removed and a 10 ff femoral sheath was placed with proglide sutures already in situ proglide sutures were tightened to help downsize the arteriotomy around the 10ff sheath. no significant blood loss. case continued with out impella as right femoral artery was not a viable option. cardiologist has pictures to show device and complications witnessed. | [5 empty] | This report is being submitted to correct the h6 medical device problem codes and to document the results of the completed product investigation. corrections made: h6 medical device problem code: added a150206; removed a150205 and a0501. h6 component code, type of investigation, investigation findings, and investigation conclusions have been updated to reflect completion of the investigation. d9 has been updated to indicate that the product was returned for evaluation. investigation summary. the cause of the difficult insertion was interaction with the introducer kink caused by vessel tortuosity. the cause of the product damage was interaction with the introducer kink as evidenced by the returned product and clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2915 | 1220648-2025-48919 | 23771435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23771435&pc=OZD | 11/11/2025 |  |  | 12/10/2025 |  |  | 11/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699949 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The canadian complainant reported that there was ischemia, major bleed, positioning issues, cardiac arrest, and sepsis during impella patient support. while on support, there was less than adequate flow to the distal limb. a mini femoral to femoral bypass was done to increase flow to the right leg where the impella was. the sheaths were changed out to pediatric ecmo sheaths with larger tubing to increase blood flow. the popliteal pulse had returned but not the dorsal pedal pulse. there was bleeding at the access site, which was successfully managed through repositioning the repositioning sheath. the alarm for ¿impella in aorta¿ alerted. the team repositioned the pump. pulseless electrical activity arrest was successfully managed through compressions. the impella was removed due to lack of distal flow to the leg. a fasciotomy was performed. fasciotomy leg was oozing serosang and dusky in color. the patient was also septic and on multiple pressors. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2916 | 1220648-2025-31941 | 23774259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774259&pc=OZD | 07/17/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 11/20/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED INC. |  |  | US |  | 1635221 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that the aorta placement signal consistently drifted above the calibrated a-line during patient support. support was maintained with the aic and implanted pump. no patient harm has been reported. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation for the reported issue has been completed. console logs, obtained from the cloud, show large number of positioning and suction alarms, which started occurring approx. 12 hours since pump start, at which time optical signal snr became low and erratic. console was not returned for investigation due to lack of maintenance contract and customer not responding to po. root cause of reportedly false positioning alarms could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
2917 | 1220648-2025-48910 | 23774306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774306&pc=OZD | 10/22/2025 | 12/10/2025 |  | 12/10/2025 |  |  | 11/12/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718844 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient underwent percutaneous coronary intervention with impella placement for acute myocardial infarction and was treated in the intensive care unit. during treatment, neurological abnormalities were noted, revealing cerebral hemorrhage. the impella was removed urgently and surgery was performed. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2918 | 1220648-2025-31942 | 23774617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774617&pc=OZD | 08/18/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 08/21/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1369445 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | A physician reported issues with placement and left ventricle signals; both were not visible during procedure, and all other values were displayed. the device was not removed. the pump was eventually successfully weaned and explanted, and no patient harm has been reported. | [5 empty] | Data analysis: console logs from the reported day of the event are consistent with the complaint. the pump was connected to the aic. after case start wizard step 6, the error "case start: optical sensor system error: disconnect cable" appeared. the pump was removed and reconnected, but the issue persisted. the aic was rebooted; after reconnecting the pump, calibration was canceled due to continued error. approximately one minute later, the alarm "placement signal not reliable (opm invalid signal, optical pump connected) #1036" was issued. the pump was run for a few minutes just before the aic was shut down. device analysis: console (b)(6) was received at field service in germany. the aic was not checked in for this complaint. the investigation was conducted by a technician without prior consultation with the assigned investigation engineer. the root cause was identified as the air gap, which is the case as the rt log data supports this finding with the lamp value remained at 100% while snr nearing 0 throughout the case. root cause: the root cause for the issue with placement and lv signal, both were not visible was due to the air gap the aic and the patient cable plug.
2919 | 1220648-2025-48920 | 23774624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774624&pc=OZD | 09/11/2025 |  |  | 12/10/2025 |  |  | 09/11/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1382144 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 and supported by an automated impella controller experienced an impella in aorta alarm due to optical signal loss. the pump position was reconfirmed. additionally, the patient experienced access site bleeding. in response, two units of packed red blood cells were transfused. later, the patient expired on support. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data analysis: console logs from the reported day of event are consistent with complaint and show large number of positioning alarms. device history review: the device passed all post sterile inspection checks.
2920 | 1220648-2025-48921 | 23774630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774630&pc=OZD | 11/13/2025 | 12/08/2025 |  | 12/10/2025 |  |  | 11/13/2025 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750931 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a 68 yr old male was implanted with a impella cp for percutaneous coronary intervention support (pci). impella kinked while insertion in common femoral artery after passing the peel away sheat. calcification was not detected in the fluoroscopy before.impella cp was after pci successfull weaned, explanation and closure without any problems. a few minutes later, the patient unexpectedly required cardiopulmonary resuscitation and expired in the cath lab. | [5 empty] | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. unable to deliver or advance: returned complaint impella cp visually inspected. both catheter kink and cannula kink observed near outlet observed. the root cause of delivery issue most likely was patient condition related due to calcified artery preventing successful delivery of impella. product damage (cannula kink): the root cause of cannula kink most likely was patient condition related due to calcified artery. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2921 | 1220648-2025-31943 | 23774637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774637&pc=OZD | 06/16/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 06/16/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432818 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that their aic showed that the optical sensor was not supported during case start. roughly two days later, an impella in aorta alarm was triggered with pulsatile motor current and a pulsatile placement signal. positioning was verified and patient support continued with the utilized devices. there was no patient harm. | [5 empty] | Analysis: the complaint was confirmed through log analysis. on (b)(6), before the complaint date, the imc logs show cp pump (b)(6) was connected, the system recognized an optical system sensor error, the pump was reconnected and was able to start up without error. the rt logs show at first there is a 0 log during the initial plug in before the rt log for the pump number began with -3276.8 optical sensor and 3000 mmhg placement signal with near zero snr. after the pump is reconnected as instructed by the console, the 5s are as expected. the same circumstances happen on the complaint date, (b)(6), when the first plug in with the rt log split over two logs, then after the console prompts for pump reconnection, the 5s was fine. fsr before and after show good 5s performance and do not indicate any repeat malfunction of the console. root cause: the optical signal issue was most likely caused by an air gap due to the -(b)(6) mmhg optical sensor and near zero snr but with optical signal performing as expected later in the cases. the cause of the repeat air gaps could not be determined since console was not returned.
2922 | 1220648-2025-31944 | 23774661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774661&pc=OZD | 07/14/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 07/14/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1120452 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal was not reliable following impella implant for patient support. despite the noted issue, the device was left in place and was subsequently managed by monitoring the motor current and hemodynamics. there was no patient harm. | [5 empty] | Analysis: on the complaint date, there are constant placement signal not reliable alarms (ic2001 imclog). the lt logs confirm low snr as placement signal was out of range at 3000 mmhg (ic2001 ltlog low snr). the console went in for pm shortly afterward and the console had the optical bench replaced due to non-conformance. root cause: the cause of the optical signal issue was most likely a defective optical bench that was replaced during pm (b)(6) 2023.
2923 | 1220648-2025-31945 | 23774694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774694&pc=OZD | 07/20/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 07/20/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic had an unreliable placement signal. investigation concluded that there was a software bug that manifested with the combination of use error. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that after previously running pump (b)(6) was connected to (b)(6), console presented with placement signal not reliable alarm (#1048 - signal out of range). rtlog data shows that while raw optical data from pump (b)(6) on (b)(6) was valid and within expected range, the raw optical data from same pump on (b)(6) had unusual values of 3276.8mmhg (all other optical parameters - snr, contrast, lamp, gain - were within specification). within several seconds since pump connection and start on (b)(6), calculated placement signal railed at 3000 (signifying unreliable signal - not an actual value), but raw optical data remained the same, implying that the optical bench was sending the unusually high values, without identifying the signal as invalid - in which case it would set its values to -3276.8 (0×8000 signed integer). further analysis of imc log revealed that just prior to connecting pump (b)(6), operator accessed ¿optical bench service¿ screen upon which the optical bench was reconfigured for 5s parameter measurement (as per design) where raw optical data reports values of cavity length (in nm), not placement (in mmhg). the service screen was promptly exited, however after pump (b)(6) was connected to console and its eeprom read (including value of calibrated g0), the optical bench had not been re-configured for g0 and units of ¿mmhg¿. as the result all values measured and transmitted from optical bench to epc corresponded to cavity length, not pressure, triggering ¿placement signal not reliable¿ alarm due to signal out of range (#1048).
2924 | 1220648-2025-31946 | 23774752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774752&pc=OZD | 10/09/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 12/06/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330532 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in japan had a cp support ongoing in which there was a sensor failure. there was no harm or injury. the pump logs point to the console (aic) being the issue. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data analysis: imc logs of console (b)(6) during a case with pump (b)(6) showed there was one instance "ossiopsens set invalid_bad_snr_sample_mask" before the event ¿placement signal not reliable (opm invalid signal, optical pump connected]) - alarm¿, event #1036. imc logs also show one instance ¿controller error (opm comm stopped) [optical pump connected] - alarm¿, event #1043. both event #1036 and #1043 resolved automatically after 35 minutes and 59 seconds. rt logs on the complaint date confirmed the reported failure mode, given that the contrast oscillating between -(b)(6). rt logs also showed unrealistic raw optical sensor signal of -(b)(6) unrealistic lamp signal of 1638.4%, and unrealistic placement signal data of 3002 mmhg. device analysis: console (b)(6) was sent back fs for further investigation. the issue was not reproduced during testing; however, device data is consistent with intermittent loss of light in the optical system due to optical bench malfunction. root cause: reported controller error and ¿placement signal not reliable¿ alarms on (b)(6) during a case with pump (b)(6) were caused by intermittent malfunction of the optical bench. all pertinent information available to abiomed has been submitted at this time.
2925 | 1220648-2025-31947 | 23774786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774786&pc=OZD | 11/02/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 11/02/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1699285 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the aic alarmed that the impella position was not reliable. waveforms were all present and within normal limits, flows were excellent and patient hemodynamics were satisfactory. after about an hour, the alarm resolved with no change in any metrics or patient condition. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ¿ on the complaint date, (b)(6) 2023, 11 minutes after the case was initiated with pump sn: (b)(6), the console displayed ¿placement signal not reliable (optical signal-to-noise ratio),¿ event #130. the alarm occurred intermittently throughout the case, confirming the reported failure mode. ¿ the rt log indicates that the snr dropped below 250. the lt log shows that the snr remained low throughout the case. device analysis: this console was tested after arriving at fs. although the snr was within specification, it was low at 3810. the voltage supplied to the optical bench from the pbm connector j3 was 4.9v as intended. the lamp resistance was measured at 5 ohms. however, the average voltage to the lamp at the optical bench connector p4 was 4.3v, whereas in a known good optical bench, it measures 4.8v. this indicates a likely optical bench malfunction. root cause: the root cause of the 'placement signal not reliable' issue was an optical bench malfunction. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
2926 | 1220648-2025-31948 | 23774869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774869&pc=OZD | 08/23/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 08/23/2023 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1151537 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | After the initial start up the impella was functioning as it should. after several minutes the placement signal alarm unreliable alarm came on. the placement signal looked fine and matched the patient¿s blood pressure. then the signal itself would come and go intermittently. then once the catheter was weaned to p2 the signal stabilized and returned to normal. the complainant reported past issues with the pump, but the pump was not removed in this instance. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs from the reported day of event are consistent with complaint and show large number of error messages related to invalid placement signal samples due to low snr (ossiopsens set invalid_bad_snr_sample_mask), followed by a ¿placement signal not reliable (optical signal-to-noise ratio) - alarm issued¿ event #130. additionally, the logs revealed two instances of the same error occurring with different pumps, happened approximately 30 days prior to the complaint date. see attachment section for imc_logs_ic2941. rtlog data from ic2941 for pump 449925 shows unrealistic value of raw optical signal oscillating between 0 and -3276.8, and snr below 200 (ic2941_rtlog_449925_detail.png) device analysis: this console was tested after arriving in fs. console inspection revealed that the 5s parameter has very low contrast and low snr, depending on test cavity length: 5s - pass cl = 14340nm 5s - fail cl = 15940nm cavity length: 14330 nm snr: 5514 contrast: 59.5% lamp: 100.0% gain: 1.30 mano: 758.4 mmhg light: 2.68 v cavity length: 15941 nm snr:222 contrast:15.3% lamp: 66.6% gain: 1.28 mano: 758.3 mmhg light: 2.70 v monitoring of the output of the optical bench ccd detector revealed distortion of the analog signal, which was negatively affecting placement signal calculation. console inspection also revealed evidence of corrupt microsd card on the rlm (rl05736 partition switching.png), which prevented the unit from connecting to wifi and streaming data (unrelated to this complaint), and a kinked impellatronic ribbon cable (unrelated to this complaint). root cause: placement signal not reliable issues were caused by a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
2927 | 1220648-2025-31949 | 23774874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774874&pc=OZD | 02/29/2024 | 12/10/2025 |  | 12/10/2025 |  |  | 02/29/2024 | 12/10/2025 |  |  |  |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1113514 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | A complainant reported that the aic showed the error message "optical sensor not supported." this was immediately resolved by switching aics. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: at the customer site on (b)(6) 2024, once the pump (sn (b)(6)) was connected, after 8 seconds, the console received an ¿optical sensor not supported ¿ alarm¿ event set #127. still, the pump was used for 4 hours and 30 minutes. event set #127 was cleared only after pump disconnection. on the complaint date ((b)(6) 2024), the user connected the demo pump four more times, and every time, the console received the event set #127 persistently, which confirms the reported failure mode. note: prior to this complaint, (b)(6) underwent service on (b)(6) 2023, as per fsr notification (b)(4). during the service, although the sau_motor_1 hardware revision was set to ¿257¿, it was accidentally reset to ¿1¿ before the console was released back to the customer. device analysis: this console was tested after arriving in fs. the aic ¿fwcheck¿ confirmed that the sau_motor_1 hardware revision was set to ¿1¿ as received. after resetting it back to ¿257¿, there was no event set #127 upon the optical pump connection. during testing, the log recorded a piston block event. a minor amount of glucose was visually confirmed underneath the gasket, which is unrelated to the reported failure mode. root cause: the root cause for the optical sensor not supported alarm was incorrect sau_motor_1 hardware revision. all pertinent information available to abiomed has been submitted at this time.
2928 | 1220648-2025-31950 | 23774875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23774875&pc=OZD | 06/20/2023 | 12/10/2025 |  | 12/10/2025 |  |  | 06/20/2023 | 12/10/2025 |  |  | 03/31/2025 |  | 12/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024321573 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a 19 day impella 5.5 support ongoing when the optical signal was lost. the pump remains on for support despite the loss of signal. no harm or injury noted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the product was not returned for investigation. data logs were returned for analysis. the patient was on ecmo during impella support, was given blood, and was having tremors. data log analysis showed that placement signal not reliable alarms were triggered on 20 june 2023 between 12pm and 11:30pm. there were intermittent placement signal spikes with no known pattern. gain values spiked to 500 and contrast values spiked to 5000 intermittently during the case. snr values were noisy with random drops to 0. no placement signal low alarms or trending placement signal values were seen which could confirm patient condition to be responsible for the placement signal not reliable alarm. optical signal issue: the root cause of the optical signal issue in the form of placement signal not reliable alarms was not determined due to lack of sufficient clinical details, unreturned product, and lack of conclusive evidence from data analysis. all pertinent information available to abiomed has been submitted at this time.
2929 | 1220648-2025-48845 | 23755623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23755623&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026754198 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with restrictive cardiomyopathy was in need of a heart transplant. the patient was status 3 on milrinone but ultimately needed escalation to an impella 5.5. while on support the patient tested positive for heparin-induced thrombocytopenia and thrombosis so they were switched to bivalirudin. the patient was successfully bridged to a heart transplant. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
2930 | 1220648-2025-48846 | 23755709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23755709&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671063 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 100 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in was implanted with an impella cp device for mechanical circulatory support. the patient was presented to the emergency room by emergency medical services (ems) after experiencing a sudden cardiac arrest at home. the patient¿s husband immediately began cardiopulmonary resuscitation while awaiting ems. ems arrived approximately 15 minutes after collapse. the patient was successfully defibrillated by ems and return of spontaneous circulation was achieved. after arrival to the hospital, the patient was admitted to the intensive care unit (icu). the patient experienced another episode of cardiac arrest in the icu and it was noted that the patient¿s troponin levels were elevated. the patient was brought to the cardiac catheterization laboratory (ccl) for further treatment. in the ccl, the patient required an increased amount of chemical hemodynamic support, and it was noted that the patient¿s hemodynamics continued to be subpar. the medical team decided to deploy extracorporeal membrane oxygenation (ecmo) to support the patient, however while awaiting ecmo cannulation, the decision was made to place an impella cp for mechanical circulatory support. an impella cp was prepped and inserted via the left femoral artery in standard fashion without complication. left heart catheterization was performed and left coronary arteries were noted to be clean. right heart catheterization was performed, and right atrial pressures were noted to be high, with a fick cardiac output of 5.51, a cardiac index of 2.67, pulmonary artery pulsatility index of 0.56, and copious pulmonary edema was noted. ecmo was then successfully cannulated, and the patient¿s left ventricle was successfully unloaded, vasoactive medications were weaned, and the patient¿s pulmonary edema decreased. it was noted that in the ccl, the medical team was unable to obtain bilateral dorsalis pedis pulse and the patient¿s lower extremities were cold and clamped. the patient was transferred to another hospital for further treatment. on impella cp support day 2, the patient was escalated from the impella cp to an impella 5.5. following impella cp explant from the femoral artery, a thrombectomy was performed. additional information regarding the impella 5.5, and the final patient outcome was not available in the complaint record accessible for assessment at time of reporting. | [5 empty] | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary no product was returned for investigation. ischemia/ thrombosis: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review ==> device with sn: (b)(6) passed all post sterile inspection checks.
2931 | 1220648-2025-48847 | 23755767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23755767&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728764 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had multivessel coronary artery disease presenting to the lab for impella cp assisted percutaneous coronary intervention. multiple attempts to deliver the impella cp were unsuccessful. the impella cp was aborted due to vascular anatomy. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy and resistance at aorta preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2932 | 1220648-2025-48848 | 23755853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23755853&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730193 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 66 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented to the hospital with shortness of breath and chest pain. and was found to have non-st-elevation myocardial infarction. left heart catheterization was done showing heavily calcified proximal left anterior descending artery and was turned down by surgery. a decision was made to proceed with impella cp assisted primary percutaneous coronary intervention (pci). prior to the case, the patient was intubated due to worsening shortness of breath and inability to lie flat. during intubation, the patient became hypotensive requiring significant levophed to maintain pressure. the right femoral artery was accessed and the groin angiogram showed tortuosity in the vessel. due to history of heparin-induced thrombocytopenia, angiomax bolus was administered. it was report that a hematoma was noted with a small ooze at the impella site. angiomax was held and activated clotting time was 235. manual pressure was held and the hematoma resolved. hemoglobin dropped to 6.7 from 7.5 preprocedural. one unit of blood was ordered. it was also reported that there was suspicion that the patient is septic from pneumonia. the patient was febrile and antibiotic treatment continued. improved hemodynamics were noted. the bleed and hematoma at the impella site were resolved. later, it was reported that hospice was consulted. the treatment team discussed options with the family and the decision was made to wean and remove the impella to proceed with hospice. the impella removal was uncomplicated however, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned. the cause of the sepsis and fever was not determined due to insufficient clinical details. the cause of the bleed and hematoma was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2933 | 1220648-2025-48849 | 23755958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23755958&pc=OZD | 11/09/2025 |  |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764608 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted because the impella cp could not cross the aortic valve. a new pump was implanted to support the patient. the patient was in stable condition. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and unable to cross aortic valve after multiple attempts preventing successful delivery of impella. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2934 | 1220648-2025-48850 | 23756180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23756180&pc=OZD | 11/09/2025 | 01/07/2026 |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714660 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced occlusion of the right coronary artery with a clot. the patient was treated with angiomax. the impella then alarmed for high purge pressure. the purge cassette was changed. no change was noted. flow was 1.1 - 1.7 milliliters. the physician did not want to exchanged the device. the physician believed it was sue to patient coagulopathy issue. the physician decided to explant the impella cp and not replace. the patient's systolic was 200's at the time on 20 of levophed. levophed was stopped and pressure remained systolic at 160s. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Investigation summary: cp pump sn615113 and 2 purge cassette returned on 11/18/2025. purge pressure high - after successful insertion of the device, clotting occurred and purge pressure high alarmed. team checked for kinks, no kinks found, purge cassette was changed. no changes, flow 1.1-1.7ml. physician believed it was due to patient coagulopathy issue. data log review shows purge pressure rises over 20 minutes before hpp alarms after case start. the pump was returned and inspected where a thin film of biomaterial was found in the purge gap. the pump was run with both returned purge cassettes and high purge pressure recreated with both. the cause of the high purge pressure was biomaterial build up due to the biomaterial found in the purge gap of the returned pump and field logs showed purge pressure rising over 20 minutes. patient codes - thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1955291. device history batch: subcomponent lot: n/a. device history review: pump s/n 615113 passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
2935 | 1220648-2025-48852 | 23756453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23756453&pc=PYX | 11/09/2025 |  |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026741941 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced a placement signal not reliable alarm. the patient survived. | [5 empty] | No product returned for investigation. data logs are not available. the cause of the placement signal alarm was not determined due to no product or data logs returned for investigation. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2936 | 1220648-2025-48851 | 23756471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23756471&pc=OZD | 11/09/2025 | 12/09/2025 |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733803 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 74 year old female was implanted with impella 5.5 to support heart bypass procedure. during insertion wire support was lost while tracking device down ascending. multiple unsuccessful attempts made to navigate tortuous anatomy. decision made to remove and place impella cp. | [5 empty] | H6: corrected the investigation findings and investigation conclusion codes. h11: added information regarding the results of this investigation. investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult and tortuous anatomy preventing successful delivery of impella. device history lot: device lot: 1967964. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2937 | 1220648-2025-48853 | 23756737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23756737&pc=OZD | 11/09/2025 | 12/16/2025 |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694506 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Heart Block | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 79-year-old male was implanted with a impella cp for cardic arrest support. it was reported that patient went into 3rd degree block with transvenous pacemaker on backup. patient had low grade fever and possible infection of unknown etiology, with cultures pending. in addition, site oozing was also noted and site was redressed with angle matching. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Correction: e4 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. access site adverse event: the cp pump visually inspected, no issue observed on the cannula or the repo sheath unit. no founding issue. the cause of the oozing issue most likely was use related. heart block and fever: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2938 | 1220648-2025-48858 | 23756952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23756952&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733132 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issue and product damage during insertion attempt of impella 5.5. it was noted that implant of pump was aborted due to vascular kinking. no impella was inserted as a result. there was no reported patient harm. additional information was received. no device is available for investigation. the described "kinking" was related to the vascular anatomy of the patient. the computed tomography showed approximately 110 -120 degree angle in the subclavian artery/truncus. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | No product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. product damage (catheter kink): the cause of product damage was most likely patient condition related due to patient's anatomy. device (sn: (b)(6) passed all post sterile inspection checks.
2939 | 1220648-2025-48859 | 23757014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757014&pc=OZD | 11/06/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026745229 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient encountered ventricular tachycardia during impella 5.5 insertion and during usage in the intensive care unit. the patient remained on support and patient outcome after explanted was noted as survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 and h6 have been updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation investigation summary device analysis: 5.5 complaint pump was returned cut. pump visually inspected, no issue or abnormality found. peri procedural adverse event, ventricular tachycardia (vt): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review the pump passed all post sterile inspection checks.
2940 | 1220648-2025-48856 | 23757022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757022&pc=OZD | 11/09/2025 |  |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. - impella cp placed independently without incident. on arrival to unit, marked bleeding noted at impella access site. repo sheath had slipped out of skin and sutures were loose. repo sheath repositioned in skin, replaced previous sutures with new forward tension sutures. fem stop still in place, bleeding well controlled. hematoma present, but not actively growing. no pressors or inotropes in use, starting vasodilators for blood pressure control. (hemolysis, hematuria, bleeding, hematoma). | [5 empty] | Complaint coding was updated to better reflect the reported event. consequently, section h6 medical device problem codes were corrected and repopulated accordingly. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data log was not provided or could not be retrieved from the remote server. access site adverse event (major bleed, hematoma): clinical details indicate marked bleeding at the impella access site upon arrival to the unit. the repo sheath had slipped out of the skin and sutures were loose. the repo sheath was repositioned and new forward tension sutures were placed. femstop remained in place and bleeding was well controlled. a hematoma was present but not actively growing. the cause of the access site bleeding and hematoma was most likely use-related, associated with sheath management and suture placement. pump suction: the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data logs. mechanical interaction with blood (hemolysis): the cause of the issue was not determined without returned product investigation and sufficient clinical details, including data logs. clinical symptom (hematuria): the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2941 | 1220648-2025-48860 | 23757113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757113&pc=OZD | 11/09/2025 | 12/09/2025 |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718537 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. ao placement signal was lost during case | [5 empty] | The investigation was completed. impella cp returned. no data logs returned. the clinical details state that after placement of the pump, shockwave was used and the pump lost ps. the returned product did not have any kinks in the catheter. the laser continuity revealed light leaking out of the sensor face. the 5s parameters taken were indicative of a damaged sensor face. based on the clinical details and returned product analysis, the root cause of the optical sensor issue is due to a material crack due to an interaction with shockwave. per capa-013581, instructions for use cp: 0048-9007 was updated to recommend physician users should ensure the shortest distance from the shockwave coronary ivl device to the impella optical sensor is =20 mm. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2942 | 1220648-2025-48861 | 23757272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757272&pc=OZD | 11/09/2025 |  |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731033 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient experienced runs a vtach last pm. amio and lido boluses and drips started. patient was sedated and cardioverted. amio drip was stopped this am. patient still experiencing short runs of vtach. impella p level was increased to p6. expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: tachycardia: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2943 | 1220648-2025-48855 | 23757522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757522&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697927 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was brought to the cardiac catheterization lab for high-risk percutaneous coronary intervention (pci) with impella cp support. failed perclose deployment during initial access was reported resulting in loss of arterial access. manual pressure was held for an appropriate amount of time, then the right femoral artery was re-accessed. the medical doctor spoke with vascular surgery and a decision was made to leave the 14 french sheath in place after pci and pump removal. successful impella placement and protected pci was noted. hemodynamics were maintained during shockwave and stenting of distal left main artery/ostial left anterior descending artery. the pump was weaned and explanted without further issues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | G11: added additional information regarding the results of the investigation. the product was not returned for investigation. data log review was not required for this case. the cause of the vascular damage issue was not determined without returned product investigation and sufficient clinical details.
2944 | 1220648-2025-48863 | 23757650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757650&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500868 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and hematuria. an echocardiogram was performed and revealed the pump was positioned towards the mitral valve. the pump could not be repositioned. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2945 | 1220648-2025-48862 | 23757666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757666&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759597 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to patient anatomy. the impella cannula could not pass through the aortic valve. the aortic valve was probed without difficulty using a pigtail catheter and a terumo guidewire. the 0.018" impella guidewire was inserted into the left ventricle without difficulty. despite repeated pushing and tensioning of the guidewire, the impella cannula could not pass through the aortic valve. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 has been updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
2946 | 1220648-2025-48864 | 23757915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23757915&pc=OZD | 11/10/2025 | 12/09/2025 |  | 12/09/2025 |  |  | 09/27/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731182 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced anemia and was transfused two units of packed red blood cells. a ct revealed a hemothorax so a chest tube was placed and anticoagulants were withheld. later, the patient experienced a discrepancy in non-invasive pressure and intermittent placement signal not reliable alarms. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. optical sensor issue: clinical details report a discrepancy between the pressure system (ps) reading and the non-invasive pressure reading on 10/nov. psnr alarms were also triggering at the time. data logs were reviewed and the reports were confirmed. returned product was investigated and optical signal metrics were within expected range, and the pump passed laser continuity testing. the pump was connected to a console and run in a bucket of water and even though optical signal 5s remained stable, the raw optical sensor was reading 400 mmhg and a psnr alarm triggered. based on the findings from data log review and the returned pump, the cause of the optical sensor issue was determined to be sensor wear. the optical sensor wear occurred prior to the device's intended 60 day design life. major bleed: clinical details report that the patient received 2 units of packed red blood cells on 26/sept/2025. the following day, computed tomography (ct) scan revealed a hemothorax. the cause of the injuries could not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
2947 | 1220648-2025-48865 | 23758139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23758139&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026744750 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support that the placement signal sensor was disrupted during intravascular lithotripsy of the right coronary artery. the pump flows remained stable, and the impella was explanted as expected at the end of the procedure. no patient harm was reported. | [5 empty] | It was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
2948 | 1220648-2025-48842 | 23758176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23758176&pc=OZD | 11/12/2025 | 12/02/2025 |  | 12/09/2025 |  |  | 11/09/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757316 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Malfunction | No | Yes | Fracture | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user faclity reported an impella cp in a 40 year old male for heart failure support. the pump appears to be broken off at transition from cannula to the outlet. it also torn in the middle. according to the intensive care physician, no force was required to remove the impella. the other half of the pump was discarded. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2949 | 1220648-2025-48866 | 23758376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23758376&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656822 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and resistance at axillary artery preventing successful delivery of impella. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2950 | 1220648-2025-48867 | 23758542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23758542&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller was unable to recognize the purge disc despite attempting multiple purge tubings. a different controller was used to support the patient. no patient harm was reported. | [5 empty] | H6 code c20 was reported incorrectly on the initial report and code c21 should of been reported. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. | Patient information was not provided. section a was left blank. the impella device was received from the customer. the investigation of the device underway. upon completion of the investigation, a supplemental mdr will be filed.
2951 | 1220648-2025-31940 | 23758691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23758691&pc=OZD | 09/01/2021 | 12/09/2025 |  | 12/09/2025 |  |  | 09/01/2021 | 12/09/2025 |  |  | 06/30/2023 |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022106253 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella support was experiencing placement signaling issues due to allegedly a damaged sensor. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
2952 | 1220648-2025-48868 | 23758930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23758930&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765577 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma at the access site. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the hematoma was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2953 | 1220648-2025-48870 | 23759077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23759077&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722864 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a suction alarm. one unit of packed red blood cells and two units of platelets were transfused. additionally, a suction alarm occurred. there was no adverse impact to patient management. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: the cause of the injury was not determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. data log was not provided or could not be retrieved from the remote server. clinical details indicate that suction alarms occurred overnight and again during a vagal episode. increased output from the jackson-pratt drain required transfusion of prbcs and platelets. the cause of suction issue could not be determined without returned product investigation and sufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2954 | 1220648-2025-48869 | 23759121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23759121&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  | 02/28/2026 |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493288 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 100 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after implantation of an impella cp the patient experienced ectopy. it was suspected that the patient's small left ventricle contributed to this condition. the pump was repositioned without resolution. the p level of the pump was reduced to resolve the ectopy. no patient harm was reported. | [5 empty] | The cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A5 and a6 are unknown. the impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2955 | 1220648-2025-48872 | 23759436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23759436&pc=OZD | 11/10/2025 | 12/09/2025 |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728058 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 64 yr old male was implanted with a impella cp for percutaneous coronary intervention to support cardiogenic shock. a hematoma was reported with radiofrequency ablation, manual pressure and fem-stop. peel away sheath wasn¿t peel away inside the body. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: hematoma: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks.
2956 | 1220648-2025-48873 | 23759595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23759595&pc=OZD | 11/10/2025 | 01/27/2026 |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1734346 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 68 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The patient was on impella 5.5 support with an automated impella controller (aic) (serial number(b)(6) for 25 days. the bedside nurse stated (b)(6) had ¿controller error¿ alarm, with patient maintaining normal flows during the alarm. the bedside medical doctor instructed the nurse to change aic controller. the nurse stated the patient was stable before and after aic controller change. (b)(6) labeled and sent to biomed department at (b)(6) hospital of usc. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
2957 | 1220648-2025-48874 | 23759813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23759813&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 68 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that while on impella 5.5 support there was an automated impella controller (aic) change. the nurse stated that the aic (serial number (b)(6)) gave the message ¿purge cassette not detected¿, and said there was a black piece covering the purge disc slot inside of the controller door. the nurse changed to a different aic controller without issue. the nurse states there was no patient harm during this event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2958 | 1220648-2025-48875 | 23760090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23760090&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693497 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 68 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was brought into the operating room for upgrade from impella cp to impella 5.5 in setting of cardiogenic shock. it was reported that while on support of the impella 5.5 the patient has a fever/cough and was put on antibiotics. the pump was later explanted and the patient survived. | [5 empty] | Investigation summary: no product was returned for investigation. fever: the cause of the injury was not determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2959 | 1220648-2025-48876 | 23760289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23760289&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 |  |  |  |  | 12/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122390 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller generated a controller error. another automated impella controller was used to continue patient support. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mc logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. device analysis: console (ic2022) was sent back fs for further investigation. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | A4 is unknown. the impella device was received from the customer, upon completion of the investigation, a supplemental mdr will be filed.
2960 | 1220648-2025-48877 | 23760788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23760788&pc=OZD | 11/10/2025 |  |  | 12/09/2025 |  |  | 11/10/2025 | 12/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561186 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 had bleeding at the anastomosis site. the patient was transfused multiple units of red blood cells and platelets. additionally, the patient coded multiple times. the patient expired on impella support. | [5 empty] | Complaint coding was updated to better reflect the reported event. consequently, section h6 health effect clinical codes were corrected and repopulated accordingly. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: no product was returned. pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. access site adverse event: the cause of the bleed was not determined due to insufficient clinical details. device history lot: device lot: 1842983. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
2961 | 1220648-2025-31939 | 23736709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23736709&pc=OZD | 03/17/2022 | 12/06/2025 |  | 12/08/2025 |  |  | 03/17/2022 | 12/08/2025 |  |  | 10/31/2023 |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159341 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal monitoring was disabled due to ongoing false position alarms. support was maintained throughout the incident. no patient harm has been reported. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
2962 | 1220648-2025-31938 | 23736880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23736880&pc=OZD | 01/03/2022 | 12/06/2025 |  | 12/08/2025 |  |  | 01/03/2022 | 12/08/2025 |  |  | 09/30/2023 |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022144796 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella support experienced some intermittent falsely triggered position in aorta alarm. the rn reviewed the patient on bedside and found no impella issues. no patient harm has been reported. | [5 empty] | No product was returned. the logs confirm occurrence of 'impella position in aorta' alarms. the root cause of the optical signal issue was patient condition related since there was no motor current spike observed in the data logs prior to alarms and clinical report mentions alarms were falsely occurring.
2963 | 1220648-2025-48789 | 23739148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739148&pc=PYX | 11/06/2025 | 01/22/2026 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025580729 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | No |  | No | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis after 16 days on impella rp support, which is now affecting kidney and liver function. therefore, impella rp explantation is planned for today, followed by ecmo support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Investigation summary mechanical interaction with blood (hemolysis): the pump was returned for investigation. biomaterial was found on the returned product. no other physical damage was noted on the returned product. the pump was started in a bucket of water and the biomaterial was flushed out of the cannula. the pump then ran as expected per specifications. the cause of the hemolysis issue was most likely related to biomaterial ingestion, based on returned product investigation. clinical symptoms (acute kidney failure): the cause of the injury was not determined due to insufficient clinical details. device history lot device batch: device history batch subcomponent lot: na device history review the pump sn passed all post sterile inspection checks.
2964 | 1220648-2025-48805 | 23739155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739155&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569788 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Low Blood Pressure/ Hypotension | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hypotension and atrial fibrillation requiring cardioversion and medication. the patient expired. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2965 | 1220648-2025-48801 | 23739171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739171&pc=OZD | 03/28/2025 | 12/03/2025 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624635 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 61 year old male was implanted with impella cp for support during percutaneous coronary intervention (pci). hematoma was noted with no evidence of active bleeding post pci with ecchymosis. patient was found to have a reduced ejection fraction and underwent cardiac cath which showed normal left main, normal left anterior descending artery a chronic total occlusion (cto) of his left circumflex and cto of his right coronary artery. he was maintained in the coronary care unit with initiation of amiodarone drip given his rapid atrial fibrillation. he was started on metop. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the hematoma could not be determined since product was not returned and insufficient clinical details were provided.
2966 | 1220648-2025-48794 | 23739182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739182&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 89 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the access site. the patient was transfused one unit of packed red blood cells. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the major bleeding is determined to be use issue due patient movement because insertion site oozed while the patient struggled in icu. the device passed all post sterile inspection checks.
2967 | 1220648-2025-48791 | 23739209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739209&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658305 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 99 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria, pleural effusion, and pericardial effusion. the patient is in stable condition. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the hematuria and pericardial effusion was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2968 | 1220648-2025-48806 | 23739225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739225&pc=OZD | 11/07/2025 | 12/03/2025 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760868 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 55-year-old male was implanted with impella cp for mitral valve repair for right ventricle support. it was reported that the surgical team placed purse string sutures to impella site soon afterwards noted arterial blood dripping from seam between red impella plug and grey section of plug. decision made to replace cp emergently. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Red handle leak: the cause of the issue was not established as no product was returned and limited clinical information was provided. major bleed/ access site adverse event: the cause of the issue was not established as limited clinical information was provided. corrected information has been provided. the initial report inadvertently had the h1 ¿summary report¿ field selected. a follow-up mdr has been submitted to correct this administrative error, and the summary report checkbox has been deselected.
2969 | 1220648-2025-48790 | 23739264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739264&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552943 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma at the access site, and hematuria. compression was applied to achieve hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematoma/minor bleed : the cause of the injury is most likely use related since angle match dressing applied and bleeding stopped. hematuria : the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2970 | 1220648-2025-48793 | 23739267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739267&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714710 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after explantation of an impella 5.5 the patient experienced a cerebrovascular accident. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2971 | 1220648-2025-48803 | 23739282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739282&pc=OZD | 11/07/2025 | 12/03/2025 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622660 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 83 year old female was implanted with impella 5.5 for support for cardiac arrest .following a 5.5 insertion, patient lost pulses in right arm - same side as 5.5 subclavian implant, ultrasound revealed clot in brachial artery, patient brought to operating room, 5.5 pocket reopened, impella explanted, thrombectomy to subclavian artery, brachial artery cutdown - thrombectomy to brachial artery, sites closed, unclear if impella was cause or related to embolism. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the returned product was investigated. traces of dried blood were seen on the inflow and outflow cages, suture hub, and sterile sleeve. no other abnormalities were seen. ischemia, thrombosis: clinical details mention that the patient lost pulses in the right arm following impella insertion. ultrasound revealed a clot in the brachial artery. the root cause of the injury was unable to be determined due to insufficient clinical details. | Correction e4. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
2972 | 1220648-2025-48800 | 23739285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739285&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768669 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 77 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Tachycardia | Hematuria | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and anemia. the patient was transfused one unit of packed red blood cells. the patient also experienced ventricular tachycardia. lidocaine was administered and the patient was defibrillated. the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the anemia and hematuria was not determined due to insufficient clinical details. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2973 | 1220648-2025-48802 | 23739286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23739286&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697048 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 86 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed a hematoma at the access site requiring surgical intervention. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. access site adverse event/hematoma: the cause of hematoma was most likely use issue related due to patient movement since hematoma developed at graft site after working with physical therapy. device history review: the device passed all post sterile inspection checks.
2974 | 1220648-2025-48792 | 23740850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23740850&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026757647 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Pulmonary Edema | Tachycardia | Respiratory Failure | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring cardioversion. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the respiratory failure and pulmonary edema was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2975 | 1220648-2025-48796 | 23740860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23740860&pc=OZD | 11/06/2025 |  |  | 12/08/2025 |  |  | 11/06/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708007 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella cp experienced suction alarms and decreased flow when the patient turned on his side. the pump was repositioned to resolve the issue. additionally, the patient experienced hemolysis. in response, the p level of the pump was reduced. | [5 empty] | This follow-up report is being submitted to add an h6 device problem code that was omitted from the initial report. h6 medical device problem code a01 (patient¿device interaction problem) has been added. | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | Correction : the follow up #1 (mwr-08122025-0001987682) report was sent stating the device was retuned but that information was sent inadvertently in error . the pump was discarded. only aic was retuned . product complaint # (b)(4). investigation summary: no product was returned. hemolysis: gross hemolysis was noted. no logs are available. the cause of hemolysis cannot be determined since insufficient clinical details were provided. device in wrong position: the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position when the patient turned on his side. suction alarms noted after that. device history lot: device lot: 1936559. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks.
2976 | 1220648-2025-48804 | 23740885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23740885&pc=OZD | 11/07/2025 |  |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739597 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an impella in aorta alarm. repositioning of the pump was unsuccessful so the pump was explanted and replaced with a new impella 5.5 pump. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was most likely repositioned incorrectly as pump alarmed ¿impella in aorta.¿ shortly after impella was repositioned by resident physician and was unable to be readvanced at bedside with echo guidance. the device passed all post sterile inspection checks.
2977 | 1220648-2025-48809 | 23741121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741121&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708299 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. while on support, it was reported that there was significant access site bleed. a purse stitch was applied, forward tension at the hub, vertical angle matching perpendicular to site, and gauze surrounding the repositioning sheath. 14 units of packed red blood cells were given. later, the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6: added code c13 to the investigation findings. h11: added the results of the investigation investigation summary no product was returned. access site adverse event: the cause of the bleed was most likely use issue related since hemostasis was achieved by adding additional purse stitch along with angle matching. device history lot device lot: 1951051 device history review device sn: (b)(6) passed all post sterile inspection checks.
2978 | 1220648-2025-48811 | 23741137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741137&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734237 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 96 KG |  |  | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Atrial Fibrillation | Pain | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, it was reported that there was slight soreness at the insertion site. no medication was required. the patient converted to atrial fibrillation and systemic anticoagulant was started. weeks later, a placement signal not reliable alarm was reported and the alarm was disabled. no further information was provided. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2979 | 1220648-2025-48808 | 23741238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741238&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651711 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated from the initial impella cp to a new impella cp for mechanical circulatory support. while on the second impella cp, the patient was brought to the operating room to escalate to an impella 5.5 when there was a large left ventricle thrombus on transesophageal echocardiogram. the following day, the family decided to withdraw care and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states) , ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation; section: impella cp with smart assist catheter insertion; section: axillary insertion of the impella cp with smart assist catheter; section: use of impella with transcatheter aortic valves; ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Medical safety/clinical reviewed the event. an impella cp device (pump 1, sn (b)(6)) was percutaneously inserted via the right femoral artery in a 61-year-old male for cardiogenic shock secondary to acute myocardial infarction (nstemi). the patient¿s past medical history and comorbidities are unknown. the patient initially presented with non-st segment elevation myocardial infarction and underwent percutaneous coronary intervention (pci) to the left anterior descending artery with placement of one drug-eluting stent. during hospitalization, the patient¿s left ventricular ejection fraction declined from 45% to 10%. the patient developed ventricular tachycardia requiring intubation and vasopressor support. labs demonstrated a lactic acid level of 2.5 mmol/l. upon arrival at the cardiac catheterization lab, left heart catheterization was unremarkable. right heart catheterization revealed a cardiac index of 2.1 l/min/m² and left ventricular end diastolic pressure of 30 mmhg. the patient was classified as scai stage e cardiogenic shock at the time of impella initiation. pump 1 was implanted without difficulty. immediately following initiation of support, purge flow was 3.4 ml/hr., increased to 3.8 ml/hr., and subsequently trended downward to below 3 ml/hr. with saturated pump flow alarms. concurrently, decreased pump flows and decoupling of aortic and left ventricular waveforms were observed. due to concern for device malfunction and inadequate support, pump 1 was removed and replaced with a second impella cp device (pump 2, sn (b)(6)). pump 2 was successfully implanted and support was reinitiated. the patient was prepared for transfer to another facility for continued management. the following day, during evaluation for escalation to an impella 5.5 device, transesophageal echocardiography revealed a large left ventricular thrombus. due to the presence of thrombus, the decision was made to initiate venoarterial extracorporeal membrane oxygenation (va-ecmo) instead of impella 5.5 support. the subsequent day, the patient¿s family elected to withdraw life-sustaining measures, and the patient expired. additional clinical details prior to withdrawal of care are unavailable. this report describes a decreased purge flow event associated with pump 1, resulting in device exchange. the device replacement procedure was completed without complication. this is a serious injury attributable to impella cp pump 1. although pump 2 will be reported in association with the patient¿s death, the patient¿s underlying conditions, acute myocardial infarction (nstemi), severe scai stage e cardiogenic shock, profound reduction in ejection fraction (10%), ventricular tachycardia and overall hemodynamic instability, contributed most to the demise outcome. correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: remove serious injury/ illness/ impairment (f12) and insert additional device required (f2301). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. note: h6. component and investigation type, findings, and conclusion remain unchanged as previously reported. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
2980 | 1220648-2025-48814 | 23741403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741403&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763573 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 98 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma and hypotension. the patient was treated with ionotropes. despite applying pressure, the hematoma did not resolve and the patient's hemoglobin continued to drop. the patient was transfused six units of packed red blood cells. the patient became unstable and coded. despite cardiopulmonary resuscitation, the patient expired. | [5 empty] | The cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the cause of the hypotension and cardiac arrest was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2981 | 1220648-2025-48812 | 23741504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741504&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768783 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 93 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. this impella cp had some position issues and had to be repositioned a couple of times. the last time it was repositioned, was at beside and urine had started to turn red. it was repositioned and a discussion was had about recommending turning the p level down a couple p levels to see if that cleared the red urine. this was agreed upon and the impella was turned to p-5. the urine cleared up in the last couple of hours. no drop in hemoglobin was noted and no blood products were given. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: data logs showed pump ran without issue during support. there were positioning & suction alarms triggered in early of the case but resolved. there were no mc spikes, abnormal ps or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2982 | 1220648-2025-48813 | 23741649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741649&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1751872 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a controller failure alarm on the supporting automated impella controller. | [5 empty] | Data analysis: the case began on october 20, 2025, and the epm hardware revision was verified as correct, which is a known good revision. however, after 18 days of pump support, the console triggered a controller failure alarm, event #1029, which persisted for the remainder of the case, leading to an aic exchange. device analysis: the console (b)(4) was returned to field service for investigation. the reported failure mode could not be reproduced. the console was tested with a known good pump and purge cassette, and no issues were observed. a firmware code was extracted from the impellatronic pca board and compared against a known good version. the extracted code was found to be corrupted. root cause: the root cause for the controller failure was due to a impellatronic pca firmware corruption. however, the cause of the firmware corruption was not been determined. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2983 | 1220648-2025-48816 | 23741663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741663&pc=OZD | 11/09/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764111 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 87 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ectopy. the patient was treated with magnesium and the pump was repositioned. the patient developed bacteremia and the pump was explanted. the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the arrhythmia and infection was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
2984 | 1220648-2025-48817 | 23741906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23741906&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in was implanted with an impella cp for mechanical circulatory support. the nurse reported the patient developed persistent ventricular tachycardia requiring 35 shocks in total. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. tachycardia : in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
2985 | 1220648-2025-48820 | 23742041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23742041&pc=OZD | 11/09/2025 |  |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766082 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 62 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost pulses in the lower leg. the pump was explanted. later, the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
2986 | 1220648-2025-48821 | 23742232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23742232&pc=OZD | 11/09/2025 |  |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026748997 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 71 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a gastrointestinal bleed. the patient was treated with a colonoscopy and is in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: gi bleed going for colonoscopy today. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
2987 | 1220648-2025-48822 | 23742391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23742391&pc=OZD | 11/08/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026739645 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 74 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. - impella 5.5 did not advance entirely due to the patient¿s vascular anatomy. after multiple wire exchanges using abiomed 0.018 and bos sci v-18, in addition to a buddy wire and use of rotaglide, the surgeon opted to abort the procedure. ct angio showed that the artery narrowed to approximately 4.5 mm near the junction of the innominate and subclavian artery. patient to remain on iabp and gtts. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy and narrowed artery preventing successful delivery of impella. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
2988 | 1220648-2025-48823 | 23742839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23742839&pc=OZD | 11/08/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026756826 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 102 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. on support day 5, it was reported that the patient experienced supraventricular tachycardia and/or ventricular tachycardia, and a lidocaine drip was discontinued and changed to procainamide, which was increased to 3 mg. the impella 5.5 was repositioned, however the following day it was noted that the patient continued to experience ventricular tachycardia. additionally, it was noted that the patient expired while on impella support. | [5 empty] | A5 is unknown. based on the information available, the impella 5.5 cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was discarded by the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2989 | 1220648-2025-48824 | 23743136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23743136&pc=PYX | 11/08/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026767438 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Device Sensing Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated to another impella rp flex device for continued mechanical circulatory support. upon insertion/device start of the 2nd impela rp flex it was reported that there was a ¿placement signal not reliable¿ and ¿flow calculation disabled¿. it was also reported that the patient decompensated yesterday while being positioned, she had arrhythmia and heart rate dropped to 20¿s. the patient expired. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data logs not available. regarding the arrhythmia the patient decompensated and had arrhythmia and drop in hr. the root cause of the injury was unable to be determined due to insufficient clinical details. placement signal issue alarm reported at case start with flow calculation disabled. no product or logs returned. the cause of the placement signal issue was not determined due to no product or data logs returned for investigation. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. revised section g pma as it was entered incorrectly on the initial report that was submitted. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. a 78-year-old female with right ventricular failure secondary to in-stent restenosis of a right coronary artery stent underwent placement of an impella rp flex via the right femoral vein. at the time of insertion, the patient was in cardiogenic shock classified as society for cardiovascular angiography and interventions (scai) stage e and had a known history of coronary artery disease. after two days of support, the patient was taken to the operating room for removal of impella rp flex pump 1 due to significant bleeding. a decision was made to place a new impella rp flex (pump 2) via the right internal jugular (rij) vein. upon insertion and device initiation of pump 2, console messages indicating ¿placement signal not reliable¿ and ¿flow calculation disabled¿ were observed. the issue was addressed, and the patient remained on support. two days later, while being repositioned, the patient acutely decompensated. the patient developed arrhythmia with heart rate decreasing into the 20s. following discussion, the family elected to withdraw care, and the patient subsequently expired. the bleeding complication associated with impella rp flex pump 1 is being reported as a serious injury. the impella rp flex pump 2 will be conservatively reported as death. the patient experienced an arrhythmic event followed by bradycardia while the patient was being repositioned. the patient¿s underlying condition contributed most to the demise outcome (cardiogenic shock scai stage e secondary to right ventricular failure from in-stent restenosis of the right coronary artery in the setting of underlying coronary artery disease). correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: inserted hemodynamic instability (e2343) and bradycardia (e060104). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | Section d catalog number was corrected. placement signal issue: the cause of the placement signal issue was not determined due to no product or data logs returned for investigation. arrhythmia/bradycardia/hemodynamic instability: the cause of the injuries was unable to be determined due to insufficient clinical details.
2990 | 1220648-2025-48810 | 23743193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23743193&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669918 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. the impella cp was placed independently in standard fashion with no reported issues. percutaneous coronary intervention was performed and no issues were reported. it was reported that after swapping the peel-away sheath for the repo sheath, the pump began experiencing suction alarms non-responsive to typical troubleshooting measures. the team turned fluoroscopy back on to visualize the pump and saw ¿an obvious kink¿ in the cannula near the motor housing. the pump was removed and a new impella cp was placed independently in standard fashion with no reported issues. the patients outcome was noted as survived. | [5 empty] | The investigation was completed. the pump was not returned for investigation. lt log shows the pump was utilized for around 2 hours. the log confirms a motor current spike to 1400, followed by a drop in pump flow and an active suction alarm at the end of the case run. rt log was not available to confirm the timing of the reported motor current spike. clinical details indicate that after swapping the peel-away sheath for the repo sheath, the impella pump began experiencing persistent suction alarms that did not respond to standard troubleshooting measures. fluoroscopy revealed an obvious kink in the cannula near the motor housing. pump suction: clinical details and available data log were not able to justify the cause of the suction. the cause of the suction issue was not determined without returned product investigation and sufficient clinical evidence. product damage (cannula kink): clinical details indicate that after swapping the peel-away sheath for the repo sheath, the impella pump began experiencing persistent suction alarms. troubleshooting was unsuccessful, and fluoroscopy revealed an obvious kink in the cannula near the motor housing. the cannula kink issue was most likely use-related, as it occurred during placement of the repositioning sheath. however, the exact damage to the cannula could not be verified without returned product investigation. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
2991 | 1220648-2025-48818 | 23743763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23743763&pc=OZD | 11/09/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026768459 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Ventricular Fibrillation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 82-year-old male was implanted with impella cp for support of heart failure. it was reported that during support patient had episode of ventricular tachycardia overnight. patient had oozing at insertion site which was reported to have resolved after dressing change. no current oozing during rounding. 2 units of packed red blood cells were given, and impella activately weaning with possible explant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This follow-up report is being submitted to document that the product has been returned for investigation and to add an h6 device problem code that was omitted from the initial report. h6 medical device problem code a01 (patient¿device interaction problem) has been added. section d9 has been updated to indicate that the product was returned for investigation. the h6 component code, type of investigation, investigation findings, and investigation conclusions have been updated to reflect that the product has been returned and that the investigation is currently in progress.
2992 | 1220648-2025-48825 | 23743825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23743825&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701072 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. hit positive. gave 1 unit of platelets overnight. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2993 | 1220648-2025-48826 | 23743980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23743980&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733544 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 143 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. (b)(6) 2025 (b)(4) renal failure, increased creatinine, crrt was initiated last night (b)(4) (b)(6) 2025 lactic and alt have worsened overnight/today. concerns for sepsis with elevated wbc, but wbc is also down trending (26.8 - 22.2) and no confirmed source of infection. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H5: added information. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. sepsis: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
2994 | 1220648-2025-48827 | 23744274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23744274&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705259 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. on (b)(6) 2025 1 unit prbcs given for low hemoglobin without signs or symptoms of bleeding. | Additional information received. it was noted there was ongoing, significant bleeding from bilateral groin access sites. patient remains on dobutamine, vasopressin, and norepinephrine. levophed requirements have escalated overnight to 35 mcg/min. patient has received multiple blood products as well over the last 24 hour course. a dialysis catheter is currently in the left groin and plans in place to transition access to subclavian. vascular surgery evaluated the patient and was considering placement of an additional suture to assist with hemostasis on the right groin where the impella cp is. discussed the case with the physician and vascular surgery. given worsening vasopressors needs and persistent bleeding, recommendation was made to escalate support to an impella 5.5. the following day the patient was extubated post right groin vascular repair. patient receiving 1 unit of packed red blood cells (prbc) and platelets. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b5. additional information was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleeding was determined to be patient condition because of the bleeding form both groin access sites. device history review: the device passed all post sterile inspection checks.
2995 | 1220648-2025-48828 | 23744534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23744534&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580454 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 99 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Multiple Organ Dysfunction Syndrome | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. post procedure: significant bleeding postoperatively. (b)(4) (b)(6) 2025 coded and reopened at bedside d/t to significant bleeding for 90 minutes. the patient ultimately expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. significant bleeding postoperatively. patient coded and reopened at bedside d/t to significant bleeding for 90 minutes. patient developed rv failure overnight. cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed/organ failure: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1856087. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
2996 | 1220648-2025-48830 | 23745161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23745161&pc=OZD | 11/08/2025 | 12/02/2025 |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647015 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp in a 74-year-old male for heart failure support. a leaking pressure reservoir was reported and fluid loss is low, pp 506 mmhg, pf 13 ml/h. recommendations and best practices shared. | [5 empty] | Investigation summary: purge leak - pump: rn reported leaking pressure reservoir. the cause of the issue was not established as no product was returned and limited clinical information was provided. device history lot: device lot number: 1911815. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
2997 | 1220648-2025-48831 | 23745209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23745209&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704128 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria, thought to be due to traumatic insertion of the pump. the patient also had oozing at the access site. the patient was in stable condition. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: available data logs were found to be insufficient with data missing from 1/7-1/9 and the hemolysis was reported to occur on 1/8. insufficient data logs showed pump ran without issue during support. there were no alarms triggered during the case. there were no mc spikes, abnormal ps or purge issues observed during support. mechanical interaction with blood/hemolysis: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. minor bleed: the cause of minor bleed was unable to be determined as limited clinical details were provided and no product was returned for review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
2998 | 1220648-2025-48795 | 23745890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23745890&pc=OZD | 11/07/2025 | 01/07/2026 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613687 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Renal Failure | Sepsis | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. thrombocytopenia, blood transfusion. aware: 7 nov 2025. occurrence: (b)(6) 2025. (b)(4): anticoagulation will be started tomorrow due to low platelets. patient received 2 unit prbc, 2 cryo, 2 platelets, 2 ffp. renal failure, additional device required. aware: 9 nov 2025. occurrence: (b)(6) 2025. (b)(4): patient is going on crrt at 1400 today. labs have all inclined in the past 24hrs. amio has been turned off in concerns for further damage to kidneys. medication required, sepsis. hit came back positive and patient switched to bival. patient on crrt lactic clearing slowly. care team discussing spesis, on broad spectrum antibiotics. values trending. paroxysmal atrial fibrillation. aware: 11 nov 2025. occurrence: (b)(6) 2025. (b)(4): patient experiencing afib/rvr and amio got started again. anemia aware: 25 nov 2025, occurrence: (b)(6) 2025. (b)(4): hgb low 1 unit prbcs. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient thrombocytopenia and received 2 units of packed red blood cell (prbc), 2 units of cryoprecipitate (cryo), 2 units platelets, and 2 units of fresh frozen plasma (ffp) as treatment. 2 days later the patient was placed on continuous renal replacement therapy (crrt) and the amio drop was stopped due to renal failure. sepsis was a concern and the patient was started on broad spectrum antibiotics. two days after stopping the amio drip the patient developed atrial fibrillation and rapid ventricular response (rvr), the amio drip was re-started. 10 days late the patient's hemoglobin can back low and was given 1 unit of prbc. the patient was reported to still be on support. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: added the explant date. | Correction: b5 re-written to appropriate formatting and without abbreviations. the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombocytopenia/ renal failure/ anemia: the cause of the injury was not determined due to insufficient clinical details. sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
2999 | 1220648-2025-48797 | 23745894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23745894&pc=OZD | 11/07/2025 | 12/07/2025 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026732537 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 113 KG |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 75 year old male was implanted wit impella 5.5 for heart failure support. second impella 5.5 was advanced using best practices however after entry into the axillary artery the impella failed to advance on further flouro inspection it appeared there was a hazy area in the axillary artery possibly indicating a calcified lesion. multiple additional attempts were made without success utilizing various arm positions and angles of entry as well as lubricating the impella 5.5 with propofol. a buddy wire was added for support without success and decision was made to abort insertion of 5.5 and insert impella cp. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the deliver issue was determined to be patient condition as the patient had calcification of vessels and resistance near the axillary artery preventing successful delivery of impella. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3000 | 1220648-2025-48798 | 23745958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23745958&pc=OZD | 11/07/2025 | 12/07/2025 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744158 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 113 KG |  | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 75 year old male was implanted with impella 5.5 for support during posterior fusion laminectomy.impella was introduced through the graft and into the axillary artery. however, the impella buckled on insertion to the axillary artery. all subsequent attempts to advance the pump resulted in cannula buckling in the same location. due to inability of catheter to advance without buckling the impella 5.5 was aborted. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the deliver issue was determined to be patient condition as the patient had calcification of vessel and resistance at axillary artery preventing successful delivery of impella. the cause of product damage, cannula kink, was most likely patient condition related since patient had calcification of the vessels. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3001 | 1220648-2025-48799 | 23746002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23746002&pc=OZD | 12/12/2024 | 12/07/2025 |  | 12/08/2025 |  |  | 11/07/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577764 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported in the long-term outcome and quality indicator (loqi) that a (b)(6) year old male was implanted with impella cp for support during high risk cardiac procedure. it was reported that after the procedure, the patient had worsening anemia with symptoms of lethargy and decreasing hemoglobin. the patient was transfused 1 unit red blood cells with an appropriate response and their appearance improved. the patient was awake and alert without acute distress. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the anemia could not be determined since product was not returned and insufficient clinical details were provided.
3002 | 1220648-2025-48829 | 23746181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23746181&pc=OZD | 11/08/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735834 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 60-year-old male was implanted with impella 5.5 for support during percutaneous coronary intervention for cardiac arrest. placement signal was noted to be significantly high after initiation despite optimal positioning & optimization of extracorporeal membrane oxygenation. possible sensor damage from mishandling of pump during insertion by st segment was noted and ultimately deemed culprit. pump flowing appropriately and left ventricle beginning to decompress so decision was made to leave pump in place and disable placement monitoring by physician. patient required x2 cardioversions overnight for ventricular tachycardia but remained stable on ecpella support throughout. | [5 empty] | The impella device has not been returned by the customer. if the device or additional information is received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. it was stated that the placement signal was noted to be significantly high after initiation despite optimal positioning & optimization of ecmo possible sensor damage from mishandling of pump during insertion by the surgical tech was noted and ultimately deemed culprit. due to lack of data logs or product returned, the root cause of the optical sensor issue cannot be established.
3003 | 1220648-2025-48832 | 23747083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23747083&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721361 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | Increase in Pressure | Device Sensing Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a flow issue and a clot was suspected. the pump was removed. the patient is in stable condition. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue - the cause of the optical sensor issue was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. high or saturated pump flow - the cause of the saturated flow was biomaterial ingestion due to motor current (mc) spike seen in the data logs that resulted in an increase in purge pressure. this then resulted in saturated pump flow of 3.5l/min at p-2 where the anticipated range should be no higher than 1.9l/min. additionally, noticeable clot entrainment at outlet of impella was noted upon explant. purge pressure high - the cause of the purge pressure high was biomaterial ingestion due to mc spike seen in the data logs that resulted in an increase in purge pressure and noticeable clot entrainment at outlet of impella was noted upon explant. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3004 | 1220648-2025-48833 | 23748047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23748047&pc=OZD | 11/08/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743939 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | False Alarm | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 63 year old male was implanted with a impella 5.5 for support of heart failure. an alarm indicating the pump position in the aorta was raised on the aic, and although the position was confirmed with an ultrasound, there was no problem (approximately 4.0 cm), and the flow rate seemed a little unusual. when the situation was checked with impella connect, a thrombus obstruction in the cannula was suspected. when anticoagulation management was confirmed, it was found that the patient had undergone thoracotomy to stop hemostasis due to cardiac rupture (woozing) after an myocardial infarction and had only undergone mylon purging, with no systemic anticoagulation. the bleeding had apparently now subsided, so the doctor decided to start systemic heparin and the call ended. impella connect then confirmed that the pump had been removed. the decision was made to remove the pump because a thrombus was suspected and the patient's hemodynamics were stable, so it was expected that there would be no problems with removal. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. false alarm: clinical details report false "position in aorta" alarms on 8/nov/2025. ultrasound verified pump position to be normal, and impella connect appeared to suggest thrombus as the cause. additionally, the patient's appts were in the low 30s at the time due to the fact that the patient was only on meylon purge post-thoracotomy. the acts were not reported at the time of the complication, but the starting act was 170, below the ifu recommendation of 250. data log review confirmed the report. around 08:30 on 8/nov, there was a sudden mc spike to 251 ma with a ms deviation at p-4. this was followed by a drop in mc and pump flows and some variability in the ps. leading up to the mc spike, the ps waveform was a typical aortic pressure trace, well aligned with the mc waveform. additionally, optical signal metrics were within expected ranges. however, after the mc spike, the ps and mc waveforms came slightly out of phase, and false positioning alarms began to trigger shortly after. there was another subsequent mc spike to 818 ma with a ms deviation 3 hours after the initial spike. this was followed by an increase in mc and pump flows. all the signatures noted in data logs support a biomaterial ingestion event. product was not returned for investigation. based on the signatures noted in data log review, the cause of the false alarm was determined to be biomaterial ingestion. the ingestion event, however, could not be definitively linked to the low starting act four days prior. thrombosis: clinical details report suspicion of thrombus at the time of false positioning alarms. the root cause of the thrombosis was unable to be determined due to insufficient clinical details.
3005 | 1220648-2025-48834 | 23749660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23749660&pc=OZD | 11/07/2025 | 02/04/2026 |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758770 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 93 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 53-year-old male was implanted with impella cp for support of cardiac arrest. following impella placement, patient running at p-9, 4lpm, red colored urine per staff. no labs drawn to detect hemolysis. patient upgraded to 5.5. | [5 empty] | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. correction. the following fields were corrected as part of this follow-up report. section h6. health effect -impact code and medical device problem code.
3006 | 1220648-2025-48835 | 23750358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23750358&pc=OZD | 11/07/2025 | 12/08/2025 |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758770 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 53 year old male was implanted with impella cp for support for cardiac arrest.following impella placement, it was reported that patient had blood in urine. there seems to be a bleed somewhere in the access site. it does not seem to be related to the graft. the cutdown site was reopened in the operating room to find the bleeding.hemoglobin & hematocrit dropped and patient was given1 unit of red blood cells. in addition the patient was successfully upgraded to the impella 5.5. | [5 empty] | The impella device discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This mdr serves as a notification, in which mdr 1220648-2025-48835 was erroneously filed as a duplicate of mdr 1220648-2025-48834. please refer to mdr 1220648-2025-48834 for all reports filed for this device.
3007 | 1220648-2025-48836 | 23750714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23750714&pc=OZD | 11/09/2025 |  |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026769625 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient started to have severe 10/10 chest pain. the patient called emergency medical services and was transported to the hospital. an electrocardiogram revealed anteriolateral st-elevation myocardial infarction. the patient was taken to the cath lab. an angiogram revealed occluded left anterior descending artery at the ostium with heavy thrombus burden. there was left ventricular ejection fraction around 15%. left ventricular end-diastolic pressure is 41 mmhg. decision made for impella cp. the impella cp was placed in left ventricle without issue. while on support, it was reported that the patient had dark urine, with hemolysis. an echocardiogram was done at bedside with initial measurement 5.5 and was moved back to 4.2, pig-tail toward apex. the patient remained on support for two days until explanation and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This follow-up report is being submitted to add an h6 device problem code that was omitted from the initial report. h6 medical device problem code a01 (patient¿device interaction problem) has been added. | E1 added zip code extension as it was omitted from the previously submitted reports. h6 b13 code was added as it was omitted from the previously submitted reports. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. hemolysis: no logs are available. the cause of hemolysis cannot be determined since insufficient clinical details were provided. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details.
3008 | 1220648-2025-48837 | 23750844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23750844&pc=OZD | 11/09/2025 |  |  | 12/08/2025 |  |  | 11/09/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708164 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Ventricular Fibrillation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was having runs of ventricular tachycardia through the night. started the patient on lidocaine. the patient had a slight fever. it was also reported that the patient expired. no further information provided. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: h6 health effect - clinical code 060110 was inadvertently omitted on the initial manufacturer device report . the investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3009 | 1220648-2025-48819 | 23752942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23752942&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583035 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 79 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Bradycardia | Low Blood Pressure/ Hypotension | Cardiac Tamponade | Multiple Organ Dysfunction Syndrome | Renal Impairment | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced atrial fibrillation requiring increasing does of vasopressors and cardioversion. the patient then experienced cardiac tamponade which resolved following pericadiocentesis. the patient continued to deteriorate with pump suction events, bradycardia, hypotension, and multiorgan failure. the patient was treated with extracorporeal membrane oxygenation and hemodialysis. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The pump was not returned for investigation, and no data log was provided or available for review. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical symptoms (cardiac tamponade, hypotension, acute kidney failure): the cause of the injury was not determined due to insufficient clinical details. clinical symptoms (atrial fibrillation, bradycardia): the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: clinical details indicate that the patient experienced persistent suction alarms, some of which were unresponsive to repositioning. without the returned product and sufficient clinical details, the cause of the suction issue could not be determined. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3010 | 1220648-2025-48843 | 23752962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23752962&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693768 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 175 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. heparin still was held per the surgeon as the patient was bleeding and blood products were given. it was also reported that flows were one liter above baseline. purge pressure was 900, up from 600. purge flow was 2.5, down from 6 milliliters. the left ventricle waveform was above the aortic. the patient remained on support until successful explanation and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3011 | 1220648-2025-48844 | 23752999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23752999&pc=OZD | 11/08/2025 |  |  | 12/08/2025 |  |  | 11/08/2025 | 12/08/2025 |  |  |  |  | 12/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763954 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Heart Block | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in non-st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the hospital with chest pain and was taken to the cardiac catheterization laboratory (ccl) for a percutaneous coronary intervention (pci) where two drug-eluting stents were placed in the left anterior descending coronary artery. following the pci, the patient was transferred to the intensive care unit (icu) where the patient¿s cardiac output continued to drop, required increased dosages of pressors and blood lactate was noted to be increasing. the medical team decided to take patient back to the ccl for an impella cp implant. the impella cp was placed with no complications, however after the impella crossed the aortic valve, the patient immediately went into sustained complete heart block. a temporary transvenous pacemaker (tvp) was inserted and completely controlled the patient¿s heart rate and rhythm. the impella was noted to be slightly deep in the ventricle at 6.3 centimeters, however the physician preferred to not reposition the impella due to the significant patient agitation and risk for accidentally pulling the impella too shallow. it was noted that there were suction alarms above performance (p) level 7, but at p7 the pump flowing appropriately and patient well supported on current inotropes, pressors and impella flows. overnight, the patient¿s ph continued to drop and placate continued to rise. the patient required additional pressors after arriving to the icu. it was noted at 02:00, the impella flows decreased to 2.5 liters per minute, and at 02:23 the patient became pulseless. cardiopulmonary resuscitation was initiated, and the patient was placed on continuous renal replacement therapy. the medical team exhausted all interventions, and the code was called at 06:12. it was also noted that the patient had reported a new onset of backpain after arrival to the icu. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. heart block: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
3012 | 1220648-2025-48779 | 23727166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23727166&pc=OZD | 11/06/2025 |  |  | 12/05/2025 |  |  | 11/06/2025 | 12/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390491 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Shipping Damage or Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | When this console was turned on for education today, an error screen appeared on the console that said ¿system self check failed¿ ¿change console immediately¿. no patient was involved. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
3013 | 1220648-2025-48781 | 23727503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23727503&pc=OZD | 11/06/2025 |  |  | 12/05/2025 |  |  | 11/07/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026748366 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, pump saturation was noted overnight. echocardiogram (echo) showed a clot on the tip on catheter. the pump stopped shortly after echo. this is one of two related reports. this report represent the impella 5.5 and another report will be submitted to represent the impella cp. refer to section h10 for the related report number. | [5 empty] | Investigation summary the pump and data logs were not returned for investigation thrombosis: as per clinical, echo showed clot on the tip of catheter on (b)(6) 2025. in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. saturated flow: as per clinical saturated flow was observed overnight on (b)(6) 2025. the cause of saturated pump flow was not determined since no pump/data logs were returned for investigation. pump stop: as per clinical pump stopped next day morning on nov 7th since pump saturation on nov 6, 2025. the cause of pump stop was not determined since no pump/data logs were returned for investigation. device history lot device lot: 1980884 device history batch subcomponent lot: n/a device history review the pump sn (b)(6) passed all post sterile inspection checks. | The impella device is unavailable for return from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
3014 | 1220648-2025-48782 | 23727643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23727643&pc=OZD | 11/07/2025 |  |  | 12/05/2025 |  |  | 11/07/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768676 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that as the impella cp ramped up in auto, flows were only 0.5 liters per minute position on fluoroscopy looked perfect and the patient received an angiomax bolus and drip so the medical doctor thought something was wrong with the pump. the medical doctor explanted this pump and placed a new impella. the patients outcome was noted as survived. | [5 empty] | The investigation was completed. low or blocked pump flow: the pump was not returned for investigation. total pump use was about three minutes. data log shows that during pump ramp-up, a small motor current spike was reported without impacting any other pump parameter. pump flow was low with an active suction alarm at the start of the case. the cause of the low pump flow issue was not determined without returned product investigation and sufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3015 | 1220648-2025-48783 | 23727707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23727707&pc=OZD | 11/07/2025 |  |  | 12/05/2025 |  |  | 11/07/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614178 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 92 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the purge fluid was leaking at the red plug on both ends of the red plug. later, the pump was explanted and the patient survived. | [5 empty] | H6: type of investigation add code b11. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. red handle leak: clinical reported purge was leaking at red pump plug on both ends. the cause of the issue was not established as no product was returned for analysis and limited clinical information was provided. device history lot: device lot number: 1888699. device history review: the pump sn (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3016 | 1220648-2025-48784 | 23728622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23728622&pc=OZD | 11/05/2025 |  |  | 12/05/2025 |  |  | 11/05/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618419 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with impella cp for percutaneous coronary intervention support. the impella cp was weaned and explanted without difficulty. however, three additional perclose sutures were added. protamine was administered, and manual compression was done for 45 minutes for adequate hemostasis. the patient also received hydralazine for some hypertension which helped the oozing. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/major bleed: the cause of the access site adverse event is use related since three additional perclose sutures were needed in order to achieve adequate hemostasis. hypertension: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3017 | 1220648-2025-48785 | 23731320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23731320&pc=OZD | 11/07/2025 |  |  | 12/05/2025 |  |  | 11/07/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754181 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support pre-percutaneous coronary intervention (pci) to the left main and circumflex arteries. during the support, purge pressure was high. the line was flushed and the cassette was changed without success and performance level at p-2. consequently, the impella cp was weaned and explanted with a survived outcome. the patient was reported to be stable following the event. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. high purge pressure: product was returned and biomaterial was observed around impeller and at purge gap. no catheter kinks/purge line kinks were seen. the pump and purge cassette were connected to console and initiated for priming. high purge pressure reproduced but resolved soon after pump was started at p-9 and biomaterial shifted/kicked off from the purge gap. the purge parameters were back in normal range later. data logs were not available for review. the cause of the high purge pressure was due to biomaterial but it is unknown if the event was a buildup or ingestion without data logs. device history review: the pump passed all post sterile inspection check. | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
3018 | 1220648-2025-48786 | 23732133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23732133&pc=OZD | 11/06/2025 |  |  | 12/05/2025 |  |  | 11/06/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739133 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was primed and waiting to be inserted. a placement signal not reliable alarm occurred after priming for 20 minutes. it had not entered the body so the white cable was tried to be pulled out and light check and cable was reinserted. the alarm went off again. the decision was made to not use that catheter as this patient may be on support for a long time waiting a transplant. this is one of three related reports. other reports will be submitted to represent the other impella 5.5 pumps. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
3019 | 1220648-2025-48787 | 23732522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23732522&pc=OZD | 11/06/2025 | 01/21/2026 |  | 12/05/2025 |  |  | 11/06/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742199 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 59 KG | Hispanic | White | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp. the medical doctor reported not to place a-line and deferred to the intensivist to place in the intensive care unit. dopamine was weaned off. there was no inotrope as the patient had arrhythmias with dobutamine and did not tolerate. the plan was to titrate down levophed for a mean arterial pressure goal of 65. had to have the impella repositioned due to a depth of 5.8 centimeters. there were no suction alarms noted. the impella was pulled back from 91 centimeters to 88.5 centimeters on the exterior catheter and 5.4 centimeters within the left ventricle. the pigtail of the impella cp caught on the papillary and was unable to be released. the impella was left as is. the same day new labs were drawn and pending. there was questionable internal bleeding as no external bleeding was noted. heparin drop was on hold. abdominal and right groin site were soft. there was minimal drainage from the right chest tube. hemoglobin dropped from 14 to 7. one unit of packed red blood cells was ordered. the patient expired on impella support. | [5 empty] | Investigation summary: no product was returned. arrhythmia: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. major bleed: questionable internal bleeding as no external bleeding noted. the cause of the bleed was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): impella reposition, due to depth 5.8cm, no suction alarms noted. impella pulled back from 91cm to 88.5cm on exterior catheter and 5.4cm within lv. pigtail of cp caught on papillary and unable to release. impella left as is. the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot. device lot: 1970109. device history batch. subcomponent lot: n/a. device history review. device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3020 | 1220648-2025-48788 | 23732689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23732689&pc=OZD | 11/06/2025 |  |  | 12/05/2025 |  |  | 11/06/2025 | 12/05/2025 |  |  |  |  | 12/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740383 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 62 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced diminished flow down the right femoral artery. an antegrade sheath was placed for distal flow down the right leg. the patient has positive pedal pulses via doppler. the next day, one unit of blood was infusing. impella flows were good and there were no alarms overnight. heparin was on hold as last partial thromboplastin time was 110. the patient was successfully weaned. | [5 empty] | No product was returned for investigation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3021 | 1220648-2025-48763 | 23714523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23714523&pc=OZD | 11/04/2025 |  |  | 12/04/2025 |  |  | 11/04/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757708 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp. when using the access port, it was noted that there was no flow distal to the repositioning sheath and no pulses were noted in the right lower extremity. the decision was made to explant the impella related to this decreased flow. the impella cp was weaned to p2 and the patient was not on any pressors/inotropes. the patient survived the procedure. | [5 empty] | The cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3022 | 1220648-2025-48765 | 23714806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23714806&pc=OZD | 11/06/2025 |  |  | 12/04/2025 |  |  | 11/06/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026749036 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. multiple unsuccessful attempts were made by the surgeon to pass the wire into the aortic arch. an aortogram was performed with visible obstruction at the level of the innominate artery, before the aortic arch. anesthesia notified the doctor of a possible thrombus versus dissection. previous computed tomography images were reviewed with no abnormal anatomical findings. multidisciplinary decision was made to abort the case. the introducer was removed and the axillary graft was ligated at the level of the skin. additional information was received. the implanting care team reported after additional diagnostic testing an axillary artery dissection with thrombus to the innominate artery. the implanting physician sent for a heparin-induced thrombocytopenia and thrombosis panel as a possibility for thrombus formation. the patient was on prior therapy of venoarterial extracorporeal membrane oxygenation and anticoagulant medication. this is one of two related reports. this report represents the impella 5.5. another report was submitted to address the impella cp. refer to section h10 for the related report number. | [5 empty] | No product was returned for investigation. the cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the vascular dissection was unable to be determined due to insufficient clinical details. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy at innominate artery, before the aortic arch preventing successful delivery of impella. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3023 | 1220648-2025-48766 | 23714845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23714845&pc=OZD | 09/08/2025 |  |  | 12/04/2025 |  |  | 09/13/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654471 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 173 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with bipella (impella cp and impella 5.5) for mechanical circulatory support (mcs). the patient experienced complications and subsequently expired. the patient was admitted with 4 out of 10 back pain waxing and waning in severity. an electrocardiogram demonstrated inferior st segment elevation myocardial infarction with hypotension and cardiogenic shock. the patient was emergently transported to the catheterization lab and animpella cp was placed for mcs. approximately two days later, an attempt was made to escalate support to an impella 5.5; however, insertion was aborted due to an axillary artery dissection. the patient was subsequently placed on veno-arterial extracorporeal membrane oxygenation (va ecmo), and the impella cp remained. during support, it was reported that the impella cp console (automated impella controller-aic) displayed a suction alarm and placement signal low alarm. however, there was no audible alarm sound. the nurse was advised to switch to an alternate aic, which resolved the issue. the following day, the patient returned to the operating room for venous drainage, and continuous renal replacement (crrt) therapy was initiated due to renal insufficiency. additional note indicated sputum positive for candida. two days later, the patient underwent surgical repair of the right axillary artery and successful insertion of an impella 5.5. ecmo, impella cp, and impella 5.5 were concurrently supporting the patient. approximately six days later, ecmo was converted from va to veno-venous (vv) configuration. both impella devices remained in place. the patient continued support and subsequently expired after a total of approximately 15 days of impella mcs. care was withdrawn. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the event was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. medical safety review. medical safety reviewed the event and noted the following: on (b)(6) 2025, an impella cp was placed in a patient presenting with cardiogenic shock for left ventricular (lv) unloading. on (b)(6) 2025, an attempt was made to escalate support to an impella 5.5; however, insertion was aborted due to an axillary artery dissection. the patient was subsequently placed on va ecmo, and the impella cp remained in place for lv unloading. during support, the bedside rn reported that the impella cp console displayed a suction alarm and placement signal low alarm, but there was no audible alarm sound. the rn was advised to switch to an alternate aic, which resolved the issue. on 09/01/2025, the patient returned to the or for venous drainage, and crrt was initiated due to renal insufficiency. on (b)(6) 2025, the patient underwent surgical repair of the right axillary artery and successful insertion of an impella 5.5. the physician elected to maintain ecmo, impella cp, and impella 5.5 concurrently for enhanced lv unloading. approximately six days later, ecmo was converted from va to vv configuration, while both impella devices remained in place. the patient continued on support until (b)(6) 2025, when care was withdrawn. device involvement and assessment: ¿ impella aic: device malfunction occurred due to lack of audible alarms on the initial aic, despite visual alarm display. this is considered a reportable device malfunction, though it is unlikely to have contributed to the patient¿s death. impella cp: renal insufficiency noted during support is likely related to patient acuity rather than device performance. ¿ first impella 5.5: initial insertion attempt resulted in axillary artery dissection requiring surgical repair. this constitutes reportable patient harm. ¿ death: patient was supported by impella cp and impella 5.5 at the time of death. the death will be conservatively reported for both devices, though it is most likely attributable to the patient¿s critical condition rather than device malfunction. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 (bipella) device report is one of four devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp (bipella) and automated impella controller. note: the medical device report for the first impella 5.5 is pending follow up and will be subsequently submitted.
3024 | 1220648-2025-48764 | 23714861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23714861&pc=OZD | 08/29/2025 | 01/23/2026 |  | 12/04/2025 |  |  | 08/31/2025 | 12/04/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720166 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Renal Failure | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received stating that the patient had low or blocked flow with frequent intermittent suction alarms. | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with bipella (impella cp and impella 5.5) for mechanical circulatory support (mcs). the patient experienced complications and subsequently expired. the patient was admitted with 4 out of 10 back pain waxing and waning in severity. an electrocardiogram demonstrated inferior st segment elevation myocardial infarction with hypotension and cardiogenic shock. the patient was emergently transported to the catheterization lab and animpella cp was placed for mcs. approximately two days later, an attempt was made to escalate support to an impella 5.5; however, insertion was aborted due to an axillary artery dissection. the patient was subsequently placed on veno-arterial extracorporeal membrane oxygenation (va ecmo), and the impella cp remained. during support, it was reported that the impella cp console (automated impella controller-aic) displayed a suction alarm and placement signal low alarm. however, there was no audible alarm sound. the nurse was advised to switch to an alternate aic, which resolved the issue. the following day, the patient returned to the operating room for venous drainage, and continuous renal replacement (crrt) therapy was initiated due to renal insufficiency. additional note indicated sputum positive for candida. two days later, the patient underwent surgical repair of the right axillary artery and successful insertion of an impella 5.5. ecmo, impella cp, and impella 5.5 were concurrently supporting the patient. approximately six days later, ecmo was converted from va to veno-venous (vv) configuration. both impella devices remained in place. the patient continued support and subsequently expired after a total of approximately 15 days of impella mcs. care was withdrawn. |  |  |  |  | Additional information was received and h6 codes updated. | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. medical safety review. medical safety reviewed the event and noted the following: on (b)(6) 2025, an impella cp was placed in a patient presenting with cardiogenic shock for left ventricular (lv) unloading. on (b)(6) 2025, an attempt was made to escalate support to an impella 5.5; however, insertion was aborted due to an axillary artery dissection. the patient was subsequently placed on va ecmo, and the impella cp remained in place for lv unloading. during support, the bedside rn reported that the impella cp console displayed a suction alarm and placement signal low alarm, but there was no audible alarm sound. the rn was advised to switch to an alternate aic, which resolved the issue. on (b)(6) 2025, the patient returned to the operating room for venous drainage, and crrt was initiated due to renal insufficiency. on (b)(6) 2025, the patient underwent surgical repair of the right axillary artery and successful insertion of an impella 5.5. the physician elected to maintain ecmo, impella cp, and impella 5.5 concurrently for enhanced lv unloading. approximately six days later, ecmo was converted from va to vv configuration, while both impella devices remained in place. the patient continued on support until (b)(6) 2025, when care was withdrawn. device involvement and assessment: impella aic: device malfunction occurred due to lack of audible alarms on the initial aic, despite visual alarm display. this is considered a reportable device malfunction, though it is unlikely to have contributed to the patient¿s death. impella cp: renal insufficiency noted during support is likely related to patient acuity rather than device performance. first impella 5.5: initial insertion attempt resulted in axillary artery dissection requiring surgical repair. this constitutes reportable patient harm. death: patient was supported by impella cp and impella 5.5 at the time of death. the death will be conservatively reported for both devices, though it is most likely attributable to the patient¿s critical condition rather than device malfunction. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp (bipella) device report is one of four devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device reports on the impella 5.5 (bipella) and automated impella controller. note: the medical device report for the first impella 5.5 is pending follow up and will be subsequently submitted. | Investigation summary: no product was returned. data logs were returned for analysis. optical sensor issue: clinical details mention that there were frequent intermittent placement signal low and suction alarms at p4 during the case. placement signal low alarms were confirmed by data log analysis. it was also seen that the placement signal was trending up around noon on (b)(6) 2025, which coincided with suction alarms, and placement signal not reliable alarms. snr was low towards the beginning of the case but stabilized after the placement signal low alarms resolved. other 5s parameters showed no abnormal trends. it was also noted from the clinical details that the patient was on va ecmo and crrt during support. the root cause of the placement signal alarms was not determined due to inconclusive evidence from the clinical details and data log analysis, and unreturned product for further evaluation. low or blocked pump flow: clinical details mention suction alarms occurring throughout the case. suction alarms were confirmed to be intermittent throughout the case, with low pump flows occurring during the alarms. no purge pressure spikes were seen outside of purge bag changes. no motor current spikes were seen outside of p-level changes. impella position in aorta and impella position unknown alarms were also seen during suction alarms. suction alarms coincided with placement signal low alarms, with trending placement signal observed. it was noted from the clinical details that the patient decompensated from respiratory status overnight on (b)(6) 2025, and had significantly diseased vasculature which prevented 5.5 escalation, in addition to being on va ecmo and crrt support. the root cause of the suction alarms was not determined due to inconclusive evidence from the clinical details and data log analysis, and unreturned product for further evaluation. renal failure: clinical details mention that the patient was on ecmo and crrt support during impella support, and crrt was net negative towards the end of the case. the root cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1958540. device history batch: subcomponent lot: n/a. device history review: pump#: (b)(6) passed all post-sterile inspection checks.
3025 | 1220648-2025-48767 | 23715038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23715038&pc=OZD | 11/06/2025 |  |  | 12/04/2025 |  |  | 11/06/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711266 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male | 85 KG | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Cardiac Arrest | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hematuria | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. while on support, limb ischemia was reported requiring external bypass. no blood volume needed. later, it was reported that there was hematuria, now clearing. additionally some hemolysis formed overnight. the p level was reduced and position was assessed. hemolysis seems to be resolving. it was also noted that the patient coded. furthermore, there was bleeding from the extracorporeal membrane oxygenation access site. heparin was discontinued and manual pressure was held and the complication was successfully managed. the pump was later explanted and the patient survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. blood/hematuria/hemolysis: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia/cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3026 | 1220648-2025-48768 | 23715457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23715457&pc=OZD | 11/05/2025 | 01/12/2026 |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743332 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 was suspected of having a clot. the pump was explanted and the patient is in stable condition. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Investigation summary: no product was returned for investigation. thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1976098. device history batch: subcomponent lot: n/a. device history review. device with sn: (b)(6) passed all post sterile inspection checks.
3027 | 1220648-2025-48769 | 23715690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23715690&pc=OZD | 11/05/2025 |  |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586062 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 55 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced going in and out of runs of ventricular tachycardia. when this happens the patient becomes laminer. flows are stable. however, pressure drops. the position was evaluated and it was fine in the mid ventricle off the septum. however, the patient was very sensitive to stimulation. the decision was made by the medical doctor to put on lidocaine., consult and paralyze to control the ectopy. the patient was very stable when leaving the intensive care unit. the patient had increased computed tomography output overnight. blood products were being given. the next day, the patient had frequent ventricular tachycardia storms on amiodarone and pacing wires. p9 causes suction but the position was stable and echocardiogram showed ejection fraction of around 20/25 percent. the nurse reported two days later that the patient was hemodynamically stable on dobutamine and milrinone. the team would be pulling the continuous renal replacement therapy (crrt) today. mean arterial pressures were in the 80's. the patient remained paralyzed. two days after, no systemic heparin was noted. the patient was subcutaneous for two days. they may start as chest tube output has decreased significantly. an echocardiogram was done at bedside and the pump was turned down to p4. ejection fraction was not improving and was still at out 20 percent. support was resumed at p7. they may trach and peg and then wean the impella. only a drip change has been done to decrease the milrinone dosage. the next day the patient was headed to the operating room for a trach and peg. the patient received trach and peg. the next day the patient was weaned off paralytics onto light sedation to assess neurology. there were no pressors and the patient remained on crrt. dressing was changed on the axillary site and the swan was swapped out. the next day impella flows were changed from p7 to p6. the plan was to continue to wean depending on the patient's tolerance. support was continued. the patient complained of gastrointestinal issues, possible ileus. no issues with the impella. care was withdrawn and the patient expired. additional information was received. the pump has been thrown away. the patient's outcome was attributed to a sick heart and open heart surgery. | [5 empty] | The root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
3028 | 1220648-2025-48770 | 23715820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23715820&pc=OZD | 11/05/2025 |  |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733793 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a new placement signal not reliable alarm. the audio was disabled. | [5 empty] | The impella device was not saved by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
3029 | 1220648-2025-48771 | 23715914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23715914&pc=OZD | 11/05/2025 |  |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026753179 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to patient anatomy. an impella cp was implanted instead. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had multi vessel disease and moderate as preventing successful delivery of impella. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3030 | 1220648-2025-48772 | 23716271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23716271&pc=OZD | 11/04/2025 |  |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624497 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient had continuous suction while on the impella. the next day, the patient was transferred for advanced care. groin dressing was changed at bedside. the patient started to produce urine output and was pink tinged. hemolysis was reported. a formal echocardiogram was pending. the echocardiogram showed the impella directed towards the anterior mitral leaflets and the pump flipped. the medical team at bedside attempted to reposition. repositioning was unsuccessful due to the pigtail caught in the structure. p level was reduced to p5 and there were no alarms. flows were 2.5. the multidisciplinary team discussed plan of care overnight. the plan was to monitor overnight and adjust inotropes for surgery reassessment for device placement. extracorporeal membrane oxygenation will be initiated if requiring additional support. the patient was being scheduled for impella cp removal and placement in the operating room later. dialysis and support was continued. the patient's labs have been hemolyzed, leading to abnormal values. | [5 empty] | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary: the pump was not returned for investigation. data log review was not provided and could not be retrieved from the remote server. pump suction / device in wrong position / mechanical interaction with blood (hemolysis): the cause of issues was not determined without returned product information and sufficient clinical details. device history review: the pump (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3031 | 1220648-2025-48773 | 23716710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23716710&pc=OZD | 11/05/2025 |  |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725964 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Air/Gas in Device | [6 empty] | Air Embolism | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. upon starting the impella, air was noted in the aorta via transesophageal echocardiogram. they are unsure what the source was or how it happened, since the pump was properly de-aired. air finally stopped with clamping of the aortic root vent. it was also noted that the patient required multiples shocks for ventricular tachycardia due to air emboli. svg was de-aired with needle. aortic root vent was clamped with final resolution of air. the medical team believed to possibly be from the aortic root vent, but not certain. the patient was stabilized and the impella was able to flow at p6 without issue or additional air. the medical team proceeded to finish the case with moderate diffuse bleeding that took some time to dry up after coagulation labs normalized. the chest was closed, final position on transesophageal echocardiogram showed good impella position with adequate waveforms. post operative bleeding from the chest tubes was noted. graft anastomoses were not the source. coagulation was corrected and four unit of red blood cells were given. the patient was hypotensive at times but was able to have quick rebound in pulse pressure with constant volume resuscitation. point-of-care ultrasound shows adequate right ventricle and left ventricle function and proper impella position. the patient was sent back to the operating room for surgical interventions. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation ventricular fibrillation: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined due to insufficient clinical details. air embolism/air in device: upon starting the impella, air was noted in the aorta via transesophageal echocardiography (tee). air finally stopped with clamping of the aortic root vent. the data logs do not show any air in purge alarms. the imc logs were reviewed and there were no air in purge alarms. de-airing and priming steps were completed successfully. the cause of the air in device issue was not determined as no product was returned for analysis and no log abnormalities were observed. device history review the pump passed all post sterile inspection checks
3032 | 1220648-2025-48774 | 23716933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23716933&pc=OZD | 11/05/2025 |  |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768773 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient experienced complete heart block and rapidly decor sated. the patient was brought back to catheterization lab to properly place transvenous pacing, stabilize and return to the unit. the goal will be to wean drops as tolerated and restabilize. two days later, the patient went into atrial fibrillation with rapid ventricular response requiring cardioversion x 1, amiodarone bolus and continued drops with increasing pressor requirements. epinephrine was weaning to off and holding off on impella wean. the patient's outcome was stable. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. heart block/ arrhythmia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: device with passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3033 | 1220648-2025-48775 | 23717032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23717032&pc=OZD | 11/04/2025 | 01/07/2026 |  | 12/04/2025 |  |  | 11/04/2025 | 12/04/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766525 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | Renal Impairment | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and hematuria. the pump was repositioned to resolve the issue. the patient was upgrade to an impella 5.5. the patient is in stable condition. | New information was received. the third adverse event site was confirmed and the event was not related to the impella study device. |  |  |  |  | Additional information was received from the reporter. the patient experienced acute kidney injury and expired. b2, h1, and h6 have been updated. | Investigation summary: product & data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria/acute kidney failure: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot #: 1982185. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3034 | 1220648-2025-48776 | 23717617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23717617&pc=OZD | 11/07/2025 |  |  | 12/04/2025 |  |  | 11/07/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695519 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 89 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. while on support bleeding/hematoma at access site was reported. manual pressure was held and lidocaine with epinephrine was injected around the site. perclose suture was tightened slightly and the site was redressed. after four days of support, the patient was made comfort measures only and expired. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned for investigation. hematoma/major bleed: bleeding/hematoma at access site after transfer to intensive care unit (icu) was noted. hemostasis was achieved through manual pressure. the cause of the bleeding/hematoma are determined to be use issue due to patient movement because the bleeding was caused due to transfer to icu. device history review; the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3035 | 1220648-2025-48777 | 23718093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23718093&pc=OZD | 11/05/2025 | 12/04/2025 |  | 12/04/2025 |  |  | 11/05/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649375 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 72 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | Multiple Organ Dysfunction Syndrome | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was implanted with impella cp for support for cardiac arrest. an impella cp was placed utilizing best practices. the pump was started in auto mode and immediately improved mean arterial pressure. right heart catheterization was performed revealing a severely dilated right heart. discussion for escalation to an impella 5.5 did not take place before the patient coded while on support. the patient was attempted to be weaned off epinephrine drip leading to ventricular fibrillation arrest. advanced cardiac life support was performed for ten minutes without any return of rhythm. the patient expired. additional information was received. the pump was not saved. they did not allege that the pump was the cause of the outcome. the family was unsure that they wanted to continue with care and were not wanting escalated care for the patient at the time of the code due to concerns with the patient's wishes and quality of life. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. ventricular fibrillation/ ischemia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. organ failure: the cause of the injury was not determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3036 | 1220648-2025-48778 | 23719611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23719611&pc=OZD | 11/04/2025 | 12/04/2025 |  | 12/04/2025 |  |  | 11/04/2025 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759816 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was implanted with a impella cp for support for cardiac arrest. an impella cp was placed via radiofrequency ablation in usual fashion without issue. purge system blocked alarm unresolved after initial troubleshooting. tissue plasminogen activator purge protocol was used and successfully resolved the alarm. high purge pressure was noted. the impella was successfully weaned and explanted. one unit of packed red blood cells was being transfused. the physician held pressure for 25 minutes and a femstop was placed for an additional forty minutes. distal pulses were present with doppler. additional information was received. the patient received blood but their was no bleeding complications in this case that were impella related. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3037 | 1220648-2025-31937 | 23720631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23720631&pc=OZD | 02/19/2024 | 12/04/2025 |  | 12/04/2025 |  |  | 02/19/2024 | 12/04/2025 |  |  |  |  | 12/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1576116 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella support, the automated impella controller (aic) showed the impella was located in the aorta even though the impella's position was checked and confirmed to be correct at all times. the same scenario was observed with a different cp with the same console. there was no reported patient harm. | [5 empty] | Data analysis: the console ic8898 was returned to fs on march 13, 2024. the last case performed before the return ran from (b)(6) 2024, with pump sn (b)(6). after 22 hours and 9 minutes from pump start, the console displayed ¿impella position in aorta (for up to 14s) - alarm,¿ event #128. this alarm was later resolved but recurred throughout the case, which confirms the reported failure mode. this confirms the reported failure mode. note: the disposables team confirms that the reported complaint pertains to disposables. device analysis: this console was tested after arriving at fs. ic8898 underwent a successful pm procedure per the fsr. root cause: not applicable, the reported failure mode was confirmed to be attributable to disposables.
3038 | 1220648-2025-48751 | 23703520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23703520&pc=OZD | 11/03/2025 |  |  | 12/03/2025 |  |  | 11/04/2025 | 12/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731186 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy in stage c cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (msc). the patient was escalated to an impella 5.5 after approximately 22 hours after start of impella cp support as the impella cp was unable to be repositioned and hemolysis developed. additionally, the patient needed more long-term support and bridge to decision. the impella 5.5 device was successfully placed, and impella mcs continued. the patient remained stable. on day six of support, the patient decompensated over the night and became hypotensive and mixed venous oxygen saturation slightly dropped. epinephrine was added and dialysis was placed on hold. later that night, the patient had runs of ventricular tachycardia. subsequently, the decision was made to go to comfort care. the following day, with family at bedside, all life sustaining measures were discontinued and the patient expired. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the ventricular tachycardia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Medical safety/clinical reviewed the event. a 59-year-old male with a medical history significant for poorly controlled diabetes mellitus (per documentation, patient reported not routinely monitoring blood glucose levels) presented with chest pain. electrocardiogram (ecg) performed by emergency medical services (ems) demonstrated st-segment elevation in the field, and a st-segment elevation myocardial infarction (stemi) alert was activated. coronary angiography revealed an acute total occlusion of the mid-right coronary artery. one drug-eluting stent was successfully deployed. additional imaging of the left coronary demonstrated diffuse multivessel coronary artery disease. the patient became hypotensive and required endotracheal intubation. due to ongoing cardiogenic shock, an impella cp was implanted for mechanical circulatory support. the society for cardiovascular angiography and interventions (scai) cardiogenic shock classification was stage c at the time of impella cp implantation. percutaneous coronary intervention (pci) was performed to the proximal left anterior descending (lad) artery, with documentation of chronic total occlusion of the remaining lad, as well as disease involving the proximal ramus and proximal left circumflex (lcx) arteries. following the procedure, the patient was transferred to the unit on impella cp support. on post-implant day one, the physician documented that the impella catheter appeared to be positioned behind the papillary muscle; however, no device alarms were present, flows were adequate, and the catheter position (91 cm at the hub and 3.1 cm from the aortic valve) was left unchanged. approximately two hours later, brief ¿impella in aorta¿ alarms occurred following a coughing episode. the alarms resolved spontaneously, and device waveforms normalized. nursing staff indicated the coughing episode was believed to have precipitated the alarms. later that day, the patient developed laboratory evidence of hemolysis, with plasma free hemoglobin (pfhgb) of 150 mg/dl. repositioning of the impella cp was unsuccessful. due to hemolysis and the need for longer-term mechanical circulatory support as a bridge-to-decision strategy, the patient was escalated to an impella 5.5. approximately seven days later, while on impella 5.5 support, the patient experienced runs of ventricular tachycardia. bedside echocardiography measured the impella 5.5 position at 4.5 cm across the aortic valve. the patient¿s scai cardiogenic shock classification had progressed to stage e. on that same day, the patient elected to transition to comfort-focused care. with family present, life-sustaining therapies were withdrawn, and the patient expired. hemolysis when device is malpositioned is a known potential complications of impella cp support and are being reported as serious injury events attributed to the impella cp. the death is being associated with the impella 5.5 as the device was in place at the time of death and the patient experienced ventricular tachycardia while on support. the patient¿s underlying condition was the most likely contributing factor to the demise outcome (poorly controlled diabetes mellitus, diffuse chronic coronary artery disease with chronic total occlusions, and progression from scai stage c to stage e shock). note: h6. component code and investigation type/findings/conclusion coding remain unchanged as previously reported. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | Date of death: the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. device status: the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
3039 | 1220648-2025-48752 | 23703613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23703613&pc=OZD | 11/06/2025 |  |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713661 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop. the impella cp was explanted. no patient harm was reported. | [5 empty] | Correction: the explant date was inadvertently left out in the initial report which has been added to this report. investigation summary: impella cp sn (b)(6) returned. pump stop - after 14 days on support, pump stop occurred. no clot found at explant. data log review showed data logs were not recovered that showed pump stopping. lt logs go through 01nov25 and rt logs recovered end at 1219 on 06nov25. the returned pump was found to have a large purge gap, indicative of bearing wear. additionally, biomaterial was found in the excess gap. pump attempted to be started and pump stop occurred immediately, even after biomaterial was removed. the cause of the pump stop was bearing wear due to the large impellor purge gap seen on the returned pump. as the pump stop occurred on day 14, and per design trace matrix 0046-9910 the cp c9 pump has an intended design life of 8 days, this is an end-of-life problem. device history review: pump s/n (b)(6) passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3040 | 1220648-2025-48753 | 23704032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23704032&pc=OZD | 11/06/2025 |  |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026738980 |  | 000000000001000100 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was to be escalated from an impella cp to an impella 5.5 but the impella 5.5 delivery was aborted due to the patient's anatomy. the patient had small vessels and the pump could not be advanced past anastomosis. this is one of two related reports. this report addresses the impella 5.5. another report will be submitted to address the impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and small vessel size preventing successful delivery of impella. the device passed all post sterile inspection checks.
3041 | 1220648-2025-48754 | 23704122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23704122&pc=OZD | 11/06/2025 |  |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759358 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female | 80 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. laboratory results were pending. the impella cp functioned without problems and without any suction events. imaging was used to check the pump position and the pump was in good position. care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3042 | 1220648-2025-48755 | 23704770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23704770&pc=OZD | 11/01/2025 | 01/07/2026 |  | 12/03/2025 |  |  | 11/01/2025 | 12/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Arrhythmia | Fever | Hematoma | Hemoptysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Thrombocytopenia | [29 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | The patient experienced hemoptysis and fever overnight. the patient was taken to ct to look for potential pancreatitis or source of bleeding. a small right non-occlusive pulmonary embolism (pe) was seen with pulmonary hemorrhage vs. infarction. the patient was also being worked up for potential heparin-induced thrombocytopenia. there were no pf4 results yet. the impella device was explanted from the patient. | It was reported that a patient implanted with an impella 5.5 became hypotensive and nonpulsatile, and experienced persistent ventricular tachycardia. the patient was defibrillated. three days later, the patient developed a hematoma at the access site. it started to occur overnight since the patient had gotten up and moved for the first time. the medical team was made aware. the patient is in stable condition. | Additional information was received stating that the patient was explanted. |  |  |  | Additional information received b5. added additional codes based on updated information. h6. health effect - clinical code. correction to b5: the impella device was explanted from the patient. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. asae/hematoma: clinical states hematoma present at the incision site and rn stated it started to occur overnight since the patient had gotten up and moved for the first time. the cause of the access site adverse event is use related due to patient movement because the bleeding started to occur overnight after the patient had gotten up and moved for the first time. ventricular tachycardia/arrhythmia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1976112. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | B5 additional information was received and d6 updated.
3043 | 1220648-2025-48757 | 23705362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23705362&pc=OZD | 11/04/2025 |  |  | 12/03/2025 |  |  | 11/04/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768376 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. vascular surgery was consulted for mono-phasic signals. the patient was taken to the hybrid operating room where vascular surgery performed a thrombectomy on the impella leg. the care team was stabilizing that patient. another round of plex was planned. the nurse reported that the impella cp had lower than normal flows. the car team had ruled out positioning and filling pressures. the medical doctor stated that poor flows may be due to severe bi-ventricular failure. the impella cp was explanted and closure was performed. the leg ended up amputated. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia / thrombosis: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3044 | 1220648-2025-48756 | 23705388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23705388&pc=OZD | 10/27/2025 | 01/05/2026 |  | 12/03/2025 |  |  | 11/04/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692946 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 61 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 received one unit of packed red blood cells for hemoglobin at 7.1 labs were pending. nitroglycerin was started for after load reductions. a few days later, the patient experienced premature ventricular contractions/self-limiting ventricular tachycardia after increased to p7 yesterday. transthoracic echocardiogram showed impella close to septum and six centimeters across atrioventricular. a discussion with surgery regarding pulling back the impella. the patient had ventricular fibrillation in the operating room prior to the impella. additionally, the patient suffered gastrointestinal bleeding. anticoagulants were withheld. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary. no product was returned. ventricular tachycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the gi bleed was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot device lot: 1932237 device history batch subcomponent lot: n/a device history review device (sn: 603364) passed all post sterile inspection checks.
3045 | 1220648-2025-48758 | 23705573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23705573&pc=OZD | 11/02/2025 |  |  | 12/03/2025 |  |  | 11/04/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739115 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 80 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient has a cerebrovascular accident. a computed tomography confirmed left main coronary artery infarct and right frontal hemorrhage. the patient's outcome is stable. additional information was received. the deficits did improve and there were no interventions. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. stroke: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the stroke was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3046 | 1220648-2025-48759 | 23705801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23705801&pc=OZD | 11/05/2025 | 01/07/2026 |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727350 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that there was an issue during the loading of the impella cp on the 0.018". the 0.018" broke in the impella. the physician had not gone though the aortic valve before breaking the guidewire. the pigtail was still in the aorta so it was not possible to push the pump in the left ventricle. the pump was removed and the patient survived. this is one of two related reports. this report represents the pump and another report will be submitted to represent the guidewire. refer to section h10 for the related report number. | Additional information received that the most likely explanation is that a nurse inadvertently activated the impella while the physician was advancing it into the descending aorta. this would explain why the 0.018" guidewire became entrapped within the impella. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5 updated with additional information. | Investigation summary: no product was returned for evaluation. unable to deliver: clinical details noted that when pump was being inserted over the guidewire, the guidewire broke and the pump was unable to be inserted across the aortic valve and into the lv. the cause of the delivery issue could not be determined as limited clinical details were provided and no product was returned for evaluation. device history lot. device lot: 1970104. device history batch. subcomponent lot: n/a. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
3047 | 1220648-2025-48760 | 23705915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23705915&pc=DQX | 11/05/2025 | 01/07/2026 |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during the loading of the impella cp on the 0.018", the 0.018" broke in the impella. the physician had not gone though the aortic valve before breaking the guidewire. the pigtail was still in the aorta so it was not possible to push the pump in the left ventricle. the pump and the guidewire were removed. this is one of two related reports. this report represents the guidewire and another report was submitted to represent the pump. refer to section h10 for the related report number. | Additional information received that the most likely explanation is that a nurse inadvertently activated the impella while the physician was advancing it into the descending aorta. this would explain why the 0.018" guidewire became entrapped within the impella. |  |  |  |  | The guidewire was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5 updated with additional information | Investigation summary: no product returned for evaluation. product damage (guidewire fracture): clinical details noted that the guidewire fractured when the pump was being loaded over the guidewire. no additional details were provided. the cause of the product damage (guidewire fracture) could not be determined as no product was returned and limited clinical details were provided. device history lot: device lot: 8854085. device history batch: subcomponent lot: n/a. device history review: 0.018" guidewire lot #8854085 passed all inspection checks.
3048 | 1220648-2025-48761 | 23706054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23706054&pc=OZD | 11/06/2025 |  |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622897 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 and then was escalated to an impella cp for mechanical circulatory support. while on the impella cp, there was concern for right leg ischemia. the device was replaced with another device and the patients outcome was noted as survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the ischemia was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3049 | 1220648-2025-48762 | 23706523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23706523&pc=OZD | 11/06/2025 |  |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 |  |  |  |  | 12/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026756774 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient had to go back down to surgery for bleeding and washout. later, the device was explanted and the patient survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the bleed could not be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3050 | 1220648-2025-48749 | 23712428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23712428&pc=OZD | 11/06/2025 | 01/22/2026 |  | 12/03/2025 |  |  | 11/06/2025 | 12/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744197 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 in a 69 year old female for chest pain support. patient had 3 stents placed to the right coronary artery. during left ventricle gram, acute ventricular septal defect was observed. the decision was made to place an impella 5.5 through the right axillary artery, and also contact outside facility to discuss further therapy/options. due to the subsequent decline, the decision was made to place the patient on peripheral venoarterial extracorporeal membrane oxygenation. the patient was then flown to outside facility. | [5 empty] | Investigation summary: no product was returned for investigation. hematoma: hematoma around site of impella 5.5. noticed while moving pt to flight stretcher. the cause of hematoma is determined to be use issue because hematoma was caused while moving patient to stretcher. device history lot: device lot: 1976109. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3051 | 1220648-2025-48718 | 23691668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23691668&pc=OZD | 11/02/2025 |  |  | 12/02/2025 |  |  | 11/02/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734287 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient had been resting on impella support and was listed status two for orthotopic heart transplantation. the nurse reported false position in aorta alarms. the medical doctor confirmed via echocardiogram. due to elevated plasma d=free hemoglobin and unable to reposition the device after multiple attempts, the decision was made to replace the impella 5.5. the patient was stable. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. no issues were found on the returned product. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3052 | 1220648-2025-48719 | 23692150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692150&pc=OZD | 11/04/2025 |  |  | 12/02/2025 |  |  | 11/05/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665799 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced suction alarms. the patient was treated with pressors and infusion of albumin. a hematoma at the left groin was identified when the pump was explanted. the patient was in stable condition. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs revealed multiple suction and impella position in aorta alarms triggered corresponded with the time the clinical team reported. both alarms then were resolved within 6 hours. no other issues were found on the logs. pump suction: the cause of pump suction issue was unable to be determined as limited clinical details were provided and no product was returned for review. hematoma: the cause of the hematoma can not be determined as insufficient clinical details were provided and product was not returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3053 | 1220648-2025-48720 | 23692236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692236&pc=OZD | 11/04/2025 | 12/01/2025 |  | 12/02/2025 |  |  | 11/04/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760524 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 70-year-old patient for chest pain support. high purge pressures after device started physician elected to remove and place new device. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. high purge pressure/purge system blocked: the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3054 | 1220648-2025-48721 | 23692341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692341&pc=OZD | 11/04/2025 |  |  | 12/02/2025 |  |  | 11/04/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1346474 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that there was a automated impella controller failure red alarm upon start up. there was no patient harm. | [5 empty] | The investigation was completed. data review: the console logs indicated a controller failure that issued alarm #418 and identifying the purge pressure sensor voltage as to be out of specified operating range. this confirming the reported "controller failure red alarm start up" error. device analysis: upon bootup, the console displayed a controller failure with a red alarm message. a functional test of the purge pressure sensor signal output to be at 0.404v where as the minimum should be operating above 0.410v. this confirming that the sensor signal output was not within the normal operating range when compared to a known good sensor at 0.498v under the same testing conditions. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3055 | 1220648-2025-48724 | 23692519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692519&pc=OZD | 11/02/2025 |  |  | 12/02/2025 |  |  | 11/02/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696997 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient went to the operating room this morning for a washout. the doctor pulled back two centimeters then readvanced two centimeters and after that the purge system started alarming high purge pressure and purge system blocked. the position looks good but a clot was found in the left atrium. the right atrium and right ventricle were distended. a washout was performed and the patient sent to the cardiovascular intensive care unit. the medical doctor did not want to start tissue plasminogen activator so they ordered heparin in the purge system and started systemic heparin. the nurse was walked through changing the purge cassette and solution but no change with the purge alarms. the medical doctors all were aware that the pump could stop. they were advised to keep an eye on the motor current and watch for a spike. the patient may be taken back to the operating room tomorrow for another wash out and possible pump swap out. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. purge pressure high/purge system blocked: data logs were reviewed and the log at time of "high purge pressure" alarms showed a spike and step up in motor current with small spike in motor speed and rise and full saturation of pump flow for a few minutes after motor current spike. the cause of the high purge pressure/purge system blocked was due to biomaterial ingestion based on data log analysis and clinical details. device history review: the device passed all post sterile inspection checks | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
3056 | 1220648-2025-48723 | 23692698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692698&pc=OZD | 11/04/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/04/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741841 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. plan to start amio for tachycardia. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3057 | 1220648-2025-48725 | 23692794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692794&pc=OZD | 11/02/2025 |  |  | 12/02/2025 |  |  | 11/02/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698681 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the impella was placed and the reposheath was introduced and sutured. no pulse was able to be doppled. the decision was made to do a bilateral runoff thru the right femoral artery sheath. the runoff revealed no distal flow to left extremities. the patient required minimal pressor support. femoral bypass was discussed and possible axillary placement for alternative access with the medical doctor. ultimately, the decision was to remove the pump and support with pressors. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3058 | 1220648-2025-48726 | 23692887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23692887&pc=OZD | 11/04/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/04/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766097 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction had an unsuccessful attempted implant with an impella cp for mechanical circulatory support (msc). the patient was transported to the catheterization lab from a floor code. the impella cp was attempted to be urgently, placed. however, the impella cp device was unable to cross aortic valve. the impella was removed, and an intra-aortic balloon pump (iabp) was placed instead for mcs with a survived outcome. the patient was noted to be in stable following the event. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the inability to deliver or advance (delivery issue), the cause was determined to be patient condition as the patient had a difficult anatomy and resistance at aortic arch along with known coronary artery disease preventing successful delivery of impella. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3059 | 1220648-2025-48727 | 23693020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23693020&pc=OZD | 11/02/2025 | 01/05/2026 |  | 12/02/2025 |  |  | 11/02/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734950 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient got an impella 5.5 as escalation. the impella 5.5 was placed utilizing a double barrel technique. the impella 5.5 flows increased and impella cp flows decreased. the impella cp was explanted. while leaving the operating room, purge pressure high alarm was presented. no kinks in the line were noted. the purge cassette was replaced. no alarm post purge change was noted. no motor current spikes were reported. purge pressure remained high and flows were low. the decision was made to monitor at this time. the medical doctor wanted to utilize tissue plasminogen activator protocol, and was initiated by the care team. discussion of alarms and close monitoring were performed with multiple care teams. this is one of two related reports. this report represents the pump and another report was submitted to represent the purge cassette. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | Investigation summary no product was returned. purge pressure high: clinical details note that a 5.5 was placed utilizing a double barrel technique. as leaving the or, purge pressure high alarm presented. no kinks in the line noted. purge cassette replaced and csc notified. no alarm post purge change. purge pressure remained high and flows low. md wanted to utilize tpa protocol, which was initiated by care team. data log review showed purge pressure around 600 mmhg after priming. pp then slightly decreased after the pump was started but then increased to ~900 mmhg in approximately 4 minutes after pump startup. an hour later, there was a small mc spike and upward shift in mc and pump flow. at this time, purge pressure increased above 1000 mmhg. the purge cassette was replaced, which did not resolve the hpp. a bag change followed a couple hours later and the pp decreased from there. the cause of the high purge pressure was biomaterial as there were clear changes in mc and pump flow at the time of the issue but the type of biomaterial is unknown since the log characteristics are inconclusive. device history lot device lot: 1964705. device history batch subcomponent lot: n/a. device history review the pump sn(b)(6) passed all post-sterile inspection checks.
3060 | 1220648-2025-48729 | 23693280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23693280&pc=OZD | 11/04/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/04/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696885 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Inaccurate Flow Rate | Decrease in Pressure | Increase in Pressure | Pumping Stopped |  |  |  | Low Blood Pressure/ Hypotension | Pulmonary Edema | Thrombosis/Thrombus | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient wishes to withdraw care. pump coincidentally also stopped. the pump was left transvalvular while on extracorporeal membrane oxygenation (ecmo). the patient expired while on ecmo. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient had a drop in bp over night (levo added). this morning the purge flow went up to 30 and purge pressure decreased. according to staff there was no leak in the system. unsure how long this stayed active. purge cassette was changed and got alarms for high purge pressure and low flows; motor current spiked and lv waveform and flows are saturating. purge flows and pressures are normal now. pump has been in since (b)(6) 2025. repeat echo done yesterday to evaluate lv thrombus. |  |  |  |  | Additional information received. b1, b5, d6b, h1, and h6 updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3061 | 1220648-2025-48731 | 23693414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23693414&pc=OZD | 11/02/2025 |  |  | 12/02/2025 |  |  | 11/02/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722807 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. it was reported that the impella flow was sluggish past the sheath on fluoroscopy image. doppler pulse was present. the doctor was assessing peripheral lower extremity and start acute coronary syndrome protocol heparin post gastrointestinal procedure. the next day it was reported that there was profuse bleeding from the mouth and heparin was turned of. later, the impella was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3062 | 1220648-2025-48733 | 23694085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23694085&pc=OZD | 04/06/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/03/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555437 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Low Blood Pressure/ Hypotension | Ischemia | Heart Block | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An long-term outcome and quality impella registry was received. it was reported that an impella cp was placed in a patient for heart failure support. the left femoral artery was accessed using best practice. the impella was inserted via 14 french x13 centimeters peel away sheath. percutaneous coronary intervention was performed to the left main left anterior descending coronary artery and left circumflex artery with a total of two stents placed and balloon angioplasty. the patient remained stable through the procedure. the pump was explanted without complication. three adverse events were reported. the patient presented to the emergency department with chest pains and tightness that started the night before and earlier that morning. th patient had a stress test today and was sent to the emergency department for evaluation. the chest pains worsened with exertion and were relieved with nitroglycerin. the patient also complained of shortness of breath. the patient was admitted for chest pain/possible acute coronary syndrome. cardiology was consulted, medications were adjusted and cardiac biomarkers were drawn. stress test showed normal function/low risk. the patient felt the chest pains could have been due to acid reflux. the patient was to follow up with gastro post-discharge. approximately four months later, the patient presented to the emergency department with a two minute syncopal episode, tachycardic, hypotensive and febrile. the patient had previously passed out in the shower. the patient was covid positive when admitted. the patient was on non-st-segment elevation myocardial infarction (nstemi) pathway. nstemi could be type two demand ischemia due to covid infection or due to the patients cardiac history. the patient wished to continue on medical therapy. the patient reported to have chronic/stable class ll-lll angina. the patients chest pain and shortness of breath improved and that patient wanted to go home. the patient declined an angiogram. the patient was discharged and was chest pain free and stable on room air. home medications were adjusted. approximately one month later, the patient presented to the emergency department with a complaint of unstable angina. per the patient, it has been intermittent for a few weeks and was relieved by nitroglycerin. the patient was admitted and an angiogram was performed. post angiogram, the patient was discouraged. the patient has mulitivessel coronary artery disease but otherwise is comfortable. options were discussed. the patient was hesitant about surgical options. an electrocardiogram showed first degree atrioventricular heart block as before, anterior and lateral down sloping st-depression (worse than prior). the angiogram showed severe multivessel coronary artery disease. the patient was angina free at discharge and did not want surgery. the lesion would make it difficult to stent. the plain is to consider higher-risk percutaneous coronary intervention versus medical management. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. ischemia/ heart block/ arrhythmia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension/ fever: the cause of the injury was not determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3063 | 1220648-2025-48735 | 23694207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23694207&pc=OZD | 11/02/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/03/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026767578 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female | 95 KG |  | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed in a patient for chest pain support. the decision was made to place the impella cp via the right femoral artery due to an end-diastolic pressure of 44 and low ejection fraction. the impella cp was placed without issue, and single access was obtained using the companion sheath. prenumbra was used to remove a thrombus in the left anterior descending artery and the left anterior descending artery was stented x 2. at the end of the case, the peel away sheath was removed, slack was pulled back and the repo sheath was advanced and sutured in place. the patient weaned off of support. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the injury, thrombosis, was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3064 | 1220648-2025-48736 | 23694465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23694465&pc=OZD | 11/05/2025 |  |  | 12/02/2025 |  |  | 11/05/2025 | 12/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739636 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced heparin induced thrombocytopenia. the patient was treated with argatroban. the patient also had a thrombosis that was treated with bilateral lower leg thrombectomy with a drain in the right leg. impella support continued. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. thrombosis: the cause of the injury was unable to be determined due to insufficient clinical details. device history lot. device lot: 1973996. device history review.the pump s/n: passed all the post sterile inspection checks. | D6b: added explant date | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3065 | 1220648-2025-48737 | 23694736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23694736&pc=OZD | 11/05/2025 | 12/26/2025 |  | 12/02/2025 |  |  | 11/05/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743984 |  | 000000000001000459 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. no impella access complications were noted. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after cabg, the patient was unable to wean off cpb, so an impella 5.5 was inserted postoperatively and the patient was returned to the icu. a cerebral hemorrhage was subsequently discovered and the patient passed away a few days later. the relationship between the cerebral hemorrhage and impella is unknown. the attending physician's opinion was that the patient had pre-existing coagulation abnormalities, and the cerebral hemorrhage may have been caused by the effects of intraoperative cardiopulmonary bypass and heparinization. dr. comment below added by (b)(6). death due to postoperative cerebral hemorrhage. the relationship with impella is unlikely. the physician believes that the cerebral hemorrhage was likely caused by the effects of intraoperative heparinization, etc., due to the patient's pre-existing coagulation abnormalities. |  |  |  |  | E1 revised initial reporter first and last name as they were entered incorrectly on the initial report that was submitted. added initial reporter postal code as it was omitted from the initial report that was submitted. g1 manufacturing site name should of been left blank on the initial report that was submitted. investigation summary: the investigation has been completed. peri-procedural adverse event (stroke): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3066 | 1220648-2025-48739 | 23694823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23694823&pc=OZD | 11/05/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/05/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737031 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was escalated from an impella cp to an impella 5.5. during support chest radiography revealed bilateral pneumonia. cultures sent. later there was run of ventricular tachycardia, on amiodarone drip. later care was withdrawn and the patient expired. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. infection/ arrhythmia: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. minor bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3067 | 1220648-2025-48738 | 23694991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23694991&pc=OZD | 10/06/2025 |  |  | 12/02/2025 |  |  | 11/21/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Cross-functional review identified that the console was the potential cause of the oscillating contrast issue on (b)(6) 2025. the user facility reported that a patient was implanted with an impella 5.5 experienced a placement signal not reliable alarm from the supporting automated impella controller. no patient harm was reported. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to h10 for the related report number. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3068 | 1220648-2025-48740 | 23695099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23695099&pc=OZD | 11/04/2025 |  |  | 12/02/2025 |  |  | 11/05/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727388 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Perforation | Pericardial Effusion | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp became hemodynamically unstable with placement signal low and impella flow reduced alarms. a left ventricle perforation was suspected, and pericardial effusion was detected on echo. a pericardiocentesis was performed and the impella was explanted. the patient survived. | [5 empty] | Correction: b1. the investigation of the reported failure mode has been completed. no product was received for evaluation. low or blocked pump flow: data log reviewed, and no indications observed to confirm root cause. the root cause of the impella low flow issue cannot be determined since the complaint device not returned for analysis and insufficient clinical details provided. cardiac perforation/ pericardial effusion/hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3069 | 1220648-2025-48742 | 23695123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23695123&pc=OZD | 11/05/2025 |  |  | 12/02/2025 |  |  | 11/05/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749574 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female | 60 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed oozing at the access site. also, retroperitoneal bleeding was identified in a ct. two units of blood were transfused. due to the patient's coagulation disorder and multiple organ failure care was withdrawn and the patient expired. | [5 empty] | No product returned for investigation. the cause of the access site bleeding was patient related due to the patient's coagulation disorder. the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3070 | 1220648-2025-48743 | 23695525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23695525&pc=OZD | 11/06/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/06/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026666100 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. hematoma and bleeding to impella site causing removal of device and surgical wound closure to rcfa. | [5 empty] | B7 added information as it was omitted from the initial report that was submitted. g1 manufacturing site name should of been left blank on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. major bleed/hematoma: hematoma and bleeding to the impella site causing removal of device and surgical wound closure to right common femoral artery. the root cause of the hematoma and bleeding were unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3071 | 1220648-2025-48744 | 23696276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23696276&pc=OZD | 11/06/2025 | 12/30/2025 |  | 12/02/2025 |  |  | 11/06/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811668 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 91 KG | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller in a 53-year-old male patient. decision to place patient on impella support with plan to send patient to left main coronary for advanced heart failure therapies. upsized to 14 fr impella short sheath. heparin administered, dilator removed, left ventricle (lv) accessed w/0.035 and pigtail. 0.035 exchanged for 0.018 guide wire. diagnostic catheter removed, impella backloaded onto 0.018 wire, sheath flushed, pump advanced and placed into lv. support initiated in auto then changed to p-9 per ic request. peel away sheath removed, repo sheath advanced, slack removed, tuohy-borst and t-lock secured. forward tension sutures to blue suture hub and site dressed w/angle matching. pt sent to intensive care unit for continued management while awaiting transfer. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data review: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer -see attached image). the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. es2023-0090 documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly.
3072 | 1220648-2025-48732 | 23696481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23696481&pc=OZD | 11/04/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/04/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696982 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. at this time, a repositioning sheath was inserted, and the pump was fixed. subsequently, while engaging a guiding catheter into the rca, the culprit lesion, the patient developed severe bradycardia, and the decision was made to start cpr and introduce ecmo. during the cpr process, the pump catheter became completely dislodged into the aorta. pci continued with ecmo support, and the patient was admitted to the ccu after treatment was completed. after a cardiovascular surgeon arrived, the cp was removed. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3073 | 1220648-2025-48745 | 23696546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23696546&pc=OZD | 11/06/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/06/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673813 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 65 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Mitral Valve Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. during support bleeding and oozing from impella site after exchange to repo sheath. forward tension sutures and angle matching dressing applied. required manual pressure. it was further reported that there were possible hemolysis, urine changed from clear yellow to dark red. pump repositioned under echo, p level decreased. the patient was required escalation to extracorporeal membrane oxygenation (ecmo) due to labile hemodynamics. after ecmo placed urine cleared and became yellow. repeat echo done and no further repositioning required. there were concerns of worsening mitral regurgitation. plan to intubate and transesophageal echocardiogram (tee). later the impella was removed due to he sudden severe onset of suspected hemolysis, lactate dehydrogenase increased from 400 to approximately 2300, with very dark almost ¿black¿ urine, and rising creatinine over the previous 24hrs. also mentioned while impella was in they dealt with ongoing pulmonary edema and that 24 hrs after removal of the impella the chest x-ray dramatically cleared. he was wondering if the pump was affecting the mitral valve. also after removal urine cleared back to clear yellow. the patient remained on ecmo and survived. | [5 empty] | H6 code a24 was incorrectly reported. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3074 | 1220648-2025-48734 | 23697239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23697239&pc=OZD | 11/03/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/03/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659430 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female | 62 KG | Non Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp in a 26-year-old female for support for cardiac arrest. it was reported that during support, the pump measured at 4.3cm on echo, requiring repositioning under formal tte guidance as soon as possible. additionally, the patient had a large mca stroke with significant cerebral edema requiring decompressive craniotomy. the stroke occurred at an outside facility prior to cannulation to (b)(6). however, it was further reported that care was withdrawn the patient subsequently expired. | [5 empty] | B1: added product problem. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. stroke: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3075 | 1220648-2025-48747 | 23697462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23697462&pc=OZD | 11/06/2025 | 12/31/2025 |  | 12/02/2025 |  |  | 11/06/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713180 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a patient for high-risk percutaneous coronary intervention support.impella inserted and activated at 1204 per instructions for use. single access using terumo destination. coronary stent infection to left anterior descending artery, patient became restless, decreased maps and pulseless electrical activity arrest.return of spontaneous circulation achieved and intervention continued. impella weaned to p2 but decision made to leave in for the night. peel away removed and repo inserted. site sutured and dressed. due to iliac anatomy - cardiothoracic surgery,following with plans for femoral cut down on removal tomorrow(part of pre case plan per dr howard). after discussion and suspicion of disseminated intravascular coagulation impella and all access sites bleeding. impella removed. dry closure attempted via contralateral access with cardiothoracic surgery present as backup. perclose successful.intra-aortic balloon pump placed in left femoral artery for continued hemodynamic support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed could not be determined.
3076 | 1220648-2025-48746 | 23697500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23697500&pc=OZD | 11/05/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/06/2025 | 12/02/2025 |  |  |  |  | 12/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696296 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  |  | White | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 53 year old female for cardiac arrest support. impella cp was prepped in standard fashion using dextrose solution. cp was back loaded and placed into mid left ventricle space under fluoro. wire was removed and impella started into auto mode with flows of 3.6 to 3.7 liters per minute noted. peel away sheath removed and repo sheath advanced, forward tension sutures applied and angle matched with dressing cardiac distress inventory. cardiothoracic surgeon consulted for venoarterial extracorporeal membrane oxygenation, eventually sats improved so extracorporeal membrane oxygenation deferred with low threshold for initiating. low pulsatility noted and high body surface area with henodynamics beginning to fall so conversation about escalation was had, dr phadke and el ashry agree she will probably need 5.5 support. transferred to unit in stable condition. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3077 | 1220648-2025-48748 | 23701399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23701399&pc=OZD | 11/06/2025 | 12/02/2025 |  | 12/02/2025 |  |  | 11/06/2025 | 12/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522168 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was brought to the cath lab in cardiogenic shock on multiple inotropes and pressors. shock team initiated and opted for impella placement at east. right common femoral artery access and angio shot which showed adequate anatomy. artery dilated up and 14 f x 13 cm sheath placed. heparin given. left ventricle (lv) access was obtained. impella loaded over 0.018 wire and entered in to the lv. distal flow showed past impella sheath. peel away removed and repositioning sheath placed. forward tension sutures placed x 2. angle matched. groin was soft without any bleeding or hematoma. the patient was transferred. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. major bleed/access site adverse event: oozing noted from all access sites. the cause of the issue was likely patient condition related as evidenced by the multiple site bleeding optical sensor issue: insufficient logs returned for analysis. the date/time on the logs returned are from 26-may-25 at 11:52am till 26-may-25 at 15:46pm. the logs returned had no impella position unknown alarms or other optical sensor related abnormalities. the 5s parameters were within range and stable but on pump start there was a decrease in snr and contrast. the cause of the issue was not established as insufficient data logs were returned and no product was returned for analysis. device history review: the pump passed all post sterile inspection checks.
3078 | 1220648-2025-48680 | 23679131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679131&pc=OZD | 11/01/2025 | 02/23/2026 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758086 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information has been provided upon request: "regarding this case, the patients low native pulsatility falsely triggered a position in aorta alarm when it should have been position unknown. the pump was not saved for return. the issue was resolved after disabling the false alarm." | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and on day three of support, the patient became hemodynamically unstable. there was also a report of device sensing issues. the patient would subsequently expire. after successful placement of the impella, the patient settled in the intensive care unit. critical care placed a central line with plans to float a swan catheter. the patient was supported at performance level p-5 due to continuous suction. the following day, discussion of escalation was made. under transesophageal echocardiography, the patient's heart looked underfilled. plans to give fluid and reassess later in the afternoon was noted. some suction alarms occurred which resolved with lowering the performance level. the team was hoping for that to resolve once fluid resuscitated. per the report, there were no other device related concerns. now day three of support, the patient began decompensating and had a position in aorta alarm around. the impella was repositioned bedside with the critical care team. the impella position was noted as okay at this time. despite the impella being repositioned, the patient was still having intermittent placement alarms. plan was to make the patient comfort care per the family and critical care team. placement signal alarm silenced. the support continued until family arrived and then care was withdrawn. the patient expired. |  |  |  |  | B5 updated. the investigation is ongoing. | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device report. medical safety/clinical review. medical safety/clinical reviewed the event. the patient presented to the emergency department via emergency medical services with st-segment elevation myocardial infarction and a several-day history of burning chest pain. while in the emergency department, the patient experienced a pulseless electrical activity (pea) cardiac arrest. cardiopulmonary resuscitation was initiated and return of spontaneous circulation was achieved with the emergency physician and interventional cardiologist present. an impella device was implanted for mechanical circulatory support without reported difficulty. single vascular access was obtained, and percutaneous coronary intervention to the left anterior descending artery was performed. during ongoing impella support, the patient demonstrated progressive clinical deterioration. at approximately 0000 hours, a ¿position in aorta¿ alarm was noted. the impella device was repositioned. despite repositioning, intermittent placement alarms continued to occur. given the patient¿s continued hemodynamic decline and poor prognosis, the family and critical care team elected to transition the patient to comfort care. the impella placement signal alarm was silenced. support was withdrawn, and the patient subsequently expired.
3079 | 1220648-2025-48681 | 23679203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679203&pc=OZD | 11/02/2025 |  |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750932 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | False Alarm | Inaccurate Flow Rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced an impella in aorta alarm, incorrect placement signal, and saturated flows. catecholamines were reduced and anticoagulation was increased, and the alarms resolved. impella support continued. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b, explant date, has been added. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. saturated flow: since neither product nor data logs were returned for investigation, the cause of the saturated flows could not be determined. optical sensor issue: since neither product nor data logs were returned for investigation, the cause of the optical sensor issue could not be determined. false alarm: since neither product nor data logs were returned for investigation, the cause of the false alarm could not be determined. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3080 | 1220648-2025-48682 | 23679246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679246&pc=OZD | 11/02/2025 |  |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532775 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had reduced blood flow to the leg with increasing edema. pedal pulses were present. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4, a5, and a6 are unknown, the impella device is not possible to obtain from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3081 | 1220648-2025-48683 | 23679357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679357&pc=OZD | 11/02/2025 |  |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706493 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced anemia and thrombocytopenia. the patient was transfused two units of packed red blood cells and two units of platelets. the patient was in stable condition. | [5 empty] | No product was returned for investigation. the cause of the anemia and thrombocytopenia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this is one of two related reports. this report represents the impella 5.5 and another report will be submitted to represent the impella cp.
3082 | 1220648-2025-48684 | 23679410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679410&pc=OZD | 11/02/2025 |  |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734970 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 98 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Tachycardia | Hemodynamic instability | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an elevated white blood cell count. the patient was treated with antibiotics. the impella alarmed for placement signal so the pump was repositioned. in response to oozing at the access site, the patient was transfused packed red blood cells, a new dressing was placed, and heparin was withheld. the patient also experienced ventricular tachycardia and ectopy, and anemia and hemodynamic instability. the p level of the pump was adjusted and the patient was transfused an additional unit of packed red blood cells. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The root cause of the infection, arrhythmia, and tachycardia was unable to be determined due to insufficient clinical details. the cause of the bleeding, hemodynamic instability, and anemia was not determined due to insufficient clinical details. the cause of the optical sensor issue was not determined as no product or data logs were returned for analysis. the cause of the access site bleeding was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3083 | 1220648-2025-48685 | 23679600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679600&pc=OZD | 11/01/2025 | 01/05/2026 |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026708808 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an incorrect placement signal. the pump successfully supported the patient to transplant. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary - optical sensor issue on the 1st of november, it was stated that the ao ps was incorrectly low. due to a lack of clinical details, and no product or logs returned, the root cause cannot be established. device history lot - device lot: 1939868. device history batch - subcomponent: n/a. device history review - impella 607705 passed all post sterile inspection checks.
3084 | 1220648-2025-48686 | 23679697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679697&pc=OZD | 11/03/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684550 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a high purge pressure alarm and bleeding from the access site. the pump alarm resolved without intervention. the patient was treated with protamine, coagulants, blood transfusions, and additional sutures to obtain hemostasis. impella support continues. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device nor data logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the bleed, clinical details noted that patient was administered protamine due to bleeding. the cause of the injury was not determined due to insufficient clinical details. high purge pressure: clinical details noted that high purge pressure occurred about a half hour after implantation, after patient was given protamine due to bleeding. no tissue plasminogen activator (tpaz0 was given and heparin was not added to purge. high purge pressures resolved on its own a few hours later. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3085 | 1220648-2025-48687 | 23679785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23679785&pc=OZD | 11/03/2025 |  |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 |  |  | 01/31/2026 |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263730 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Death | Yes | No | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during patient support the impella automated impella controller (aic) fell off the stator and onto the floor. there was visible damage to the plastic housing of the aic. the aic was exchanged for a new one to continue patient support. later, care was withdrawn and the patient expired. | [5 empty] | A4 and a5 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3086 | 1220648-2025-48688 | 23680378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23680378&pc=PYX | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026718233 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support. hemolysis was noted and validated with lactate dehydrogenase. the impella cp was repositioned without improvement three hours later. the impella rp flex appears the inlet may be in contact with tv and be the source of hemolysis. low flows with high pulmonary artery pressure reading on automated impella controller were noted. clinical picture worsened. the device was explanted for concern of clot ingestion. the patient was escalated to veno-arterial extracorporeal membrane oxygenation. the patient was noted to survive. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D9 the device was returned and the date received was added. h6 codes updated to reflect the device being returned. h6 medical device problem code was also updated to be better reflect what was reported.
3087 | 1220648-2025-48689 | 23680815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23680815&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026746967 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, there were concern for a gastrointestinal (gi) bleed. the patient was also bleeding from medial right leg fasciotomy site. currently, no anticoagulation was ordered but the team considering adding bivalirudin. arterial blood gas analysis improved plan to decrease sweep. an echocardiogram was not officially read but rounding team estimates ejection fraction at 40% this morning. the plan was to decrease dobutamine to 2.5 mcg and discontinue amiodarone. later, it was reported that the patient went into ventricular fibrillation arrest overnight and was taken to the catheterization lab. no blockages were noted but appeared to be having global spasms. the impella increased to p6 and amiodarone was restarted, nitroglycerin and calcium channel blocker were added. the gi bleed and bleeding from leg resolved. one unit of blood was given overnight. the patient was later successfully weaned and explanted. the patient outcome after explant was noted as survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the arrhythmia/ ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: concern for gi bleed, patient also bleeding from medial right leg faciotomy site. amio restarted, nitro and calcium channel blocker added. gi bleed and bleeding from leg resolved. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3088 | 1220648-2025-48690 | 23681402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23681402&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695271 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Sepsis | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, lidocaine was given for ectopy overnight. the impella was at p5 and was attempted to increase to p6 to give more flow from the impella before going up on pressors but suction alarm on p6. echocardiogram showed the impella in good position. later, it was noted that there were questions for mixed shock presentation, maybe septic? blood cultures were pending at this time. last white blood cell was 20. additionally, there was a concern of upper gastrointestinal (gi) bleeding yesterday so the team stopped systemic heparin. gi was consulted and per gi note, endoscopy for upper gi would be unfavorable due to critical condition but no further bleeding from gi tube. the patient was later weaned and explanted. the patient's outcome at explant was noted as survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. arrhythmia/ sepsis: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: pt with gi bleed, noted when checking tf residual. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: device passed all post sterile inspection checks.
3089 | 1220648-2025-48691 | 23681473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23681473&pc=OZD | 11/15/2025 | 12/24/2025 |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655390 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | [6 empty] | Stroke/CVA | Renal Failure | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | It was reported that a patient implanted with an impella 5.5 suffered an optical sensor issue, hematuria, and a cerebral vascular accident. the patient is in stable condition and impella support continues. | [5 empty] | D6b, explant date, has been added. | The patient also experienced renal failure. the patient was treated with dialysis and transfusion of platelets. | B3 has been updated. | D6b, explant date, has been added. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3090 | 1220648-2025-48692 | 23681623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23681623&pc=OZD | 11/01/2025 | 02/20/2026 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026746891 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was on bipella support for mechanical circulatory support. while on support, hemolysis was noted and the impella cp was repositioned without resolution. the pump was later explanted and the patient survived. this is one of two related reports. this report represents the impella cp and another report was submitted for the impella rp flex. refer to section h10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3091 | 1220648-2025-48694 | 23681846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23681846&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747636 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 72 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Inflammation | Pain | Renal Failure | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, it was reported that the patient had arm swelling and pain to the right upper extremity (impella 5.5 cutdown). it was also reported that the patient has a hematoma at the impella site. plans for a wash out in the operating room. furthermore, there was a gastrointestinal (gi) bleed reported by the nurse. heparin drips were discontinued for the gi bleed management. renal failure was noted which an additional device was required. days later, there was a pump suction issue. the care team reported lower flows on the impella with some suction alarms overnight. they were unable to increase support. sever right ventricle dysfunction seemed to be the cause. the pump was later explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3092 | 1220648-2025-48693 | 23681854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23681854&pc=OZD | 11/04/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689342 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced suction alarms and reduced flows. the medical staff suspected the patient had a thrombus. repositioning of the pump was unsuccessful and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the cause of the injury was not determined due to insufficient clinical details. the cause of the product issue was not determined without returned product investigation and sufficient clinical details, including data logs.
3093 | 1220648-2025-48695 | 23681913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23681913&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756751 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 73 YR |  | Male | 109 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) pre-percutaneous coronary intervention (pci). a failed attempt to pass the guidewire for pci led to the intervention being aborted with surgery consult. the patient was transferred to the unit on impella mcs. on arrival to the unit, the access site, right groin site, was oozing off and on. the oozing was resolved with manual compression and dressing change with proper angle matching. partial thromboplastin time was greater than 120. three days later, after successful revascularization of the left anterior descending artery and circumflex artery, the patient was weaned from all pressor support and impella cp. heparin was held and activated clotting time in range to explant the impella cp, which was successfully completed with a survived patient outcome. the patient remained stable post explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the bleeding, the cause is most likely use related since on arrival to intensive care unit, the right groin site oozing off. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3094 | 1220648-2025-48698 | 23682428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23682428&pc=OZD | 11/03/2025 | 01/05/2026 |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ischemia of the right leg. a patchplasty was performed to resolve the issue. | Resuscitation was performed during the impella implant. |  |  |  |  | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. investigation summary: ppae (ischemia): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot. device lot: unknown. device history batch. subcomponent: n/a. device history review a post sterile inspection was unable to be conducted as the serial number of the complaint device is not known.
3095 | 1220648-2025-48697 | 23682684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23682684&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740413 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 159 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. audio alarms for suction were noted and placement signal monitoring remained off per the doctor. lactate was already down trending. the next day, ongoing placement signal was low and suction alarms were noted. ac was held due to diffuse bleeding from the impella yesterday. the next day, thrombocytopenia continued. two days later, the patient was transfused one unit of packed red blood cells and one unit of platelets overnight. the patient was escalated to an impella 5.5 about a week later. this is one of three related reports. this report represents the impella cp and another report will be submitted to represent the automated impella controller and another report will be submitted to represent the impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3096 | 1220648-2025-48696 | 23682990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23682990&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1131197 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. aortic placement signal disappeared and then controller fault error message appeared. the console was swapped. this is one of three related reports. this report represents the automated impella controller and another report will be submitted to represent the impella cp and another report will be submitted to represent the impella 5.5. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3097 | 1220648-2025-48699 | 23684171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23684171&pc=OZD | 11/03/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713802 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 67 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. ischemic limb distal to impella cp. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3098 | 1220648-2025-48700 | 23684583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23684583&pc=OZD | 11/01/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706505 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 and experienced a gastrointestinal bleed during support. the patient was admitted to the hospital with hypotension and was receiving bumex and dobutamine. the patient was transferred to the cardiac catheterization laboratory, and the medical team attempted to wean the patient from dobutamine without success. the patient¿s blood lactate results were 8.4 mmol/l and subsequently dobutamine was increased and lasix was initiated. due to clinical worsening, the patient was intubated and cannulated for veno arterial extracorporeal membrane oxygenation. subsequently, an impella 5.5 was placed. the impella 5.5 was inadvertently pulled back into the aorta, and the physician was unable to readvance the impella back into the ventricle. the decision was made to place a new impella 5.5. on support day 7 with the replacement impella 5.5, it was reported that the patient experienced a gastrointestinal bleed. a medical plan was made for a gastrointestinal scope for bleeding. no further information was provided. the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. this complaint involves two impella devices: pump 1: impella 5.5, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other device associated with this complaint. | The user facility reported an impella 5.5 in a 45 year male patient for heart failure support. the patient was hypotensive and was on bumex and dba 5 and then transferred to the coronary care unit. they tried to wean the dobutamine with no success,. repeat labs showed lactate at 8.4,. they increased dobutamine to 7.5 and started lasix drops. due to clinical worsening, cardiothoracic surgery was consulted for extracorporeal membrane oxygenation cannulation. the patient arrived to the cardiothoracic intensive care unit intubated and then cannulated for venoarterial extracorporeal membrane oxygenation. the impella 5.5 was placed and was at 5.5 centimeters in the left ventricle on echocardiogram. the impella went into the aorta and the physician was unable to advance so the decision was made to place a new impella 5.5. |  |  |  |  | B5 is being updated to include all related device information not included in the initial report. the revised b5 provides a complete event narrative. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other impella 5.5. investigation summary: no product was returned for investigation. major bleed: plan to go get a lower gi scope today for gi bleeding. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history lot: device lot: 1942004. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3099 | 1220648-2025-48701 | 23684764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23684764&pc=OZD | 11/04/2025 |  |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1622811 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after the maintenance work was completed at the head office, the equipment was brought into the hospital, powered on, and when the ce checked it, an error message appeared. when the control device was turned on, a message appeared saying, "self-diagnosis failed. please replace the control device immediately." there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3100 | 1220648-2025-48702 | 23685761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23685761&pc=PYX | 11/02/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551600 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 127 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Renal Failure | Hematuria | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock (stage e) from an acute myocardial infarction was implanted with an impella rp flex for mechanical circulatory support (msc) along with an intra-aortic balloon pump (iabp). post impella implant, urine output was greater than 30 cc/hour and pink-red. the following day it was noted that the patient was failing on the current therapy. renal was consulted; the patient was not clearing potassium. the patient was transported back to the catheterization lab for impella rp flex repositioning and a line for continuous renal replacement therapy (crrt) was placed. this day, the patient¿s status was changed to dnr (do not resuscitate). the next day the patient unexpectedly became severely bradycardic and expired. the patient was a dnr and no rescue was provided. it was noted possibly potassium (k+) related and the team was currently waiting for crrt to be resumed. potassium level was over 5. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the hematuria/ renal failure, the cause was not determined due to insufficient clinical details. regarding the bradycardia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Medical safety reviewed the event and noted the event is a death type of reportable event. however, it is noted underlying condition due to abnormal lab values and that this was a high-risk intervention supported by impella mechanical circulatory support (mcs). severe right ventricular dysfunction was noted post procedure by echocardiography. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3101 | 1220648-2025-48703 | 23685778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23685778&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759811 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 63 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support (msc) and experienced thrombus, access site bleeding and limb ischemia. the patient presented to the emergency department with diffused chest pain for two hours with radiation to neck. st-segment elevation myocardial infarction was noted, and the patient was emergently taken to the catheterization lab. initially, an intra-aortic balloon pump (iabp) was placed. the percutaneous coronary intervention (pci) began utilizing the right radial. the patient continued to decline, and the decision was made to exchange the iabp to an impella cp device which was implanted without complication. the pci was with two stents to the right coronary artery. the patient had reoccluded the right coronary artery requiring further intervention. a large clot was noted, and pneumbra was utilized as well. of note, the physician reported moderate to severe right ventricular dysfunction and considered right-sided impella support. there is no information that indicates right-sided support was initiated. the pci was completed. following, it was noted there was bleeding around the blue suture hub was noted. the site remained oozy. aggrastat was continued and activated clotting time was 385. the scrub tech was educated on proper groin dressing. however, there is continued disregard for the angle of insertion. the angle was lost and bleeding worsened. the bleeding was successfully managed with manual pressure, perclose synch, angle matching, and pressure dressing. the patient had received two units of packed red blood cells overnight. additionally, the previous night, a distal perfusion catheter was placed as the patient lost pulses after impella cp manipulation. impella mcs continued. a couple of days later, the patient had staged pci to the left anterior descending artery, and the impella was removed post procedure in the catheterization lab with a survived outcome. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the bleeding around blue suture hub, activated clotting time (act) values reported during the time of bleed was 385. the cause of the bleeding is determined to be use issue due to anticoagulation management because the act values at the time of the bleed are higher than recommended. for ischemia/ thrombosis, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3102 | 1220648-2025-48704 | 23686120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23686120&pc=OZD | 11/01/2025 |  |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026747722 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Thrombocytopenia | [7 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, it was reported that the patient had many arrythmias overnight and defibrillator shocks. it was also reported that renal failure and the plan was for continuous renal replacement therapy. furthermore, bivalirudin was paused for concern of gastrointestinal (gi) bleed. platelets have dropped, so the plan to give one unit of blood. later, care was withdrawn and the patient expired. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. arrhythmia/ tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. renal failure/ anemia/ thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3103 | 1220648-2025-48705 | 23686721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23686721&pc=OZD | 11/02/2025 | 02/20/2026 |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612915 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Fever | Tachycardia | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. arrhythmias overnight and this am w/ afib and vt; confirmed placement of impella on pocus and formal tte, no c/f positioning contributing. patient also was febrile, with blood cultures drawn. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. fever: the cause of the injury was not determined due to insufficient clinical details. paroxysmal atrial fibrillation/ ventricular tachycardia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks.
3104 | 1220648-2025-48706 | 23686960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23686960&pc=OZD | 11/02/2025 | 01/29/2026 |  | 12/01/2025 |  |  | 11/02/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593770 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 120 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hematoma | Hemolysis | Low Blood Pressure/ Hypotension | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. red urine, concern for hemolysis, hematoma and bruising at impella site. i met pt in icu. impella on auto with flows around 2lpm, waveforms were separated so i recommended echo to check placement. on echo impella measured deep in lv but seemed to jump with attempts to reposition (suspected pigtail under pap). we got it to an ok position with better flows. all the while the patient was with complaints of being uncomfortable and hot, she was diaphoretic, tachycardic and cold. her bps were low so pressors were started (levo and vaso). hr began to steadily decrease and flows on impella were also decreasing. epi, bicarb and atropine pushed and cpr was started and patient was intubated. rosc achieved after 6 minutes. echo to verify placement showed impella shallow so advanced for final placement at ~3cm in lv. | [5 empty] | Investigation summary: data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. logs also showed multiple suction alarms triggered during the case but resolved. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. hematoma: the cause of hematoma was unable to be determined as limited clinical details were provided and no product was returned for review. tachycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot#: 1862709. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3105 | 1220648-2025-48707 | 23688804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688804&pc=OZD | 11/01/2025 | 01/03/2026 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699676 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a (b)(6) male patient for support of chest pain. the decision was made by physician to wean and explant impella after unable to detect right pedal pulses via doppler. popliteal pulses were present, right lower extremity remained warm with adequate capillary refill, and no mottling noted. patient taken to cardiovascular laboratory and runoff of right leg revealed peripheral artery disease and sluggish flow. levophed used as bridge to explant to tolerate procedure, narcotic analgesia and sedation, as per physician. patient started on dobutamine and levophed being actively weaned. hemostasis achieved using contralateral balloon followed by manual compression for a total of one hour. patient stable and taken back to intensive care unit with pedal pulses present via doppler. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ischemia could not be determined.
3106 | 1220648-2025-48708 | 23688810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688810&pc=OZD | 11/01/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679738 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 in a 65 year old female patient. patient was previously placed on impella cp support early this morning. drips were increased over the course of the morning and patient began to worsen significantly. decision was made to escalate support to impella 5.5. right aortic arch to place graft and insert impella 5.5. 8mm graft was placed and peel away was inserted into graft site. impella was inserted over .018 wire at axillary site while impella cp was pulled back and removed at groin site. flouro and transesophageal echocardiogram used to visualize placement and insertion of both pumps. groin site was closed successfully and impella 5.5 was placed without issue. wire was removed and impella 5.5 was gradually increased to p8. plan for patient to remain on impella support until point-of-care testing can be determined. planning to bridge to either left ventricular assist device or transplant is possible. axillary site was sutured and dressed appropriately. 3 point fixation was placed slack was removed. | The complainant reported that a patient implanted with an impella 5.5 experienced a loss of placement signal waveforms on the device. it was reported that the patient was previously placed on impella cp support but subsequent to the patient¿s worsening conditions, the decision was made to escalate support to an impella 5.5. the impella 5.5 was placed without reported complications, and the impella cp was explanted. on support day 2, it was reported that the placement signal was erratic and reading falsely high pressures of 270/112 mmhg. the nurse observed that the placement signal would occasionally dash out and then re-appear with elevated pressures. it was noted that the patient¿s hemodynamics were stable and unchanged and there were no alarms observed on the automated controller unit. no signs or symptoms of patient injury were reported, there was no reported harm associated with the event. the device was later explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 is being updated to include related device information not included in the initial report. the revised b5 provides a complete event narrative. d9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. the manufacturer was notified that the event reported in the initial mdr 1220648-2025-48708 was also reported to the food and drug administration through the medsun program submitted nov-2025 with medsun report number 2300380000-2025-8049. this submission is to document the receipt of the medsun notification and to link the medsun report number to the previous submitted mdr.
3107 | 1220648-2025-48709 | 23688814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688814&pc=OZD | 11/01/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736457 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 111 KG | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 in a 53 year old male patient. prior to cut down, impella cp was weaned and explanted from the radiofrequency ablation and surgically closed. impella 5.5 placed via right axillary artery, imaging reviewed prior to case. successful cut down and graft attachment of 8mm graft. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the clinical states that the ps were falsely negative but then all placement signal was lost. the data logs show that the ps was stable until 10/31 and had run for about 736 hours until the ps declines from 80mmhg to about 40mmhg and slowly declines to 20mmhg over 43 hours until it rapidly declines about 60mmhg to -40 over triggering ps alarms over the course of an hour. no ps adjustments were made during this time. snr decreases to 0 but the rest of the 5s parameters remain relatively stable. based on the pattern in the data logs as well as the clinical details, the root cause of the optical sensor issue is due to material wear. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. d6 explant information updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3108 | 1220648-2025-48710 | 23688817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688817&pc=PYX | 11/03/2025 | 12/23/2025 |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026636132 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 102 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon explant of an impella rp a clot was observed on the pump inlet. the clot was removed. no patient harm was reported. | [5 empty] | The impella rp flex was returned for investigation. no issues were found on the returned product. the cause of the thrombosis and arrhythmia was unable to be determined due to insufficient clinical details. the cause of the access site bleed was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3109 | 1220648-2025-48711 | 23688821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688821&pc=OZD | 10/02/2025 |  |  | 12/01/2025 |  |  | 11/03/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722311 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Cardiac Arrest | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hypotension and cardiac arrest. chest compressions were given. the patient is in stable condition. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the hypotension, anemia, and cardiac arrest was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
3110 | 1220648-2025-48712 | 23688824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688824&pc=OZD | 11/04/2025 | 01/21/2026 |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026760755 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to the patient¿s anatomy. no patient harm was reported. | [5 empty] | Investigation summary: no product was returned. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had abnormal vessel anatomy and resistance at axillary artery preventing successful delivery of impella. device history lot: device lot: 1971334. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3111 | 1220648-2025-48714 | 23688862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688862&pc=OZD | 11/04/2025 |  |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1909656 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller showed a controller failure upon start up and before patient support. a different controller was used to successfully support the patient. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3112 | 1220648-2025-48715 | 23688875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688875&pc=OZD | 11/03/2025 |  |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 |  |  |  |  | 12/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1665435 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella was not recognized by the console. additionally, optical sensor error messages appeared. troubleshooting was unsuccessful. the automated impella controller was exchanged. the patient was noted as stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3113 | 1220648-2025-48716 | 23688942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688942&pc=OZD | 11/04/2025 |  |  | 12/01/2025 |  |  | 11/04/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736960 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Cardiac Tamponade | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | It was reported that a patient implanted with an impella 5.5 suffered multiple strokes and pericardial tamponade that the physician stated were not related to impella use. impella support continues. | Additional information was received stating that the patient was explanted. |  |  |  |  | B: added product problem. h6: updated codes based on the results of the updated investigation. h11: updated based on the results of the updated investigation. investigation summary: pump stop: the cause of the pump stop was an electrical open circuit failure due to mc spikes reaching above 32000ma. however, the cause of the electrical open was not determined due to no product returned and insufficient clinical details. patient codes - stroke: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac tamponade: the cause of the injury was not determined due to insufficient clinical details. clinical review: impella 5.5 support was initiated via axillary access in a 52-year-old male with severely reduced left ventricular ejection fraction and advanced underlying cardiac disease who underwent elective placement of an impella 5.5 prior to high-risk coronary artery bypass graft surgery. the patient presented with cardiogenic shock classified as society for cardiovascular angiography and interventions stage c and required inotropic therapy, vasopressors, and noninvasive positive pressure ventilation with continuous positive airway pressure and bilevel positive airway pressure for respiratory support. postoperatively, multiple cerebral infarcts were identified on imaging, and the patient experienced episodes of hypotension requiring cardiopulmonary resuscitation prior to coronary artery bypass graft surgery. neurologic events such as stroke are known risks in patients undergoing high-risk percutaneous and surgical coronary interventions and may occur due to atheroembolization and hemodynamic instability in patients with severe atherosclerotic disease and cardiogenic shock. per the treating providers, no correlation to impella therapy was identified. the patient subsequently developed recurrent pericardial tamponade requiring multiple thoracotomies, which was attributed to the complexity of the surgical course and patient condition and not considered related to the impella 5.5 device. the patient later required additional hemodynamic support with venoarterial extracorporeal membrane oxygenation. following a prolonged duration of mechanical support in the setting of multisystem failure and a palliative clinical course, a pump stop with inability to restart was reported after a performance level adjustment. pump stop is addressed in the instructions for use with guidance for clinical assessment and appropriate patient management when pump function cannot be restored. the patient expired shortly thereafter. the reported events are attributed to the patient¿s severe pre-existing cardiac dysfunction, high-risk surgical intervention, prolonged critical illness, and associated complications. cardiac tamponade, neurologic events, and hemodynamic instability are known risks described in the instructions for use for the impella 5.5. the patient codes associated with this device include stroke, cardiac tamponade, surgical intervention, and death; however, the death is being recorded conservatively. comprehensive review did not identify evidence suggesting a causal relationship between the impella 5.5 device and the reported patient event. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5. additional information was received and d6. updated. | The root cause of the stroke was unable to be determined due to insufficient clinical details. the cause of the cardiac tamponade was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | B2: updated reportability type and date of death. h1: updated reportability type. h6: added codes based on information in the complaint file.
3114 | 1220648-2025-48713 | 23688948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23688948&pc=OZD | 10/31/2025 | 12/01/2025 |  | 12/01/2025 |  |  | 11/01/2025 | 12/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026738766 |  | 000000000001000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Decrease in Pressure | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. nurse informed me that yesterday 10/31 pt developed gi bleed. gi doctor consulted and ptt range dropped on heparin gtt. pump also had a decrease in pressure. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: medical device problem code-coded incorrectly in the initial report. omit code a24 and a141101. replaced with code a141102. h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product returned. purge pressure high - after 9 days on support, purge flow <2 and pressure >1200. nurse practitioner made decision for tpa protocol. one day later, purge system back to normal flows and pressure. data log review showed pp rise begins around 0600 on 01nov and continues for 3 hours before high pp alarms triggered. high pp remains for over 24 hours before recovering. once purge metrics recover, they do not return to pre-issue levels. no mc spike seen outside of p-level changes. the cause of the purge pressure high was biomaterial build up due to pp rising over 3 hours and the issue resolving within 2 days after the use of tpa. patient codes - major bleed: the cause of the injury was not determined due to insufficient clinical details. device history review: pump s/n (b)(6) passed all post sterile inspection checks.
3115 | 1220648-2025-48662 | 23677007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677007&pc=OZD | 10/31/2025 | 11/29/2025 |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026740923 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. holding systemic heparin due to gastrointestinal bleeding. optical sensor issue, major bleed, arrhythmia, tachycardia, cardiac arrest. hemostasis, medication required, resuscitation | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrythmia/tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the injury was not determined due to insufficient clinical details. optical sensor issue: due to insufficient clinical information and no data logs or product returned, the root cause cannot be established. device history review: the device passed all post sterile inspection checks.
3116 | 1220648-2025-48667 | 23677009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677009&pc=OZD | 10/31/2025 | 11/30/2025 |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026719767 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. frequent ectopy noted upon placing impella 5.5. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product was returned for investigation. arrhythmia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history review device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3117 | 1220648-2025-48668 | 23677015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677015&pc=OZD | 10/31/2025 | 11/28/2025 |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699612 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 68-year-old male patient for support of chest pain. patient was taken to operating room to remove the impella cp due to diminishing left distal pulses. upon removal, vascular surgery did a thrombectomy where there were multiple clots removed from the iliac. artery was repaired and palpable pulses were noted upon return to the cardiovascular intensive care unit. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. ischemia / thrombosis: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review ==> the pump s/n: (b)(6) passed all the post sterile inspection checks.
3118 | 1220648-2025-48669 | 23677019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677019&pc=OZD | 10/31/2025 | 11/30/2025 |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731041 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 65 year old female patient for support of chest pain. the decision was made to place impella cp sinoatrial node (sa). successful insertion of impella cp sa to radiofrequency ablation. impella catheter at 83 cm, palpable positive distal pulses. patient to be transferred to unit. a thrombus was noted during explant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted.
3119 | 1220648-2025-48670 | 23677022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677022&pc=OZD | 10/31/2025 | 11/30/2025 |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734368 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 70 year old male patient for st-elevation myocardial infarction for support. impella was inserted pre percutaneous coronary intervention utilizing best practices. patient had great hemodynamic support for remainder of case, including rotational atherectomy to left anterior descending artery. at the conclusion of the case the left ventricular end-diastolic pressure trend was 40 - decision made to leave pump in overnight. peel away sheath removed and repositioning sheath was inserted and secured utilizing best practices. small hematoma noted to left groin, but was controlled with manual pressure and seatbelt pressure with fem stop. patient transported to imtensive care unit (icu). +emesis upon arrival to icu with subsequent groin bleeding and hypotension. epinephrine started and pressure held to left groin. decision made to explant impella. site closed with preclose x2. pt stable post explant. | [5 empty] | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product was returned for investigation. hematoma: there was a small hematoma which resolved with manual pressure. major bleed: on the transport to the intensive care unit (icu) and had emesis upon arrival to the icu with subsequent bleeding at the impella insertion site. the cause of the bleeding was determined to be use issue due to patient movement because the bleeding was noted at the insertion site after transport to icu. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history review device with passed all post sterile inspection checks.
3120 | 1220648-2025-48671 | 23677025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677025&pc=OZD | 10/31/2025 |  |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1208376 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. aic cord appears frayed and only charges at certain positions, aic to aic transferred performed. aic removed from service and sent to biomed. | [5 empty] | Device received, inv in progress: the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. no problems were found with the aic related to the reported issue of the power cord appearing frayed and charging only in certain positions.
3121 | 1220648-2025-48672 | 23677037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677037&pc=OZD | 10/31/2025 |  |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725329 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. pre procedure narrative: impella 5.5 placed with dr. (b)(6). phone support provided- had to re tunnel access site, rotate shoulder and use viper slide and buddy wire to achieve successful placement. impella measures at 5cm, did want to flip towards mitral, pump rotated towards apex and catheter re-straightened to eliminate turning. waveforms and flows look good. impella not interfering with any structures on tee. | [5 empty] | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. difficult to place or position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3122 | 1220648-2025-48673 | 23677042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677042&pc=OZD | 10/31/2025 | 11/30/2025 |  | 11/30/2025 |  |  | 10/31/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026768114 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 64-year-old male patient for st-elevation myocardial infarction support. the decision was made to place impella cp for support. impella cp was prepped and inserted per instructions for use protocol via radiofrequency ablation. percutaneous coronary intervention to left main performed. after the procedure the patient was sent to intensive care unit on support. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3123 | 1220648-2025-48674 | 23677046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677046&pc=OZD | 11/01/2025 | 11/30/2025 |  | 11/30/2025 |  |  | 11/01/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730834 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | False Alarm | [6 empty] | Respiratory Failure | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support after aortocoronary venous bypass (acvb) developed acute respiratory distress syndrome (ards) and the 5.5 alarmed impella in aorta. the impella was checked several times via transthoracic echocardiogram, and it was always in good position. the doctor stated that the 5.5 was working effectively and there was no issue with therapy. the impella position alarm was disabled. it was further reported that the patient care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. false alarm: no further details were provided. data logs nor product were available for investigation. since limited clinical details were provided and data logs nor product were returned for investigation, the cause of the false alarm could not be determined. respiratory failure: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3124 | 1220648-2025-48676 | 23677059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677059&pc=OZD | 10/31/2025 | 11/30/2025 |  | 11/30/2025 |  |  | 11/01/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026748527 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Device Damaged by Another Device | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support (msc) and after four days of support during removal, the impella 5.5 device go caught on the valve causing the valve to shift resulting in redo thoracotomy and aortic valve replacement. therapies used after mcs were listed as respiratory and extracorporeal membrane oxygenation (ecmo). the patient survived the explant with an outcome status of critical. the following additional information was noted: there were no issues with the placement position. the insertion was performed under direct vision using direct aorta, so there were no problems with the placement, etc. the impella 5.5 caught on the persival valve during removal. there was no decrease in flow during support and after removal, the patient was managed with central ecmo. | [5 empty] | Patient-specific information a4. weight is unknown. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. after 4 days on support without issue, the pump was explanted. during explant, the pump got caught on the perceval valve, causing the valve to shift, resulting in redo thoracotomy and aortic valve replacement. the patient was then managed on central extracorporeal membrane oxygenation without issue. the cause of the device interaction and heart valve incompetence could not be determined due to no product return and insufficient clinical details.
3125 | 1220648-2025-48677 | 23677065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677065&pc=OZD | 11/01/2025 |  |  | 11/30/2025 |  |  | 11/01/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683637 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ischemia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experience access site bleeding, hematoma, limb ischemia and positioning issue requiring intervention. the patient presented to the emergency department and found to have non-st segment elevation myocardial infarction and also had hypoxia. the patient was intubated, angiogram completed and impella cp recommended which was successfully placed via the right femoral artery. percutaneous coronary intervention was staged. one stent was placed in the left main artery and planned later for the right coronary artery. after the sheath exchange, due to the patient¿s extensive right femoral artery calcification, bleeding was not properly controlled. femstop was placed and the patient was transferred to the unit with the femstop in place. close was attempted but failed. during patient settling in the intensive care unit, the patient¿s pressure dropped and the patient went into pulseless electrical activity arrest with one round of cardiopulmonary resuscitation that achieved return of spontaneous circulation. the patient had bleeding from site and mottled right lower extremity. purge was changed and echo confirmed deep position of the impella cp. however, due to changed status: do not resuscitate, repositioning of the impella cp device was on hold. the patient continued support, and a couple of days later, the patient was noted to be restless with some concern for delirium. seroquel increased. the patient remained off vasoactive infusions with plans to wean impella versus escalation care and transfer to higher level of care. discussion made regarding the purge fluid, currently heparin with change to sodium bicarb and start systemic heparin. one unit of packed red blood cells was transfused overnight for low hemoglobin. hematoma on right groin was noted as stable. the following day the impella cp was explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia, in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. additionally, for hypotension, cardiac arrest, and the resuscitation, the cause was unable to be determined due to insufficient clinical details. however, regarding the access site adverse event, the cause of bleed was most likely patient condition since patient had calcification in the vessels. for the device in wrong position, the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3126 | 1220648-2025-48678 | 23677071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677071&pc=OZD | 11/01/2025 | 01/12/2026 |  | 11/30/2025 |  |  | 11/01/2025 | 11/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759368 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a kink in the pump. the impella was explanted and replaced with a new pump to support the patient. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: pump (b)(6), the purge cassette, and a 14 fr x 25 cm introducer were returned for investigation. the cannula was noted to be kink near outlet cage and also found cannula fracture at the inlet cage. the long introducer also had a severe kink. product damage (cannula kink): clinical details indicate the patient had a tortuous vessel condition, which likely contributed to the kinking and damage observed. the cause of the cannula kink damage was most likely related to challenging patient anatomy (tortuous vessels). device history lot: device batch: 1990606. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks.
3127 | 1220648-2025-48679 | 23677073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23677073&pc=OZD | 11/01/2025 |  |  | 11/30/2025 |  |  | 11/01/2025 | 11/30/2025 |  |  |  |  | 11/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737491 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 101 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. impella placement signal dashed out and placement signal unreliable alarm on. flows still calculating. thrombus present. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. no data logs returned. optical sensor issue: the returned product was returned cut and there was blood/biomaterial around the outflow cage. no kinks were noted in the catheter. due to the product being returned cut, and no data logs available, the root cause of the optical sensor issue cannot be determined. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3128 | 1220648-2025-48643 | 23675786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675786&pc=OZD | 10/30/2025 | 11/28/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720854 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 69 year old male patient for support of shortness of breath. patient in atrial fibrillation and chemically cardioverted. impella site pre-closed with pro-glide x2. impella placed using best practices. angio via repo port showed adequate flow. pulses palpable. angle matched dressing, forward placed sutures. knee immobilizer placed in intesnsive care unit. the plan is for patient to remain on support. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the ischemia was unable to be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3129 | 1220648-2025-48644 | 23675787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675787&pc=OZD | 10/30/2025 | 01/09/2026 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601546 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 in a 65 year old male. upon closing post impella 5.5 insertion, graft anastomosis site began to bleed. physician unable to control bleeding at site with repair. vessel damage occurred at graft site. case was aborted and additional impella 5.5 was inserted via an alternate site. patient then had to undergo vascular repair surgery to repair previous impella site. | [5 empty] | Investigation summary: major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. vessel perforation : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot : 1873913. device history batch: sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3130 | 1220648-2025-48645 | 23675790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675790&pc=OZD | 10/30/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743585 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 69-year-old male patient for st-elevation myocardial infarction support. successful percutaneous coronary intervention of lm and right coronary artery with stents x2. sheath exchanged. groin is clean dry and intact. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: infection: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3131 | 1220648-2025-48646 | 23675793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675793&pc=OZD | 10/30/2025 | 02/24/2026 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1380109 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation | [10 empty] | Malfunction | No | Yes | Display or Visual Feedback Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Controller error and screen goes blank. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 updated; the product has been returned. investigation summary: (b)(4) had a controller error alarm caused by a communication anomaly between the optical pressure measuring unit and the main computer.
3132 | 1220648-2025-48647 | 23675794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675794&pc=OZD | 10/30/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686736 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 82 year old male patient for st-elevation myocardial infarction support. patient was brought to the coronary care unit for aright heart catheterization. decision was made to place an impella cp using best practice techniques. peel away sheath was removed and repositioning sheath secured in place with sutures and tegaderm. knee immobilizer placed. continuous cardiac output swanz cath noted in left femoral vein. pt transferred to the intensive care unit in stable condition. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the hematuria was not determined due to insufficient clinical details.
3133 | 1220648-2025-48648 | 23675797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675797&pc=OZD | 10/29/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755806 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Asystole | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. runs of vt observed echo done and impella pulled back. | [5 empty] | Correction d4, h6 the investigation of the reported failure mode has been completed. no product was received for evaluation. asystole/ tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3134 | 1220648-2025-48649 | 23675801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675801&pc=OZD | 10/30/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680859 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 72 year old male patient for cardiac arrest support. impella was backloaded on the wire and advanced into the ventricle without issue. wire was removed and support was initiated. veno-arterial extra-corporeal membrane oxygenation cannulation was followed with venous and arterial cannula placed in the radiofrequency ablation and right femoral vein antegrade sheaths were also placed for peripheral perfusion. this was followed by removal of the peel away sheath and advancement of the repositioning sheath, which is sutured and dressed. following initiated of mechanical circulatory support levo was able to be weaned to 15mcg/min. patient was then transferred to atrium cmc for advance therapy workup. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. hematoma/ minor bleed: the cause of the injury is most likely use related since large hematoma with active bleeding on arrival from transferring facility. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3135 | 1220648-2025-48650 | 23675804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675804&pc=OZD | 10/30/2025 | 11/25/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026734277 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support, in or post 5.5 insertion, md trying to reposition the pump and sterile sleeve broke off. advised md to cover sterile sleeve and handoff to staff managing the patient. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation code b11 h11: added information regarding the results of the investigation. investigation summary pump was not returned for investigation. device in wrong position: the cause of the positioning issue was not determined without returned product investigation and sufficient clinical information. product damage (anticontamination sleeve split/torn): the cause of the sterile sleeve damage issue was not determined without returned product investigation and sufficient clinical information. device history lot device batch:1964716 device history review the pump (b)(6) passed all post sterile inspection checks.
3136 | 1220648-2025-48651 | 23675809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675809&pc=OZD | 10/30/2025 |  |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548058 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. pt had thready to unidentifiable pulses the whole admission, but r leg which was the impella leg was looking paler this am. us duplex obtained and little to no flow was shown alongside with clot, so patient went down to or for thrombectomy, fasciotomy, and dpc placement. during the surgery, the patient's vascular anatomy was severely calcified and heavily diseased. per the vascular surgeon "the impella is innocent; this injury was not caused by it." during the surgery, there was a fair amount of blood loss, so 3 uprbc and 1 plts given. | [5 empty] | A review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis/ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3137 | 1220648-2025-48652 | 23675811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675811&pc=OZD | 10/30/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698903 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 67-year-old male patient for support of chest pain. patient brought down to cardiac catheterization laboratory for emergent impella and percutaneous coronary intervention. right radial aborted, with focus on radiofrequency ablation. impella backloaded and inserted in normal fashion. while crossing valve, wire pulled back and unable to cross with impella. impella removed and steps repeated with impella insertion in normal fashion. started on p-auto and companion sheath used for single access technique. patient hemodynamically stable throughout case. patient taken to or for escalation to impella 5.5. 5.5 placed and impella cp removed from radiofrequency ablation (rfa). issues due to patient body habitus contributed to hematoma at impella cp site. corrected site issue with covered stent placed in rfa at site of impella insertion. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device or logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/hematoma: clinical states issues due to patient body habitus contributed to hematoma at impella cp site. no further details known about what specific patient issue/condition contributed to the hematoma. the cause of the access site adverse event was not determined due to lack of clinical information. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3138 | 1220648-2025-48655 | 23675818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675818&pc=OZD | 10/30/2025 | 11/28/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026659499 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 77 year old male patient for non-st-segment elevation myocardial infarction support with ultrasound guided radiofrequency ablation access obtained using micro venous puncture kit. decision was made to place impella. impella loaded onto wire and inserted into left ventricle position confirmed. impella started @ on auto, flow 4.0 lpm. single access technique with companion sheath. performed thrombectomy to om2 and cirx, balloon and additional drug-eluting stent (des) placed x2. moved to right coronary artery stented area which is also clotted. slack and peel away sheath removed. repo sheath advanced. tb locked at 90cm, angle matched and site dressed. small track ooze noted, surgicel applied and physician aware. pt transferred to cardiovascular intensive care unit. pt continued to require increasing pressor support overnight. had an episode of pulseless electrical activity (pea), was resuscitated. family was called in and patient was made do not resuscitate. early this morning patient(pt) had another episode of pea and then went asystole. pt expired on support. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. major bleed: the cause of the injury is most likely use related since the stitch was placed yesterday and no further oozing. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3139 | 1220648-2025-48656 | 23675820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675820&pc=OZD | 10/30/2025 | 12/31/2025 |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679380 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 in a 42-year-old male. initial 5.5 pump attempted unsuccessfully for escalation due to patient anatomy. | [5 empty] | The investigation for the reported delivery issue has been completed. the root cause of the delivery issue was determined to be patient condition as the patient had difficult vasculature anatomy and resistance at innominate artery preventing successful delivery of impella. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3140 | 1220648-2025-48657 | 23675822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675822&pc=OZD | 10/30/2025 |  |  | 11/29/2025 |  |  | 10/30/2025 | 11/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694917 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. pre procedure note: this was a no cross 5.5 due to axillary anatomy. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult vasculature anatomy and resistance at axillary artery preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3141 | 1220648-2025-48658 | 23675897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675897&pc=OZD | 10/31/2025 |  |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 |  |  | 12/31/2025 |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468862 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 had placement signal issues. the impella position was verified to be accurate. no patient harm was reported. impella support continues. | [5 empty] | D6b: updated explant date. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. optical sensor issue: clinical details report that the ps map was ~50 mmhg below the patient's bp reading 5 hrs into support (70 mmhg vs 19 mmhg). no further details were provided. data logs were reviewed and the report was confirmed. initially, the ps was reading ~80/40 mmhg when the pump was first inserted in the body and started. the optical sensor calibration was checked, and the g0 went from 16607 to 16629, which is not abnormal. within the first 5.5 hours of starting the pump, the ps map had trended down from ~50 mmhg to 20 mmhg even though the p-level/mc remained constant or increased. in reviewing the first rt log, there was also an instance of the ps stepping down without any change in mc or p-level within the larger trend, however, the pressure trace remained aortic. product was not returned for investigation. since limited clinical details were provided, data logs were inconclusive, and product wasn't returned for investigation, the cause of the optical sensor issue could not be determined.
3142 | 1220648-2025-48659 | 23675902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675902&pc=OZD | 10/31/2025 |  |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751941 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to the patient¿s anatomy. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult vasculature anatomy preventing successful delivery of impella. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3143 | 1220648-2025-48660 | 23675912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675912&pc=OZD | 10/31/2025 |  |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731707 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Gastrointestinal Hemorrhage | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma and gastrointestinal bleed. the patient was transfused one unit of blood. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. major bleed/renal failure: gi bleed and renal failure reported on impella support. the cause of the injury was not determined due to insufficient clinical details. access site adverse event/hematoma: the cause of the access site adverse event was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3144 | 1220648-2025-48663 | 23675938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675938&pc=OZD | 10/31/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756856 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 53-year-old male patient for non-st-elevation myocardial infarction support. impella cp placed per standard procedure without difficulty and into proper position in the left ventricle under fluoroscopy without issues. wire removed and impella cp started in auto. due to hypertension, it was requested for impella to be at p5. percutaneous coronary intervention performed with drug-eluting stents to left anterior descending artery. impella sheath was then removed from the body, broke and peeled away. repositioning sheath placed into groin and sutured with forward tension maintaining angle of insertion. patient transferred to intensive care unit on impella support. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation minor bleed: the cause of the injury is most likely use related since the patient sat up with projectile emesis causing slight oozing at groin site. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review ==> the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3145 | 1220648-2025-48664 | 23675943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675943&pc=OZD | 10/31/2025 | 02/20/2026 |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681737 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. right chest hematoma noted, sandbag applied by rn. | The complainant reported additional information upon request: "this hematoma in this case resolved after sandbag application. the pump is not available for return- the hospital withdrew care on this patient on (b)(6) 2025." |  |  |  |  | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated.
3146 | 1220648-2025-48665 | 23675952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675952&pc=OZD | 10/31/2025 |  |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 |  |  |  |  | 11/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730405 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 56 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. - pt went into vfib, shocked 16 times, cpr, unable to obtain rosc. tod called. | [5 empty] | Correction d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3147 | 1220648-2025-48661 | 23675954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675954&pc=OZD | 10/31/2025 | 11/29/2025 |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766013 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. bleeding at left groin. manual pressure held. bleeding controlled. unable to get flows above p-7 without suction alarms. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Device repositioning and slow flow rate occurred during treatment with the impella 5.5 serial number (b)(6) on related report# 1220648-2025-48634. these issues did not occur with the impella cp serial number (b)(6), therefore, report 1220648-2025-48661 was submitted in error.
3148 | 1220648-2025-48666 | 23675959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23675959&pc=OZD | 10/31/2025 | 11/28/2025 |  | 11/29/2025 |  |  | 10/31/2025 | 11/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610283 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 70-year-old male patient for support of shortness of breath. patient was on levo alone to start, that was increased and dobutamine was also added. patient was taking cardzem and that was stopped. amio was also added and this is when patient became hypotensive and rapid was called. patient was taken to cardiac catheterization laboratorywhere right heart cath was performed and impella was placed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. mechanical interaction: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3149 | 1220648-2025-48633 | 23671231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23671231&pc=OZD | 10/24/2025 |  |  | 11/28/2025 |  |  | 10/24/2025 | 11/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688254 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | [6 empty] | Fever | Low Blood Pressure/ Hypotension | Ischemia | Hematuria | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella cp for mechanical circulatory support (mcs). the patient experienced adverse events, device positioning issue. the impella cp device was explanted and replaced with an impella 5.5 for escalation in care. the patient experienced additional adverse events and subsequently expired. the patient presented to the emergency department with complaint of chest pain, shortness of breath, and hypotension. while in the emergency department, the patient ventricular fibrillation arrested. return of spontaneous circulation (rosc) was achieved, and the patient was taken to the catheterization lab. three shocks were delivered. the impella cp was successfully placed. however, the patient required initiation of veno-arterial extracorporeal membrane oxygenation (va ecmo). percutaneous coronary intervention (pci) was completed with two stents placed to the left anterior descending artery and three stents to the circumflex. thereafter, the patient was transferred to another hospital for a higher level of care and ecmo management. later this same day, hypotension was still noted. 100mcg doses of epinephrine were given ivp due to the hypotension and lack of cardiac pulsatility. additionally noted this day from an echocardiography was the impella cp was at 5cm. the physician elected not to reposition at this time. the following day, no pulses were found bilaterally on lower extremities in which the physician attributed the patient's core temperature of 33 degrees and had no pulse pressure on ecpella (ecmo and impella). the patient was rewarmed. four days later, hematuria was noted; urine output turned red and the patient was febrile. later this day, the patient was transported to the operating room for the impella cp removal and impella 5.5 escalation with continued ecmo support which was completed without incident. the following day the patient experienced epistaxis, pericardial effusion and after the patient arrived back from interventional radiology after having a pericardial drain placed (550mls were removed), heparin was placed on hold due to the effusion and nosebleed. ecpella support continued and approximately three days later ecmo was decannulated and the patient went into atrial fibrillation with rapid ventricular response and needed to be recannulated the same day with possibly placing a right ventricular assist device the following day. however, it appears the rvad was not placed and three days after the ecmo decannulation/recannulation, the patient was noted to be on dexmedetomidine for pneumonia and was in liver failure. two days later the patient expired as care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The product was received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. no issues were found with the returned product. regarding the ischemia, in order to make a cause determination, all of the clinical details would have to be provided. therefore, the cause was unable to be determined due to insufficient clinical details. additionally, hypotension, hematuria, fever and the positioning issue, the cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
3150 | 1220648-2025-48632 | 23671233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23671233&pc=OZD | 10/29/2025 |  |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731171 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 82 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Multiple Organ Dysfunction Syndrome | Pericardial Effusion | [8 empty] | Life Threatening | Required Intervention | Death |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella cp for mechanical circulatory support (mcs). the patient experienced adverse events, device positioning issue. the impella cp device was explanted and replaced with an impella 5.5 for escalation in care. the patient experienced additional adverse events and subsequently expired. the patient presented to the emergency department with complaint of chest pain, shortness of breath, and hypotension. while in the emergency department, the patient ventricular fibrillation arrested. return of spontaneous circulation (rosc) was achieved, and the patient was taken to the catheterization lab. three shocks were delivered. the impella cp was successfully placed. however, the patient required initiation of veno-arterial extracorporeal membrane oxygenation (va ecmo). percutaneous coronary intervention (pci) was completed with two stents placed to the left anterior descending artery and three stents to the circumflex. thereafter, the patient was transferred to another hospital for a higher level of care and ecmo management. later this same day, hypotension was still noted. 100mcg doses of epinephrine were given ivp due to the hypotension and lack of cardiac pulsatility. additionally noted this day from an echocardiography was the impella cp was at 5cm. the physician elected not to reposition at this time. the following day, no pulses were found bilaterally on lower extremities in which the physician attributed the patient's core temperature of 33 degrees and had no pulse pressure on ecpella (ecmo and impella). the patient was rewarmed. four days later, hematuria was noted; urine output turned red and the patient was febrile. later this day, the patient was transported to the operating room for the impella cp removal and impella 5.5 escalation with continued ecmo support which was completed without incident. the following day the patient experienced epistaxis, pericardial effusion and after the patient arrived back from interventional radiology after having a pericardial drain placed (550mls were removed), heparin was placed on hold due to the effusion and nosebleed. ecpella support continued and approximately three days later ecmo was decannulated and the patient went into atrial fibrillation with rapid ventricular response and needed to be recannulated the same day with possibly placing a right ventricular assist device the following day. however, it appears the rvad was not placed and three days after the ecmo decannulation/recannulation, the patient was noted to be on dexmedetomidine for pneumonia and was in liver failure. two days later the patient expired as care was withdrawn. | [5 empty] | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review: medical safety/clinical reviewed the event: the patient presented to the emergency department with chest pain, shortness of breath, and hypotension. the patient experienced ventricular fibrillation cardiac arrest, underwent defibrillation three times, and achieved return of spontaneous circulation. the patient was emergently taken to the cardiac catheterization laboratory. an impella cp device was successfully implanted for mechanical circulatory support. due to ongoing hemodynamic instability, venoarterial extracorporeal membrane oxygenation (va ecmo) was initiated (ecpella support). percutaneous coronary intervention was performed with placement of two stents to the left anterior descending artery and three stents to the left circumflex artery. the patient was pending transfer to a higher level of care facility for ecmo management. following impella cp implantation, the patient required intermittent intravenous push doses of epinephrine (100 mcg) for hypotension and lack of cardiac pulsatility. the following day, bilateral lower extremity pulses were absent. per the physician, the absence of pulses was attributed to core temperature approximately 33°c and lack of pulse pressure while on ecpella support. the patient was rewarmed. a positioning concern was later identified with echocardiography demonstrating the impella cp positioned at approximately 5 cm. cardiothoracic surgery elected not to reposition the device at that time. approximately four days later, the patient developed hematuria and fever. the impella cp was explanted, and the patient was escalated to an impella 5.5 device for continued mechanical circulatory support. during impella 5.5 support, the patient experienced multiple complications, including epistaxis, pericardial effusion, infection, arrhythmias, and organ failure. the patient underwent pericardial drain placement in interventional radiology, with approximately 550 ml of fluid removed. anticoagulation with heparin was held due to pericardial effusion and a bloody nose. the patient developed pneumonia requiring sedation with dexmedetomidine and was noted to be in liver failure. va ecmo had been decannulated three days prior; however, the patient developed atrial fibrillation with rapid ventricular response and required recannulation on the same day. additional advanced therapies, including consideration of right ventricular assist device placement, were discussed. the patient's condition continued to deteriorate, and the patient expired following withdrawal of care. the event story involves two product complaints: impella cp - mdr 1220648-2025-48633: serious injury. impella 5.5 - mdr 1220648-2025-48632: death. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, epistaxis/ pericardial effusion/ organ failure, the cause was not determined due to insufficient clinical details. infection/ arrhythmia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
3151 | 1220648-2025-48634 | 23671319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23671319&pc=OZD | 10/31/2025 |  |  | 11/28/2025 |  |  | 10/31/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722843 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs). post implant same day, suction alarms were encountered a performance level p-5. 1000ml bolus fluids were given with no change. the interventional cardiologist repositioned the impella 5.5 device pulling back 3cm. the patient was unable to get flows above p-7 without suction alarms. an echocardiography showed akinesis left ventricle. later overnight, suction alarms triggered. the physician arrived at bedside to reposition the impella device using ultrasound guidance. the patient was evaluated throughout the day with hopes to reduce vasopressors. however, approximately four days later the patient expired. | [5 empty] | Patient-specific information a4. weight is unknown. additionally, the exact date of death is unknown; however it has been captured as the date of explant. device status. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. additionally, data log was not provided or available on the remote server. however, the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, suction issue, the cause was not determined without returned product investigation and sufficient clinical details, including data log. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Medical safety/clinical reviewed the event. a 72-year-old male patient was transferred from another hospital status post st-elevation myocardial infarction (stemi) with an intra-aortic balloon pump (iabp) in place and multivessel coronary artery disease. past medical history included known coronary artery disease with placement of two bare-metal stents (bms) to the right coronary artery (rca) in 2024. an impella cp was percutaneously inserted via the left femoral artery for mechanical circulatory support. during impella cp support, bleeding at the left groin access site was observed and controlled with manual pressure. after approximately 17 hours of impella cp support, the patient demonstrated inadequate hemodynamic support requiring three vasopressors. the clinical team elected to escalate support to an impella 5.5 device, which was successfully implanted. the impella cp was removed prior to completion of the procedure. hemostasis at the access site was achieved with manual pressure, and three-point fixation was applied. on post-implant day 0 with the impella 5.5, suction alarms were noted at performance level p-5. a 1000 ml fluid bolus was administered without resolution of the alarms. the device was repositioned by withdrawing approximately 3 cm; however, flows above p-7 could not be achieved without recurrent suction alarms. echocardiogram reportedly demonstrated left ventricular akinesis per nursing documentation. the patient remained on multiple inotropes and was in cardiogenic shock, stage e. the clinical team continued monitoring and supportive care. approximately four days later, care was withdrawn, and the patient expired. the patient's underlying condition contributed most to the demise outcome. the patient required multiple inotropes and remained in advanced (stage e) cardiogenic shock, indicating hemodynamic compromise. the recurrent suction alarms may be associated with low left ventricular preload, left ventricular dysfunction (akinesis), and ongoing cardiogenic shock, which are conditions noted in the instructions for use as potential causes of for suction. h6. investigation type/finding/conclusion remains unchanged.
3152 | 1220648-2025-48635 | 23672126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23672126&pc=OZD | 10/29/2025 |  |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754281 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 80 KG |  |  | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implante with an impella cp experienced saturated flow and flat motor current. no patient harm was reported. the patient was successfully weaned and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for investigation. saturated/high pump flow: the returned device showed no biomaterial evidence or bearing wear defect. pump was ran for around 22 hours with no saturated flow issue reproducing. there was no defect found with return product. data logs were reviewed and saturated flows, intermittent low flows seen at respective p-levels during the cp pump support. the cause of saturated flow was likely due to biomaterial buildup based on log analysis. device history review: the pump passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3153 | 1220648-2025-48636 | 23672128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23672128&pc=OZD | 10/29/2025 | 11/21/2025 |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653724 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematoma | Hemoptysis | Sepsis | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had sepsis, anemia, hematoma and hemoptysis during impella 5.5 support. patient showed signs of sepsis during support. additionally, patient's hemoglobin levels trended down for which 1 unit of packed red blood cells was given. the next day, it was noted that the day prior, patient's right chest showed a small hematoma. patient also had significant bloody secretions from trachea. pump was explanted a few days later. patient was stable. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. sepsis: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hemoptysis/ anemia: the cause of the injury was not determined due to insufficient clinical details. hematoma: ultrasound of r chest yesterday showed small hematoma, 2cmx1cm. the cause of the hematoma is determined to be patient condition because the hematoma is noted at the non-impella site. device history review: device with passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3154 | 1220648-2025-48639 | 23673921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23673921&pc=OZD | 10/29/2025 | 11/25/2025 |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026739626 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient bleeding postoperatively. required 3 units rbcs to reach hgb 8.1 urine is red. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. no product was returned for investigation. data logs were returned for analysis, but were only available till the afternoon of 12 nov 2025. pump suction: clinical details mention suction alarms occurring overnight on 29 oct 2025. cvp setting was 0. the alarms resolved after prbc transfusions. the patient was on ecmo support during the case. data log analysis confirmed suction alarms occurring during the reported timeframe. placement signal was fluctuating but showed no abnormal trends, while there were drops in pump flow for the entire case. impella position unknown alarms were seen during the suction events. no other abnormalities were seen. the root cause of the pump suction was most likely patient condition related based on the provided clinical details and data log analysis. major bleed, hematuria: clinical details mention that the patient was bleeding postoperatively after impella and ecmo intubation and required (b)(4) units rbcs to reach hgb 8.1. red urine output was also observed during the case. the root cause of the injury was not determined due to insufficient clinical details.
3155 | 1220648-2025-48640 | 23673966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23673966&pc=OZD | 10/29/2025 |  |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757947 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. hematoma formed at the groin site of the impella cp. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d4, g1. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
3156 | 1220648-2025-48641 | 23673978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23673978&pc=OZD | 10/03/2025 | 11/28/2025 |  | 11/28/2025 |  |  | 10/30/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108097 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller in a 70 year old female patient. aic failure noted during priming and z.eroing of impella 5.5. switched to aic after pump inserted and started pump failure noted on aic. decision made by surgeon to explant current 5.5 and replace with new 5.5. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the console (b)(4) was tested, but the reported issue could not be reproduced. no performance irregularities were observed during testing. additionally, the console was manually power-cycled several times without any issues. no power supply interruptions were detected between the pbm, 4-in-1, and impellatronic pca assembly. the console was tested with both the pump and purge cassette for several hours, and no abnormalities were observed. furthermore, the pump was repeatedly disconnected and reconnected during testing, and each time the pump was successfully recognized. the root cause of the issue ¿device insertion and startup resulted in a 'device failure' alarm on aic¿ could not be determined, as the reported behavior could not be reproduced during testing. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
3157 | 1220648-2025-48638 | 23673981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23673981&pc=OZD | 10/03/2025 | 11/28/2025 |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED INC. |  |  | GM |  |  |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. alarm cannot be turned off, minimum signal for complete cure, low value alarm frequency. therefore, even if the pump position monitoring stop is performed, an alarm will be generated afterwards. the incident was confirmed at the (b)(6) hospital, where he was transferred. | [5 empty] | The device was returned. data analysis: there were significant number of placement signal alarms at the latter portion of the case. root cause: returned product performed as expected and no problems were found. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | Device received, inv in progress: the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3158 | 1220648-2025-48642 | 23674002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23674002&pc=OZD | 10/30/2025 |  |  | 11/28/2025 |  |  | 10/30/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722505 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female | 78 KG |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during pump preparation an impella cp had a purge pressure low alarm. the alarm resolved without intervention. the patient was successfully supported by the pump and no patient harm was reported. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. purge cassette returned for evaluation. pump not returned for evaluation. the returned cassette was cut open and attempted to run through de-air but a blockage was observed resulting in the purge disc to pop. there was dried yellow purge fluid beyond the purge disc so the pressure line with the purge disc was then cut off. then the system was purged with colored fluid and the system was closed and pressurized to >1050mmhg for 10 minutes and no leak was observed. the cause of the purge issue was not determined since the pump was not returned for review and no defect found on the cassette.
3159 | 1220648-2025-48637 | 23674004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23674004&pc=OZD | 10/03/2025 | 11/28/2025 |  | 11/28/2025 |  |  | 10/29/2025 | 11/28/2025 |  |  |  |  | 11/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597983 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. sensing issue, no s/symptoms, explant to lvad/revision/replace console: unable to silence alarm for ps, no s/symptoms, false alarm/buzzer. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H3: updated. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. optical sensor issue: clinical details report that roughly 2 weeks into support the low pressure (ps) was diverging from the patient's radial artery blood pressure (bp) measurements. ps low alarms began to trigger. the patient remained on support and was bridged to a left ventricular assist device (lvad) 23 days later. data logs were reviewed and the reports were confirmed. returned product was investigated and the 5s parameters were within expected ranges and the pump passed laser continuity testing. the pump was connected to a console and run in a bucket of water for 5 hours. the pump reproduced what was noted in the field: optical 5s parameters remained stable while the ps gradually trended from -25 mmhg to -35 mmhg. based on the evidence from field data logs and returned product evaluation, the cause of the optical sensor issue was determined to be sensor wear. the first indication of sensor wear occurred 23.75 days into support, which is within the design life of the product. | Correction: d4. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
3160 | 1220648-2025-48613 | 23655810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23655810&pc=OZD | 09/25/2025 |  |  | 11/26/2025 |  |  | 09/25/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697167 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 58 KG | Non Hispanic | White | Death | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  | Cardiac Perforation | [12 empty] | Disability | Death | Life Threatening | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported a patient in cardiogenic shock from an acute myocardial infarction implanted with an impella cp device for mechanical circulatory support experienced a left ventricle perforation and subsequently expired. the patient was admitted to telemetry floor for non-st-segment elevation myocardial infarction. new st changes occurred in the afternoon with subsequent complete heart block and cardiac arrest. the patient was in cardiogenic shock (stage e) and emergently brought to catheterization lab for an impella cp device placement. the impella cp placement was noted to be very difficult due to ongoing cardiopulmonary resuscitation and intubation. the patient continued to decompensate with low impella flows thought to be related to right ventricular failure from massive the right ventricular infarct. the decision was made by multiple interventionalists to proceed with the impella rp flex placement. upon bipella support, the patient remained "very" unstable. echocardiogram showed effusion; pericardiocentesis done. imaging showed that impella cp had perforated through the pericardium. the patient continued to decline, cardiac surgery was consulted and declined intervention. the patient ultimately expired; time of death was called. | [5 empty] | Additional information was received and noted the following: initially the low flows were presumed to be related to rv failure as due to the urgent nature of the situation, an echo was not yet completed so a decision was made to implant an rp flex. retrospectively, the md now assumes the initial low flows were related to the perforation. the site of the perforation was thought to be due to friable tissue due to the infarct. the cause of the death is the perforation, rather than care withdrawal. after the impella rp flex was implanted, an echo was done showing the perforation and cardiac surgery was consulted. the patient was deemed not a candidate for further escalation by 3 interventional cardiologists, a cardiac surgeon, an intensivist and a heart failure physician. all heroic measures were stopped; the impella was turned off and the patient expired in the catheterization lab. it was determined that the patient would not have survived even a trip up to the intensive care unit to be with family. medical safety review. medical safety reviewed the event. this event describes a patient who had an impella cp implanted during active cpr which is believed to have led to lv perforation during the insertion. prior to confirmation of the lv perforation an impella rp flex was implanted for right ventricular infarct with right ventricular failure. upon imaging it was noted that the impella cp had perforated the left ventricle. intervention to correct the lv perforation was declined which led to the patient's death. since the impella cp did perforate the left ventricle there is not enough information to dissociate the impella cp from the demise of the patient. due to the perforation and the decision to decline intervention the impella rp flex can be dissociated from the patient's death. device status: the impella cp device was not received from the customer at the time of this mdr writing; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Complaint coding was updated to better align with the reported details. the updated impacted the medical device problem coding. therefore, section h6 medical device problem code was repopulated to reflect the update. note: h6 heath effect ¿ clinical/impact coding remains unchanged as previously reported. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, data logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. impella cp placement was very difficult due to ongoing cardiopulmonary resuscitation and intubation. the patient was critically ill, with low impella flows thought to be related to right ventricular failure from a massive right ventricular infarct. despite escalation to bipella support, the patient remained unstable. imaging revealed pericardial effusion and perforation, and the patient ultimately expired after cardiac surgery declined intervention. regarding the low or blocked pump flow, the pump was not returned for investigation. however, data log was analyzed and shows the pump was unable to achieve flow greater than 2 l/min, with active suction alarms. there were some pump positioning/ventricular alarms triggered while the pump was set to p-9 and not generating enough flow. no motor current spike was reported on the data log. the cause of the low flow issue was most likely use-related, since the presence of ventricular alarms on the data log and confirmation that the pump was positioned deeply support that the pump was in an improper position. mechanical interaction with tissue-cardiac perforation cause was most likely use-related since placement of the pump was performed during active cpr, and difficult to place or position the pump cause was most likely related to patient movement as placement was initiated during active cpr. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3161 | 1220648-2025-48614 | 23655986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23655986&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711206 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support and while on support the patient was having episodes of ventricular tachycardia. later care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b11 and b13. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. ventricular tachycardia: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3162 | 1220648-2025-48615 | 23656079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23656079&pc=OZD | 10/10/2025 | 11/25/2025 |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713994 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella 5.5 in a 82 year male patient for support of dyspnea on exertion. axillary sheath inserted into graft and secured with graft lock x2. once in ascending aorta, catheter exchanged to 65cm 5fr pig and crossed into left ventricle. wire exchanged to impella 018 wire. catheter removed and impella pump backloaded and advanced over wire into position. transesophageal echocardiogram showed impella well positioned, angled towards apex and measuring 4.5cm. jackson-pratt drain to site. impella secured and locked at 52cm. blood products started before being taken to operating room. impella weaned to p2 and successfully explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B3 updated date of event as it was entered incorrectly on the initial report that was submitted. e1 multiple fields were added as the information was omitted from the initial report that was submitted. e4 added as selection was omitted from the initial report that was submitted. g3 date should of been 10-oct-2025 on the initial report that was submitted. h6 revised health effect - clinical code from e0303 to e1302. medical device problem code was reported incorrectly on the initial report that was submitted.
3163 | 1220648-2025-48617 | 23656596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23656596&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685635 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Great Vessel Perforation | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. pt w/ known peripheral vascular disease. amputated toes on both extremities prior to admission. absent distal flow to bilateral lower extremities. unsure pulses prior to admission at outside hospital. discussed w/ dr. (b)(6). reducing vasoconstricting medications. plan to have vascular team evaluate bilateral lower extremities. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | 6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. ischemia/thrombosis: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. vessel perforation (peripheral): impella removed¿leaving wire in place with device pre-closed with perclose x2. a sheath was inserted into preclosed site and iliac artery perforated. the cause of the vessel perforation is determined to be use issue because iliac artery perforated after the sheath was inserted. device history review: device with passed all post sterile inspection checks.
3164 | 1220648-2025-48616 | 23656628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23656628&pc=OZD | 10/27/2025 | 11/26/2025 |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026722339 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. placement signal not reliable and absent placement signal alarms present. patient outcome, unknown. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller. investigation summary ==> suction: clinical details report ongoing suction alarms throughout support. on (b)(6), it was noted that the patient's bp was low. on (b)(6), the placement signal (ps) was reported to be lower than the patient's blood pressure (bp), potentially contribution to the alarms. no troubleshooting methods were reported. data logs were reviewed and the alarms were confirmed. there were over 2000 triggers of diastolic suction throughout the case. the ps did trend down through the first week of support, however, from that point forward it was generally steady around 75/35 mmhg. the user did not make any ps adjustments either. though signal not reliable (snr) became widespread as the case went on, the ps pressure trace appeared normal when reviewing rt logs. suction alarms on (b)(6) were reviewed and the criteria for the trigger were being met. however, the alarms were triggering when mcmin values were will within the expected range for p-8. the alarms resolved for a period when there was a step up in the ps without a change in mc values. as for (b)(6), the alarms were confirmed to be meeting the triggering criteria again, however, in this case, there was a trend down in mcmin values toward the lower end of the expected range for p-8. since no troubleshooting methods were provided, data logs were inconclusive, and product wasn't returned for investigation, the cause of the suction could not be determined. optical sensor issue 1 (psnr alarm): on (b)(6), a placement signal not reliable (psnr) alarm was reported. data logs were reviewed and the psnr alarm was confirmed. on (b)(6) the psnr alarm triggered when contrast began to oscillate between +/-3200, snr, gain, and light all dropped, and lamp increased. the alarm resolved after 9 hours and the ps returned for the rest of the available logs. this pattern is indicative of a console issue. upon review of data logs, it is apparent that the console is the potential cause of the psnr alarm. optical sensor issue 2 (ps reading lower): clinical details report that the ps was reading lower than the patient's bp on (b)(6). no values for the patient's actual bp were provided. as mentioned above, there was a slight trend in the ps during the first week of support, however, it remained generally stable from that point forward. additionally, there were no ps adjustments made by the user. on the reported date of the discrepancy, the ps was reading 89/46 mmhg. there were no bp readings provided to compare to. the real time event log at case start appears to show the ps being zeroed from 14 mmhg, however, event logs weren't available to review g0 values. product wasn't returned for investigation. since limited clinical details were provided, event logs weren't available, and product wasn't returned for investigation, the cause of the optical sensor issue (ps reading lower) could not be determined. device history lot ==> device lot: 1958862. device history review ==> pump sn (b)(6) passed all post-sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b explant date updated.
3165 | 1220648-2025-48618 | 23656870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23656870&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757758 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 89 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient decompensated at an outside hospital and was transferred to the complainant hospital for escalation of care. upon arrival, the patient was struggling to breathe and was immediately intubated. the patient was then implanted with an impella cp due to the patient's surgical aortic valve stenotic would not accommodate an impella 5.5. the impella cp was successfully implanted and support initiated without complications. on support day 2, the patient was noted to have decreased and concentrated urine output. the patient was place on continuous renal replacement. it was noted that the patient had an acute kidney injury related to late presentation cardiogenic shock. on support day 4, the pump had alarms for suction and placement signal unreliable. the performance (p) level was decreased from p-8 to p-6 and resolved the suction alarms. an echocardiogram was used to verify impella positioning. on support day 5, it was noted that the impella was slightly deep when imaged with echocardiography. the family decided to withdraw care from the patient. the patient's care was withdrawn, and the patient expired. | [5 empty] | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed | Section e1 address was omitted in initial report and has been updated. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. optical signal issue: clinical reported placement signal not reliable alarm noted. the cause of the issue was not established as no product or data logs were returned for analysis. pump suction: clinical reported suction alarm noted. the cause of the issue was not established as no product or data logs were returned for analysis. minor bleed/access site adverse event: minor oozing reported. the cause of the issue was not established as limited clinical information was provided.
3166 | 1220648-2025-48619 | 23657092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23657092&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731319 |  | 000000000001000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. bleeding at insertion site, plan to take patient back to or to find bleeder. | [5 empty] | Correction b1, d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. access site adverse event: the cause of the bleed was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3167 | 1220648-2025-48620 | 23657303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23657303&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713329 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 89 KG | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. on review of alarm history, ¿placement signal unreliable¿ alarm noted. rn stated no change in placement signal waveforms or impella flows/motor current. no change in pt condition at time of alarm. happened after suction alarm. dr. (b)(6) turned down flow from p-8 to p-6 and alarm resolved. discussed that ps and motor/flows are independent of one another and reminded not to use ps as hemodynamic monitor. staff verbalized understanding. | [5 empty] | After further review the reportability changed to malfunction. b1, h1, and h6 were updated. the death will be reported in mwr-24112025-0001979439. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3168 | 1220648-2025-48622 | 23657665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23657665&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728718 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp after coding in the intensive care unit. the impella cp was implanted via the right femoral artery without complications; support was initiated and the patient stabilized. on support day 2, the patient experienced a few hypotensive events during impella cp support and required an increase in pressors. additionally, the purge solution was changed to bicarbonate and heparin was withheld. the patient was transfused with 7 units of packed red blood cells due to cardiotomy bleeding post cardiopulmonary resuscitation following cardiac arrest the previous day, prior to impella use. of note, there was no bleeding or oozing from the impella access site, and the impella was noted to have good flows. on support day 3, heparin continued to be withheld due to cardiotomy and surgical site bleeding. the patient was noted to have improved hemodynamically and was being weaned from pressor and impella support. additional blood products were administered overnight and throughout the morning. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A2, a4, a5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: the device was not returned for evaluation. hypotension: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
3169 | 1220648-2025-48624 | 23658449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23658449&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026746429 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was put on impella cp support to unload and for left heart catheterization. the impella was implanted in left femoral artery (multiple extracorporeal membrane oxygenation (ecmo) cannulation attempts in this groin as well). it ran in auto and was then turned to p4. the team was unable to keep ecmo flows above one liters near end of case. a large hematoma was noted at left groin site. pressure was held. the team decided to implant the impella using a larger non-impella sheath. the team is aware and understands that this is not recommended. the patient's lactic continued to climb and increasing pressor support. they were unable to keep ecmo flows up and the decision was made to send the patient to the intensive care unit for family to withdraw care. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3170 | 1220648-2025-48623 | 23658824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23658824&pc=OXD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OXD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026744001 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. washout today at impella site- clot evacuated with no active bleeding noted. holding anticoagulation for 24hrs. continued milrinone therapy and aggressive mobility as bridge to advanced therapies. hematoma noticed as well. | [5 empty] | H6: added type of investigation codes b11 and b13. h11: added the results of the investigation. investigation summary: no product was returned for investigation. thrombosis: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3171 | 1220648-2025-48625 | 23658844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23658844&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744143 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 147 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Multiple Organ Dysfunction Syndrome | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient spiked a temperature overnight so he was placed on airborne precautions and full panel of cultures were ran including covid rule out. furthermore, it was reported that the patients kidney and liver were not functioning well. additionally, the patient was reported having nose bleeds and was off anticoagulation. the patient was successfully weaned and survival outcome was noted as survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6: added type of investigation codes b11 and b13. h11: added the results of the investigation. investigation summary: no product was returned. fever, organ failure, epistaxis: the cause of the injury was not determined due to insufficient clinical details. hematoma: the cause of the injury was not determined due to insufficient clinical details. device history review: device (sn: (b)(6) passed all post sterile inspection checks.
3172 | 1220648-2025-48627 | 23659581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23659581&pc=OZD | 10/27/2025 | 02/25/2026 |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026636898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Renal Failure | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. post procedure: after completing thrombectomy, the patient continued to be unstable and dr. (b)(6) elected to remove the iabp and place impella cp for hemodynamic support in the setting of cgs. upon placement of abiomed 0.018 wire into the lv the patient converted into a lethal rhythm and cpr was initiated. md backloaded the impella cp onto the abiomed 0.018 wire and attempted to place the device. md lost wire support for placement and elected to completely remove the impella device and rewire the device with a v-18 wire to get the cp placed across the av. once the device was placed and wire was removed the impella ramped up appropriately, patient continued to have cpr and multiple shocks until rosc was achieved. (b)(4) - on (b)(6) 2025 hemolysis and renal failure- patient remains in severe cardiogenic shock. concerns for hemolysis with dark urinary output and high ldh. attempting to place lines for crrt now but having difficulty finding a vein. patient is sedated but follows commands. cv surgery is aware and is discussing escalation options such as 5.5/ecmo. bedside is attempting to wean pressers today. p-6 maintained for now, 3.1 lpm, with no alarms or issues at this time. patient expired. | Additional information has been provided upon request: "i am not certain as to what interventions were performed regarding the alleged complaint events, but typically, echo is obtained to assess for malpositioning in the setting of hemolysis concerns. if warranted, echo guided repositioning is typically attempted at bedside by ic or surgeon (if 5.5). i believe crrt was maintained throughout the treatment course. this pump is unfortunately unavailable for return." |  |  |  |  | Correction: h1. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. | H6: added type of investigation code b13, b11. investigation summary: the device was not returned for evaluation. clinical symptoms (renal failure, cardiac arrest, hypotension, major bleed): the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood (hemolysis): the pump was not returned for investigation. clinical details indicate the patient developed hemolysis, evidenced by dark urine output and elevated ldh, particularly noted on (b)(6) 2025, while in severe cardiogenic shock and on impella and crrt support. data log was available until (b)(6) 2025, 3:30 pm, showing low pump flow during the case, which resolved within 30 minutes and then ran without any reported abnormalities. the cause of the hemolysis issue was not determined without the returned product and sufficient clinical details, including full data logs. device history lot: device batch: 1888611. device history review: the pump (b)(6) passed all post sterile inspection checks.
3173 | 1220648-2025-48628 | 23659844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23659844&pc=DQX | 10/27/2025 | 02/23/2026 |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 | [6 empty] | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported additional information upon request: "i am not certain as to what interventions were performed regarding the alleged complaint events, but typically, echo is obtained to assess for malpositioning in the setting of hemolysis concerns. if warranted, echo guided repositioning is typically attempted at bedside by ic or surgeon (if 5.5). i believe crrt was maintained throughout the treatment course. this pump is unfortunately unavailable for return." | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. post procedure: after completing thrombectomy the patient continued to be unstable and dr. gilbert elected to remove the iabp and place impella cp for hemodynamic support in the setting of cgs. upon placement of abiomed 0.018 wire into the lv the patient converted into a lethal rhythm and cpr was initiated. md backloaded the impella cp onto the abiomed 0.018 wire and attempted to place the device. md lost wire support for placement and elected to completely remove the impella device and rewire the device with a v-18 wire to get the cp placed across the av. once the device was placed and wire was removed the impella ramped up appropriately. patient continued to have cpr and multiple shocks until rosc was achieved. hemolysis and renal failure- patient remains in severe cardiogenic shock. concerns for hemolysis with dark urinary output and high ldh. attempting to place lines for crrt now but having difficulty finding a vein. patient is sedated but follows commands. cv surgery is aware and is discussing escalation options such as 5.5/ecmo. bedside is attempting to wean pressers today. p-6 maintained for now, 3.1 lpm, with no alarms or issues at this time. |  |  |  |  | B5 updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 death and date of death were inadvertently selected. patient death has been captured in report mwr (b)(4). the investigation of the reported failure mode has been completed. no product was received for evaluation. clinical symptoms (arrhythmia): the root cause of the injury was unable to be determined due to insufficient clinical details. unable to deliver or advance: the root cause was most likely use-related, due to patient movement and instability during active cpr. device history review: the device passed all post sterile inspection checks.
3174 | 1220648-2025-48626 | 23660122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23660122&pc=OZD | 10/27/2025 | 11/25/2025 |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656712 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5in a 71-year-old female patient for support of cardiac arrest. she decomensated and coded after she was in the intensive care unit. the patient was taken back to the operating room where the doctor (md) placed another vein graft. the patient was placed on central veno-arterial extra-corporeal membrane oxygenation and the decision was made to also place another vein graft impella 5.5 axillary. the md was not able to get the impella 5.5 to track through the axillary artery so the decision was made to place the impella cp axillary. impella 5.5 would not track through axillary artery. decision was made to move to a cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult vasculature anatomy and resistance at axillary artery preventing successful delivery of impella. device history review: device (sn: (b)(6) passed all post sterile inspection checks.
3175 | 1220648-2025-48629 | 23660192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23660192&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744162 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 69 KG | Non Hispanic |  | Injury | Yes | No | Malposition of Device | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella 5.5 support. it was reported unable to titrate above p4 without ventricular tachycardia. intensivist attempted repositioning but struggled. elected to keep device as is for now because the patient is perfusing well. the patient remains on support. additional information 14-jan-2025: the patient was successfully weaned from the impella 5.5 during the transplant procedure and underwent a successful heart transplant. the patient survived the procedure. | The complainant reported a patient on impella 5.5 support. it was reported unable to titrate above p4 without ventricular tachycardia. intensivist attempted repositioning but struggled. elected to keep device as is for now because the patient is perfusing well. the patient remains on support. |  |  |  |  | This report is being submitted to update explant date and to report patient outcome. b5 is being updated to include new information that was not included in the initial report. the revised b5 provides a complete event narrative. d6b is being updated to include explant date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product was returned. ventricular fibrillation: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review device (b)(6) passed all post sterile inspection checks.
3176 | 1220648-2025-48630 | 23660285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23660285&pc=OZD | 10/27/2025 | 11/25/2025 |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026707474 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Anemia | Bradycardia | Hematoma | Pain | Hematuria | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 in a 56 year old male patient for support of dyspnea, fatigue, and lower extremity edema. patient noted to have tricuspid annular plane systolic excursion mild mitral regurgitation. doctor cut bevel on 10mm graft and placed on axilary using standard technique, short tunnel. pump initially in ventricle and then deflected towards septum. doctor successfully repositioned impella, with transesophageal echocardiogram. successful implant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: updated explant date h6: added code e1302 and a0709 based additional information in the complaint file. additional information: the patient experienced hematuria. urinary output - greater than 30 cc/hr and concentrated/amber. additionally, the device exhibited intermittent placement signal unreliable alarms. the team disabled placement signal notifications as they were intermittently getting ¿placement, signal unreliable¿alarms. they conducted more frequent echocardiograms as needed; otherwise, team was unconcerned with device moving.
3177 | 1220648-2025-48631 | 23660533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23660533&pc=OZD | 10/27/2025 |  |  | 11/26/2025 |  |  | 10/27/2025 | 11/26/2025 |  |  |  |  | 11/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704264 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 86 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in non-st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support for a percutaneous coronary intervention (pci). the patient decompensated during dialysis treatment and became increasingly tachypneic and was taken to the emergency room. the impella was successfully implanted via the right femoral artery without complications. during the pci, it was reported that the pump had an alarm for placement signal not reliable. the impella position was imaged and confirmed to be in proper position. the alarm was disabled. this was followed by a loss of placement signal and left ventricle signals. impella flows remained as expected at 3.6l/min throughout the procedure. the patient was noted to be hemodynamically stable throughout the procedure. no patient harm was reported in relation to this event. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The device and logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the optical sensor issue: the clinical stated that initially after starting the impella, there were no alarms and there were appropriate waveforms. during a pci procedure, the impella began alarming placement signal unreliable with a loss of ao and lv signals. due to a lack of clinical information, product or data logs returned, the root cause of the optical sensor issue cannot be established. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3178 | 1220648-2025-48591 | 23643871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23643871&pc=OZD | 10/27/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1637336 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in japan had an automated impella controller (aic)/impella cp and extracorporeal membrane oxygenation supporting a patient admitted in with ventricular fibrillation. during the support of over one week, there was issues with the display on the aic. there were flashing lights but the console was not swapped out and support was continued till death from heart failure. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3179 | 1220648-2025-48592 | 23644191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23644191&pc=OZD | 10/26/2025 | 11/24/2025 |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760932 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic |  | Injury | Yes | No | False Alarm | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp in an 82-year-old male for mechanical circulatory support for limb ischemia. patient remains on bicarb, lidocaine, dobutamine, vasoconstrictor and high dose levosimendan. vascular following patient due to decreased flow to lower extremities. on continuous renal replacement therapy for acute kidney injury. still creating a lactate but down trending. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3180 | 1220648-2025-48593 | 23644780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23644780&pc=OZD | 10/26/2025 | 11/24/2025 |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657546 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a a 67 year old male for st-elevation myocardial infarction support. a golf ball sized hematoma was noticed at insertion site. manual pressure held, however bleeding worsened. angio was taken from repo sheath port and small amount of contrast was visualized at arteriotomy. shock call was placed, escalation was declined due to bleeding complication. vascular surgery was called but ultimately could not support until impella is removed. patient required rotational thromboelastometry of blood and blood product. after approximately one hour of manual pressure and with the stop of bivalirudin, bleeding slowed and dsg was applied. disseminated intravascular coagulation was suspected and patient continued to deteriorate. care withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. asae/major bleed/hematoma: noted golf ball sized hematoma at insertion site. manual pressure held, however bleeding worsened. after approximately one hour of manual pressure and with the stop of bivalirudin, bleeding slowed. the cause of the access site adverse event was user related as the bleeding was slowed by stopping anticoagulation. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3181 | 1220648-2025-48594 | 23645026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23645026&pc=PYX | 10/26/2025 |  |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026639583 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. plan is to monitor hr as the patient has intermittent pauses while continuing to support the patient with pressors. patient has heart block. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. heart block : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3182 | 1220648-2025-48595 | 23646206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23646206&pc=OZD | 10/26/2025 | 11/25/2025 |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717701 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 71 year old male patient for support during cardiac arrest. during cardiopulmonary resuscitation (cpr) was in progress and return of spontaneous circulation was achieved. the impella sheath was able to place, and while advancing the impella, the patient went back into pulseless electrical activity. cardiopulmonary resuscitation was performed but patient subsequently expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3183 | 1220648-2025-48596 | 23646503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23646503&pc=OZD | 10/25/2025 | 11/25/2025 |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760499 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in a 65 year old female for seizure with st elevations support, which was noted on electrocardiogram. impella cp implanted post percutaneous coronary intervention of the left anterior descending artery. nurse described bleeding at the site with a drop in hemoglobin and blood given. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: bleeding at the site with a drop in hgb. multiple mds consulted, and blood given. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3184 | 1220648-2025-48597 | 23646583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23646583&pc=OZD | 10/26/2025 | 11/25/2025 |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729614 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp in an 84-year-old patient for primary percutaneous coronary intervention support. during support of impella cp the patient was off heparin for bleeding at impella groin site. positive distal pulses, no hematoma. echocardiogram done today, not yet resulted. per rn possible plan to remove impella monday. the pump was explanted. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: bleed around the sheath was resolved by applying manual pressure. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3185 | 1220648-2025-48598 | 23647051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23647051&pc=OZD | 10/26/2025 |  |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595749 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. - slight oozing still noted at impella site with transfers. | [5 empty] | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data logs were returned for analysis. low or blocked pump flow: data log analysis confirmed low flows at the beginning of the case, which coincided with impella in aorta alarms. impella position unknown and suction alarms were seen towards the end of the case. impella flow reduced alarm was also triggered, which resolved with the impella in aorta alarm. clinical details mention that ¿impella flow reduced¿ alarms were shortly seen after initiation of support, along with "impella in aorta" alarm. the root cause of the low pump flow was most likely user related owing to wrong position of the pump. device in wrong position: data log analysis confirmed impella in aorta alarms. clinical details mention that the pump was advanced in small increments under fluoroscopy till it was brought to the proper position after the alarms triggered. the peel-away sheath was not removed till that point, after which the sheath was replaced with a repo sheath, sterile sleeve was attached and the tb valve was locked. the root cause of the device being in the wrong position was most likely user related as the wrong position was confirmed under fluoroscopy, and the sheath was not locked in place prior to the positioning alarms. access site adverse event: clinical details mention that slight oozing was noted at the impella site when the patient was transferred to the bed and later to another hospital, with activated clotting time being above 180 at 247. the root cause of the access site bleeding was user-related as the act was above the recommended range during transfer of the patient. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3186 | 1220648-2025-48600 | 23647427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23647427&pc=OZD | 10/26/2025 |  |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764048 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp and impella rp support had ventricular tachycardia and was converted back to normal rhythm by cardioversion. the patient continued with impella support until the impella was successfully explanted as planned. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the tachycardia : in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3187 | 1220648-2025-48602 | 23647709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23647709&pc=OZD | 10/27/2025 | 12/31/2025 |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736485 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella 5.5 support. while on support the patient has been on and off anticoagulation due to bleeding issues. the automated impella controller started alarming purge system blocked. tissue plasminogen activator started with plan to exchange pump on if purge flow did not return. the patient remains on support. | [5 empty] | D6b updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined due to insufficient clinical details. purge system blocked: the cause of purge system blocked was due to biomaterial buildup based on log and clinical review. device history review: the device passed all post sterile inspection checks.
3188 | 1220648-2025-48603 | 23647892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23647892&pc=DQX | 10/27/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Injury | Yes | No | Malposition of Device | Deformation Due to Compressive Stress | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  | Pericardial Effusion | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. buddy wire used to straighten anatomy. 20 minutes after procedure was successfully completed & while waiting for icu space patient decompensated. bp dropped and cvp raised. echo showed pericardial effusion. md decided to perform pericardial window. after removing 300cc of bright red blood, decided to convert to sternotomy and explore cavity. no active bleeding found, only 2 small clots. 1 prbc & cell saver given during the procedure. md conclusion was that most likely injury was caused by one of the wires. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction h6. | H6: updated codes for type of investigation, investigation findings, investigation conclusions. h11: additional information regarding the results of the investigation. investigation summary no product was returned. thrombosis: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. pericardial effusion: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an adverse event related to the anatomy of the patient since patient had tortuosity in innominate and curved axillary artery at innominate take off would not let impella advance. product damage (guidewire kink): the cause of the product damage was most likely patient condition related since patient had tortuosity in innominate and curved axillary artery at innominate take off would not let impella advance. device history review guide wire 0.018" batch #8854090 passed all post sterile inspection checks.
3189 | 1220648-2025-48604 | 23648710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23648710&pc=OZD | 10/26/2025 | 11/25/2025 |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737061 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Cardiogenic Shock | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an imella 5.5 in a 62 year old male patient for acute myocardial infarction and cardiogenic shock support. he was alert on presentation but then brought back and became unresponsive and pulseless. rhythm shown to be pulseless electrical activity. cardiopulmonary resuscitation started and return of spontaneous circulation achieved. patient was intubated and hypoxemic. impella cp placed. the device experienced pump stop at bedside while awaiting definitive surgical planning- unable to restart with standard troubleshooting steps. device explanted at bedside and converted to intra-aortic balloon pump as bridge to 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the stroke: the cause of the injury was unable to be determined due to insufficient clinical details. cardiogenic shock: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3190 | 1220648-2025-48607 | 23649179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23649179&pc=OZD | 10/26/2025 | 11/25/2025 |  | 11/25/2025 |  |  | 10/26/2025 | 11/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Low Blood Pressure/ Hypotension | Tachycardia | Thrombosis/Thrombus | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella cp was placed in a 64 year old male patient for cardiac arrest support. an impella was inserted with ease, was pre-closed x 2 prior to impella sheath insertion. coronary arteries were rechecked using single access with companion sheath. no changes from previous left heart catheterization. the patient was transported to the intensive care unit. the patient foley has dark urine with some blood clots. labs have been sent and no hemolysis was noted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: the device was not returned for evaluation. tachycardia / thrombosis: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. hypotension/ fever: the cause of the injury was not determined due to insufficient clinical details. device history review the pump sn: (b)(6) passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. a 64-year-old male with a significant past medical history of hypertension, former tobacco use, and gastroesophageal reflux disease (gerd) experienced cardiac arrest at home and was subsequently admitted to a cardiac catheterization laboratory. coronary angiography demonstrated multivessel coronary artery disease. cardiothoracic surgery was consulted, and the patient was transferred to another facility for further evaluation. during hospitalization, the patient developed increased work of breathing and required reintubation, and was then transferred to a tertiary care facility for placement of an impella cp for mechanical circulatory support. at the time of implantation, the patient was classified as society for cardiovascular angiography and interventions (scai) stage e cardiogenic shock. the impella cp was inserted without reported difficulty. two perclose devices were deployed prior to impella sheath insertion for vascular closure preparation. the patient was transferred to the intensive care unit (icu) post-procedure. following placement of a foley catheter, urine was noted to be dark with visible blood clots. foley irrigation yielded large clots described as appearing consistent with old blood per the nursing staff. labs were obtained and did not demonstrate evidence of hemolysis. two days after initiation of impella cp support, the patient underwent aggressive deep suctioning by respiratory therapy, after which the patient aspirated and experienced oxygen desaturation into the 70 percent range. emergent bronchoscopy was performed. mechanical ventilation settings were increased. the patient subsequently developed tachycardia and hypotension with elevated pulmonary artery pressures. epoprostenol (veletri) and norepinephrine (levophed) were initiated for hemodynamic support. the patient was medically paralyzed. white blood cell count was 21.5. the patient had been febrile the prior night but was afebrile the next morning. the following day, mechanical circulatory support was escalated to an impella 5.5 for continued hemodynamic support. after approximately nine days on impella 5.5 support, the patient developed atrial fibrillation with rapid ventricular response and was treated with amiodarone. despite impella support, the patient¿s condition remained a concern by the family. two days later, the patient expired following withdrawal of care. tachycardia, hypotension, and fever occurring during impella cp support are being reported as serious injury events. the impella 5.5 is being associated with the death as it was in place at the time of demise and the patient experienced atrial fibrillation with rapid ventricular response requiring antiarrhythmic therapy while on support. however, underlying condition contributed most to the demise outcome (including scai stage e cardiogenic shock and severe coronary artery disease). note: investigation type/findings/conclusion remain unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3191 | 1220648-2025-48606 | 23649304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23649304&pc=OZD | 10/27/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026686374 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was on impella cp support. while o support it was reported that the patient was escalated to impella 5.5 after impella cp was unable to be repositioned and hemolysis - plasma free hemoglobin 150. the patient was escalated to the impella 5.5 and patient outcome was noted as expired. this report is for the impella cp and a additional report will be sent for the impella 5.5 (serial number (b)(6)). | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: data logs showed pump ran without issue during support. there were positioning alarms triggered during the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3192 | 1220648-2025-48608 | 23649431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23649431&pc=OZD | 10/27/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765504 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella cp support. while on support the impella console displayed placement signal not reliable alarm. previous to alarm, impella in aorta alarm appeared, device was repositioned at bedside by advancing ~4 cm. just prior to alarm, purge bag switched from 5% dextrose in water to 5% dextrose in water with sodium bicarbonate. device still functioned after alarm, care team continued to monitor pump overnight and weaned device from p-6 to p-2. later it was reported that there was a thrombus noted in left ventricle. the impella cp was removed and patient was noted as stable with a survival outcome. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. optical sensor issue: the pump was returned for investigation, however, it was cut at the catheter, so standard testing of the optical sensor could not be completed. the only abnormality noted on the pump was a kink in the catheter distal to the red plug, however, it was determined to be unrelated to the reported failure based on the fact that data logs show the optical signal/5s parameters returning to normal values during the case. upon review of data logs, it is apparent that the console is the potential cause of the issue. device in wrong position: returned product was investigated and there were no abnormalities noted on the repositioning unit. since insufficient clinical details were provided to determine how the pump came out of position, the cause of the device in wrong position could not be determined. thrombosis: returned product had evidence of biomaterial in the outflow cage and purge gap. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3193 | 1220648-2025-48601 | 23649620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23649620&pc=OZD | 10/28/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026732625 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 119 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella 5.5 support. while on support a small left ventricle thrombus was noted on echocardiogram. no other information provided other then the patient remained stable and survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis/peri-procedural adverse event: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3194 | 1220648-2025-48611 | 23649791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23649791&pc=OZD | 10/22/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674281 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp and cannulated for extra corporeal membrane oxygenation after being admitted to the hospital for a bentall procedure. on support day 5, the patient experienced an extensive cerebral infarction, and the neurological prognosis was determined to be poor. after obtaining informed consent from the family, the patient¿s code status was changed to not resuscitate. the physician noted that the relationship between the cerebral infarction and the impella use was unknown and noted that ecmo was also in use at the time of the incident. the patient expired. | [5 empty] | A2, a4 are unknown. b2. date of death was not available in the complaint record accessible for mdr assessment at time of reporting. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the stroke was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | This follow-up report is being submitted to update a2 (patient age) with newly obtained demographic information. no additional clinical or device-related updates are available at this time
3195 | 1220648-2025-48610 | 23649792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23649792&pc=OZD | 10/27/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734972 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Cardiac Arrest | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support during surgery for an aortic valve replacement (avr). the impella 5.5 was placed without complications after balloon aortic valvuloplasty was performed due to the patient¿s severe aortic stenosis. the patient was taken to the cardiovascular operating room for avr, and the patient was placed on cardiopulmonary bypass (cpb). after cpb was initiated, the physician reviewed the placement of the impella and decided to increase the depth of the pump. it was noted that the device was not very mobile under the surgical drapes, and it was pulled significantly to free the catheter. the pump was advanced further into the left ventricle. after repositioning, the impella 5.5 abruptly stopped pumping. the device was placed into surgical mode and medical team decided to place a cross clamp and give retrograde cardioplegia. following the procedure, the device was unable to be restarted. the decision was made to explant the impella 5.5 as the medical team determined that the patient did not require mechanical circulatory support. after removal, it was noted that there was a bend in the catheter near the red plug, where the device was fixed to the patient¿s body. patient outcome is unknown as outcome information was not available in the complaint record accessible for mdr assessment at time of reporting. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3196 | 1220648-2025-31936 | 23650158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23650158&pc=OZD | 02/21/2025 | 11/25/2025 |  | 11/25/2025 |  |  | 02/21/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612249 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm was encountered during impella cp support. the alarm was silenced and support continued uninterrupted. the pump was eventually weaned and explanted. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and evaluated. impella cp returned for evaluation. upon visual inspection, no kinks in catheter present. pump was connected to console and placement signal not reliable (psnr) alarm reproduced. optical parameters of the pump were measured, and values were consistent with a broken fiber. pump did not pass laser continuity testing and light was emitting from patient cable side of red handle. further examination of the red handle found that a broken fiber was present on the patient cable side of the red handle. additionally, bond between patient cable and kink protector was not intact. patient cable was able to easily rotate and translate. wires exiting from patient cable into red handle were very tightly twisted around themselves. complaint pump kink protector was compared against a control pump. upon further inspection of the bond of the patient cable kink protector, complaint patient cable showed thin glue layer present. kink protector was able to easily slide up patient cable without resistance. control patient cable had thick layer of glue and bits of kink protector still attached after kink protector was removed. complaint kink protector shows smooth internal canal, no tearing of kink protector present. very minimal glue residue on kink protector. control kink protector torn after removing from patient cable. some glue residue present on inside of control kink protector. no data logs were returned for analysis. clinical details noted that a psnr alarm occurred on the third day of support a few hours after troubleshooting with the touhy-borsht lock was performed. the root cause of the optical signal issue was most likely due to a damaged fiber line on the patient cable side of the red handle as a result of insufficient adhesive in the manufacturing process. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3197 | 1220648-2025-48612 | 23651347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23651347&pc=OZD | 10/27/2025 |  |  | 11/25/2025 |  |  | 10/27/2025 | 11/25/2025 |  |  |  |  | 11/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764855 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in non-st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced ventricular fibrillation in the emergency room and was brought emergently to the cardiac catheterization laboratory. an impella cp was placed for support during a percutaneous coronary intervention. on day 2 of support, the patient was noted to be febrile and had questionable neurological status. the medical team suspected the patient had anoxia from multiple arrests prior to impella insertion. the patient¿s prognosis was poor. the medical team and patient¿s family decided to withdraw care of the patient. the patient¿s outcome was expired. | [5 empty] | A4 is unknown. b2 date of death was not available in the complaint record accessible for mdr assessment at time of reporting. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device and logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the fever: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e1 added initial reporter facility name and address as it was omitted from the initial report that was submitted.
3198 | 1220648-2025-48564 | 23630099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630099&pc=OZD | 10/26/2025 | 11/21/2025 |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729635 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Fever | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, there was blood-tinged urine on arrival from transport from (b)(6) hospital to (b)(6) hospital. suction events after moving the patient to the hospital bed. no issues per transport team. the patient was weaned to p-4. hemolysis was noted. an echocardiogram was performed and the inlet was shallow just past the aortic valve. the pump was advanced and was measuring at 3.7 centimeters with improved flows and waveforms. the plan was to leave at p-5, until a new pulmonary artery catheter (pa) was placed. the current pa catheter clotted off and was they were unable to get filled and pressures current output and current intake. the patient was brought to the cardiac catheterization lab for impella repositioning and new pa catheter placement. 500 cubic centimeters of albumin bolus were given. lasix was given. hemolysis cleared. and urine output was adequate and yellow. confidence interval (ci) was 1.9. the pump was increased to p-8. the plan was to trend lactates and ci¿s. two day later, the patient was running, temperatures overnight. a culture was sent and antibiotics were started. the patient's outcome was unknown. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in proper position and hemolysis was improved after repositioning of the pump. the cause of positioning issue was unknown. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. pump suction: the cause of pump suction was most likely patient condition related. fever: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3199 | 1220648-2025-48569 | 23630277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630277&pc=OZD | 10/27/2025 | 11/21/2025 |  | 11/24/2025 |  |  | 10/27/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725108 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pain | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during impella cp support, a patient complained of pain at the insertion site. gabapentin was added for pain at the impella insertion site. the patient continues on current support while waiting for a transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b, date of explant, has been added. | D6b, date of explant, has been added.
3200 | 1220648-2025-48570 | 23630348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630348&pc=OZD | 10/27/2025 |  |  | 11/24/2025 |  |  | 10/27/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719732 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support per the bedside registered nurse, a patient began having cerebrovascular accident symptoms early in the morning and was taken down for a head computed tomography. significant bleed with midline shift was found. heparin was discontinued. the decision was made to withdraw care and the patient expired. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3201 | 1220648-2025-48571 | 23630467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630467&pc=OZD | 10/27/2025 |  |  | 11/24/2025 |  |  | 10/27/2025 | 11/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756508 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced motor current high alarm followed by a pump stop. attempts to restart the pump were unsuccessful. the pump was explanted and the patient remains in critical condition. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. pump stop: clinical details report a sudden high motor current pump stop on 27/oct/2925. the pump was unable to be restarted, so it was explanted. salesforce reports a 207 hour case duration. data logs were reviewed, however, they were only available from 18/oct to 26/oct (180 hours). during this period, all pump metrics were generally stable, and there were no pump stop alarms. returned product was investigated and visual inspection confirmed a enlarged purge gap. the pump passed electrical testing. when the pump was connected to a console and attempted to be run in a bucket of water, high motor current pump stops reproduced. based on the inspection of returned product, the cause of the pump stop was determined to be due to bearing wear. based on the clinical details provided, the pump was run longer than the pump's design life of 8 days, per design trace matrix 0046-9910. thrombosis: clinical details report that the crrt circuit had clotted off 2 times during the case. the root cause of the injury was unable to be determined due to insufficient clinical details. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3202 | 1220648-2025-48572 | 23630473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630473&pc=OZD | 10/27/2025 |  |  | 11/24/2025 |  |  | 10/27/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731276 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Decrease in Pressure | Increase in Pressure | [5 empty] | Hematuria | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and a low platelet count. heparin was discontinued and argatroban was started. the patient also experienced sudden low purge flow. tissue plasminogen activator was added to the purge to resolve the issue. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Correction: h6 medical device problem code 141101 was erroneously entered on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombocytopenia / hematuria: the cause of the injury was not determined due to insufficient clinical details. high purge pressure: the cause of the high purge pressure was not determined due to lack of product and data logs returned for investigation. device history review: the device passed all post sterile inspection checks.
3203 | 1220648-2025-48573 | 23630563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630563&pc=OZD | 10/27/2025 |  |  | 11/24/2025 |  |  | 10/27/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649478 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to patient anatomy. the impella cp was prepped and backloaded on to the wire and was unable to pass innominate. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: "dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve." section: warnings & cautions: warnings: "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance." | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had calcification and resistance at innominate axillary artery preventing successful delivery of impella. the device passed all post sterile inspection checks.
3204 | 1220648-2025-48574 | 23630747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23630747&pc=OZD | 10/29/2025 | 01/05/2026 |  | 11/24/2025 |  |  | 10/29/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674404 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced very poor peripheral perfusion to both lower limbs. the impella cp was explanted due to severely compromised lower limb perfusion. a central extracorporeal membrane oxygenation was subsequently initiated. the patient was stable on impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Investigation summary: no product was returned for investigation. ischemia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1921936. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3205 | 1220648-2025-48575 | 23631586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23631586&pc=OZD | 10/25/2025 |  |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725992 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported while a patient was on impella 5.5 support, the patient went into atrial fibrillation during a walk but soon converted back to normal sinus rhythm. the patient was given volume bolus and lasix was held. urine output was adequate. later, it was reported that the impella site was unremarkable. one unit of packed red blood cells were given. hemoglobin was 7.7. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. anemia: the cause of the injury was not determined due to insufficient clinical details. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details.
3206 | 1220648-2025-48576 | 23631785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23631785&pc=OZD | 10/25/2025 |  |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731036 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 69 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient required 200 joules of defibrillation for sustained torsade's. the impella was removed and hemostasis was obtained. the patient's outcome is stable. additional information was received. the pump was not saved after explant. | [5 empty] | Investigation summary: no product was returned for investigation. arrhythmia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history review device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
3207 | 1220648-2025-48578 | 23632391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23632391&pc=OZD | 10/25/2025 |  |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026746357 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 86 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Anemia | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Renal Impairment | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella cp support due to acute myocardial infarction and cardiogenic shock. while on support, two units of packed red blood cells were given for hemoglobin of 5.8. central venous pressure (cvp) dropped to 5, and hemolysis was noted on labs. post blood administration, cvp was up to 15. a large chest wall hematoma was noted. a chest tube was placed with well over three liters out. the patient was anuric, continuous renal replacement therapy was started. the patient was stable post procedure and was sent to the intensive care unit. additional information was received. the pump is not available for return. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemolysis: data log shows low pump flow with suction alarms from the start of the case, with some improvement in pump flow while running on steady p-level p4. later, the pump was even able to tolerate higher p-levels for some time before being turned back to p4. there was no motor current spike or positioning issue noted during the case. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. clinical symptoms (hypotension, major bleed, hematoma, anemia, acute kidney failure): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3208 | 1220648-2025-48579 | 23632974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23632974&pc=OZD | 10/25/2025 |  |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026742703 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with acute myocardial infarction and cardiogenic shock was being escalated from an impella cp to an impella 5.5 for support. after connecting the impella 5.5 to the automated impella controller, an "optical sensor system error" was presented. the pump was disconnected, troubleshooting was performed, and reconnected. the console presented a ¿controller error: switch to back up controller¿. a new pump and console were used. there was no patient harm. this is one of two related reports. this report represents the pump and another report will be submitted to represent the automated impella controller. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. since the pump was not returned for investigation, the cause of the optical sensor issue could not be determined.
3209 | 1220648-2025-48580 | 23633226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23633226&pc=OZD | 10/24/2025 |  |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766027 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient experienced renal failure and ana additional device was required. the p became anuric overnight and was started on continuous renal replacement therapy. the plan is to escalate to an impella 5.5. this is one of four related reports. this report represents the impella cp. other reported will be submitted for the two related impella 5.5 pumps and the automated impella controller. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. renal failure : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3210 | 1220648-2025-48581 | 23633740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23633740&pc=OZD | 10/25/2025 | 02/12/2026 |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1649323 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with acute myocardial infarction and cardiogenic shock was being escalated from an impella cp to an impella 5.5 for support. after connecting the impella 5.5 to the automated impella controller (aic), an "optical sensor system error" was presented. the pump was disconnected, troubleshooting was performed, and reconnected. the console presented ¿controller error: switch to back up controller¿. a new pump and aic were used. there was no patient harm. this is one of four related reports. this report represents the automated impella controller and other reports will be submitted for the impella cp and the two impella 5.5 pump. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D9 updated; the product has been returned. investigation summary: the cause of the optical issue is due to defective optical component.
3211 | 1220648-2025-48582 | 23633869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23633869&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025611054 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Decrease in Pressure | Increase in Pressure | [5 empty] | Pulmonary Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 while supporting the patient had purge flow decreased. the purge system was changed, and lysis was administered. it was further reported that the impella was kept at p-4 level by mistake which led to the patient having some pulmonary edema. the p-level was raised. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6: omitted code a141101 and added code a141102 per the results of the investigation. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary purge pressure high: the pump was not returned for investigation. clinical details indicate that after impella 5.5 s2 implantation, the team monitored purge flow and purge pressure closely. on (b)(6) 2025, purge pressure values were high (pp: 793 mmhg and 768 mmhg) while purge flow remained within range. multiple interventions were performed without any improvement in purge pressure, including changing the purge system, initiating tpa lysis therapy, and later switching to bicarbonate purge. no red purge alarm occurred. data log shows the purge pressure was within specification range (approx. 775) during the entire case run without any noted trend. on the reported event on (b)(6) 2025, there were four instances of purge volume low alarms, which correlated with bag/cassette change procedures. based on the data log, no high purge pressure issue was found during the case run. clinical symptoms (pulmonary edema): the cause of the injury was not determined due to insufficient clinical details. device history review the pump passed all post sterile inspection checks.
3212 | 1220648-2025-48583 | 23633963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23633963&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666090 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemoptysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. cardiopulmonary resuscitation was performed prior and throughout the procedure. heparin and cangrelor were off due to hemoptysis. the patient outcome is unknown. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemoptysis: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3213 | 1220648-2025-48585 | 23635366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23635366&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756641 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient was placed onto impella cp support due to post-cardiotomy cardiogenic shock and low cardiac output syndrome and coronary artery bypass graft/valve replacement. while on support, the staff expressed concern for right pleural effusion with 600 milliliters of serosanguineous drainage. pericardial effusion was also noted. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3214 | 1220648-2025-48584 | 23636910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23636910&pc=PYX | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026614464 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Low Blood Pressure/ Hypotension | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella rp flex for mechanical circulatory support. the patient was brought to the cardiovascular operating room on (b)(6) 2025 for a right minithoracotomy with mitral valve repair and annuloplasty. the following day, the patient experienced right ventricular heart failure and then continued to decompensate overnight requiring additional vasoactive support. the decision was then made to place an impella rp flex. the impella rp flex was successfully placed, and support was initiated with no complications. on day 4 of support, the patient experienced atrial fibrillation with heart rate in the 120s beats per minute. it was noted that the medical team was considering treating the arrhythmia with amiodarone, however they were reluctant to treat the patient with it as the patient became hypotensive the previous time it was delivered. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> paroxysmal atrial fibrillation: the cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device lot: 1885537. device history review ==> the pump s/n: (b)(6) passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3215 | 1220648-2025-48577 | 23637830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23637830&pc=OZD | 10/25/2025 |  |  | 11/24/2025 |  |  | 10/25/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631643 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support had bleeding from the impella access site. all best practices were shared with the team. blood products were administered. two hours later, bleeding was successfully resolved. the patient outcome is unknown as outcome information was not available in the complaint record accessible for mdr assessment at time of reporting. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. access site adverse event: the cause of bleed cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
3216 | 1220648-2025-31935 | 23637976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23637976&pc=OZD | 07/03/2025 | 11/24/2025 |  | 11/24/2025 |  |  | 07/09/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713329 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient came as a transfer from another facility on the implanting center's automated impella controller (aic). due to patient's instability at time of transfer, the decision was made to not switch consoles. the patient had stabilized, and the decision was made to switch to the current facility¿s aic. white cable switched, and immediate placement signal not reliable alarm. purge cassette and disc transferred over, and purge system failure alarm noted. purge bag changed while waiting for new cassette without resolution of alarm. purge cassette changed, and alarm resolved. no changes in impella flows during these alarms. white cable connection confirmed as secure/snug, resolution of placement signal not reliable alarm, and all waveforms returned to aic screen. motor current remained pulsatile, impella flows remained at 2.1 lpm throughout. no hemodynamic instability at any time during alarms. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. analysis: cloud logs were able to be retrieved for both consoles. (b)(6) performed as expected until (b)(6) when cp pump 691621 was transferred to (b)(6). (b)(6) had two independent failure modes that were captured in the complaint. the imc logs show that a placement signal not reliable alarm was triggered at 10:28 resolving at 10:41. the optical signal started with placement signal frequently going to 3000mmhg until 10:41 when the placement signal normalized for the duration of the case. for the purge, the imc logs show that after purge cassette 207580 was inserted into (b)(6), there was a controller failure for the piston being blocked. this was only resolved by swapping purge cassette for sn (b)(6) at which point the purge was able to perform as expected from 10:39 on. the rtlogs confirm that at the time cassette 207580 was inserted, the purge flow ticks could not go through the full range that are available after the purge cassette swap. root cause: the cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug as the complaint clarified that making that connection more secure resolved the issue. there was no problem found with (b)(6) priming as the piston block controller failure was resolved by swapping out purge cassettes.
3217 | 1220648-2025-48586 | 23638043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23638043&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761362 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was implanted in a patient. on day 2 of support, the patient coded, and it was noted that the impella was not in the correct position and there were continuous alarms for suction. the impella was found to be too shallow and was repositioned. later, it was noted that the patient¿s potassium was elevated to >7 mmol/l and nephrology was consulted. the patient¿s epinephrine requirements increased, and it was reported that the impella alarmed and the patient then coded. a return of spontaneous circulation was achieved, however the family decided to not escalate care. it was noted that the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. clinical symptoms (cardiac arrest): the cause of the injury was not determined due to insufficient clinical details. pump suction / device in wrong position: the pump was not returned for investigation. data log shows several pump suction alarms, with a related drop in flow and fluctuating placement signal, while running on steady p-level. there were also suction events when the pump was turned to p2, with dampened motor current noted. there were a few "pump position in ventricle" alarms triggered during the suction event around. the cause of the suction issue was most likely use-related, specifically due to suboptimal device positioning, and was resolved with appropriate repositioning. however, the cause of the pump position issue was not determined without returned product investigation and sufficient clinical details. device history review: the device passed all post sterile inspection checks. | A. patient information is unknown. b2. date of death information was not available in the complaint record accessible for mdr assessment at time of reporting. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3218 | 1220648-2025-48589 | 23638297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23638297&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729510 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Impella went into the aorta and unable to advance it in so was replaced with a new impella 5.5. there was no patient harm. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the device in wrong position (positioning issue): impella went into the aorta and unable to reposition it. the cause of the positioning issue was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3219 | 1220648-2025-48588 | 23638301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23638301&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756561 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp that a hematoma at impella site related to track ooze per team, manual pressure held with resolution of hematoma. the impella in aorta alarm. self-resolved but point of care ultrasonography done and cardiologist pushed cp in a couple centimeter. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the clinical stated there was an impella position in aorta alarm that self resolved but the pump was pushed deeper a couple cm. access site adverse event: the cause of hematoma/ bleed was most likely use issue related since the bleed and hematoma occurred after the transfer, in the icu and the act was 300 well above the recommended range. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3220 | 1220648-2025-48587 | 23638454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23638454&pc=OZD | 10/26/2025 |  |  | 11/24/2025 |  |  | 10/26/2025 | 11/24/2025 |  |  |  |  | 11/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026763706 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp and during support the pump became malpositioned. transthoracic echocardiogram revealed that the device had migrated underneath a papillary muscle. the doctor made several unsuccessful attempts to reposition, which resulted in continuous suction alarms. the status began deteriorating, and the decision was made to explant the cp and escalate to impella 5.5. | [5 empty] | The impella device was not returned from the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered impella position in ventricle & suction alarms. no other issues were found on the logs. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in proper position since the physician confirmed pump was deep under tte & caused hemolysis as reposition of the pump was not successful. device in wrong position: the cause of device in wrong position was most likely the patient movement related since the pump was in improper position after being rolled in icu by the clinical team. pump suction: the cause of pump suction was most likely due to pump was not in proper position since the physician confirmed pump was deep under tte that triggered suction alarms. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3221 | 1220648-2025-48529 | 23615865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23615865&pc=PYX | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026741942 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [5 empty] | INTRA-AORTIC BALLOON PUMP | [22 empty] | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp flex device for mechanical circulatory support (mcs). the patient was admitted and taken to the intensive care unit on multiple drips, lactic elevated. echocardiogram showed a constricted left ventricle with a small cavity, possible cardiac amyloidosis. the right ventricular was dilated with tricuspid regurgitation. the heart team decided to place an impella rp flex for right ventricle unloading and intra-aortic balloon pump (iabp). the patient was intubated prior to procedure. an 8 french/50 cubic centimeter iabp was placed in the left femoral artery, and the impella rp flex was then placed in the right femoral vein. one mattress suture was placed. post implant, the physician was pleased with the impella position and the patient's hemodynamics. the patient was transferred back to the unit on impella mcs. later that day, the impella rp flex was viewed and position was unchanged. support continued with impella rp flex and iabp. drips were weaned as tolerated. international normalized ratio (inr) and activated clotting time (act) continued to be elevated. the following day, a placement signal not reliable (psnr) alarm was noted. the team was educated on how to monitor the impella running without these numbers visible. pressors were continued to be weaned as tolerated. however, the patient's prognosis was noted as "poor." the impella access site was clean, dry and intact and there were no impella questions or concerns. however, approximately three hours later, it was noted the patient expired on support. the patient went into cardiac arrest and return of spontaneous circulation was unable to be achieved. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the hemodynamic instability, the cause was not determined due to insufficient clinical details. for the placement signal issue, data logs show pulmonary artery pulsatility (pap) values trending down from 0 to -55mmhg in less than 1 minute and then returns to a positive value before quickly trending back down. values continue to fluctuate in this manner for the remainder of support. data logs confirm the placement signal not reliable (psnr) alarms starting on oct-23. psnr alarms dp signal out of range and epm sensor failure was observed for the remainder of support with flow calculation disabled. optical 5s parameters stable. no motor current spikes outside of p-level changes. without the returned product, the pump cannot be tested for dp drift nor can the dp electrical values be tested. the cause of the placement signal issue was not determined due to no product returned and inconclusive data log analysis. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A4 weight is unknown. medical safety review: medical safety reviewed the event and noted the following: a patient with worsening heart failure symptoms underwent successful implantation of an impella rp flex and an intra-aortic balloon pump (iabp). post-implant, vasoactive drips were weaned as tolerated. on the following day, a ¿placement signal not reliable¿ alarm was noted on the impella rp flex. this alarm did not affect pump function. approximately three hours later, the patient experienced cardiac arrest and expired while on support. patient outcome: the patient expired while on support. the death will be conservatively reported as associated with the rp flex; however, the most likely contributing factor appears to be the patient¿s worsening clinical condition rather than device performance. manufacturer assessment: the impella rp flex displayed a ¿placement signal not reliable¿ alarm. although the alarm did not impact pump performance, it is classified as a device malfunction per reporting requirements. the patient's death is conservatively attributed to the impella rp flex, as the patient was on support at the time of death. however, the patient's clinical deterioration is most likely primary cause of death. no evidence suggests that the malfunction contributed directly to the patient's death. device status: the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3222 | 1220648-2025-48530 | 23616203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616203&pc=OZD | 10/25/2025 |  |  | 11/21/2025 |  |  | 10/25/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670074 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp, the pump was pushed into the left ventricle until the impella prolapsed. the impella was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was pushed into lv until pump prolapsed while peeling away sheath and then removed after provider pulled it out of ventricle. the device passed all post sterile inspection checks.
3223 | 1220648-2025-48531 | 23616249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616249&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/28/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026753250 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 78 KG |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced an optical sensor failure. no patient harm was reported. | [5 empty] | H6: updated codes for type of investigation, investigation findings, cause not established. h11: updated based on the results of the investigation. investigation summary: purge cassette returned. no data logs returned. the clinical details state that there was an optical sensor failure. the returned product did not have any damages noted, 5s were within specification and it passed laser continuity. the pump was run in a bucket of water and no alarms were noted. due to a lack of data logs provided and limited clinical information, the root cause cannot be determined. device history review: impella 634110 passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3224 | 1220648-2025-48532 | 23616351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616351&pc=OZD | 10/28/2025 |  |  | 11/21/2025 |  |  | 10/28/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261012 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that when setting the clock on the automated impella controller, a system self-check failure occurred. there was no patient involvement. | [5 empty] | The reported event did not occur with patient use. therefore, section a was left blank. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | Product complaint # ==> (b)(4). investigation summary ==> data analysis: the data logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the console was tested in aachen field service. upon bootup, the console showed the "system self check failure" screen. fwcheck was performed (see ic4658_fw), and the sau_powermod_1 hardware revision displayed as ¿n/a¿ and the firmware version also showed ¿n/a.¿ this occurred due to a communication issue between the pbm and the 4-in-1. visual inspection confirmed pbm contamination (pbm_contamination.jpg) which had impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
3225 | 1220648-2025-48534 | 23616464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616464&pc=OZD | 10/26/2025 |  |  | 11/21/2025 |  |  | 10/26/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730953 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the groin site and suspected hemolysis. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. minor bleed: the cause of minor bleed was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks.
3226 | 1220648-2025-48533 | 23616684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616684&pc=OZD | 10/26/2025 |  |  | 11/21/2025 |  |  | 10/26/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706432 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed a thrombus in the left ventricle. the patient had been anticoagulated with bivalirudin for most of the impella support. the plan was to place the patient on a left ventricular assist device (lvad) by the end of the week to prevent an accidental clot ingestion of the pump. the patient was changed to heparin instead as a result of a clot finding. the patient was successfully bridged to a lvad. the patient tolerated the lvad and left sided support well. | [5 empty] | The cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
3227 | 1220648-2025-48535 | 23616834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616834&pc=OZD | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/28/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611084 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the attending physician suspected a positioning error of an implanted impella 5.5. attempts to reposition the pump were unsuccessful. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the difficult to place/ position was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3228 | 1220648-2025-48536 | 23616928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23616928&pc=OZD | 10/22/2025 |  |  | 11/21/2025 |  |  | 10/28/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698691 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp suffered a stroke. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3229 | 1220648-2025-48537 | 23617241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23617241&pc=OZD | 10/22/2025 |  |  | 11/21/2025 |  |  | 10/22/2025 | 11/21/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638023 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation | Distributor | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was implanted via the left femoral artery without complications. on day two of impella support, the patient experienced oozing from the impella access site and the central venous catheter. the site was redressed with angle matching, and the patient was transfused with a unit of blood products. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The investigation for the reported event, major bleed, has been completed. no product was returned for investigation. the root cause of the injury could not be determined as product was not returned and insufficient clinical details were provided. | A4 is unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
3230 | 1220648-2025-48538 | 23617541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23617541&pc=OZD | 12/11/2025 | 12/22/2025 |  | 11/21/2025 |  |  | 10/28/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734060 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male | 100 KG |  |  | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Unspecified Infection | Sepsis | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella 5.5 developed an infection (hospital germ) unrelated to the impella and experienced a high purge pressure alarm with low purge flow when the cassette was changed. zero flow of the purge system was noted at the time of the exchange. the issue self resolved. impella support continued. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Explant date, section d6b, has been updated. | No product returned for evaluation for investigation. the root cause of the infection and sepsis was unable to be determined due to insufficient clinical details. the cause of the high purge pressure could not be determined as no product and insufficient data logs and clinical details were returned for analysis. the cause of the optical signal issue could not be determined as no product and insufficient data logs and clinical details were provided. the device passed all post-sterile inspection checks. | Sections b3, date of occurrence, and d6a, explantation date have been updated.
3231 | 1220648-2025-48539 | 23617655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23617655&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731099 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had hematuria during impella cp support. it was noted that the patient had concentrated dark red urine. the patient did not have a foley catheter. the team was recommended to assess the position and decrease p-level. however, impella weaning was initiated and explanted shortly afterwards. the patient remained stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: type of investigation, added b11 and b13. h11: added the results of the investigation. investigation summary: no product was returned for investigation. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1960398. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3232 | 1220648-2025-48541 | 23617828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23617828&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735943 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding that could not be controlled. the pump was explanted. after removal, there was some bleeding and a drop in blood pressure, but after hemostasis, blood pressure remained stable. the puncture site was stopped with a percutaneous transluminal angioplasty balloon from the contralateral side and vascular treatment was performed surgically. no decrease in blood pressure was observed, but there were signs of hypoperfusion. the patient in this case was large and had a lot of fat, so there was resistance during insertion of the peel-away sheath. the gap between the sheath and the dilator became caught, creating a larger gap, which likely led to a larger puncture hole than usual. this likely resulted in a larger gap with the repositioning sheath, preventing the bleeding from stopping. | [5 empty] | The cause of the injury is most likely patient condition related since the patient in this case was large and had a lot of fat, so there was resistance during insertion of the peel-away sheath. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3233 | 1220648-2025-48540 | 23617829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23617829&pc=OZD | 10/24/2025 |  |  | 11/21/2025 |  |  | 11/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765646 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding and cardiac arrest. the patient received massive transfusion protocol and the impella was exchanged for a new pump. the patient also received surgical repair of the iliac artery. the patient survived. | The complainant reported that a patient implanted with an impella cp experienced bleeding that required blood transfusion, hypotension, cardiac arrest that required cardiopulmonary resuscitation (cpr), impella alarms for suction, an impella cannula that folded during use, and an impella that was positioned shallow in the ventricle. it was reported that during a coronary procedure, the patient¿s left main coronary artery was perforated. subsequently, the patient sustained cardiac arrest requiring cpr and the decision was made to place an impella cp to stabilize the patient. the impella cp was placed without reported complications, however the patient¿s aortic arch was noted to be very short and caused folding of the impella cp below the motor housing and suction alarms. it was noted that the catheter was placed posterior in the patient¿s small left ventricle. the decision was made to leave the impella in a shallow position and allow the patient to stabilize for further treatment the following day. on support day two, it was reported the patient had a pericardial washout and lung lacerations reported with impella cp support. following the procedure, the patient¿s condition was noted to be improving and the medical team planned on weaning and explanting the device the following day. the impella cp was successfully weaned and explant. immediately following the explant, the patient experienced significant bleeding and sustained cardiac arrest. massive transfusion protocol was initiated, and a new impella was placed to stabilize the patient. the patient was taken to the operating room to repair the patient¿s iliac artery and interventional radiology to repair small arteries. there were no complications reported with the replacement impella cp. the replacement device was later explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. |  |  |  |  | A4 is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 is being updated to include new and corrected information that was not included in the initial report. the revised b5 provides a complete event narrative. a4 weight of patient was not included in the initial mdr and is now correctly added.
3234 | 1220648-2025-48543 | 23617988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23617988&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1259143 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) power cable had loose connection. multiple sockets were replaced but the aic did not charge. the aic started charging again and the issue was resolved. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3235 | 1220648-2025-48542 | 23618046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618046&pc=OZD | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026743355 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Inflammation | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced access site bleeding upon explant. the patient was treated with albumin. the next day a hematoma developed at the access site and down the patient¿s right arm. the patient also had premature ventricular contractions and bigeminy. one unit of packed red blood cells was transfused. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of the inflammation and anemia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3236 | 1220648-2025-31931 | 23618058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618058&pc=OZD | 11/10/2023 | 11/21/2025 |  | 11/21/2025 |  |  | 11/10/2023 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481065 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal not reliable alarm. the nurse disabled audio and monitored motor current flows. the decision made to withdraw care, and the patient expired. the patient's death is captured in the sister record of this record. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console log shows placement signal not reliable alarm (#1036, opm signal invalid), and multiple psnr events. ltlog data shows that less than 27 days after pump 486754 was transferred from another console to (b)(6), optical placement signal became invalid and did not recover for the duration of the case. rtlogs indicated that during that time snr was very low, and contrast had occasionally unrealistic values (below 0% or above 100%). these characteristics of the optical signal are consistent with fiber break. device analysis: console testing showed that the system - including optical bench - was operating within specification. complaint issue was not reproduced. root cause: no problem was found with the console. pump investigation 23-19937-1 points to optical signal issues being related to the kinked fiber. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3237 | 1220648-2025-48544 | 23618224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618224&pc=OZD | 10/24/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757877 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 106 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Unspecified Infection | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed onto impella support for acute myocardial infarction complicated by cardiogenic shock. while on support, the patient experienced hematuria and hemolysis. antibiotics were prescribed for elevated white blood cells. additional information was received. the device is not available for return. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: product and data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria: the cause of the injury was not determined due to insufficient clinical details. infection: in order to make a cause determination, the clinical details outlined would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: device with passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3238 | 1220648-2025-31932 | 23618488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618488&pc=OZD | 04/16/2024 | 11/21/2025 |  | 11/21/2025 |  |  | 05/31/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1544031 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was placed onto impella 5.5 support for acute myocardial infarction / cardiogenic shock. while on support, the pump experienced optical signal issues that were resolved by disabling the positioning alarm and utilizing alternate means for monitoring position. the loss of placement signal did not prompt any intervention or pump replacement. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show multiple impella position in aorta alarms and large number of impella position unknown alarms triggered by low pulsatility algorithm. most of the alarms occurred on (b)(6) 2025. ltlog data shows that throughout most of the case, optical snr was noisy and fluctuated between 1000 and 8000. rtlog data confirm periods of low pulsatility of the placement signal, and degradation of the raw optical data signal due to lower snr. when console arrived for its annual pm it was sent for investigation. during its inspection in danvers some contamination of blue receptacle fiber was observed, as well as contamination of the internal optical connector between optical bench and fiber pigtail. console in its original configuration failed 5s testing, due to low contrast: test cavity 1 test cavity 2, measurement 1 test cavity 2, measurement 2 cavity length: 14335 nm snr: 12281 contrast: 43.8% lamp: 70.8% gain: 1.29 mano: 762.4 mmhg light: 2.60 v. cavity length: 16495 nm snr: 4014 contrast: 20.4% lamp: 58.8% gain: 1.29 mano: 762.5 mmhg light: 2.66 v. cavity length: 16496 nm snr: 5058 contrast: 45.8% lamp: 79.1% gain: 1.33 mano: 762.5 mmhg light: 2.71 v. after replacing optical pump connector with a known good one, optical parameters improved, but contrast value would still fail: test cavity 1, measurement 1. test cavity 1, measurement 2. test cavity 2, measurement 1. cavity length: 14336 nm snr: 14895 contrast: 60.0% lamp: 83.0% gain: 1.36 mano: 762.7 mmhg light: 2.67 v cavity length: 14336 nm snr: 14507 contrast: 58.8% lamp: 83.0% gain: 1.37 mano: 762.7 mmhg light: 2.71 v. cavity length: 16496 nm snr: 5954 contrast: 35.0% lamp: 71.9% gain: 1.36 mano: 762.7 mmhg light: 2.68 v. measurement of the optical bench ccd detector output did not show any distortions, however there was some noise located in the center of pattern, corresponding to longer test cavities (lower pressure values), correlating with lower values of snr. reportedly false positioning alarms were most likely caused by degradation of the optical signal due to failed optical pump connector and malfunction of the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3239 | 1220648-2025-31933 | 23618514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618514&pc=OZD | 05/01/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 06/05/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811284 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an implanted impella 5.5 pump became inaccurate during patient support. support continued uninterrupted despite the noted inconsistencies with the optical signal. upon further review by the investigating engineers, it was believed that the automated impella controller (aic) could have contributed to the noted sensor issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: a review of the imc logs reveals the occurrence of frequent suction alarms throughout the case. additionally, one instance of ¿calculator placement signal 1045¿ is found, and the rt log during this time shows that all optical 5s parameters are reported as a variation of -16384, with an invalid placement signal of 3000. this is a known condition that results from a transient optical bench firmware anomaly. the condition does not persist long enough to generate an alarm. beginning on (b)(6), frequent ¿impella position unknown¿ alarms begin to occur. a review of the rt logs for this period shows that all optical 5s parameters that are stable. device analysis: (b)(4) was not returned for this investigation. the cause of the aic issue was a 3rd party hardware/software issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3240 | 1220648-2025-48547 | 23618534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618534&pc=PYX | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026764947 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced an unreliable placement signal and absence of flow calculations. the patient had mediastinal bleeding and bleeding from chest tubes. the patient was transfused three units of packed red blood cells, fresh frozen plasma, platelets, and cyroprecipitate. next, the pump displayed a high purge pressure alarm. no kinks or external causes for the alarm were identified. the purge cassette was changed but the alarm did not resolved. tissue plasminogen activator was added to the purge, and vasopressors were administered. systemic heparin was withheld. the decision was made to withdraw care and the patient expired. additional information was received. the pump was not retained. there are two related reports. this report represents the pump and another reported was submitted to represent the automated impella controller. refer to section h10 for the related report number. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. major bleed: the cause of the injury was not determined due to insufficient clinical details. placement signal issue: the data logs confirm placement signal not reliable alarms (#1051 and #1053). the logs show a motor current spike down with a small speed deviation shortly before a gradual decrease in placement signal while pump flows increased. pump was at p9. the placement signal drift continued until placement signal went out of range and flow calculation was lost. it is unknown what led to the downwards motor current spike and subsequent drift. the cause of the issue was unable to be established as no product was returned for analysis and log analysis was not conclusive. purge pressure high: the data logs confirm purge pressure high alarms and show a gradual rise in purge pressure over about an hour with a decrease in purge flows. there were no motor current spikes outside p-level changes. clinical mentioned systemic heparin was delayed due to patient's bleeding. tpa was added to the purge. elevated purge pressure remained for about 3 hours before decreasing towards normal values. the cause of the issue was determined to be biomaterial build up as evidenced by log analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3241 | 1220648-2025-48548 | 23618611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618611&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759143 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient¿s urine pink, suspected hemolysis. the impella was repositioned with transthoracic echocardiogram, p-level reduced to p-6, will reassess throughout day for possible upgrade. the pump was explanted the following day. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary no product was returned for investigation. no data logs were returned for analysis. the device was in the wrong position: clinical details mention that the impella inlet was 4.5cm from the atrioventricular (av) and was repositioned with the help of the transthoracic echocardiogram (tte), and the inlet was then 4cm from av. the root cause of the positioning issue was not determined due to lack of sufficient conclusive evidence from the provided clinical details and unreturned product for further evaluation. mechanical interaction with blood: clinical details mention hemolysis was suspected due to pink urine. the root cause of the hemolysis was not determined due to lack of sufficient conclusive evidence from the provided clinical details, unreturned data logs for further analysis, and unreturned product for further evaluation. device history review ==> pump #629015 passed all post-sterile inspection checks.
3242 | 1220648-2025-48546 | 23618776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618776&pc=PYX | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251194 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp experienced a placement signal not reliable alarm on the supporting automated impella controller. there are two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to section h10 for the related report number. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3243 | 1220648-2025-48549 | 23618887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23618887&pc=OZD | 08/22/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693227 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. during the first day of support, it was reported that the purge system was leaking. the purge cassette was replaced, and the nurse reported that this resolved the issue. however, the following day it was again reported the purge system was leaking. the device was explanted successfully after weaning. the pump and the second cassette were sent for investigation for the purge leak. on (b)(6) 2025, investigation findings revealed that the purge leak issue was due to the damaged sidearm portion of the pump. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | Investigation summary: according to the clinical details, a purge leak was reported on 22-aug and the issue was resolved. a second purge leak was reported on 24-aug, and the product was sent for investigation. data log review was not required for this case run. purge leak ¿ pump: the pump and purge cassette were investigated, and no issue was found related to the purge cassette. damage was identified on the pump sidearm, with a lengthwise crack down the yellow luer, where the purge leak was reproduced. the cause of the purge leak was a crack in the yellow luer but the cause of the crack was unable to be determined. the supplier invetigation of similar complaint indicated possible high tension in combination with disinfectants and certain drug ingredients like for example oil, alcohol, lipids but was unable to confirm. device history lot: device batch: 1940910. device history review: the pump (b)(6) passed all post sterile inspection checks.
3244 | 1220648-2025-48550 | 23619091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619091&pc=OZD | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695421 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Decrease in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced suction alarms. the position of the pump was confirmed to be correct but the suction continued and the patient developed hematuria. tissue plasminogen activator was added to the purge with minimal improvements. the pump was explanted out of concern for a thrombus. the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction/low or blocked pump flow: the cause of the issue was most likely related to ingested biomaterial, based on returned product investigation.
3245 | 1220648-2025-31934 | 23619220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619220&pc=OZD | 01/24/2024 | 11/21/2025 |  | 11/21/2025 |  |  | 01/24/2024 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133102 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support with an automated impella controller due to acute heart failure. while on support, there were placement signal unreliable alarms, however support was continued. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imclogs show (b)(6) had 5.5 pump (b)(6) connected on (b)(6). the imc log show persistent bad snr entries. the lt log shows that for the duration of the case, snr was below 1200 and placement signal frequently was unreliable. (b)(6) was returned for investigation as part of (b)(4) for a similar complaint and the console was found to have failing snr and a defect in the ccd array. the repair performed at that time, replacing the optical bench, would have addressed this complaint's symptoms as well. root cause: the cause of the optical signal issues was a defective optical bench with very low snr and a distorted ccd array. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3246 | 1220648-2025-48545 | 23619278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619278&pc=OZD | 10/25/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 10/25/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731189 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed in shock setting of post-cardiotomy cardiogenic shock. it was noted that the placement of cp was successful. on echocardiogram to assess placement it was noted that right atrial (ra) was collapsed. decision made to take patient back for washout. clots removed and compression relieved from ra. during support there was poor flow down leg due to tortuosity and blockages. decision to leave peel-away in place & placed antegrade perfusion sheath down left leg to refuse. additionally, the patient continued to bleed with heparin. decision to turn heparin down by half. ultimately the cp was explanted and inserted intra-aortic balloon pump. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: patient continues to bleed in chest with heparin. the cause of the bleeding is determined to be patient condition because of the bleeding from non-impella site. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3247 | 1220648-2025-48551 | 23619475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619475&pc=OZD | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680835 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 94 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Stroke/CVA | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation and ventricular tachycardia, and a cerebrovascular accident. there was no adverse impact to patient management reported. | [5 empty] | D6b, explant date, has been added. | A5 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the paroxysmal atrial fibrillation, tachycardia, arrhythmia, and stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
3248 | 1220648-2025-48554 | 23619577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619577&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741074 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hematuria | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. while on support, the patient lost pulsatility on placement signals and required initiation of epinephrine and norepinephrine infusions. the patient went into bradycardia with heart rate in the 40's. the patient had hematuria, urinary output greater than 30 cubic centimeters per hour and was pink-red in color. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary no product was returned for investigation. cardiovascular insufficiency/ hematuria: the cause of the injury was not determined due to insufficient clinical details. bradycardia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the bradycardia was unable to be determined due to insufficient clinical details. device history review ==> device with sn: (b)(6) passed all post sterile inspection checks.
3249 | 1220648-2025-48552 | 23619712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619712&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623961 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient began to experience high afterload and suction alarms while on the route back to the unit. vasopressin was stopped. the patient also experienced very labile swings from great pulsatility to none and appropriate blood pressure to hypotension within a very short period. at this point pulsatility was returned and remained and waveforms and flow improved on device. an echocardiogram (echo) was done showing the impella 5.5 was mispositioned waiting on echo to return to reposition. urine was red tinged most likely due to malposition. anticoagulation was still off, due to gastrointestinal bleeding. urine was clearing up but volume was down, so the team may try to use diuretics. the impella was going to be removed as a code stroke was called on the patient. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump suction: the cause of the suction issue was most likely related to patient condition, based on data log review and provided clinical details. mechanical interaction with blood (hemolysis, hematuria): the cause of the hemolysis issue was not determined without sufficient clinical information and returned product investigation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3250 | 1220648-2025-48553 | 23619778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619778&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient experience decreased distal perfusion to impella leg. providers are concerned about distal perfusion in impella leg. foot is warm and there¿s a slight flow signal on doppler. heparin has been held due to internal bleeding concerns. urine is pink or red, plasma-free >100. the patient was taken to cath lab for impella explant due to hemolysis. the patient outcome is unknown. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary data review: data logs showed pump ran without issue during support. there were positioning & suction alarms triggered during the case but resolved. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. hemostasis: the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device lot #: 1967764 device history batch ==> subcomponent lot#: n/a device history review ==> pump sn: (b)(6) passed all post sterile inspection checks.
3251 | 1220648-2025-48555 | 23619800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619800&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765712 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient had groin bleeding. the physician was going to look at the groin and possibly tighten the perclose sutures to aid in bleeding. the team will re-dress site at that time. no vasoactive drips currently and holding on heparin until partial thromboplastin time is within range. the patient was taken to operating room for impella removal. the impella was removed without issue and the patient was stable. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. impella cp sn (b)(6) returned the returned product was investigated. traces of dried blood were seen on the inflow and outflow cages, around the suture hub, and on the sterile sleeve. no other abnormalities were seen. access site adverse event: clinical details mention that the patient suffered from groin bleeding on (b)(6) 2025 while on impella support. the root cause of the access site bleeding was not determined due to lack of sufficient clinical details and inconclusive evidence obtained from the investigation of the returned product. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3252 | 1220648-2025-48558 | 23619821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23619821&pc=OZD | 10/25/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 10/25/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697137 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during impella 5.5 explantation, the impella cannula was torn off just above the blood outlet. the malfunction of cannula detachment, if it were to recur, is likely to cause or contribute to a serious injury or death. additional information received: all components of the pump were removed from the patient¿s body, and no additional intervention was needed. the patient was successfully bridged to left ventricular assist device | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3253 | 1220648-2025-48559 | 23620124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23620124&pc=OZD | 10/25/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 10/25/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719956 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pulmonary Edema | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient presented to the emergency department with groin pain. the patient had a past medical history of diabetes and meth use. the patient was admitted to rule out infection and found to be in cardiogenic shock. a decision was made to take the patient to ccl for impella cp placement. on impella support the patient developed bleeding from the impella access site. two units of packed red blood cells were required. repo sheath re sutured and angle matching dressing reapplied. bleeding was resolved. the patient also had flash pulmonary edema and required intubation. it was further reported that a decision was made to withdraw care and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: type of investigation, added codes b13 and b11. investigation findings, added code c13. investigation summary: the device was not returned for investigation. pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the injury is most likely use related since repo sheath re sutured and angle matching dressing reapplied resolved bleeding. device history lot: device lot: 1956248. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3254 | 1220648-2025-48561 | 23620228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23620228&pc=OZD | 10/25/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 10/25/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756801 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was successfully placed on a non-st-elevation myocardial infarction patient without issues. during support it was noted that the patient was not having much ventricular tachycardia. the patient remained on support for a total of 116.82 and was successfully weaned and explanted. | [5 empty] | Investigation summary: no product was returned for investigation. ventricular fibrillation: in order to make a root cause determination, the clinical details have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3255 | 1220648-2025-48562 | 23620260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23620260&pc=OZD | 10/25/2025 |  |  | 11/21/2025 |  |  | 10/25/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744191 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced ectopy. the dose of milrinone was reduced and epi was administered. the patient also experienced hypotension and was treated with albumin. impella support continues. | [5 empty] | D6b, date of explant, has been updated. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3256 | 1220648-2025-48563 | 23620571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23620571&pc=OZD | 10/26/2025 | 11/21/2025 |  | 11/21/2025 |  |  | 10/26/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026742686 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Unspecified Infection | Sepsis | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp the patient did not tolerate attempt to wean, lactates bumped to 4.1 and hypotensive. bumped to p 7. elevated white blood cells - sepsis workup (w/u) ongoing - + urinary tract infection (uti). the patient went into comfort care. support was turned down to p2 and patient eventually passed away. pump was turned off and left in. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned for investigation. sepsis/infection: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3257 | 1220648-2025-48556 | 23620574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23620574&pc=OZD | 10/24/2025 |  |  | 11/21/2025 |  |  | 10/24/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758792 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was fever during impella cp patient support. while on support, the patient was febrile at 38.2 and blood culture panels were being drawn. the impella cp was surgically removed in operating room after successful insertion of right axillary impella 5.5 and veno-arterial extracorporeal membrane oxygenation maintained. the patient expired while on impella 5.5 (serial number: (b)(6) therefore an additional report will be sent. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned for investigation. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3258 | 1220648-2025-48565 | 23622006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23622006&pc=OZD | 10/27/2025 |  |  | 11/21/2025 |  |  | 10/27/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720792 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Pulmonary Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced atrial fibrillation requiring cardioversion. one unit of packed red blood cells was transfused along with amio and digoxin. the patient also exhibited flash pulmonary edema requiring reintubation. the patient outcome is unknown. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the cause of the pulmonary edema was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3259 | 1220648-2025-48566 | 23622122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23622122&pc=OZD | 10/27/2025 |  |  | 11/21/2025 |  |  | 10/27/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756548 |  | 000000000001000080 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a loss of placement signal lost. the issue resolved without intervention. no patient harm was reported. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: clinical details indicate placement signal was lost, with no position monitoring. the root cause of the optical sensor issue was not determined due to unreturned logs, product, and insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3260 | 1220648-2025-48567 | 23622260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23622260&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640538 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. the patient was admitted to the hospital for a percutaneous coronary intervention (pci). the patient underwent a successful pci to the mid left anterior descending coronary artery and was noted to be stable during the procedure. following the procedure, the patient later declined, and the decision was made to bring the patient back to the cardiac catheterization lab (ccl) for right heart catheterization and a possible impella implant. upon arrival at the ccl, the patient¿s heart rate was 40 beats per minute. a temporary pacemaker was placed, followed by an impella cp. during the initiation of support, the patient¿s heart was noted to not be moving on fluoroscopy even though the pacemaker was capturing. impella support was initiated with a peak pump flow of 2.8 l/min and then began to decline. abiomed clinical support advised the physician that this was likely due to a drained ventricle with no refill occurring, possibly due to pulseless electrical activity (pea) arrest. compressions were initiated and epinephrine was administered. after 5 minutes, cardiopulmonary resuscitation was paused and return of spontaneous circulation was achieved. a femoral arteriogram revealed there was no flow past the impella sheath. the peel-away sheath was removed, and a repeat angiogram revealed flow was re-established down the leg. on day 2 of impella cp support, the medical team reported that the anticontamination sleeve separated from the distal touhy lock. the medical team was advised to clean the portion of the catheter that had been exposed and then place tegaderm around the sleeve and touhy lock. on day 6 of impella cp support, it was noted that the patient¿s lactate was trending up, however it was unclear if the patient had options for advanced support or escalation of support. the following day the patient was weaned from the impella cp to receive and magnetic resonance imaging of the brain. it was noted that the patient was intubated throughout the entirety of impella support without recovering cognitive function. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two automated impella devices: impella cp with serial number (b)(6). 14fr introducer with serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Investigation summary: clinical symptoms (cardiac arrest): the cause of the injury was not determined due to insufficient clinical details. product damage (anticontamination sleeve split/torn): the pump was not returned for investigation. data log review was not required for this case run. the cause of the sterile sleeve damage issue was not determined without returned product investigation and sufficient clinical details. device history lot: device batch: 1904174. device history review: the pump (b)(6) passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3261 | 1220648-2025-48568 | 23622340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23622340&pc=OZD | 10/23/2025 |  |  | 11/21/2025 |  |  | 10/23/2025 | 11/21/2025 |  |  |  |  | 11/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734842 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support for a percutaneous coronary intervention. the impella cp was successfully implanted, and support was initiated, but adequate flows were unable to be achieved. fluoroscopy was performed and revealed that the impella appeared to be kinked. the physician attempted to reposition the impella, however due to the patient¿s severe aortic stenosis repositioning was unsuccessful. the decision was made to replace the impella cp. no patient harm was reported in relation to this event. follow-up efforts were unsuccessful in obtaining additional information on the replacement impella device, and the final patient outcome is unknown. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. product damage (cannula kink): the cause of product damage was most likely patient condition related since patient had severe aortic stenosis. low or blocked pump flow: no product was returned. patient had inadequate flows. after reviewing the logs, multiple suction alarms were observed. the cause of low pump flow was most likely patient condition related since the pump was activated even after it had kinked and since patient had severe aortic stenosis. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an adverse event related to the anatomy of the patient since patient had severe aortic stenosis. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3262 | 1220648-2025-48493 | 23603641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23603641&pc=OZD | 10/21/2025 | 01/06/2026 |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755293 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in with known cardiac history. the pump was placed but the single access site had a hematoma and a little bit of pericardial effusion while on support. the pump was explanted and the patient survived. additional information was received. the hematoma was not at the groin. it was a hematoma due to percutaneous coronary intervention of a venous graft from a coronary artery graft out of 1997. the hematoma was in the right ventricle myocardial wall and was not related to the impella but to the percutaneous coronary intervention. the impella pump was not collected for evaluation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. hematoma: big hematoma as before, little bit of pericardial effusion. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1990598. device history batch subcomponent lot: n/a. device history review device with sn: (b)(6) passed all post sterile inspection checks.
3263 | 1220648-2025-48494 | 23603674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23603674&pc=OZD | 10/23/2025 | 02/25/2026 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136134 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation | [10 empty] | Injury | No | Yes | Display or Visual Feedback Problem | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered an automated impella controller (aic) issue. the impella cp was being inserted as per usual. after hooking up the power cable to the aic, the aic deviated from the normal screen prompts and appeared to have displayed older versions of aic prompts similar to the impella 2.5. the local team recommended to switch to a backup aic which worked without concern. there was no patient harm. | [5 empty] | H6 codes were corrected. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3264 | 1220648-2025-48495 | 23603736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23603736&pc=OZD | 10/23/2025 |  |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730353 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered pump flow issues during impella 5.5 support. it was noted that the motor current was flat, and team had turned the p-level down to p-2 to wean and possibly explant. when reviewing the motor current and aorta placement signal, it was noted to have caused an alarm of placement signal in aorta. when the intensivist changed the p-level from p-2 to p-4, the motor current remained flat with no changes. the flow did increase. it was discussed if there could be a possible ingestion at the inlet or within the cannula itself versus the motor or purge pressure area. the pump was successfully weaned and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 medical device problem code 140508 was inadvertently omitted on the initial manufacturer device report.
3265 | 1220648-2025-48496 | 23603936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23603936&pc=OZD | 10/23/2025 |  |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699720 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 125 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had minor bleeding and that they encountered positioning issues. the patient had a small hematoma at the impella cp site that increased. manual pressure was held to the site, the hematoma was expressed and the site became stable. additionally, the impella had another ¿in aorta¿ alarm that had not resolved. the impella was switched to p-2 and the doctor was to be at bedside to reposition under echo as soon as possible. the doctor performed a pocus and the pigtail appears to still be across the valve. extensive education was provided to the care team. the patient remained on impella support for two days. the patient remained stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary no product was returned. minor bleed/ hematoma: the cause of the bleed/ hematoma was most likely use issue related since the user placed the blue suture hub into the skin which is not advised with the use of cp. difficult to place or position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1947612 device history review ==> device (sn: (b)(6) passed all post sterile inspection checks.
3266 | 1220648-2025-48500 | 23604594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23604594&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689449 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed on a patient for left ventricle support. during cp support it was noted of some oozing at site. the site was redressed and preclose tightened. pedal pulses were present however heparin was discontinued. additionally, it was reported that the focus was to resolve gastrointestinal bleed. the patient remained on support for a total of 103.15 hours and was successfully weaned and explanted | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3267 | 1220648-2025-48502 | 23604736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23604736&pc=OZD | 10/21/2025 |  |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758125 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy awaiting transcatheter aortic valve replacement (tvar) was implanted with an impella cp for mechanical circulatory support. the patient presented to the emergency room for dyspnea, orthopnea, lower leg edema, an ejection fraction of 20%, and severe aortic stenosis. balloon aortic valvuloplasty was performed and the impella cp was implanted without complications, and the patient was transferred to the unit for rest and to await a tavr. on day 7 of support, the patient successfully underwent a tvar and remained on impella support post procedure. the following day, the medical team attempted to wean the patient from impella support, and the patient became hypotensive and was noted to be in heart failure. support was increased, and the decision was made to let the patient rest and recover for another day. on day 10 of support, the patient was transfused with two units of packed red blood cells for a drop in hemoglobin and melena in stool. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary no product was returned for investigation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. hypotension/ cardiac failure: the cause of the injury was not determined due to insufficient clinical details. device history lot device lot: 1979323 device history review device with sn: (b)(6) passed all post sterile inspection checks. as per qn (b)(4), for pump with sn: (b)(6), pressure accumulator set not secured in blister pack. the qn is set in process to implement the rework. inspection plan ensures that every other potentially affected lot will be found. therefore, no further immediate action is required.
3268 | 1220648-2025-48503 | 23604895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23604895&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590785 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 78 KG |  |  | Injury | Yes | Yes | False Alarm | Restricted Flow rate | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was placed on intra-aortic balloon pump due to hemodynamically unstable, requiring pressers. however, the decision was made to escalate to an impella cp. the impella cp was placed successfully. during support, hematuria was present and labs were hemolyzed. additionally, there was a clinical support call for a suction alarm at p8. they came down to p4 and were unable to go back up. there was another clinical support call for position in aorta alarms along with suction. the patient remained on support and was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause for the reported events could not be determined as the product was not returned and due to insufficient clinical details.
3269 | 1220648-2025-48501 | 23604929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23604929&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763065 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 71 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hematoma | Tachycardia | Thrombocytopenia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient(pt) presented for an outpatient cardiac output response to stress test that revealed an 80% cardiovascular imaging research center (circ) lesion that was stented. after stenting pt became bradycardic and ultimately coded. advanced cardiovascular life support initiated, pt emergently intubated and after coding for a total of 2 separate times return of spontaneous circulation was achieved. decision was made to place an impella cp. the cp was placed and started in auto. intravascular ultrasound (ivus) and plain old balloon angioplasty performed to circ. evidence of clot showering on intravascular ultrasound. p level turned to p9. pt began having suction alarms. during support it was noted that amio was started for atrial fibrillation the pt was turned to p5 due to drop in hemoglobin. ultrasound done around impella site, no bleed, however large contained hematoma was found around the swan site. 2 units of blood given. the patient continued to receive blood for low hemoglobin. the pt's platelets were also down trending. the pt was worked up for heparin-induced thrombocytopenia with thrombosis. plan was to remove impella due to concerns of heparin-induced thrombocytopenia and desire to stop systemic heparin. the patient was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. no data logs were returned for analysis. pump suction: clinical details mention that the patient began having suction alarms a few minutes after turning the p-level to p-9 on 24 oct 2025. fluid bolus was ordered. the root cause of the suction alarms was not determined due to lack of sufficient clinical details providing information on whether the suction alarms resolved after administering blood products, and unreturned data logs and product to confirm suction alarms and related trends. thrombosis, tachycardia: clinical details mention the evidence of clot showering was found under ivus. patient had an increased heart rate as well. the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma, thrombocytopenia: clinical details mention that a large contained hematoma was found around the swan site in the rfv under ultrasound. the patient had low hgb values during the case and the platelets were down trending as well. the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3270 | 1220648-2025-48504 | 23605003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605003&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750127 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was placed after a patient rapidly decompensated post percutaneous coronary intervention. after a few hours on impella cp support, the patient's urine became red-tinged and labs were hemolyzing. the impella cp was repositioned and the urine cleared up and labs were no longer hemolyzed. the patient remained on support and was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to improper positioning which is use related due to physician deciding to leave pump in wrong position and hemolysis was improved after repositioning of the pump. device history review: the device passed all post sterile inspection checks.
3271 | 1220648-2025-48505 | 23605136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605136&pc=OZD | 10/24/2025 |  |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729447 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 89 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Anemia | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Pericardial Effusion | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced pulseless electrical activity requiring cardiopulmonary resuscitation (cpr), epinephrine, and bicarbonate. during cpr, the pump migrated out of the left ventricle and into the ascending aorta. multiple attempts to reposition the pump were unsuccessful. additionally, the patient developed a small pericardial effusion that was not big enough to require intervention. the patient was treated with multiple vasoactive drips, anticoagulants, and blood transfusions due to a hematoma at the axillary site as well as ongoing nose bleeds. later, the patient was successfully weaned from the pump and survived. additional information was received. the customer discarded the pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the cardiac arrest requiring resuscitation, pericardial effusion, hypotension, anemia, and epistaxis was not determined due to insufficient clinical details. the cause of the hematoma was most likely patient condition since patient had hematoma at axillary site along with nose bleeds. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the user during cpr. the device passed all post sterile inspection checks.
3272 | 1220648-2025-48497 | 23605139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605139&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026662609 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | [6 empty] | Ventricular Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a decision was made to place an impella 5.5 on a patient with severe mitral regurgitation with multivessel diagnosis that had an intra-aortic balloon pump due to peripheral artery disease. however it was reported that the impella 0.018 wire was in the left ventricle (lv) and advanced on wire. once the impella entered the axillary artery, the patient went into ventricular fibrillation. the staff defibrillated the patient and cardiopulmonary resuscitation was started. the wire was lost from the lv and was pulled back into the aorta. the decision was made to pull the impella and place on bypass. later, the patient was converted to extracorporeal membrane oxygenation. the impella 5.5 was aborted. additional information was received. there was no issue with the impella 5.5 itself. the pump is not available for return. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation: the cause of the injury was unable to be determined due to insufficient clinical details. resuscitation: the cause of the injury was unable to be determined due to insufficient clinical details. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be use issue related as wire was lost from lv and was pulled back into aorta, decision made to pull impella and place on bypass. device history review: the device passed all post sterile inspection checks.
3273 | 1220648-2025-48506 | 23605504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605504&pc=OZD | 10/23/2025 | 02/19/2026 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686599 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Respiratory Failure | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | A patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the patient was hypoxic when transferring over in the operating room for transfer. a major bleed was noted and one unit of blood was given. the patient had arrhythmia ventricular tachycardia that resolved by pump repositioning. four days later, the patient experienced supraventricular tachycardia. the patient remains on impella support. | The complainant reported additional information upon request: "at this time, i do not recall specific details regarding the reported events from october 24, 2025 beyond what is documented in the complaint and patient updates section i am not aware of any additional device malfunction beyond what was report. the device is not available for a return or further evaluation." |  |  |  |  | The cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the cause of the respiratory failure was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b, explant date, has been added.
3274 | 1220648-2025-48507 | 23605578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605578&pc=OZD | 10/21/2025 |  |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732059 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the hospital with chest pain and a left heart catheterization revealed lesions in the left anterior descending artery (lad) and right coronary artery (rca). one drug eluting stent was placed in the lad and one in the rca. following the coronary intervention, an impella cp was placed via the right femoral artery with no complications and the patient was sent to the unit for rest and recovery. on day three of support, it was reported that the patient experienced ventricular tachycardia that self-resolved. amiodarone was administered as a result of the arrhythmia and additionally the patient was administered levophed to achieve a mean arterial pressure greater than 60 mmhg. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the device was not returned for evaluation. it was reported that the device was explanted and not saved by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | H6: type of investigation, added codes b11 and b13. h11: added the results of the investigation. investigation summary no product was returned for investigation. ventricular tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1970125 device history review ==> device with sn: (b)(6) passed all post sterile inspection checks.
3275 | 1220648-2025-48508 | 23605917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605917&pc=OZD | 10/26/2025 |  |  | 11/20/2025 |  |  | 10/26/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207017 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller generated a placement signal unreliable alarm during patient use. tracing did not appear on the monitor for approximately 10-15 minutes. the placement signal and left ventricle tracing returned and are currently displaced. impella support continues. this is one of two related reports. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h10 for the related report number. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3276 | 1220648-2025-48509 | 23605989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23605989&pc=OZD | 10/26/2025 |  |  | 11/20/2025 |  |  | 10/26/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739654 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bloody stools and low platelet counts. heparin was discontinued and the patient was successfully weaned off the pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the bleeding is determined to be patient condition because the bleeding is noted at a non-impella site. the device passed all post sterile inspection checks.
3277 | 1220648-2025-48511 | 23606201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23606201&pc=OZD | 10/26/2025 |  |  | 11/20/2025 |  |  | 10/27/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026751109 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop. the supporting automated impella controller was exchanged to continue patient support. no patient harm was reported. this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. | [5 empty] | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the newly completed investigation. investigation summary: the investigation was updated due to data logs being received. temporary pump stop: clinical information reported the pump stopped and controller failure before high motor current and pump stopped. the automated impella controller (aic) was replaced and the issue did not happen on the replacement aic. the pump was able to support the patient. the data logs confirmed temporary pump stop and show a alarm controller error pump stop. there were no motor current spikes outside p-level changes and no pump related log abnormalities leading to the pump stop. in addition, the pump was able to run without issues on the replacement console. the cause of the issue was not traced to the pump as log analysis revealed no pump related log abnormalities and pump was able to run on replacement console without issues. device history review: the pump passed all post sterile inspection checks. | The cause of the pump stop was not established as no product or logs were returned for analysis. the device passed all post sterile inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3278 | 1220648-2025-48510 | 23606318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23606318&pc=OZD | 10/26/2025 |  |  | 11/20/2025 |  |  | 10/27/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1905647 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop along with a failure of the automated impella controller (aic). a new aic was used to continue patient support. no patient harm was reported. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to section h10 for the related report number. | [5 empty] | A4, a5, and a6 are unknown. the automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3279 | 1220648-2025-48512 | 23606385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23606385&pc=OZD | 06/02/2025 |  |  | 11/20/2025 |  |  | 10/27/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646254 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through an investigational study that a patient implanted with an impella cp developed a hematoma at the groin. no intervention was performed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks.
3280 | 1220648-2025-48513 | 23606673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23606673&pc=OZD | 10/27/2025 |  |  | 11/20/2025 |  |  | 10/27/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726031 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male | 90 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced organ failure. a computed tomography scan showed the patient had severe mesenteric ischemia. care was withdrawn and the patient expired. | [5 empty] | A5 and a6 are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned for investigation. the cause of the organ failure was not determined due to insufficient clinical details. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3281 | 1220648-2025-48514 | 23606813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23606813&pc=OZD | 10/27/2025 |  |  | 11/20/2025 |  |  | 10/27/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705330 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Female | 55 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an unreliable placement signal was reading incorrectly. the issue resolved without intervention. additionally, lysis was given due to purge pressure and purge flows issues. the patient was successfully transplanted, the pump was explanted, and the patient survived. additional information was received. the impella 5.5 was not available to be returned for investigation. | [5 empty] | No product or logs were returned for evaluation. the cause of the optical signal issue could not be determined as no product or logs were returned for evaluation with limited clinical details provided. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. the device passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3282 | 1220648-2025-48498 | 23606848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23606848&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763973 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 85 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a decision was made to place an impella cp to optimize a patient for percutaneous coronary intervention (pci). the impella cp was successfully placed and the pci was performed. however, a hematoma was noted at the impella cp site. the impella was successfully weaned and explanted at the end of the case. additional information was received the impella cp was explanted and tossed out. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hematoma: the cause of hematoma was unable to be determined as limited clinical details were provided and no product was returned for review. h6 codes were corrected from a01 was corrected to a24.
3283 | 1220648-2025-48516 | 23607047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607047&pc=OZD | 01/07/2025 | 12/23/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ischemia | Respiratory Failure | Pericardial Effusion | Heart Block | Heart Failure/Congestive Heart Failure | Renal Impairment |  |  |  | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support for acute myocardial infarction complicated by cardiogenic shock. while on support, oozing noted at impella site. heparin was held. pt ventricular fibrillation arrested with cardiopulmonary resuscitation and epinephrine transthoracic echocardiogram shows pericardial effusion requiring pericardiocentesis and pericardial drain placement. heart rhythm converted to bbb. transfusion to be given for low hgb in seeing of high pericardial drain output. the patient developed new-onset acute systolic heart failure characterized by reduced left ventricular function and contributed to multi-organ dysfunction, including aki and hemopericardium. patient developed new-onset ventricular tachycardia (vt). the arrhythmia contributed to clinical deterioration and was managed with amiodarone. the patient developed acute renal failure. clinical suspicion for acute tubular necrosis (atn) was high due to prolonged hypotension and ischemic insult. the patient developed acute hypoxemic respiratory failure associated with pneumonia and esbl e. coli infection, and required intubation. | [5 empty] | The investigation for the reported major bleed, respiratory failure,tachycardia,ischemia, and infection has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed, respiratory failure,tachycardia,ischemia, and infection could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3284 | 1220648-2025-48518 | 23607095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607095&pc=OZD | 10/21/2025 | 12/30/2025 |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737105 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 82 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support during a high-risk mitral valve replacement and possible tricuspid valve repair or replacement. the patient arrived at the hospital for the scheduled and the mitral valve replacement was successfully performed with the patient on cardiopulmonary bypass (cpb). the physician decided to place the impella 5.5 to remove the patient from cpb. the impella 5.5 was implanted via the right axillary artery without complications, and the patient was successfully removed from cpb. the patient¿s chest was closed, and the patient was transferred to the intensive care unit (icu) for rest and recovery. after arrival at the icu, it was noted that the impella device required repositioning to point the impella inlet away from the mitral valve and point it towards the apex. the patient was given fluid volume and blood products and heparin was withheld. the following day it was noted that the patient¿s chest tube discharge was slowing down, and heparin remained withheld. on day four of support, the patient was transfused with an additional two units of blood products, and it was noted that the pump was having suction alarms at performance (p) level 4-5. the medical team thought the suction was likely related to the patients fluid volume and lasix was administered. an echocardiogram was ordered and the impella was repositioned. the following day, it was noted that since the pump was repositioned there were no further suction events, and the p level was increased to p7. on day six of support, the patient was brought to the operating room for a chest wash out. upon returning to the icu it was noted that the patient had significant chest tube output. heparin was withheld until the chest tube output decreased. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. major bleed: the cause of the injury was not determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. data log shows several suction alarms while running on high p-levels, especially around midnight on 23-oct and the following day. no abnormalities were reported regarding motor current spikes or positioning issues. placement signal showed a decreased trend compared to previous days with reported suction events on 24-oct. as per the clinical details, blood products were given during the time of reported events, along with some repositioning, which allowed an increase in p-level without suction. the cause of the suction issue was most likely related to patient volume status and preload, based on the given clinical details and data log review. device in wrong position: the cause of the improper position issue was not determined without returned product investigation and sufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | H6 medical device problem code added. the impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
3285 | 1220648-2025-48499 | 23607212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607212&pc=OZD | 10/21/2025 |  |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627446 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support the patient admitted in with st-elevation myocardial infarction and pulmonary edema. the pump site had an ooze that did not require intervention. the pump was explanted on the second day of support with concerns for hemolysis. | [5 empty] | A5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Additional information was received indicating that the patient did not experience oozing as initially reported.
3286 | 1220648-2025-48517 | 23607264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607264&pc=OZD | 10/23/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736328 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support for coronary artery bypass graft. while on support, patient was diagnosed with potential hit+ prior to surgery. chest was closed and up to the unit after multiple vasoactive medications started and blood products given. | [5 empty] | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary no product was returned for investigation. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history lot ==> device lot: 1969461 device history review ==> device with sn: (b)(6) passed all post sterile inspection checks.
3287 | 1220648-2025-48515 | 23607270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607270&pc=OZD | 10/26/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/27/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709702 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced migration of the impella into the aorta. the impella could not be repositioned so the pump was explanted and replaced with an impella cp. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was dislocated into aorta by the patient as the patient was in delirium phase and restless. the device passed all post sterile inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3288 | 1220648-2025-48520 | 23607402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607402&pc=OZD | 10/21/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026738801 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | Cardiac Tamponade | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient's chest was reopened for treatment of tamponade. the patient remained intubated and spiked a fever overnight. cultures were sent for possible pneumonia. the patient remained on p5 and not increasing due to right ventricle failure. pulmonary artery pulsatility index was 1.67 and creatinine was 4.82. continuous renal replacement therapy was started. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. infection: the root cause of the injuries were not determined due to insufficient clinical details.
3289 | 1220648-2025-48521 | 23607468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607468&pc=OZD | 10/21/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/21/2025 | 11/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | Hemodynamic instability | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was placed with an impella cp and was tachycardic and hypertensive. the patient was positive for anoxic brain injury, and the family made the decision to withdraw care. the patient expired soon after support was stopped. | Clinical review/rationale: an impella cp was inserted via the femoral artery for the support of a 40-year-old female patient. the patient had arrested with ventricular fibrillation at home and was coded by emergency services for 45 minutes, inclusive of defibrillations. any other medical history or comorbidities were unknown, though the patient was known to use cocaine. the patient was having tachycardia and hypertension, presenting in scai stage e shock. after presenting to the cardiac catheterization lab the team medically was managing the patient and revealed the coronary arteries were clear of disease, but the impella cp was placed for hemodynamic support. while on support the arrythmia and hypertension persisted, as did the medical management with nitroglycerine and propofol, but the diagnosis of an anoxic brain injury was made. care was withdrawn and the patient expired after 2 days of support. the death is conservatively being reported although an unlikely contributing factor to the death, and more likely the patient's presenting native critical condition and cocaine use, as presented in scai stage e. |  |  |  |  | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the tachycardia, stroke: the root causes of the injuries were unable to be determined due to insufficient clinical details. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5 revised due to additional information/clinical review/rationale received.
3290 | 1220648-2025-48519 | 23607636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23607636&pc=OZD | 10/23/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/23/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733797 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support for coronary artery bypass graft. while on support, patient went into atrial fibrillation with rapid ventricular response. amio bolus given and intravenous drip started. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was not returned for evaluation. paroxysmal atrial fibrillation : in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review ==> the pump s/n: 621506 passed all the post sterile inspection checks.
3291 | 1220648-2025-48522 | 23608144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23608144&pc=OZD | 10/24/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1648111 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 79 KG |  |  | Death | Yes | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had open heart mitral valve replacement and was unable to come off bypass. the patient was placed on extracorporeal membrane oxygenation. the patient¿s right heart pressures were elevating and upon further evaluation the impella rp flex was decided to be placed for worsening right ventricle failure. the patient was successfully placed on impella rp flex support and about five minutes later a yellow alarm displayed ¿controller error - switch to back-up controller.¿ the patient was still getting 3.2 liters of flow on p-9 with a pulsatile motor current, but the placement signal was blank. the controller was switched to a backup and there were no issues with the new console. it was further reported that the patient was unable to recover and expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Console logs from the reported day of event are consistent with complaint and show an controller error alarm controller error and controller error triggered during the clinical procedure. no alarms related to ¿pump not being detected¿ were observed in the console logs. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. root cause: the root cause for the reported controller error alarm was due to the optical bench fw communication protocol anomaly. no issues were identified in the aic regarding the reported failure mode of ¿pump not detected.¿ e1 added initial reporter state as it was omitted from the initial report that was submitted.
3292 | 1220648-2025-48523 | 23608550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23608550&pc=OZD | 10/24/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026766105 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 58 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was alarming high purge pressure, and purge system blocked. tissue plasminogen activator was added to the purge to improve flow. the patient continued with support and was explanted as planned. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. purge pressure high/purge system blocked: the cause of the high purge pressure could not be determined as no product or logs were returned for analysis and limited clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3293 | 1220648-2025-48524 | 23609237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23609237&pc=OZD | 10/24/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716049 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | Yes | Decrease in Pressure | Increase in Pressure | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient presented for percutaneous coronary intervention (pci) and impella cp placement due to severe left main lesion. pci was performed and an impella cp was placed. during impella cp support, there was a clinical call reporting purge system fluctuation. the nurse reported decrease in purge flow. no active alarms were noted. purge fluid remained 5% dextrose in water with bicarbonate. no kinks were observed. during the call, purge flow/purge pressure was returned to baseline. additionally it was noted that the pump was repositioned due to hemolysis. the patient remained on support for a total of 32.95 hours and was weaned and explanted successfully. additional information was received. the product was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hemolysis: clinical details noted that pump was repositioned under echo a few hours in to support due to hemolysis and p-level was turned down to p-5. data logs were reviewed and lt logs showed no motor current spike to indicate ingestion or positioning alarms. brief instance of suction alarms for a few minutes then resolved on its own. the cause of the hemolysis could not be determined as no product was returned and limited clinical details were provided. purge pressure high: clinical details noted that purge flows were decreased with no active alarms. purge flows and pressure returned to baseline on its own. data logs were reviewed and lt log shows a gradual but choppy rise in purge pressure over approximately 30 minutes before a singular "purge system blocked" alarm, indicating a possible gradual buildup of biomaterial on the purge gap. then a sudden resolve in purge pressure and increase in purge flow occurs with no change in motor current or pump flows, indicating possible cause of increase in purge pressure due to kink in the purge line. no motor current spike at time of purge pressure rise was observed. the cause of the high purge pressure could not be determined as no product was returned and data logs do not show a consistent cause of the increase in purge pressure. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | Section h6 codes were updated.
3294 | 1220648-2025-48526 | 23609433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23609433&pc=OZD | 10/24/2025 | 01/05/2026 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1675422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 61 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Display or Visual Feedback Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that they encountered automated impella controller (aic) issues concerning the display and an unexpected controller shutdown. it was reported that the aic console screen went blank. additionally, the aic underwent an unexpected controller shutdown. the aic abruptly shutdown and automatically restarted. there was no alarm observed prior to shutdown. the registered nurse was advised to have back up controller on standby if problem reoccurred. | [5 empty] | The device was returned. the failure mode was not reproduced during investigation on engineering lab bench. the console¿s display and associated cables were inspected, and no abnormalities were found. the root cause of the aic abruptly shutting down, automatically restarting, and triggering unexpected controller shutdown alarm was due io-graphics process failure. however, the exact reason for this process failure could not be determined. no issues related to the display were identified within the aic. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3295 | 1220648-2025-48525 | 23609450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23609450&pc=OZD | 10/24/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666039 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Ischemia | Hematuria | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a decision was made to place an impella cp after a patient experienced shortness of breath and required bilevel positive airway pressure. the impella cp was successfully placed and support began. during support, it was noted that pulses were unable to be deplored and the patient's urine had a slight red tinge. it was noted that when the patient attempted to get out of bed, he yanked on the impella cp and pulled it back into the descending aorta and that was seen on x-ray. dressing site had an old ooze but no hematoma was present. additionally, the urine turned frank red. the impella cp was advanced back across the valve without issue. percutaneous coronary interventions were performed to the left anterior descending artery. an echocardiogram was performed and showed the impella in the mid-left ventricle. ejection fraction was noted to be very poor. the patient became hypotensive on return requiring levophed to be started. 500 milliliters of albumin was also given. the patient was intubated upon arrival with concerns for him pulling the impella back again. they were now unable to doppler pulses and the leg was cold. the team discussed and planned to proceed with escalation of the impella 5.5 with surgical repair of the right groin site. the impella cp was removed and vascular surgery repaired the artery. a clot was removed at the insertion site. the patient continued to clot so the artery was patched. additional information was received that the pump was not saved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | H6 added medical device problem code as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. ischemia, thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. hematuria, hypotension: the cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: no product was returned. urine turned frank red around 4am. after reviewing the logs, multiple suction alarms were observed. the cause of hemolysis was most likely use issue related since based on patient updates hemolysis resolved/managed with pump reposition. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient. device history review: the device passed all post sterile inspection checks.
3296 | 1220648-2025-48527 | 23609610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23609610&pc=OZD | 10/24/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738732 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 94 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp upon arrival to intensive care unit, the patient¿s gauze saturated blood for groin site. site was redressed with topical hemostatic products and new gauze. physicians np contacted physician w update and current drops which included aggrastat. aggrastat discontinued, light pressure femstop applied to strap pressure. hgb was 11.8 at this time. no blood products were ordered. and oozing appeared to be controlled post femstop. the patient became a do not resuscitate. the patient continued to deteriorate. decision made to withdraw care and the patient expired. | [5 empty] | H6: added codes to type of investigation: b11, b13. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot: 1970179. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3297 | 1220648-2025-48528 | 23609965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23609965&pc=OZD | 10/24/2025 | 11/20/2025 |  | 11/20/2025 |  |  | 10/24/2025 | 11/20/2025 |  |  |  |  | 11/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713404 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient experienced limb ischemia on impella cp and was escalated to the 5.5. the plan was to let the patient rest and unload with the 5.5. during 5.5 support it was noted that the patient had several suction events, however resolved with 1 l ivf bolus and decreasing impella. the patient also was having runs of non sustained ventricular tachycardia when agitated/coughing and with endotracheal tube suctioning. the 5.5 was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3298 | 1220648-2025-48456 | 23592403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23592403&pc=OZD | 10/20/2025 | 01/05/2026 |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324766 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller would not detect purge cassettes. the console was swapped out and support continued. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to section h10 for the related report number. | [5 empty] | Investigation summary : data review: imc logs from the complaint date 10/20/2025 revealed that the console issued the purge disc not detected alarm 20/10/25 11:35:39 imcgui: event set: #402 from several times and follow with alarm 20/10/25 11:39:31 imcgui: event set: #407 from purge system open [pressure < 100mmhg for 20s or more] - alarm issued. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer - (see attached image purge flag jpg ). root cause:the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. (b)(4) documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. (see attached image glucose build up jpg). device history review : q1) service history review - are there any qns associated with the complaint device¿s most recent service report? no. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a. q3) complaint history review - have there been any other complaints against this console? no. phr summary: there were no issues related to complaint discovered when reviewing the product history of console (b)(6). | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3299 | 1220648-2025-48455 | 23592474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23592474&pc=OZD | 10/29/2025 |  |  | 11/19/2025 |  |  | 10/29/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736990 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 78 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the patient had on/off arrhythmias. care was withdrawn and the patient expired. this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. | [5 empty] | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. arrhythmia: the root cause of the injury was not determined due to insufficient clinical details. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
3300 | 1220648-2025-48451 | 23592591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23592591&pc=OZD | 10/22/2025 | 11/18/2025 |  | 11/19/2025 |  |  | 10/22/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681469 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after seven days on impella cp support without any pressure event and patient good evolution, the pressure began to increase so the team decided to explant. due to axillary access, the device was removed out of the graft. a little thrombus around the motor house was observed. although during the implantation bicarbonate was used, the team changed to heparin during patient support. the patient suffered an ischemia event in the second finger of the left hand. the issue resolved hours later. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3301 | 1220648-2025-48457 | 23592759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23592759&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708074 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost pulses in the right extremity. a femoral-femoral bypass was placed to resolve the ischemia. it was also noted that the perfusion catheter was not providing enough flow to the lower extremity. the impella cp was explanted and the patient survived. additional information was received. the impella is not available for return. the impella was removed by vascular surgery due to small vessel size. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned for investigation. the cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3302 | 1220648-2025-48458 | 23592843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23592843&pc=OZD | 10/20/2025 | 01/22/2026 |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728374 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Pain | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was implanted with an impella cp and experienced an abrupt rhythm change to wide complex bradycardia and oozing from the impella insertion site. the patient also complained of pain at multiple sites. the impella blue hub appeared to be partially under the tissue so the angle of entry was adjusted. the patient was transfused one unit of blood. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | Investigation summary: impella cp + purge cassette returned for investigation. the pump was investigated. the sterile sleeve and tuohy-borst valve had dried blood on them. no damage to the sterile sleeve, cannula, tuohy-borst valve, or the blue hub were noted. it was seen that the blue suture hub had dried blood on it from all sides. no other abnormalities were seen. access site adverse event: based on the clinical details, oozing was noted from the impella access site, which subsidized after changing the angle of entry. the blue hub was noted to appear partially under the tissue. the root cause of the access site bleeding was most likely use related as the blue hub was partially inserted in the skin and the initial angle of insertion was improper as noted from the clinical details, and product evaluation confirmed the bleeding around the suture hub. bradycardia: clinical details mention that the patient suffered from bradycardia on 20 and 21 oct 2025. in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pain: clinical details mention that the patient was complaining about pain at multiple sites while on impella support. the root cause of the injury was not determined due to insufficient clinical details. device history lot. device lot: 1967742. device history batch. subcomponent lot: n/a. device history review: pump #621663 passed all post-sterile inspection checks.
3303 | 1220648-2025-48459 | 23592957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23592957&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660460 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a gastrointestinal bleed. the patient was transfused four units of packed red blood cells and two units of platelets. the patient continued to decompensate, requiring pressors. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3304 | 1220648-2025-48460 | 23593035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23593035&pc=PYX | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1477997 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon powering on an impella automated impella controller the screen had a blue screen and generated a "system self-check failed change console immediately" alarm. no harm to patient management was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The returned console was tested after arriving in field service. the reported failure mode was reproduced during testing. visual inspection confirmed the power manager corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3305 | 1220648-2025-48461 | 23593198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23593198&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682519 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 76 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an embolic stroke post removal which is suspected to be related to the impella. the patient survived. | [5 empty] | The cause of the injury was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3306 | 1220648-2025-48462 | 23593294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23593294&pc=PYX | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675561 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 122 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced a placement signal not reliable alarm. the placement signal was absent and the impella flow calculation was disabled. impella support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The pump was not returned for investigation. the cause of the placement signal issue was not determined due to no product returned and inconclusive data log review. the device passed all post sterile inspection checks.
3307 | 1220648-2025-48463 | 23593616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23593616&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614136 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 135 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced hemolyzed labs and hematuria. the impella p level was reduced to resolve the issue. later, care was withdrawn and the patient expired. additional information was received. the family decided to withdraw care and the pump was discarded by the staff. | [5 empty] | The cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the death was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3308 | 1220648-2025-48464 | 23593924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23593924&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731173 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a suction alarm and ventricular tachycardia. the tachycardia resolved without intervention. the patient is on a tablet form of amiodarone now. later, the patient was successfully weaned from the pump and survived. additional information was received. the pump was discarded by the staff. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed a suction alarm around 12:20pm on 18-oct-2025. impella position unknown alarms were also seen during this time. no motor current spikes outside of p-level changes were seen. no other abnormalities were seen. clinical details mention that the patient experienced ventricular tachycardia for three minutes during support. one suction alarm was also noted. pump suction: the root cause of the pump suction was not determined due to insufficient conclusive evidence from the data logs and clinical details, and unreturned product. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. device history review: the pump passed all post-sterile inspection checks.
3309 | 1220648-2025-48465 | 23594064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594064&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743647 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Great Vessel Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a left main perforation. the patient was treated with epinephrine and levophed, and heparin was withheld. anticoagulation was reversed and the sheaths were sutured in place. the patient also experienced hematuria. the impella position was slightly shallow so the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. vessel perforation (peripheral)/hemostasis: the cause of vessel perforation (peripheral)/hemostasis was unable to be determined as limited clinical details were provided and no product was returned for review. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3310 | 1220648-2025-48467 | 23594111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594111&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763765 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction (stemi)/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient woke up the morning of (b)(6) 2025 with chest pain and presented to the emergency room in stemi and a 100% occluded the right coronary artery (rca). percutaneous transluminal coronary angioplasty was performed, and a drug eluting stent was placed in the rca. following the procedure, the patient remained hypotensive, and an intra-aortic balloon pump (iabp) was placed. the patient¿s requirement for medical support continued to increase, and the decision was made to escalate the patient to an impella cp. the impella cp was successfully placed via the right femoral artery, and the iabp was removed. upon initiation of impella support, low flows were reported. a right heart catheterization was performed and revealed a pulmonary artery pulsatility index of 0.4. the medical team decided to peripherally cannulate the patient with veno-arterial extracorporeal membrane oxygenation (va ecmo) and continued support with impella cp. the patient was then transferred to an outside hospital for a durable left ventricular assist device work up. upon arriving to the hospital, the patient was noted to have a large hematoma at the impella insertion side, and there were suction alarms. the patient required epinephrine to maintain a stable blood pressure. the patient received two units of blood products and was emergently taken to the operating room for removal of impella cp and to reconfigure va ecmo from peripheral to central cannulation. the patient's outcome at the time of impella explant was survived, and the final patient outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Corrections: d4, h4. no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed suction alarms. the pump was run on two consoles, and suction alarms were seen on both consoles. purge pressure spikes were seen with no known pattern. there were no motor current or placement signal spikes/trends outside of p-level changes. no other abnormalities were seen. pump suction: the root cause of the pump suction was most likely patient condition based on provided clinical details and data log analysis. access site adverse event: the cause of the large bleeding hematoma was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3311 | 1220648-2025-48466 | 23594138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594138&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730329 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant escalated a surgical patient from intra-aortic balloon pump to the impella 5.5 pump for support. the 5.5 had high purge pressures on day 2 and was remedied by the addition of tissue plasminogen activator to the purge. the pump remained on for 10 days of support til wean and explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3312 | 1220648-2025-48468 | 23594248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594248&pc=OZD | 10/23/2025 |  |  | 11/19/2025 |  |  | 10/23/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723651 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | False Alarm | Backflow | Restricted Flow rate | Pumping Stopped | Device Sensing Problem | Patient Device Interaction Problem |  | Arrhythmia | Cardiac Arrest | Stroke/CVA | Fever | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ischemia | Ventricular Fibrillation | Thrombosis/Thrombus | [27 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced bleeding and a hematoma at the access site. the dressing was changed and angle matching was confirmed. heparin was withheld. and a femostop was applied. the patient had a fever and intermittent suction events so the femostop was removed. hemostasis was achieved. then, the patient spontaneously arrested during an echocardiogram, transitioning from ventricular tachycardia to ventricular fibrillation. the p level was reduced and compressions were started. the impella was repositioned. the patient was transfused three units of blood. a pump stop occurred with an alarm of "retrograde flow¿. the impella was successfully restarted but experienced multiple pump stops. the pump was explanted and a clot was noted at the inlet of the pump¿s cannula. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | H6 added medical device problem code.
3313 | 1220648-2025-48469 | 23594411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594411&pc=OZD | 10/17/2025 | 01/06/2026 |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729885 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support for a percutaneous coronary intervention (pci). following placement of the impella cp, the catheter was externally fixed to the patient¿s body using two fixation rings and the pci was performed successfully. after the procedure, it was found that the anti-contamination sleeve on the catheter had detached from the fixation point on the catheter plug and exposed the catheter. the sterile field was maintained, so the detached area was reinforced with tegaderm. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Investigation summary: product damage (anticontamination sleeve split/torn): the pump was returned for investigation. the sterile sleeve end on the red handle side was taped to the ring and found detached from the secure ring. clinical details indicate that after the procedure, detachment of the sterile sleeve and fixation ring on the catheter plug was observed, exposing the shaft. the sterile field was maintained, and the area was reinforced with tegaderm. the physician reported that both fixation rings were tightened during placement, but visual confirmation of the sterile sleeve fixation was not documented. the damage was confirmed on the returned device; however, the cause of the sterile sleeve damage could not be confirmed without sufficient clinical details. device history lot device batch: 1957247. device history batch subcomponent lot: na. device history review the pump (b)(6) passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3314 | 1220648-2025-48471 | 23594515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594515&pc=OZD | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705921 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bradycardia, decreased blood pressure and loss of pulsatility. a temporary pacemaker was placed to restore pulsatility, and later a permanent pacemaker was placed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. bradycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3315 | 1220648-2025-48470 | 23594540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594540&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758861 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 72 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ventricular Fibrillation | Gastrointestinal Hemorrhage | [7 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation and atrial fibrillation and was treated with amiodarone, levophed, and defibrillation. the patient had an increased white blood cell count and hematemesis, so heparin was reduced. after percutaneous coronary intervention the patient developed a gastrointestinal/intra-abdominal bleed. heparin was discontinued, two units of packed red blood cells were transfused, and protamine was administered. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the arrhythmia/ ventricular fibrillation/ infection: the root cause of the injury was unable to be determined due to insufficient clinical details. anemia: the cause of the injury was not determined due to insufficient clinical details. major bleed: hematemesis post procedure- heparin dose reduced, and patient taken to radiology for stat abdominal ct. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3316 | 1220648-2025-48472 | 23594584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23594584&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725300 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 71 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia. the impella was adjusted and the ventricular tachycardia resolved. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3317 | 1220648-2025-48473 | 23595377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23595377&pc=OZD | 09/02/2024 |  |  | 11/19/2025 |  |  | 09/02/2024 | 11/19/2025 |  |  | 12/31/2025 |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025442013 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Backflow | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there was a retrograde flow alarm. no further information was provided; however, it was later reported that the patient expired. | [5 empty] | Additional information was received in the form of medical safety review and noted the following: medical safety reviewed the event and noted there is not enough information to exclude impella 5.5 pump 3 as an associated factor in the patient's outcome. impella cp pump 1 and impella 5.5 pump 2 events are serious injury. related medical device reports. this impella 5.5 pump 3 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 1 and impella 5.5 pump 2. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3318 | 1220648-2025-48474 | 23595510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23595510&pc=OZD | 09/03/2025 |  |  | 11/19/2025 |  |  | 09/03/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1176067 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella device for mechanical circulatory support. shortly after the automated impella controller was switched on, a controller error appeared. no further information was provided. there was no reported harm or injury to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3319 | 1220648-2025-48476 | 23595523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23595523&pc=OZD | 09/11/2025 |  |  | 11/19/2025 |  |  | 09/11/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734367 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 101 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Renal Failure | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp was inserted to the left ventricle with no issues and ramped up. position was unstable until multiple troubleshooting attempts were made. during support it was noted that o2 sat continued to slowly rise. bedside echocardiogram completed and impella position was acceptable, but may need further adjustments at a future time. decided to leave where it is due to difficulty of positioning. the patient was maxed out on all 5 vasopressors and was on continuous renal replacement therapy and not pulling any fluid. ultimately care was withdrawn and the patient expired. | [5 empty] | A5 and a6 is unknown ( patient's race and ethnicity). the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3320 | 1220648-2025-48475 | 23595575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23595575&pc=OZD | 10/18/2025 | 12/30/2025 |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741897 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information received indicated that the pump was in the incorrect position. | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day 10 of impella 5.5 support, the patient experienced a gastrointestinal bleed and heparin was withheld. the patient¿s activated clotting time (act) dropped from the absence of heparin and it was noted that the flows on the impella 5.5 decreased. it was reported that on (b)(6) 2025 the patient¿s act was 26 seconds and on (b)(6) 2025 the act was 22 seconds. the patient was noted to be in progredient shock, and the impella had alarms for suction and low flows. it was noted that the patient had a predisposition to form blood clots and had active clots in multiple locations. additionally, it was noted that the patient experienced signs of hemolysis and that the patient¿s urine was hemolytic. the decision was made to exchange the impella 5.5 with a new impella 5.5. upon removal of the first impella 5.5, a clot was observed in the pump¿s inlet. the new impella was successfully implanted via direct aorta insertion. the patient remained on support with the replacement impella 5.5 for four days. it was noted that the patient¿s care was withdrawn and the patient expired during impella 5.5 support. |  |  |  |  | H6 added medical device problem due to information in b5. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Impella 5.5 (cut) returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. low pump flow/suction: clinical details report that the pump was experiencing low flows and suction alarms. data logs were reviewed and the reports were confirmed. returned product was investigated, and the catheter was cut. there were no damages/abnormalities noted on the cages, impeller, or cannula. after soaking the pump, biomaterial was confirmed to be in the purge gap and around the impeller. based on the clinical details provided, the patterns noted in data log review, and confirmation of biomaterial on returned product, the cause of the low flows/suction was determined to be biomaterial ingestion. hemolysis: clinical details report that the patient had hemolytic urine. returned product did not exhibit damages/abnormalities noted on the cages, impeller, or cannula. after soaking the pump, biomaterial was confirmed to be in the purge gap and around the impeller. based on the clinical details provided, the patterns noted in data log review, and confirmation of biomaterial on returned product, the cause of hemolysis was determined to be biomaterial ingestion. device in wrong position: data logs were reviewed and the positioning alarms on the reported date were confirmed. there was a biomaterial ingestion on that date followed by the positioning alarms 2 minutes later. however, the position in ventricle alarm was likely a correct trigger as the waveform appeared more ventricular than aortic, and the pressure values corresponded with the pump being in the ventricle. therefore, there is potential that some of the alarms were false due to a biomaterial ingestion event, however, there is potential that the pump was also being repositioned at the time to troubleshoot. returned product was cut at the catheter. there were no abnormalities noted at the catheter lock/repositioning unit. since insufficient information was provided, the cause of the device in wrong position could not be determined. thrombosis: clinical details report that thrombus was noted at the time of explant. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: clinical details report that the patient had a gastrointestinal bleed on (b)(6). heparin was paused. no further details were provided. the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | Medical safety/clinical reviewed the event. a 51-year-old female with acute decompensated chronic non-ischemic cardiomyopathy presented in society for cardiovascular angiography and interventions stage c cardiogenic shock, which progressed to stage d despite medical therapy. due to progressive hemodynamic deterioration, an impella 5.5 (pump 1) was implanted for mechanical circulatory support. during support, low pump flow was observed. in response, the device was exchanged for a second impella 5.5 (pump 2) to optimize hemodynamic support. it was reported that low flow may have been related to suspected thrombus formation. at the time of exchange, it was unclear whether the replacement pump provided full circulatory support. additional clinical information indicated the patient developed gastrointestinal bleeding, necessitating discontinuation of heparin therapy. as a result of the bleeding, activated clotting time values were maintained below the intended therapeutic target range. during this period of subtherapeutic anticoagulation, a decrease in pump flow was observed. thrombus was identified on the device at explant. the patient was described as having a prothrombotic state, with thrombi identified in multiple veins. hemolysis was suspected based on discolored urine. no further details regarding the gastrointestinal bleeding were available. the patient remained on impella 5.5 support for a total of 11 days. despite mechanical circulatory support, the patient¿s cardiogenic shock continued to progress in the setting of advanced underlying cardiomyopathy. given the refractory nature of her condition, a decision was made to withdraw care, and the patient subsequently expired. note: h6. component and investigation type, findings and conclusion remain unchanged as previously reported. | Correction: h11 erroneously stated no product was returned in the initial report. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. additional information: further additional information was requested for medical safety review and noted the following response: question(s): confirm if pump 1 gave support? and did pump 2 provided better flows? response(s): clinical rep "confirm nor neglect it." pump 2 provided full flow.
3321 | 1220648-2025-48477 | 23595862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23595862&pc=OZD | 10/19/2025 |  |  | 11/19/2025 |  |  | 10/19/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729652 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 55 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced multiple episodes of ventricular tachycardia. the patient was treated with defibrillation and increased vasopressors. later, the family chose to withdraw care and patient expired. additional information was received that the pump was not retained to return for evaluation. | [5 empty] | The cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3322 | 1220648-2025-48478 | 23596043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596043&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719299 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 60 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced continuous suction alarms. the device was repositioned but the alarm did not resolve. additionally, there was a loss of pulses to the left lower extremity. the complication was successfully managed with a placed distal perfusion catheter. later, care was withdrawn and the patient expired. additional information was received. the pump was not available for return. the patient was successfully explanted. no further issues with ischemia. | [5 empty] | No product was returned for investigation. data logs were returned for analysis and confirmed the complaint. the root cause of the pump suction was not determined due to lack of sufficient clinical details, unreturned product, and inconclusive evidence from the data logs. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3323 | 1220648-2025-48479 | 23596171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596171&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759772 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in with a st-elevation myocardial infarction. it was reported that a repeat echocardiogram was done to confirm placement due to hemolysis. the pump was pulled back three centimeters. the catheter was mid ventricular and measured three and a half centimeters. systemic vascular resistance is now 2200 which was down from 3980. the team plans on re assessing with possible explant the following day. the impella cp was explanted the following day and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary: product & data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in the proper position since hemolysis was resolved by repositioning of the pump. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. device history lot: device lot #: 1980806. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
3324 | 1220648-2025-48481 | 23596360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596360&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717808 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 72 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Renal Failure | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 placed for a patient with an extensive cardiac history and getting a coronary artery bypass graft for multivessel coronary artery disease. the patient first had an intra-aortic balloon pump and then was put on the impella 5.5. the patient had periods of intermittent suction and low pump flows and the pump was repositioned a few different times. furthermore, the patient coded multiple times while on support and there were multiple rounds of advanced cardiac life support given to the patient. it was also report that the patient was given continuous renal replacement therapy for worsening renal lab results. despite aggressive measures and efforts, the patient continued to decline and require additional doses of pressors/inotropes. the patient was extubated and all infusions were stopped. the impella was pulled back in the ascending aorta/graft and flows stopped. the patient expired shortly after. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6: added type of investigation code b13. h11: added the results of the investigation. investigation summary: additional information was provided. per the physician, the patient would combat ventilator and endotracheal tube causing the patient to eventually vagal which in turn dropped the patient¿s blood pressure (bp). due to patient having very limited native cardiac function, patient would enter ventricular tachycardia/ventricular fibrillation rhythm and/or require advanced cardiovascular life support (acls)due to blood pressure. low or blocked pump flow: the pump was not returned for investigation. clinical details indicate that the patient experienced low pump flows. low flows were often associated with suction events, suboptimal device position, and ongoing hemodynamic instability. data log shows several suction alarms during the case run, with placement signal showing a fluctuating trend. there was low flow reported while running on a steady p-level, without any noted motor current spike or positioning issue. the cause of the low pump flow issue was not determined without returned product investigation and sufficient clinical details. device in wrong position: the cause of the positioning issue was not determined without returned product and sufficient clinical information. clinical symptom (cardiac arrest): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptom (renal failure): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
3325 | 1220648-2025-48483 | 23596465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596465&pc=OZD | 09/18/2025 | 11/06/2025 |  | 11/19/2025 |  |  | 09/18/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707188 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 104 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was escalated from a intra-aortic balloon pump to a 5.5 to continue with heart failure work up. however the first attempt to place a 5.5 was aborted due to suspected calcium within axillary artery. impella cp was placed. after a couple of days on cp support decision was made to escalate to 5.5 but through left axillary artery. the cp was explanted and the second 5.5 was placed successfully. during support of the 5.5 the nurse reported that there were concerns for sepsis. the patient had automatic implantable cardioverter-defibrillator interrogated, found to have frequent episodes of atrial fibrillation with rapid ventricular response requiring cardioversion from external defibrillator. additionally the patient sustained ventricular tachycardia overnight, automatic implantable cardioverter-defibrillator shock once and an external defibrillation 100j converted to a paced rhythm. the patient was given amiodarone bolus and remains on drip. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted to provide medical safety/clinical review, correct previously submitted information, and to report the product investigation. b2 is being updated to include death and date of death. date of death the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. h1. type of reportable event has been corrected to death. h6. health effect - impact code f02 (death) has been added. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot. device lot: 1941985. device history batch. sub-component lot: n/a. device history review. the pump s/n: (b)(6) passed all the post sterile inspection checks. mjedical safety/clinical review. medical safety/clinical reviewed the event. the patient was admitted for congestive heart failure exacerbation and supraventricular tachycardia, underwent cardioversion, and was initially transferred to telemetry. due to worsening cardiogenic shock, the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va-ecmo) on (b)(6) 2025 (4 days prior to impella support), subsequently weaned from vasopressor and inotropic support, had diuretics discontinued, and was supported with a femoral intra-aortic balloon pump (iabp). four days later, escalation to impella support was planned. on (b)(6) 2025, the patient was taken to the cardiovascular operating room for implantation of an impella 5.5 (pump 1). a 10-mm graft was anastomosed to the right axillary artery, and impella 5.5 advancement was attempted under fluoroscopic. the procedure was aborted due to suspected calcification of the axillary artery, and the impella 5.5 was not implanted. an impella cp (pump 2) was subsequently implanted and following this implant, the patient became hypotensive with low mixed venous oxygen saturation (svo2). dobutamine was initiated and escalated without improvement and was associated with arrhythmia, requiring multiple cardioversions. the right hand was noted to be dusky with low flow. the patient required multiple defibrillation shocks. plans were made to escalate support with a left axillary impella 5.5, and the impella cp was explanted on (b)(6) 2025. an impella 5.5 (pump 3) was implanted and maintained through (b)(6) 2025. on (b)(6) 2025, nursing staff reported increasing vasopressor requirements, administration of multiple blood products and fluids, and concern for possible sepsis. nephrology adjusted continuous renal replacement therapy (crrt). the following day on (b)(6) 2025, interrogation of the patient¿s implanted cardioverter-defibrillator revealed frequent episodes of atrial fibrillation with rapid ventricular response, requiring cardioversion with an external defibrillator. on (b)(6) 2025, the patient developed sustained ventricular tachycardia overnight, receiving one icd shock and one external defibrillation at 100 joules, resulting in conversion to a paced rhythm. the patient received an amiodarone bolus and remained on continuous infusion. labs revealed an elevated plasma-free hemoglobin of 189, increased from the prior day. central venous pressure was reported as 10 mmhg, and the left ventricular pressure on the automated impella controller was reported as ¿31 mmhg, raising concern for possible impella positioning. echocardiography was pending and possible impella pump repositioning (echo results and if the impella was repositioned are unknown; no further details were noted on this). impella 5.5 and va-ecmo support were continued. subsequently, impella support was withdrawn overnight. no additional details regarding the exact timing or circumstances of withdrawal or patient outcome were noted. based on the information at hand, the aborted implantation of impella 5.5 (pump 1) was attributed to patient anatomy, specifically suspected axillary artery calcification. this a malfunction type of reportable event. the adverse events observed during impella cp (pump 2) with no device malfunction is a serious injury. regarding impella 5.5 (pump 3) support, there were recurrent atrial and ventricular arrhythmias, rising plasma-free hemoglobin, and concern for possible device positioning were noted. there is not enough evidence to exclude the impella 5.5 pump 3 as an association factor to the patient's outcome. however, patient's underlying condition contributed most to the outcome (advanced heart failure progression). the event story involves three product complaints. mpella 5.5 - mdr 1220648-2025-47536: malfunction. impella cp pump 2 - mdr 1220648-2025-47828: serious injury. impella 5.5 pump 3 - mdr 1220648-2025-48483: death. related medical device reports: this impella 5.5 pump 3 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 2 and impella 5.5 pump 1.
3326 | 1220648-2025-48480 | 23596466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596466&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/19/2025 |  |  | 10/23/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758709 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the patient encountered positioning issues, low pump flow, pump suction and patient had hemolysis during impella cp support. patient¿s labs were reported as hemolyzed. urine output still concentrated yellow. additionally, on echocardiogram, impella was measuring shallow again at 2.5 cm from av (aortic valve) annulus but also pointed toward mitral valve. the doctor at bedside attempted to reposition with echo guidance but it appeared that the pigtail was caught and the impella either advanced too far or was pulled back too shallow. attempted to torque, but flow decreased to 0.5l/min and started getting suction alarms even on p2. doctor called to discuss taking patient to catheter lab to reposition. patient remained on p2 and hemodynamics remained stable. options of replacing impella vs explant vs impella 5.5 intra-aortic balloon pump placement discussed between the team. the decision was ultimately made to take patient to catheter lab and explant the pump. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4 and h6 medical device problem code 01 was inadvertently omitted on the initial manufacturer device report. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hemolysis: the cause hemolysis was unable to be determined due to insufficient clinical details and since it is unknown if giving fluid bolus resolved the hemolysis. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. low pump flow, pump suction: no product or logs were returned. the cause of low pump flow/ pump suction cannot be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3327 | 1220648-2025-48484 | 23596725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596725&pc=OZD | 10/22/2025 |  |  | 11/19/2025 |  |  | 10/22/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724350 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 127 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. the stitch was tightened and a femstop was applied to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the major bleed is most likely use related since z stitch tightened and femstop with very gentle pressure above site applied and hemostasis was achieved. the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3328 | 1220648-2025-48482 | 23596827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596827&pc=OZD | 10/21/2025 | 01/07/2026 |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716280 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 106 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Fever | Unspecified Infection | Ischemia | Tachycardia | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient who had anemia of an unknown origin was implanted with an impella 5.5 and required a blood transfusion. no bleeding was identified, despite diagnostic workup for a retroperitoneal and gastrointestinal bleed. milrinone was started briefly but was discontinued due to ventricular ectopy. the patient continued to have intermittent ectopy with movement, despite adequate impella position. the patient had sustained ventricular tachycardia (vt) with a rate of 178. the patient refused amiodarone and the vt self-converted before defibrillation was needed. the patient has an automatic implantable cardioverter-defibrillator (aicd) that did not shock the patient and was being investigated. another episode of sustained vt occurred the following day and the aicd initiated one shock. the patient had short runs of vt and the alarms resolved on their own. the position of the pump was confirmed to be adequate on imaging. the patient was found to have fever and urinary tract infection requiring antibiotics. there was blood drainage from the impella access site. the oozing was managed by more frequent dressing changes. the patient remains on support. | The complainant reported an aerobic drainage culture obtained from right axillary impella insertion site had a positive result for klebsiella. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B5 updated. h6 updated. the pump will not be returned. the pump s/n: (b)(6) passed all the post sterile inspection checks. | Minor bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. anemia/ fever: the cause of the injury was not determined due to insufficient clinical details. arrhythmia/ tachycardia/ infection: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details.
3329 | 1220648-2025-48485 | 23596928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23596928&pc=OZD | 10/22/2025 |  |  | 11/19/2025 |  |  | 10/22/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726920 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost pulses in the lower extremity. the pump was explanted to resolve the ischemia. later, the patient decompensated and expired. | [5 empty] | No product was returned for investigation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3330 | 1220648-2025-31929 | 23597221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597221&pc=OZD | 03/18/2024 | 11/19/2025 |  | 11/19/2025 |  |  | 04/23/2024 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126087 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm appeared on the automated impella controller during patient support. the alarm was disabled and support continued with the implanted impella cp pump. care was eventually withdrawn, and the patient expired. the death is captured on the sister records of this record. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the following communication-related error "algs suspended opm signal invalid, and "invalid bad snr sample mask". the first placement signal not reliable alarm (#1036) was recorded approximately five hours after the pump was connected, followed by repeated #1036 alarms throughout the case. additionally, on 18 march 2024, a controller error opm comm crc invalid (#1045) alarm was observed. lt and rt log data indicates that the unreliable placement signal persisted for most of the case with low snr, which led to placement signal not reliable alarms. in the subsequent case with pump sn (b)(6), the snr remained low for the entire duration as well. device analysis: console ic2158 was not returned to fs for further investigation. root cause: the root cause for the placement signal not reliable alarm was not determined since aic was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3331 | 1220648-2025-48486 | 23597238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597238&pc=OZD | 10/22/2025 |  |  | 11/19/2025 |  |  | 10/22/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712052 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Sepsis | Tachycardia | Ventricular Fibrillation | Gastrointestinal Hemorrhage |  |  |  |  | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular fibrillation and ventricular tachycardia requiring cardiopulmonary resuscitation, defibrillation, and medication. the patient also experienced oozing and a hematoma at the axillary cutdown site. the patient was transfused multiple units of packed red blood cells and platelets, and pressure dressings and an addition suture were placed. the pump was repositioned multiple times and pacing was stopped but the axillary site continued to ooze. the patient was diagnosed with sepsis and gastrointestinal bleeding. the purge was changed. the patient then had additional ventricular tachycardia and decrease in blood pressure. blood transfusions continued. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the tachycardia, ventricular fibrillation, sepsis, infection, and arrhythmia was unable to be determined due to insufficient clinical details. the cause of the hypotension was not determined due to insufficient clinical details. the root cause of the minor bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the major bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
3332 | 1220648-2025-48488 | 23597356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597356&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/19/2025 |  |  | 10/23/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 50 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed onto impella support due to post cardiotomy cardiogenic shock/low cardiac output syndrome. while on support, oozing was noted from the proximal anastomosis of otitis media graft. two units of platelets and two units of fresh frozen plasma were given in the operating room. the patient had respiratory arrested during explant attempt prior to dressing removal. the doctor and team suspected over-sedation. one round of cardiopulmonary resuscitation was performed and epinephrine was given. the patient recovered. additional information was received. the pump was explanted and the patient had a good outcome. the pump was not saved to be returned for evaluation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. respiratory failure : the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3333 | 1220648-2025-48487 | 23597359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597359&pc=OZD | 10/21/2025 |  |  | 11/19/2025 |  |  | 10/21/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679010 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support an obese cardiac patient. the pump was placed however during the support there was low placement signal and suction. it was also reported that the pump came out of the left ventricle during patient movement. the pump went to the aorta but the patient was stable. it was noted that given that the patient was morbidly obese, the interventional cardiologist pulled (one centimeter) the impella under fluoroscopy. the plan is to reposition the device under fluoroscopy. no other information was provided about the repositioning. the pump was explanted and the patient survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device in wrong position: clinical reported during transfer from ccu to icu, the patient moved suddenly, causing the impella cp to migrate into the aorta. the cause of the issue was determined to be use related as evidenced by clinical information. pump suction: clinical reported ongoing suction. no patient volume issues were noted nor was ecmo use mentioned. log analysis shows suction alarm #14 events were triggered correctly. logs showed severe change in motor current at onset of suction alarms. the cause of the suction alarms was not established as no product was returned and limited clinical information was provided. optical sensor issue: clinical reported intermittent placement signal low alarms. alarms did not coincide with suction events. the data logs confirm alarm events. placement signal at the time of alarm occurrence was noisy and around 50mmhg. further rt log analysis showed intermittent dips in placement signal towards 20mmhg verifying correct trigger of alarms. motor current values were stable and unaffected. there were no position related alarms. no relevant patient condition or abnormality noted. no volume issues noted. the 5s parameters were within specification with contrast slightly variable. the cause of the issue was not established as no product was returned and limited clinical information was provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3334 | 1220648-2025-48489 | 23597537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597537&pc=PYX | 10/20/2025 |  |  | 11/19/2025 |  |  | 10/20/2025 | 11/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026629365 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient presented for coronary artery bypass graft. protamine was given at the end of the case and the patient suffered hemodynamic collapse requiring to go back on bypass. the decision for biventricular impella support was made. it was noted that impella cp support was started with suction noted when coming off support and the impella rp flex was placed. during support, it was noted that the patient went into ventricular tachycardia prior to leaving the operating room and advanced cardiovascular life support was required (acls). volume resuscitation continued post acls. bicarbonate was started. critical care was continued. the patient required constant attention to maintain volume. bedside echocardiogram showed completed sucked down left ventricle and right ventricle with the impellas only able to run around 2 liters per minute. after chasing volume: two units of blood, two amp of bicarbonate, 500 albumin, 500 lactated ringer's solution bolus amiodarone bicarbonate drip, they were able to get flows close to three liters per minute. more blood products were given: unit of platelets, cryoprecipitate, plasma and two more bicarbonate amp pushes, two units of blood and a setup for auto transfusion with chest tubes. the patient had been stable for several hours now. after continuous volume replacement, maxed pressor support was done and the family decided to withdraw care and the patient expired. this report is for the impella rp flex. the patient's death was reported against the impella cp in another report. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation was updated/corrected to incorporate additional coding for bleeding (e0506), as appropriately documented in the follow-up 2 report. the following was noted: no product was returned for evaluation. regarding resuscitation and cardiac arrest, a cause determination would require comprehensive clinical details; however, the cause could not be determined due to insufficient clinical information. similarly, for the major bleed, the cause could not be determined due to insufficient clinical details. low pump flow: the patient had low volume issues. after reviewing the logs, multiple suction alarms were observed. the cause of low pump flow was most likely patient condition related since patient required constant attention to maintain volume and also after giving more blood products and volume, the patient has been stable for several hours with flows close to 3l/min. device history review was completed and noted the pump passed all post sterile inspection checks. h6. investigation type, findings and conclusion codes, along with the h6 component code, were repopulated/updated accordingly based on this completed investigation. | The investigation of the reported failure mode has been completed. no product was received for evaluation. resuscitation, cardiac arrest: the cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: after reviewing the logs, multiple suction alarms were observed. the cause of low pump flow was most likely patient condition related since patient required constant attention to maintain volume and also after giving more blood products and volume, the patient has been stable for several hours with flows close to 3l/min. device history review: the device passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. a 56-year-old male with a history of atrial fibrillation and end-stage renal disease on chronic dialysis (4 years) was scheduled for coronary artery bypass grafting (cabg) due to complaints of shortness of breath. eight months prior, the patient had been hospitalized for three weeks for infective endocarditis with mitral valve vegetation. the vegetation was later described as calcified lesions on the posterior leaflet. at the time of the current admission, echocardiography demonstrated 1+ mitral regurgitation (mr) with a left ventricular ejection fraction (ef) of 35%. pulmonary function testing was pending per the medical record. cabg was performed. at the conclusion of cardiopulmonary bypass, protamine sulfate was administered for heparin reversal. subsequently, the patient experienced acute hemodynamic collapse requiring return to cardiopulmonary bypass. the patient was placed on biventricular mechanical circulatory support with an impella cp and an impella rp flex. prior to leaving the operating room, the patient developed ventricular tachycardia (vt). advanced cardiovascular life support (acls) was initiated, including cardiopulmonary resuscitation (cpr). during resuscitative efforts, the impella device position appeared to advance. bedside echocardiography was performed, and the device was repositioned by the physician to 89 cm. the ventricles appeared significantly underfilled (¿grossly dry¿), and aggressive volume resuscitation was initiated following acls. upon transfer to the intensive care unit (icu), the patient required continuous volume administration to maintain hemodynamic stability. ongoing bleeding was suspected, potentially exacerbated by chest compressions performed during cpr. bedside echocardiography in the icu demonstrated severely underfilled left and right ventricles with both impella devices only able to generate flows of approximately 2.0 liters per minute. the impella cp appeared positioned deep and was repositioned; however, low flow persisted due to inadequate preload. extensive volume and blood product resuscitation was administered, including multiple units of packed red blood cells, platelets, cryoprecipitate, plasma, sodium bicarbonate boluses and infusion, albumin, lactated ringer¿s solution, and initiation of autotransfusion via chest tubes. following aggressive resuscitative efforts, device flows improved to approximately 3.0 liters per minute. the patient remained temporarily hemodynamically stable with continuous high-volume replacement and maximal vasoactive support. despite ongoing efforts to control postoperative bleeding and maintain circulatory support, the patient remained critically ill. after approximately 23 hours of biventricular impella support and continued high-dose vasoactive therapy, the family elected to withdraw care. the patient subsequently expired. the impella cp and impella rp flex are being conservatively reported as associated with a death outcome. however, the patient¿s death was most consistent with refractory cardiogenic and hemorrhagic shock following protamine administration and separation from cardiopulmonary bypass, compounded by postoperative bleeding and hemodynamic collapse requiring acls and are considered the primary factors to the patient's demise. correction. after review of the case by medical safety, it was determined the impella rp flex is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3335 | 1220648-2025-48490 | 23597653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597653&pc=OZD | 10/24/2025 |  |  | 11/19/2025 |  |  | 10/24/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced low hemoglobin. one unit of blood was transfused. additionally, antibiotics were given over concern for sepsis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | No product was returned for investigation. the cause of the anemia was not determined due to insufficient clinical details. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3336 | 1220648-2025-31930 | 23597872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597872&pc=OZD | 06/23/2025 | 11/19/2025 |  | 11/19/2025 |  |  | 06/23/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398270 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that a cardiomyopathy patient was on impella 5.5 when the intensive care unit staff reported that the placement signal was suddenly gone. there was no alarms on aic screen and no harm to the patient reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on 6/16 5.5 pump (b)(6) was connected and run until 6/23. on 6/18 there was a placement signal not reliable alarm that did not resolve but the audio was silenced. after the pump was unplugged, the lamp was found to be outputting no light (0.44v) triggering controller error #1039. the lt and rt logs confirm that starting on 6/18 the placement signal went to 3000mmhg as the contrast oscillated between +/-3200% and snr went to 1. device analysis: on console return, when the console was booted up, the optical bench firmware was read as n/a. this was not reproduced during subsequent testing and is potentially unrelated to the complaint. the bad lamp and lost placement signal was not reproduced during testing. root cause: the cause of the intermittent optical signal issue was a defective lamp because the lamp persisted to malfunction after the pump was removed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3337 | 1220648-2025-48491 | 23597972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23597972&pc=OZD | 10/22/2025 |  |  | 11/19/2025 |  |  | 10/22/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603964 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp on the patient¿s pump was repositioned last night after patient had episode of vomitting followed by position alarms. added norepinephrine drip and lr bolus for cvp-3. post esophagogastroduodenoscopy, the patient developed tachycardia and decompensated requiring addition of vasopressin, dobutamine and amiodarone. the patient was slowly weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was most likely pulled out of position by the patient when they vomited. device history review: the device passed all post sterile inspection checks.
3338 | 1220648-2025-48492 | 23598121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23598121&pc=OZD | 10/23/2025 | 11/19/2025 |  | 11/19/2025 |  |  | 10/23/2025 | 11/19/2025 |  |  |  |  | 11/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682688 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support for high risk cardiac procedure. while on support, patient coded while transferring to intensive care unit. 1mg epi given, return of spontaneous circulation. patient had angina that resolved. | [5 empty] | The impella device has been discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. cardiac arrest/ pain: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3339 | 1220648-2025-48412 | 23579405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23579405&pc=OZD | 10/20/2025 | 12/31/2025 |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675650 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed on impella cp support for a acute myocardial infarction complicated by cardiogenic shock. while on support, the patient continued to remain severely acidotic and hypotensive despite medical management. the patient went into pulseless electrical activity arrest with cardiopulmonary resuscitation for 15 minutes before there was a return of spontaneous circulation again. the family decided to make the patient do not resuscitate and ultimately the patient passed shortly after. | [5 empty] | Investigation summary: cardiac arrest : the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot : 1931409. device history batch: sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device is not available for return. therefore, an investigation of the device is not possible. should any new information be received, a supplemental report will be filed.
3340 | 1220648-2025-48414 | 23579456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23579456&pc=OZD | 10/20/2025 | 01/06/2026 |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666843 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 83 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an inferior st-elevation myocardial infarction patient presented and had a right coronary artery ballooned and stented. the left main (lm) was shot and showed significant left anterior descending artery stenosis. the decision was made for an impella cp. it was noted that perclose x 2 were placed. an impella cp was inserted and initiated in auto-mode. the peel-away sheath was exchanged for a repositioning sheath. during suture placement, a large hematoma was palpated, and large amounts of groin access bleeding was noted. at this time, the physician pulled on the perclose strings and noticed they did not take. manual pressure was held to attempt to control the bleed with no success. the patient's hemodynamics had improved substantially, so the decision was made to remove the impella. one more perclose was attempted and the groin bleed ceased. | [5 empty] | Investigation summary: product was not returned for review. major bleed/hematoma: the cause of major bleed/hematoma was use related since bleed occurred during suture placement and resolved with additional perclose. device history lot: device lot #: 1915025. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3341 | 1220648-2025-48420 | 23580561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23580561&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622849 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the nurse notified the field team that the patient continued to experience persistent suction alarms and hypotension. complete blood count panel revealed a significant drop in hemoglobin. the doctor noted new abdominal distention and the patient was transported for a computed tomography which confirmed a retroperitoneal bleed originating from the right common femoral artery, the site of the impella cp insertion. orders were placed to apply a femostop over the right femoral access site. the patient was then transferred to the catheterization lab where an angiography demonstrated cessation of bleeding. a collective decision was made to leave the impella cp in place and continue close monitoring of the site. the impella cp has appropriate angle with gauze and the nurse is going to apply the dressing soon. blue hub to the skin with no bleeding currently. heparin remains off during this time. later while attempting to float the swan at bedside, the patient required defibrillation for ventricular fibrillation. swan placement was stopped. after the patient had good pulse pressure and was looking better and overnight slowly became more hypotensive. a low dose of levophed was started. no further bleeding from anywhere was noted but one unit of packed red blood cells was given this morning. still no heparin was being used currently. the patient still has red urine but is not concerned with hemolysis. the patient experienced massive bleeding during impella cp removal, thrombectomy and impella 5.5 insertion. coagulopathic was given after case. the patient was noted to be stable and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data logs were returned for analysis. data log analysis confirmed suction alarms persistently occurring. placement signal low and placement signal not reliable alarms were triggered at the beginning of the case but were resolved within a minute. impella position unknown alarms were seen. there were also preventing retrograde flow (impella flow high) alarms during suction alarms. no motor current spikes were seen outside of p-level changes. no purge related issues were seen. pump suction: the root cause of the pump suction was most likely patient condition based on the provided clinical details and data log analysis. access site adverse event: the root cause of the access site bleeding was not determined due to lack of conclusive evidence from provided clinical details. retroperitoneal hemorrhage: the root cause of the retroperitoneal hemorrhage was not determined due to lack of conclusive evidence from provided clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. ventricular fibrillation, thrombosis: the root cause of the injuries was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3342 | 1220648-2025-48421 | 23580627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23580627&pc=OZD | 10/19/2025 |  |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026765018 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia. the patient was intubated and the impella flow was increased. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3343 | 1220648-2025-48422 | 23580781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23580781&pc=OZD | 10/19/2025 |  |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750146 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 51 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a non-st-segment elevation myocardial infarction patient presented with severe disease in the right coronary artery (rca). the rca was stented however had no reflow was noted. the patient tanked pressures and attempted to access both groins for impella to support the patient. both illiacs were 100% occluded. surgeon cutdown was done of the left axillary artery. they ended up placing the sheath in the vein instead as well as the impella. the impella was removed and the surgeon switched to the right axillary artery. once surgery was attempted to close the access site, it would not stop oozing. a 14 french short cook sheath was placed back into the left axillary vein. along with thrombin, manual pressure was held to the site. the pump went into the body and was removed for 30-45 minutes. a new pump was used at that time. it was noted that the closure method was left in and closed the cutdown site around the sheath, and protamine was given. the second impella was successfully inserted into the left ventricle. a swan was floated through the left subclavian vein. then it was noted that the swan was removed, to attempt to close the vein surgically. ph level was found to be 7.08 and hemoglobin was 6. the patient received four units of fresh frozen plasma, six units of blood, and multiple activated protein kinase of sodium bicarbonate. all sites were dressed and draped. the pressure dressings were placed on both cutdown sites. ultimately, the patient expired on support. | [5 empty] | This report is being submitted to provide medical safety/clinical review and to correct and add coding to align with clinical information previously described in the initial mdr. h6 coding for malposition of device, blood transfusion, delay to treatment/therapy, and medication required have been added, however no new event details or outcomes are being reported. medical safety/clinical reviewed the event. a 69-year-old female with a medical history significant for coronary artery disease presented with a non-st-segment elevation myocardial infarction (nstemi). cardiac catheterization revealed severe right coronary artery (rca) disease. percutaneous intervention with rotational atherectomy and stent placement was attempted but was unsuccessful. the rca subsequently developed no-reflow. at this point the patient became hypotensive and was intubated. the decision was made to place the patient on impella mechanical circulatory support (mcs). femoral arterial access was attempted; however, both iliac arteries were found to be 100% occluded. cardiothoracic surgery was consulted for left axillary artery cutdown. however, the physician ended up placing the introducer sheath and impella cp pump 1 into a vein rather than the artery. the impella cp pump 1 was removed, and access was then obtained via the right axillary artery. surgical closure of the initial access site was unsuccessful, and a 14 fr sheath was placed, left in the vein. after approximately 30¿45 minutes, an impella cp pump 2 was successfully implanted. during this period, laboratory values demonstrated ph of 7.08 and hemoglobin of 6 g/dl. the patient received four units of fresh frozen plasma, six units of packed red blood cells, and multiple ampules of sodium bicarbonate. the patient was transferred to the intensive care unit on impella mcs and additional blood products (2 units) were subsequently administered. subsequent laboratory testing showed improvement in ph to 7.28 and hemoglobin to 10.2 g/dl. the impella cp pump 2 provided circulatory support for approximately six hours; however, the patient¿s hemodynamic status continued to deteriorate. the patient expired while on impella support. based on the available clinical information, the patient¿s underlying clinical condition contributed most to the death outcome (cardiogenic shock [scai stage d] secondary to acute myocardial infarction and severe underlying coronary artery disease) rather than the performance of either impella device. for reporting purposes, the event is conservatively assigned to the impella cp pump 1 due to a delay in initiation of mechanical circulatory support associated with inadvertent venous placement, which was caused by the user. this complaint involves two impella devices: pump 1: impella cp, with serial number: (b)(6). 14fr introducer used with pump 1, with serial number: (b)(6). this mdr specifically addresses pump 1: impella cp, with serial number: (b)(6), which is the subject of this report. separate mdrs will be submitted for the 14fr introducer associated with this complaint. refer to the related manufacturer report number as noted in section h10 for the medical device report on the 14fr introducer. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned. the cause of the bleed cannot be determined since insufficient clinical details were provided. the cause of the delivery issue was determined to be patient condition as the patient had occluded illiacs along with difficult anatomy and resistance in axillary artery preventing successful delivery of impella. the device passed all post sterile inspection checks.
3344 | 1220648-2025-48424 | 23580921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23580921&pc=OZD | 10/19/2025 |  |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711514 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 63 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 encountered damage to the red impella plug resulting in a pump stop. the pump was explanted and replaced with an intra-aortic balloon pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The cause of the pump stop was not established as no product or data logs were returned for analysis and limited information was provided. the cause of the product damage was not established as no product/images were returned for analysis and limited information was provided the device passed all post sterile inspection checks.
3345 | 1220648-2025-48423 | 23580929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23580929&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700740 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support the team was concerned about pump position due to its potential interaction with mitral valve. an attempt to reposition was made but wire prolapsed on aortic valve so the pump was left shallow. however, the a few days later, patient had runs of ventricular tachycardia and suction events. the pump was repositioned but it became entrapped in mitral valve apparatus and it was unable to be rotated away. p-levels were unable to be increased above level 6. as a result, impella support was escalated to impella 5.5. additionally, due to concern for hematoma, patient was given 1 unit of packed red blood cells. | [5 empty] | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. h10 this is one of two related reports. this report represents the impella cp and another report was submitted to represent the impella 5.5. investigation summary: the investigation has been completed. no product or data logs were returned for evaluation. clinical details mention that the impella position was concerning for interaction with mitral valve (mv) and repositioning was attempted. the impella was left shallow. arrythmia and suction events were observed during the case, and the pump was trapped in the mv apparatus, unable to rotate away. unable to increase the p-level above p-6 as suction alarms were triggered. the patient was also given blood products due to concern for hematoma. pump suction: the root cause of the pump suction was not determined due to lack of conclusive evidence from the provided clinical details and unreturned product and logs to confirm reported findings. device in wrong position: the root cause of the positioning issues was not determined due to lack of conclusive evidence from the provided clinical details. hematoma: the cause of the hematoma was not determined due to insufficient clinical details. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. device history review: the pump passed all post-sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3346 | 1220648-2025-48426 | 23581071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581071&pc=OZD | 10/19/2025 |  |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced bleeding so heparin was withheld. later, the patient expired on support. | [5 empty] | The cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
3347 | 1220648-2025-48425 | 23581082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581082&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611190 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female | 69 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp while repositioning it was pulled back across aorta and unable to recross. the device was removed and new impella cp implanted. there was no patient harm. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position across aorta during transport and was unable to be recrossed aortic valve. device history review: the device passed all post sterile inspection checks.
3348 | 1220648-2025-48328 | 23581157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581157&pc=OZD | 10/12/2025 | 11/12/2025 |  | 11/18/2025 |  |  | 10/12/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758553 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Sepsis | Renal Impairment | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support the patient had an acute kidney injury requiring dialysis. the patient became septic due to having an open chest. additionally doppler ultrasound done on the left upper extremity due to heavily bruised lower limb. concerns for clot in upper arm not seen on echocardiogram. bilateral feet distal half were also purple. the care team pushed the impella cp in by 1 centimeter the patient's chest was closed but patient had a rough night and lost pulsatility requiring multiple drips. the patient's white blood cell count was high and the patient was likely septic. care was withdrawn and the patient expired. | [5 empty] | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. sepsis/ ischemia : the cause of the injury was unable to be determined due to insufficient clinical details. acute kidney failure : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3349 | 1220648-2025-48427 | 23581219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581219&pc=OZD | 10/19/2025 |  |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728812 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced oozing at the access site. hemostasis was obtained with a mitral clip. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot device lot: 1960387. device history batch sub-component lot: n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3350 | 1220648-2025-48428 | 23581929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581929&pc=OZD | 10/19/2025 | 12/31/2025 |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657621 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia and bleeding. the patient was transfused one unit of packed red blood cells. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: tachycardia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot: 1911465. device history batch: sub-component lot: n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3351 | 1220648-2025-48430 | 23581942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581942&pc=OZD | 10/19/2025 |  |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690963 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia. an echocardiogram showed the pump was positioned deep in the ventricle. resistance was encountered when the pump was pulled back into proper position, so repositioning was halted. later, the patient was successfully weaned from the pump and survived. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product was returned. tachycardia: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump most likely got out of position as the impella console alarmed for position unknown after patient experienced a run of ventricular tachycardia (vt) after coughing. evice (sn: (b)(6) passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3352 | 1220648-2025-48431 | 23581944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23581944&pc=OZD | 10/08/2025 |  |  | 11/18/2025 |  |  | 10/08/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731297 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemoptysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 coughed up blood and was bleeding from the nose. the issue was resolved with systemic anticoagulation and transfusion of one unit of packed red blood cells. additionally, the impella generated a placement signal not reliable alarm. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b: updated the explant date. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. optical sensor issue: clinical details report psnr alarms on 19/oct and 26/oct. data logs were reviewed and the alarms were confirmed. there were a total of 3 psnr triggers during the case. at all three instances, contrast oscillated between +/- 3200 and snr dropped to zero. in all three cases, the alarm resolved and the ps calculation returned. this is a known console issue. product was not returned for investigation. based on the pattern noted in data log review, there was no problem found with the pump. epistaxis/hemoptysis: clinical details report that the patient had bleeding from the nose and was coughing up blood on 8/oct. the cause of the injury was not determined due to insufficient clinical details. the pump passed all post-sterile inspection checks.
3353 | 1220648-2025-48433 | 23583891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23583891&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026754542 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had arrythmia and was resuscitated. an echocardiogram was done and the pump was pulled back two centimeters. three days later, hemolysis was suspected. the impella potentially was a little too deep in the echocardiogram. no correction was done at the moment due to the patient being unstable. it was further reported that the patient care was withdrawn due to severe multiple organ failure and the patient expired. this is one of two related reports. this report represents the pump. another report with be submitted to represent the automated impella controller. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter - ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Medical safety/clinical reviewed the event. a 75-year-old female with unknown medical history presented in cardiogenic shock (scai classification unknown) and was implanted with an impella cp (pump 1) for mechanical circulatory support (mcs). on post-implant day 0, no aortic (ao) placement signal was observed, and the device was replaced with a new impella cp (pump 2). during this support, a loose connection was identified with the automated impella controller (aic) power cable. multiple electrical outlets were attempted; however, the aic did not charge as expected. however, it was noted that the aic started charging without further issue. while on impella cp pump 2 support, the patient experienced an arrhythmic event requiring brief resuscitation. the patient stabilized thereafter, and echocardiography was performed. the device was repositioned by withdrawing approximately 2 cm. approximately three days later, elevated motor current and potential hemolysis were noted and closely monitored. repeat echocardiography suggested the device may have been positioned slightly deep within the ventricle. no repositioning was performed at that time due to the patient¿s hemodynamic instability. subsequently, the patient was noted to have developed severe multiorgan failure (mof). care was withdrawn, and the patient expired. the device performance issues associated with impella cp pump 1 (loss of ao placement signal) and the automated impella controller (power cable/charging issue) are being reported as malfunction type of reportable event. these issues are not considered to have contributed to the patient¿s death and were limited to device performance and limited user management concerns. impella cp pump 2 will be reported as a death type of reportable event. the patient experienced device positioning concerns and an arrhythmic event during support; however, based on available information, the most likely cause of death was severe cardiogenic shock complicated by ¿severe¿ multiorgan failure. correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: insert organ failure/multiple organ dysfunction syndrome (e2342) and hemodynamic instability (e2343). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the final coding that accurately reflects the reported event. related medical device reports: this is one of three devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3354 | 1220648-2025-48432 | 23584020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584020&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680789 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support due to cardiomyopathy. upon priming 5.5, high purge pressure alarm alarming. despite troubleshooting, purge pressure reading 1100 and could not resolve. decision made to not place that device and prep a new device. there was no patient harm. | [5 empty] | B7 added information that was omitted from the initial report that was submitted. d4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. d9 the device was returned and date was added. h6 added code b13 as it was omitted from the initial report that was submitted. type of investigation, investigation findings and investigation conclusions were updated as the device was returned. h11 revised additional manufacturer narrative as the device was returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3355 | 1220648-2025-48434 | 23584292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584292&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690286 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that during implantation of an impella cp a cannula kink was encountered. the pump prolapsed in the aorta after placing the repositioning sheath. a new impella cp was implanted to support the patient. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of delivery issue was most likely use issue related since the pump was accidently pulled back across the valve when removing the peel away introducer and inserting the repo sheath. the cause of product damage/cannula kink was most likely use issue related since the pump was accidently pulled back across the valve when removing the peel away introducer and inserting the repo sheath resulting in a cannula kink. the device passed all post sterile inspection checks.
3356 | 1220648-2025-48435 | 23584387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584387&pc=OZD | 10/20/2025 | 11/18/2025 |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763960 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during support of impella cp the patient became hypotensive during process of suturing and dressing to the left and right groin sites. impella increased to p8 and levo 0.1 mcg/kg/min restarted. small hematoma was noted as superior and medial to impella insertion site. slight pressure held for a minute or two and hematoma decreased in size. no active oozing noted. upon undraping patient, unequal abdomen size noted. right side larger than left and firm to touch. due to agitation upon arrival at the lab, no staff could speak to abdominal assessment prior to procedure start. though hemoglobin was stable, dr. lin was concerned due to patient¿s persistent hypotension despite multiple chemical sensitivity and large dose pressor. lab staff accompanied the patient for computed tomography angiography to rule out a potential bleed then delivered patient to intensive care unit. the pump was explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3357 | 1220648-2025-48436 | 23584420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584420&pc=OZD | 10/20/2025 | 11/18/2025 |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716843 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was successfully inserted via the left femoral artery and the pump was connected. the cannula was across the aortic valve, however when the physician attempted to advance the repositioning sheath up the catheter length, the sheath did not move. the pump was aborted and a secondary impella cp was implanted for patient support. the patient was on support with the second impella cp for a total of 142.77 hours and the decision was made to escalate to an impella 5.5. the patient was taken to the operating room for upgrade to an impella 5.5. upon induction, the patient became acutely hypotensive, suction on the impella cp, right ventricle was dilated and dysfunctional on transesophageal echocardiogram, ventricular fibrillation arrest, cardiopulmonary resuscitation was performed. multiple shocks both internal and external were given. veno-arterial extracorporeal membrane oxygenation was placed emergently. the impella cp was successfully weaned and explanted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data logs were returned for analysis. data log analysis confirmed suction alarms occurring. a purge pressure spike was noted around the time of the second suction event, with abrupt drop in purge flows. no motor current spikes were seen outside of p-level changes. impella position in ventricle alarms were seen towards the end of the case. a single purge pressure high alarm was seen which resolved within a minute. placement signal low alarms were seen towards the end of the case. no placement signal spikes or trends were seen outside of p-level changes. mean flows were variable and slightly lower than the normal range for a specific p-level, especially after the second suction event. no other abnormalities were seen. pump suction: the root cause of the pump suction was most likely patient condition based on the provided clinical details and data log analysis. hypotension: the cause of the hypotension was not determined due to insufficient clinical details. ventricular fibrillation: the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3358 | 1220648-2025-48437 | 23584601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584601&pc=OZD | 10/21/2025 |  |  | 11/18/2025 |  |  | 10/21/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660547 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a impella cp pump placed to support the acute myocardial infarction/cardiogenic shock patient. there was a kink beneath impella outflow which obstructed the blood flow, hence low flow support. there are several angiogram picks and videos demonstrating the kink. impella had to be replaced with a different pump to support the patient. there was no patient harm. | [5 empty] | The investigation was completed. impella cp #593661 returned for evaluation. no data logs were returned for analysis. clinical details indicate that a kink was noted beneath the impella outflow which obstructed the blood flow, which caused low pump flow. images and videos under fluoroscopic imaging were sent which confirmed a cannula kink during support. the pump had to be replaced with a different pump to continue support to the patient after unsuccessful attempts made by the physicians to get rid of the kink by rotating, pushing, and pulling the pump. upon investigation of the returned pump, a cannula kink was confirmed under microscopic guidance. the kink was near the outflow cage of the pump around the same location as the attached fluoroscopic imagery received. no foreign material or biomaterial was seen in or around the impeller blades, inflow, and outflow cages. the pump was successfully primed and run in the lab, with flows within the expected range. no physical abnormalities other than the cannula kink were seen. the pump functioned per specification. product damage (cannula kink): the root cause of the cannula kink could not be determined due to the lack of sufficient clinical details and inconclusive evidence from the product evaluation. low or blocked pump flow: the root cause of the low or blocked pump flow was most likely due to the cannula kink based on provided clinical details, attached images, and evaluation of the returned product. device history review was completed and passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3359 | 1220648-2025-48440 | 23584822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584822&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716051 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3360 | 1220648-2025-48438 | 23584824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23584824&pc=OZD | 10/21/2025 |  |  | 11/18/2025 |  |  | 10/21/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 59 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had a impella 5.5-automated impella controller (aic) support placed for the patient admitted in for bypass surgery. during impella setup, customer inserted purge cassette and aic did not recognize that purge cassette was inserted. customer exchanged aic and purge cassette was recognized in new aic. there was no patient harm. this complaint is for the aic and another report will be sent for the impella 5.5. | [5 empty] | B7 added information that was omitted from the initial report that was submitted. d9 added date device was returned as it was omitted from the initial report that was submitted. e1 added initial reporter phone number and zip code as they were omitted from the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. type of investigation, investigation findings and investigation conclusions were revised as the device was returned. h10 added related report number. h11 revised additional manufacturer narrative as the device was returned at the time the initial report was submitted. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3361 | 1220648-2025-48441 | 23585244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585244&pc=OZD | 10/21/2025 |  |  | 11/18/2025 |  |  | 10/21/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026744083 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 59 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5- automated impella controller support placed for the patient admitted in for bypass surgery. while completing icu check in, placement signal was lost. motor current and flow remained on the screen. this did not occur during patient movement. patient was intubated and sedated. placement signal spontaneously returned. overnight, placement signal was spontaneously lost again and returned multiple hours later. placement signal has been intermittent since then. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3362 | 1220648-2025-48442 | 23585355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585355&pc=OZD | 10/19/2025 | 11/18/2025 |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736373 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 93 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The patient was successfully bridged to transplant. | The user facility reported that during impella 5.5 support, the patient had increased ventricular tachycardia. echocardiogram verified the impella placement but did show an encapsulated mural thrombus in the left ventricle (lv) not close to the inlet. the day after the patient had a liver biopsy, so bivalirudin was on hold. lv thrombus was discussed with the doctor and the nurse will watch motor current and flow closely. three units of packed red blood cells were given. the patient then continues to have a slow gastrointestinal (gi) bleed. the patient is going for another scope. the patient got another unit of blood for low hemoglobin. no impella alarms and remains of anticoagulation were noted. the patient continues with an active gi bleed. beside colonoscopy was done. the patient was no longer experiencing a gi bleeding per scope. the patient was able to eat a normal diet. the patient was stable. day three there was no gi bleed. the plan was to restart bivalirudin. hemoglobin was dropped to six overnight. the patient received two units of packed red blood cell¿s and turned off of bivalirudin. hemoglobin was more stable so bivalirudin will be starting soon. they were unaware where the bleeding was coming from, but there was no obvious signs. platelets were stable and no more blood products were given. the patient was still on support. the doctor stated he did not attribute the gi bleed to the impella 5.5. |  |  |  |  | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the returned product was investigated. the cannula and motor assembly was returned cut from the rest of the pump. dried blood was seen on the red handle, catheter, sterile sleeve, and suture hub. no biomaterial was seen inside the cannula and around the inflow and outflow cages. no other abnormalities were seen. major bleed: patient going down to ir for additional gi bleed evaluation. the cause of the injury was not determined due to insufficient clinical details. tachycardia/ thrombosis: increased vt on (b)(6) 2025. echo verified impella placement but did show an encapsulated mural thrombus in the lv not close to the inlet. the root cause of the injury was unable to be determined due to insufficient clinical details. | Additional information received: b5 and d6b were updated. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. tachycardia/ thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | Additional information updated: d. 9 date device returned to manufacturer and is device returned to manufacturer? the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
3363 | 1220648-2025-48444 | 23585503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585503&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764012 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 84 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma. in response, pressure was held. later, care was withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3364 | 1220648-2025-48446 | 23585681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585681&pc=OZD | 10/24/2025 |  |  | 11/18/2025 |  |  | 10/24/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1758284 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 76 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a placement signal issue with the automated impella controller (aic). it was noted that there appeared to be an ao drift. aic ao placement signal was 15-20 points off aline reading. ao placement reading was 75/35(49) while aline reading was 127/48(63). there was no patient harm. this is one of two related reports. this report represents the automated impella controller and another report will be submitted to represent the pump. refer to section h10 for the related report number. | [5 empty] | D4: corrected lot number and udi. d9: updated devices return information. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3365 | 1220648-2025-48447 | 23585733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585733&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761397 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the groin site. the angle of insertion was supported, and manual pressure was held. the patient then became hypotensive and experienced pulseless electrical activity. the p level of the pump was reduced to 4 and cardiopulmonary resuscitation was provided. the patient was intubated and defibrillated. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the minor bleed is most likely use related since angle of insertion was supported, manual pressure was held and the oozing ceased. the cause of the hypotension and cardiac arrest was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3366 | 1220648-2025-48445 | 23585764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585764&pc=OZD | 10/19/2025 | 01/06/2026 |  | 11/18/2025 |  |  | 10/19/2025 | 11/18/2025 |  |  | 12/31/2025 |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025447362 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 76 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient experienced continuous pump suction issues which was resolved through standard troubleshooting numerous times. additionally, it was noted that there appeared to be an ao drift. automated impella controller ao placement signal was 15-20 points off aline reading. ao placement reading was 75/35(49) while aline reading was 127/48(63). moreover, the patient went in and out of atrial fibrillation. however, after nine days of support, care was withdrawn and the patient expired. this is one of two related reports. this report represents the impella 5.5 and another report was submitted to represent the automated impella controller. | [5 empty] | Investigation summary: suction: clinical details report persistent intermittent suction throughout support. initially, it was reported that there was concern the pump was positioned deep, however, it was later confirmed to be in proper position. stopping epi, reducing p-level, and administering volume were all troubleshooting methods cited to help resolve suction. neither data logs nor product were returned for investigation. since data logs and product were not available, the cause of the suction could not be determined. optical sensor issue: clinical details report that the aops was reading lower than the a-line (aops: 75/35(49), a-line: 127/48(63)) on (b)(6) 2025. no further details were provided. neither data logs nor product were returned for investigation. since data logs and product were not available, the cause of the optical sensor issue could not be determined. arrhythmia: clinical details report the patient was in and out of a-fib on (b)(6) 2025. in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1784973. device history batch: subcomponent: na. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3367 | 1220648-2025-48448 | 23585977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23585977&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738286 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 107 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hemolyzed labs and positioning issues. the pump was repositioned and the p level was lowered. later, care was withdrawn and the patient expired. | [5 empty] | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review. medical safety/clinical reviewed the event. a 49-year-old male with history of alcohol use disorder and nonischemic cardiomyopathy on gdmt (nonadherent) presented to the emergency department on (b)(6) 2025 with epigastric and chest pain. the patient had worsening delirium, flash pulmonary edema, and cardiogenic shock with poor hemodynamics following swan-ganz placement. on (b)(6) 2025, biventricular impella support was initiated (impella cp for lv support and impella rp flex for rv support). the impella rp flex was explanted on (b)(6) 2025. due to persistent shock, va ecmo was initiated with continued impella cp for lv unloading (ecpella). the impella cp remained in place from (b)(6) 2025 to (b)(6) 2025. during ecpella support, the patient developed laboratory evidence of significant hemolysis with multiple plasma-free hemoglobin (pfhgb) values >30 mg/dl, peaking at 160 mg/dl. the impella cp was repositioned over the weekend and subsequently reported by cardiology to be in good position. the impella cp device performance level was reduced from p-4 to p-2 for left ventricular venting. despite ongoing support, the patient¿s condition did not improve. goals of care were transitioned to comfort measures, all life-sustaining therapies were withdrawn, and the patient expired. | Data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved. there were no motor current spikes or purge issues observed during support. no product was not returned for review. the cause of the hemolysis was the patient related since position was good and ecmo was started at the time of hemolysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3368 | 1220648-2025-48449 | 23586000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23586000&pc=OZD | 10/20/2025 |  |  | 11/18/2025 |  |  | 10/20/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756579 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female | 54 KG |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Infection | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp exhibited mottled bilateral lower extremities, likely from toxic shock syndrome, and right lower extremity pulses were lost. the patient was treated with a perfusion catheter. the patient developed purpura fulminans and underwent a fasciotomy. additionally, there were placement signal alarms. later, care was withdrawn and the patient expired on impella support. | [5 empty] | No product was returned for investigation. the root cause of the purpura fulminans and ischemia was unable to be determined due to insufficient clinical details. the root cause of the optical sensor issue is most likely due to the patients condition based on the placement signal trend and clinical details. the device passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3369 | 1220648-2025-48387 | 23586180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23586180&pc=OZD | 10/17/2025 | 11/15/2025 |  | 11/18/2025 |  |  | 10/17/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760639 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 139 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient with a known cardiac history and upon admit to the cath lab the patient went into ventricular fibrillation arrest, had cardiopulmonary resuscitation and was shocked x10. the cp and extracorporeal membrane oxygenation were then placed. the cp had low flows and suction and was removed and replaced after 1/2 an hour. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3370 | 1220648-2025-48385 | 23586196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23586196&pc=OZD | 10/17/2025 | 11/15/2025 |  | 11/18/2025 |  |  | 10/17/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703391 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was a delivery issue during impella cp device implant. the patient had an anatomy issue and the impella implant was aborted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had an anatomy issue preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3371 | 1220648-2025-47342 | 23586659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23586659&pc=OZD | 09/08/2025 |  |  | 11/18/2025 |  |  | 09/08/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662820 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp suffered a cardiac arrest requiring resuscitation. the patient also suffered a perforation bleed. multiple units of packed red blood cells and platelets were transfused, and the patient was treated with sodium bicarbonate and albumin. frequent suction alarms occurred due to the bleeding. the patient entered pulseless electrical activity and return of spontaneous circulation was obtained. later, the patient expired. the patient arrived at the catheterization lab to received percutaneous transluminal coronary angioplasty (ptca) to right posterior descending artery (rpda) and drug-eluting stent to the diagonal artery. staff reported when the patient was about to be moved off of catheterization lab table the patient began having chest pain and st elevation. angiogram revealed perforation to diagonal artery and pericardial effusion was noted. the patient went into cardiac arrest. a pericardiocentesis was performed, and the decision made to place an impella cp device for mechanical circulatory support. at the start of the emergent case, the nurse turned on the automated impella controller (aic), and the yellow controller error alarm went off. the loaner aic was used instead. impella cp was inserted and return of spontaneous circulation (rosc) was achieved after impella insertion. heparin and angiomax had previously been administered; activated clotting time (act) were 272 seconds reported. multiple units of packed red blood cells (prbc) were administered. the impella peel-away sheath was removed, and repositioning sheath was advanced. the impella cp catheter secured with forward tensions sutures and angle match dressing. impella access site was clean dry and intact with no evidence of bleeding or hematoma with right distal pulse present. the patient remained while awaiting flight transport to receiving facility. patient remained intubated. patient had a central line to the left internal jugular, arterial line to right radial. volume was being administered. approximately 600 ml of output from pericardiocentesis drain. urine was noted to be clear yellow in foley bag; knee immobilizer was in place to right leg. the patient arrived via life net flight and coded as helicopter landed. resuscitation was administered for approximately 15 minutes with rosc in the emergency department. the aic transferred to the receiving hospital's aic successfully without incident. the patient was transported to unit and multiple prbcs, platelets, sodium bicarb and albumin infusing. pericardiocentesis drain was draining continuously. the patientâ¿¿s pupils were fixed and dilated at 4. discussions were made with the family. the physician consulted for percutaneous coronary intervention to the left anterior descending/diagonal arteries to try to control perforation bleed. impella site was still clean dry and intact. frequent suction alarms were noted due to bleeding issues. knee immobilizer was in place; feet were warm. no mottling was noted. the patient was severely acidotic with continued correction measures. on-site bedside assessment with the team and the patient was taken to catheterization lab. access to left femoral artery was obtained. diagnostic left heart catheterization revealed right coronary artery stent open and diagonal stent totally occluded with no bleeding from prior perforation noted. left ventricle angiogram was performed which showed no left ventricular perforation; however, left ventricle was very small with little contractility noted. end diastolic pressure was 30. the patient went into pulseless electrical activity (pea). lucus device was started. prbcs, calcium, epi and bicarb were administered. access to left femoral vein was made with swan advanced to the pulmonary artery. return of spontaneous circulation was achieved after 20 minutes of cardiopulmonary resuscitation (cpr). the patient was taken back to the unit and went into pea after arriving. lucus device was started and return of spontaneous circulation achieved after approximately 10 minutes of cpr. at this time, the family decided to assign do not resuscitate. patient went pea and impella was turned down to performance level p-2 and all drips were stopped. time of death was called. the impella and ventilator were stopped. | [5 empty] | Patient-specific information a4. weight is unknown. medical safety review. medical safety reviewed the event. this event describes a coronary artery perforation and pericardial effusion noted during pci prior to impella placement. patient went into cardiac arrest; pericardiocentesis performed and this time a decision made to place impella cp. during the support there was frequent suction alarms noted due to bleeding issues. the patient was eventually made dnr after multiple rounds of cpr. the patient would never fully recover and expired. the patient encountered suction due to small ventricle and most likely preexisting bleed. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. conservatively, the death should be reported on the impella cp device that in use at time of expiration. regarding the aic, there is no association to outcome. the impella device was never connected to this aic. the controller error, however, is a device malfunction and should be reported as a malfunction type of reportable event. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the cardiac arrest, resuscitation, and major bleed, the cause was not determined due to insufficient clinical details. regarding the pump suction, no product was returned. the patient had frequent suction alarms. after reviewing the logs, multiple suction alarms along impella in ventricle alarms were observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller.
3372 | 1220648-2025-48453 | 23589259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23589259&pc=OZD | 10/22/2025 | 11/18/2025 |  | 11/18/2025 |  |  | 10/22/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689346 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female | 118 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Pulmonary Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to cardiomyopathy. while on support, chest x-ray showed a small effusion and pulmonary edema, furosemide to be administered. pt went into afib and was treated with amio. | [5 empty] | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. product complaint # (b)(4). investigation summary no product was returned for investigation. pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. paroxysmal atrial fibrillation: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot: 1932204 device history review device with sn: (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3373 | 1220648-2025-48452 | 23589472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23589472&pc=OZD | 10/22/2025 | 12/24/2025 |  | 11/18/2025 |  |  | 10/22/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026721299 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female | 70 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support experienced a perforation in ramus interventricular is anterior (riva) procedure which led to pericardial effusion. the impella was weaned and explanted due to the pericardial effusion. the patient expired. per the physician, the complication and the death are not related with the impella pump. | [5 empty] | The investigation for the reported pericardial effusion has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the pericardial effusion could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3374 | 1220648-2025-48450 | 23589667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23589667&pc=OZD | 10/20/2025 | 02/23/2026 |  | 11/18/2025 |  |  | 10/21/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154079 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 94 KG | Hispanic |  | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information has been reported: "the device is not available for return. i cannot recall the requested information as the case is from october." | The user facility reported that during aic transfer, the console displayed unreliable placement signals with loss of aortic and lv indicators upon start-up. a second aic transfer was performed to rule out console malfunction, but placement signals remained absent. the pump continued to deliver flows of 4 lpm on p-6 with pulsatile motor current, indicating pump functionality was intact. the placement signal unreliable alarm was silenced, and staff were instructed to monitor flows, motor current, and patient hemodynamics. |  |  |  |  | B5 updated. the investigation is still ongoing. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3375 | 1220648-2025-48454 | 23589685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23589685&pc=OZD | 10/22/2025 | 11/18/2025 |  | 11/18/2025 |  |  | 10/22/2025 | 11/18/2025 |  |  |  |  | 11/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026764160 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 66 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  | Hematoma | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp the patient had suction reported overnight but given fluid (total 3l lr) and was able to come back to p-8. drop in hemoglobin from 10 to 7.6. 1-unit of packed red blood cells (prbc) infused. recheck up to 9.1. second unit ordered concern for rp bleeding after hemoglobin drop overnight, however, after results of cat scan without contrast returned positive for hematoma around right iliac extending into pelvis and retroperitoneum. non-con providing limited info on whether extravasation is current. the plan currently was to trial wean the patient at p-4 and assess for readiness for explant vs upgrade to impella 5.5. heart failure following, will consult cardiothoracic surgery pink-tinged urine also noted with lactate dehydrogenase over 1000. the pump was explanted and the 5.5 was placed. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. low or blocked pump flow: clinical details indicate frequent suction events and episodes of low flow, which resolved after volume administration and p-level adjustments. the cause of the low pump flow issue associated with suction alarms was most likely related to patient condition, based on data log review and provided clinical details. mechanical interaction with blood (hemolysis): clinical details confirm that hemolysis was resolved after volume administration. no relative pump's mechanical defect was found on the data log to justify the cause of hemolysis. the cause of the hemolysis was most likely patient condition-related, since the issue resolved with volume administration. device in wrong position: echo showed the device was slightly deep at 4.5 cm from the aortic valve, but the team was not interested in repositioning at this time. the cause of the device positioning issue could not be determined without returned product investigation and sufficient clinical details. access site adverse event (hematoma): clinical details indicate that after impella cp implantation, the patient developed a moderate-sized hematoma centered around the right iliac vasculature, extending into the pelvis and lower retroperitoneum, as seen on ct imaging. the hematoma was associated with a significant drop in hemoglobin and required transfusion of prbcs. the cause of the access site hematoma could not be determined without returned product investigation and sufficient clinical details. clinical symptom (hematuria): the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3376 | 1220648-2025-48389 | 23568835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23568835&pc=OZD | 10/18/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631332 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Inflammation | Thrombosis/Thrombus | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a hematoma in the right axillary site and a swollen face. the patient was transfused red blood cells, fresh frozen plasma, and platelets. later, the patient was given a do not resuscitate status and expired on impella support. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the inflammation was not determined due to insufficient clinical details. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
3377 | 1220648-2025-48390 | 23568974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23568974&pc=PYX | 10/18/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026719867 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced oozing from the access site. a mattress suture was placed and the patient was transfused two units of red blood cells to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the bleeding is determined to be use issue because hemostasis was achieved by adding new mattress suture. the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3378 | 1220648-2025-48391 | 23569166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23569166&pc=OZD | 10/18/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739029 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemolysis | Inflammation | Ischemia | Sepsis | Rupture | Multiple Organ Dysfunction Syndrome | [6 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced sudden cardiac arrest with papillary muscle rupture, shock with metabolic acidosis, hyperemia, right upper extremity ischemia, multi organ failure, and hemolysis. the impella was repositioned without resolution, and the patient ultimately expired. | Additional information received and stated the patient showed progressive worsening of limb ischemia over a 24¿48 hour period, ultimately requiring amputation. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. papillary muscle rupture/ inflammation/ organ failure: the root cause of the injury was unable to be determined due to insufficient clinical details. sepsis/ ischemia/ cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: log reviewed, many suction alarms observed during impella support, no mc spike observed, no positioning issue or placement signal trend seen. the root cause of the hemolysis issue cannot be determined since the product not returned for investigation and insufficient clinical data provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3379 | 1220648-2025-48392 | 23569310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23569310&pc=OZD | 10/18/2025 | 12/30/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759922 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 102 KG | Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had complaints of chest pain and presented to urgent care. at urgent care, st elevation was noted, and the patient was transferred to a hospital for care. upon arrival to the cardiac catheterization laboratory the patient was diaphoretic and hypotensive. angiogram performed and found complete loss of left anterior descending coronary artery and left circumflex coronary artery due to left main coronary artery disease. an initial percutaneous transluminal coronary angioplasty was performed but during the procedure the patient became progressively worse with cardiogenic shock symptoms, and the decision was made to implant an impella cp for mechanical support. the impella cp was prepped and placed with no issues. 2 stents were deployed successfully to address the blockages in the patient¿s coronary arteries and restored flow. impella position was reviewed with fluoroscopy, the pump was secured, and the patient was transferred to the intensive care unit (icu) for rest and recovery. upon arrival at icu, an echocardiogram was performed and found the impella cp was trapped under papillary muscle. the patient continued to decline despite increase in medication drip requirement. the patient was volume resuscitated, and medications were increased. the patient continued to decline and arrested. four rounds of cardiovascular resuscitation were performed and return of spontaneous circulation could not be achieved. the patient expired during impella cp support. | [5 empty] | Investigation summary: no product was returned. resuscitation, cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was most likely pulled out of position during the transport to icu. device history lot: device lot: 1980814. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3380 | 1220648-2025-48393 | 23569437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23569437&pc=OZD | 10/18/2025 |  |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708748 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced an apical mural thrombus in the left ventricle, and an apical aneurysm was found using an echocardiogram. the patient became hypotensive and the patient was treated with epi and levo. impella support continues. | [5 empty] | No product was returned for investigation. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3381 | 1220648-2025-48394 | 23570561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23570561&pc=OZD | 10/18/2025 |  |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730009 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 56 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to a durable left ventricular assist device (lvad). on day 5 of support, the patient¿s anticoagulation was withheld due to a drop in the patient¿s hematocrit and hemoglobin. the patient was successfully weaned from the impella 5.5 during successful surgery for an lvad implant. the patient survived the procedure. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3382 | 1220648-2025-48397 | 23570849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23570849&pc=OZD | 10/18/2025 | 12/31/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758677 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Fever | Hematuria | Hemodynamic instability | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. while on support, patient experienced hematuria and low-grade fever for which antibiotics were given. the patient noted to not have a cough or gag reflex with blown pupils. computed tomography which showed a massive bleed with mid-line shift and herniation to the brain stem. the patient's blood pressure has been very unstable. care was ultimately withdrawn and patient expired. | [5 empty] | Investigation summary: fever/hematuria: the cause of the injury was not determined due to insufficient clinical details. stroke: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1978898. device history batch: sub-component lot: n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3383 | 1220648-2025-48401 | 23571187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23571187&pc=OZD | 10/18/2025 | 02/19/2026 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635224 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Pulmonary Edema | Pericardial Effusion | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient was placed onto impella 5.5 support for off-pump coronary artery bypass. while on support, patient went into atrial fibrillation then went back in to normal sinus rhythm. a few days later it was reported that a computed tomography scan was done, scan revealed large pericardial effusion (hemopericardium), moderate pleural effusion, and hemothorax. furthermore, the patient went back for washout, evacuated 1l of blood/clots. the patient went into atrial fibrillation with rapid ventricular response, was given amio bolus and started drip. cardioverted x2 which was unsuccessful so an additional amio bolus was given. the patient was later successfully explanted and survived. | [5 empty] | Investigation summary: (atrial fibrillation, thrombosis/thrombus): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (pulmonary edema, hemorrhage/bleeding, pericardial effusion): the cause of the injury was not determined due to insufficient clinical details. peri-procedural adverse event: the cause of the peri-procedural adverse event was not determined due to insufficient clinical details. device history lot: device batch: 1894050. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3384 | 1220648-2025-48404 | 23571946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23571946&pc=OZD | 10/18/2025 | 12/12/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713948 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Application Network Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while ambulating the patient, there was an unexpected low battery warning on the automated impella console (aic) after unplugging it from ac power with a 100% charge. the alarm indicated 50% battery remaining but self-resolved while the aic remained unplugged, with the battery level subsequently displaying 95%. additionally, the aic was not streaming via impella connect (ic) during the event. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
3385 | 1220648-2025-48395 | 23572241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23572241&pc=OZD | 10/21/2025 |  |  | 11/17/2025 |  |  | 10/21/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758100 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was initially admitted on (b)(6) 2025 with acute cholecystitis and a non-st-elevation myocardial infarction with a troponin peak of 5,000, which was initially considered non-ischemic. a choley drain was placed, and his initial ejection fraction was 60%. the patient was re-presented on (b)(6) 2025, from rehab with shortness of breath, chest pain, and signs of heart failure. he was transferred on (b)(6) 2025 for a video-assisted thoracoscopic surgery (vats) to address a hemothorax. following the vats on (b)(6) 2025, the patient severely decompensated overnight into cardiogenic shock, requiring three pressors and having a lactate level of 13. an urgent echo showed a drastically reduced ejection fraction of 10%. the shock team initiated urgent mechanical circulatory support by placing veno-arterial extracorporeal membrane oxygenation (va ecmo) and an impella cp. upon initiation of va ecmo and aimpella cp (ecpella support), the patient stabilized rapidly, allowing weaning off most pressors to only a low dose of levophed. perfusion to both legs is confirmed. on day 4 of ecpella support, it was noted that the patient would be weaned from va ecmo over the weekend. on day 7 of impella cp support, the patient developed a gastrointestinal bleed (gib). the patient¿s anticoagulation therapy was withheld. an esophagogastroduodenoscopy was successfully performed, and a total of four clips were applied to an ulcer, which appears to have resolved the bleeding issue. the patient required two units of packed red blood cells for stabilization. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | H6: added type of investigation code b13 h11: added the results of the investigation. investigation summary no product was returned for investigation. major bleed: - gi bleed, pt underwent egd, two clips applied to ulcer. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review ==> device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3386 | 1220648-2025-48396 | 23572303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23572303&pc=OZD | 10/18/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581512 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 68 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that there was hypotension during impella cp patient support. while on support, there was some hypotension when fluid removal was increased. an abiomed representative noted that there was an increase in vasoactive drugs. the patient experienced worsening acidosis. care was withdrawn and the patient expired. | [5 empty] | Medical safety reviewed the event and noted the following: the patient had congenital heart disease and was receiving extracorporeal membrane oxygenation (ecmo) support. physicians attempted to insert an impella 5.5 device; however, due to vessel size and calcification, the device could not be advanced. this was classified as a non-delivery event attributable to the patient¿s known anatomy. subsequently, an impella cp device was successfully placed. while on impella cp support, the patient experienced a hypotensive episode during continuous renal replacement therapy (crrt). the hypotension was associated with excessive fluid removal and resolved after intervention. the patient¿s condition continued to deteriorate, with severe biventricular failure noted. the family elected to withdraw care, and the patient expired. device involvement and reporting: impella 5.5: non-delivery event due to anatomical limitations. impella cp: hypotensive episode conservatively reported as device-associated, though root cause was excessive fluid removal during crrt. patient death conservatively reported while on device support; however, the device is unlikely to have contributed to the outcome, which was attributed to the patient¿s underlying condition. conclusion: the events were reported in accordance with regulatory requirements. no device malfunction was identified. the inability to insert impella 5.5 was due to patient anatomy, and subsequent complications were related to clinical management and disease progression rather than device performance. device status. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hypotension, the cause was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3387 | 1220648-2025-48408 | 23572464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23572464&pc=OZD | 10/19/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/19/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755502 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was a hematoma during impella cp patient support. while on support, a femostop was placed for a small hematoma. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematoma: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3388 | 1220648-2025-48398 | 23572778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23572778&pc=OZD | 10/17/2025 |  |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712129 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 66 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support due to postcardiotomy cardiogenic shock/low cardiac output syndrome. while on support, patient received multiple blood transfusions. the patient was taken to the operating room for washout. the patient was also placed on dialysis. the patient was transferred to another hospital and no other information was provided. | [5 empty] | Explant date was provided and d6b was updated. the investigation was completed and no product was returned for investigation. major bleed: ecpella + iabp continues. patient (pt) receiving multiple blood transfusions overnight. received 7units red blood cells yesterday and 3units overnight, plt, cryo. plan to take pt to the operating room for washout and find source of bleeding. the cause of the bleeding is determined to be patient condition because of unknown location of bleed. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3389 | 1220648-2025-48399 | 23572881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23572881&pc=OZD | 10/18/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026639786 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 109 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was placed onto impella support for acute myocardial infarction complicated by cardiogenic shock. while on support, patient attempted to extubate. waveforms began to appear like the device was pushed deeper in ventricle. echocardiogram performed at bedside revealed device measuring at 6.0 cm. pump was under papillary muscle, unable to completely release without risk of explant. suction alarms began and unable to obtain flows more than 2.0 l/min. suction alarms unable to be resolved with decreasing of p-level. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data log shows a sudden drop in pump flow while running on higher p-levels, with active suction alarms throughout the case run. no sign of motor current spike or positioning alarm was noted during the case. the pump was left for another 48 hours on p-level 7 with flow less than 2 l/min and active suction alarms. pump suction: as per the call, pump was positioned at 6.0 cm, under the papillary muscle and unable to be completely released without risk of explant. despite attempts to resolve suction by pulling back the device and decreasing the p-level, suction alarms persisted and flows could not be obtained above 2.0 l/min. the cause of the suction issue was most likely use-related, based on data log review and provided clinical details suggesting that the device was under the papillary muscle and unable to be completely released. device in wrong position: the cause of the positioning issue was most likely unintentional use-related, as repositioning appeared to be needed after the patient woke up and sat up in bed. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3390 | 1220648-2025-48400 | 23572967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23572967&pc=OZD | 10/18/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697945 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support experienced access site complications, including hematoma formation and oozing around the repositioning sheath. a femstop device was applied superior to the arteriotomy/veinotomy and inflated to 60 mmhg to achieve hemostasis. patient support was escalated to impella 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: major bleed/ hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot : 1939795. device history batch: sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3391 | 1220648-2025-48402 | 23573295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23573295&pc=PYX | 10/18/2025 |  |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026741775 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella rp flex experienced a temporary pump stop. the pump restarted without intervention. additionally, a high purge pressure alarm occurred. the issue was resolved with the administration of tissue plasminogen activator was initiated. impella support continues. | The pump was successfully weaned. |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received: b5 and d6b updated corrections: d2a, d2b, g4. | The investigation of the reported failure mode has been completed. no product was received for evaluation. temporary pump stop: data logs were reviewed and the log at beginning of case shows a sudden increase in motor current and deviation in motor speed right before alarm #13 pump stop alarm. placement signal also spikes up at time of pump stop. pump was able to be restarted after multiple attempts and pump metrics were within normal limits after pump was able to be restarted. the cause of the temporary pump stop could not be determined as limited clinical details were provided and no product was returned for evaluation. high purge pressure/purge system blocked: data logs were reviewed and log at time of high purge pressure alarms noted a sudden step up in purge pressure with a small motor current spike. purge pressure remained elevated without alarms for approximately 4 hours before another step up in purge pressure with a small motor current spike. "high purge pressure" alarms were triggered. purge pressure remained elevated for approximately 4 hours before slowly resolving. the cause of the high purge pressure/purge system blocked was due to biomaterial ingestion based on data log analysis and clinical details. device history review: the device passed all post sterile inspection checks.
3392 | 1220648-2025-48406 | 23573311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23573311&pc=OZD | 10/18/2025 |  |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026720645 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient drove themselves to the hospital and upon arrival went into cardiac arrest. cardiopulmonary resuscitation was performed, and the patient was shocked twice in the cardiac catheterization laboratory (ccl). the patient was found to have an anterior myocardial infarction in the left anterior descending coronary artery. a percutaneous coronary intervention was performed and an impella cp was placed for support. the patient was transported to the cardiac intensive care unit (cicu) and approximately 15 minutes after arrival, it was reported that the pump began alarming after the patient had been turned. the impella was repositioned under echocardiogram guidance. the physician decided to wean and explant the impella as they were concerned about the alarms that occurred despite best practices being used and locking the impella in place. the physician felt that the patient was hemodynamically recovered enough that impella was no longer needed. the patient was taken back to the ccl and the pump was successfully weaned and explanted. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue/suction: data logs do confirm failure mode, however, since the clinical details confirmed correct positioning and no product was returned, the root cause of the optical sensor issue and suction issue cannot be determined. device history review: the device passed all post sterile inspection checks.
3393 | 1220648-2025-48403 | 23573639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23573639&pc=OZD | 10/18/2025 |  |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740412 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female | 58 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Malposition of Device |  |  |  |  | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Inflammation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with heart transplant rejection was implanted with an impella cp device for mechanical circulatory support. the patient presented to the emergency room (er) on (b)(6) 2025 with complaints of syncope. the patient¿s baseline ejection fraction was 54% but deteriorated to 15%. there was concern for acute transplant rejection with biopsy results positive for 1r rejection. the patient experienced pulseless electrical activity arrested the following morning with an approximate downtime of 40 minutes. the patient briefly experienced a return of spontaneous circulation during this time. the patient was urgently transferred to the complainant hospital where they were cannulated for extracorporeal cardiopulmonary resuscitation. the patient continued to wide-complex rhythm, and was transferred to the cardiac catheterization lab (ccl) in critical condition, with a baseline lactate level of 18 mmol/l. during transport and procedures, the patient remained intubated, on veno-arterial extracorporeal membrane oxygenation (va-ecmo), and required significant hemodynamic support with 4 units of epinephrine. coronary angiography identified significant underlying vascular disease in the ramus branch and an acute and occlusive thrombus in the right coronary artery. the decision was then made to place the impella cp and complete a percutaneous coronary intervention. the impella cp was placed without complications, and support was initiated in auto performance. due to suction alarms, the performance (p) level was decreased to p5, and then p3 due to ongoing suction alarms. va ecmo flows were also reduced, and the patient was given fluids and albumin. after the procedure, the patient was transferred to the intensive care unit (icu). upon arrival to the icu, the impella continued to have suction alarms and p level was dropped to p2 and more volume was administered. a point of care ultrasound was completed and the physician felt that the impella was lightly shallow, however it was still across the aortic valve and did not appear to be interacting with intracardiac structures. of note, the patient¿s lactate level decreased to 17 mmol/l after support was initiated. an additional ultrasound was performed and impella was noted to be shallow. the impella was advanced at the patient¿s bedside and suction was noted to have improved after the positioning of the pump. additionally, the patient¿s left chest wall appeared swollen and a stat chest x-ray was ordered. on support day 2 there were noted suction alarms when the patient was diuresed and the patient¿s abdomen appeared to be distended. an exploratory laparotomy was performed but no source of bleeding could be found. the patient was massively transfused with 10 units of packed red blood cells (prbcs), 8 units of fresh frozen plasma, 3 units of cryoprecipitate, 3 units of platelets, 4 liters of lactated ringer¿s solutions, and 4 liters of albumin. on support day 3, the patient received a unit of prbcs and it was noted that the patient¿s bowel movements were dark in color. the patient had some increase in entropy the previous night and the patient¿s amiodarone was increased. the impella position was imaged with point of care ultrasound and noted to be in good position. on support day 4 there were noted to be no impella concerns, however there were two purge blocked alarms that self-resolved. the patient was also noted to have been transfused with a unit of prbcs and a unit of platelet on support day 6, the patient had dark colored stool and vomited coffee colored emesis twice. anticoagulation therapy was withheld as a result, and the patient¿s hemoglobin dropped to 6.1. as a result of the blood loss, the patient received 2 units of prbcs, 2 units of platelets, 1 unit of fresh frozen plasma, and 120mg of lasix. the patient had two clips placed in the duodenum to resolve the bleeding. on support day 7, the patient was transfused with 1 unit of prbcs and 1 unit of platelets and remined off of anticoagulation therapy. on support day 9, there were no impella related alarms. the patient experienced a few runs of supraventricular tachycardia and/or atrial fibrillation. a point of care ultrasound showed the pump in a adequate position. the patient received an additional unit of prbcs. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injuries, arrythmia, inflammation, major bleed and anemia was unable to be determined due to insufficient clinical details. data logs were reviewed and intermittent suction alarms were observed during support. main instances of suction alarms occurred of first day of support. no motor current spikes observed. placement signal and motor current have low pulsatility at time of initial suction alarms. clinical details noted that suction started soon after pump was initiated. extracorporeal membrane oxygenation flows were decreased, fluid and albumin administered in order to resolve suction alarms. additionally, noted that pump was slightly shallow but not interacting with any cardiac structures. suction improved slightly after repositioning. the cause of the pump suction was patient condition related as motor current and placement signal had low pulsatility during initial suction events and patient as administered fluids in order to resolve suction. clinical details noted that pump was slightly shallow on ultrasound once patient was in unit. pump was across av and there was no interaction with any cardiac structures. pump was repositioned. the cause of the device in wrong position could not be determined as no details were provided as to how the pump came out of position. clinical details noted that pump had multiple instances of "purge system blocked" alarms during support. data logs were reviewed and first instance of high purge pressure showed a sudden increase purge pressure with a drop in purge flow. purge flow ticks also become stalled at the same time. purge pressure is able to suddenly resolve. the cause of the purge system blocked is due to a kinked purge line based on a review of returned data logs. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3394 | 1220648-2025-48405 | 23574509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23574509&pc=OZD | 10/18/2025 | 12/31/2025 |  | 11/17/2025 |  |  | 10/18/2025 | 11/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739605 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that distal perfusion to the patient¿s leg was compromised during impella cp support. attempts to place a distal perfusion catheter were unsuccessful, leading to escalation of therapy from the initial device to an impella 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: ischemia: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1970222. device history batch: sub-component lot: n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3395 | 1220648-2025-48407 | 23575567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23575567&pc=OZD | 10/20/2025 | 02/23/2026 |  | 11/17/2025 |  |  | 10/19/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655017 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 94 KG | Hispanic |  | Injury | Yes | No | False Alarm | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the patient experienced multiple hematoma sites (head, chest, right axillary, and groins) while supported on impella 5.5. the bedside nurse was unsure if these were related to the device or graft. the pump was functioning properly and optimally positioned. during transfer, the patient remained hemodynamically labile, requiring transfusion of 2 units of red blood cells and fluid resuscitation for low filling pressures. impella position was confirmed via transesophageal echocardiogram prior to leaving the chronic lymphocytic leukemia and 3-point fixation was in place. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the optical sensor issue: clinical details report that a psnr alarm triggered upon connection to new console at transfer hospital. to troubleshoot, the team connected the pump to an additional console, and the alarm persisted. the case carried on without a ps, and echo was used to confirm position. data logs were reviewed and the report was confirmed. the pump was run on (b)(4) from 18/oct to 21/oct without any abnormalities noted in the optical signal metrics or alarms. when the pump was transferred to (b)(4) on 21/oct, the ps immediately railed to 3000 mmhg and the psnr alarm triggered. however, during this time, all optical signal metrics were within expected ranges. the pump was then transferred to (b)(4) , and the same pattern persisted. the pump was run nonetheless without a ps on (b)(4) and (b)(4) through 25/oct/2025. at the time of investigation, imc logs were not available, and the pump was not returned. since neither imc logs nor the pump were available for investigation, the cause of the optical sensor issue could not be determined. please refer to pi-17629749081040463 for an investigation into the console. hematoma/major bleed: clinical details report that the patient had hematomas at the head, chest, rax, and groin sites on 20/oct. (b)(4) units of rbcs were administered the day before. the impella 5.5 was inserted via the rax. no further details were provided and product was not returned for investigation. since insufficient clinical details were provided and product wasn't returned for investigation, the cause of the access site adverse event and hematoma could not be determined. hemodynamic instability clinical details report that the patient had to remain in the ccl for an extended period of time due to unstable blood pressures. the following day, a patient update reported that the patient's pressures were still unstable. the cause of the injuries were not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3396 | 1220648-2025-48410 | 23576530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23576530&pc=OZD | 10/20/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755873 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hematoma | Hemolysis | Low Blood Pressure/ Hypotension | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp on (b)(6) 2025, a hematoma was noted to right groin, manual pressure was held for 10 minutes with placement of femstop. groin was accessed pre-impella with a generic sheath and lhc performed. plan was to use left groin for impella placement but there was too much calcium. original sheath exchange for impella peel away and hematoma noted after exchange but before peel away was removed at end of case. activated clotting time (act) was out of range/greater than 400 for the entire case. systemic anticoagulation held in intensive care unit (icu). some blood in the urine was also noted, rn to continue to monitor, medical doctor (md) aware. eventually, urine cleared, no longer blood tinged. weaned to p2 overnight, patient didn¿t tolerate, blood pressure (bp) dropped, a lot of ectopy, needed a total of 2.5l in intravenous fluids (ivf) bolus, increased dobutamine. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injuries, arrhythmia, and low blood pressure and were unable to be determined due to insufficient clinical details. as per the clinical details, the hematoma was managed with manual pressure, dressing with angle matching, and femstop. the patient was also kept off anticoagulation, and the reported act during the hematoma was 400 at the time of the event. the cause of the hematoma was most likely use-related based on the clinical details provided. mechanical interaction with blood (hemolysis): pump was not returned for investigation. data log shows no abnormalities related to motor current spike, positioning issues, or suction during the time of the hemolysis event. pump flow and placement signals had no major trend justifying the cause of the hemolysis. the cause of the hemolysis issue was not determined without returned product investigation and sufficient related clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3397 | 1220648-2025-48416 | 23576549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23576549&pc=OZD | 10/20/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697046 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 78 KG | Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The use facility reported patient had a 5.5 pump placed for the support of a patient awaiting transplant. during the 6 day support the pump was poorly positioned but not replaced. the patient was successfully transplanted and survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the device in wrong position (positioning issue): the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3398 | 1220648-2025-48418 | 23576550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23576550&pc=OZD | 10/17/2025 |  |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133541 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp and automated impella controller (aic) being prepped for use for the patient admitted in with acute myocardial infarction/cardiogenic shock. the nurse reported that the aic was showing a problem, "the purge disc was not recognized¿ and subsequently the pump was not recognized. the clinical staff replaced both the pump and the aic which resolved the issue. there was no patient harm. this report is for the aic and another report will be sent for the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3399 | 1220648-2025-48409 | 23577209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23577209&pc=OZD | 10/20/2025 |  |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740423 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 114 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a vascular dissection and major bleed during impella cp patient support. while on support, there was aortic bleeding. the patient was taken to the operating room for an aortic valve replacement and a patch repair of the ascending aorta. the patient received three units of blood. the following day the pump was successfully explanted and the patient survived. | [5 empty] | The investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the product was not returned for investigation. the root cause of the injuries, vascular dissection and major bleed, could not be determined as the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3400 | 1220648-2025-48411 | 23577772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23577772&pc=OZD | 10/20/2025 | 11/17/2025 |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722895 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 for hemodynamic support. during support it was noted that the patient's hemoglobin had dropped with dark stools, received 1 unit of red blood cells. in addition, the care team stated that 5.5 was increased, 250 of albumin, 2 packed red blood cells and 2 platelets, were given. the patient was being evaluated for heparin induced thrombocytopenia. there was also a note of thrombus on the tip of the 5.5, however after explant it was identified to not be thrombus (see attachments). the 5.5 was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3401 | 1220648-2025-48417 | 23577773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23577773&pc=OZD | 10/17/2025 |  |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608370 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp and automated impella controller (aic) being prepped for use for the patient admitted in with acute myocardial infarction/cardiogenic shock. the nurse reported that the aic was showing a problem, "the purge disc was not recognized¿and subsequently the pump was not recognized. the clinical staff replaced both the pump and the aic which resolved the issue. there was no patient harm. this report is for the pump and another report will be sent for the aic. | [5 empty] | E1 added initial reporter postal code as it was omitted from the initial report that submitted. h6 code a141203 was removed. h10 related report number was added. investigation summary: the investigation has been completed. no product returned. logs are not available for the date of the complaint. pump not detected: clinical details noted that there were pump detection issues, including an "impella defective" alarm, so a new pump was used. no other details are known. the cause of the pump not being detected was not determined since no product/logs were returned and insufficient clinical details were provided. device history review: the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3402 | 1220648-2025-48419 | 23577775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23577775&pc=OZD | 10/20/2025 |  |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 |  |  |  |  | 11/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026743489 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the patient with extensive cardiac history. during support the left ventricle placement signal was off significantly once guide catheter inserted. sometimes would settle out to normal then landed on having a decoupled waveform. eventually suction alarm. anesthesia at first said this was due to patient coughing, however suction alarm persisted. the impella was adjusted but didn¿t help. already only in p4 so the physician proceeded with percutaneous coronary intervention. the patient tolerated the procedure well. decision was made to wean impella at the end of the case and remove. physician stated that he thought that maybe ¿the first existing tavr valve was up against an outlet or possibly the guide catheter.¿ there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3403 | 1220648-2025-48415 | 23577959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23577959&pc=OZD | 10/20/2025 | 12/30/2025 |  | 11/17/2025 |  |  | 10/20/2025 | 11/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718528 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Heart Failure/Congestive Heart Failure | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that a patient presented for coronary artery bypass graft. protamine was given at the end of the case and the patients suffered hemodynamic collapse requiring to go back on bypass. decision for biventricular impella support. it was noted that impella cp support started with suction noted when coming off support. the impella rp was placed. during support it was noted that the patient went into ventricular tachycardia prior to leaving the operating room. during advanced cardiovascular life support (acls) the impella appeared to have advanced. the cp was pulled back. the heart appeared grossly dry. volume resuscitation continued post acls. bicarb started. critical care continued. the patient required constant attention to maintain volume. it was noted that cardiopulmonary resuscitation while leaving the operating room had probably led to increased bleeding. bedside echocardiogram showed completed sucked down left ventricle and right ventricle with impellas only able to run around 2 l/min. cp appeared deep and was pulled back but ultimately needed more preload. after chasing volume: 2 units of blood, 2 amp bicarb, 500 albumin, 500 lactated ringer's solution bolus amio bicarb drip, they were able to get flows close to 3l/min. more blood products unit of platelets, cryo, plasma and 2 more bicarb amp pushes, 2 units of blood and a setup for auto transfusion with chest tubes the patient had been stable for several hours now. as long as they continue to chase the volume he is losing. after continued efforts to stop post op bleeding with continuous volume replacement and on max vaso active therapy, the family decided to withdraw care and the patient expired. | [5 empty] | H6 medical device problem code updated. the impella device has been discarded, and the investigation is still underway. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the clinical symptoms, ventricular fibrillation, cardiac failure and major bleed were unable to be determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. data log shows several suction alarms with fluctuating placement signal throughout the 7.5 hours of the case run. during suction alarms, pump flow trended down while running on high p-levels. no abnormalities were reported regarding positioning issues or motor current spikes during the case run. as per the clinical details, persistent suction events occurred following impella placement, particularly after cpr and advanced cardiac resuscitation, which likely led to increased bleeding and hypovolemia. echo showed underfilled left and right ventricles, with impella flows only around 2 l/min, consistent with low preload and volume depletion. the cause of the suction issue was most likely related to patient condition, based on data log review and provided clinical details. device in wrong position: the pump was not returned for investigation. clinical details indicate that during acls, the impella cp appeared to have advanced. a bedside echocardiogram (echo) was performed, and the pump was pulled back by the doctor to 89 cm. in the icu, bedside echo showed the cp appeared deep and was pulled back, but ultimately the patient needed more preload. the cause of the device position issue was not determined without returned product investigation and sufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. clinical symptoms (ventricular fibrillation): the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms ( major bleed): the cause of the injury was not determined due to insufficient clinical details. pump suction: the cause of the suction issue was most likely related to patient condition, based on data log review and provided clinical details. device in wrong position: the cause of the device position issue was not determined without returned product investigation and sufficient clinical details. | Medical safety/clinical reviewed the event. a 56-year-old male with a history of atrial fibrillation and end-stage renal disease on chronic dialysis (4 years) was scheduled for coronary artery bypass grafting (cabg) due to complaints of shortness of breath. eight months prior, the patient had been hospitalized for three weeks for infective endocarditis with mitral valve vegetation. the vegetation was later described as calcified lesions on the posterior leaflet. at the time of the current admission, echocardiography demonstrated 1+ mitral regurgitation (mr) with a left ventricular ejection fraction (ef) of 35%. pulmonary function testing was pending per the medical record. cabg was performed. at the conclusion of cardiopulmonary bypass, protamine sulfate was administered for heparin reversal. subsequently, the patient experienced acute hemodynamic collapse requiring return to cardiopulmonary bypass. the patient was placed on biventricular mechanical circulatory support with an impella cp and an impella rp flex. prior to leaving the operating room, the patient developed ventricular tachycardia (vt). advanced cardiovascular life support (acls) was initiated, including cardiopulmonary resuscitation (cpr). during resuscitative efforts, the impella device position appeared to advance. bedside echocardiography was performed, and the device was repositioned by the physician to 89 cm. the ventricles appeared significantly underfilled (¿grossly dry¿), and aggressive volume resuscitation was initiated following acls. upon transfer to the intensive care unit (icu), the patient required continuous volume administration to maintain hemodynamic stability. ongoing bleeding was suspected, potentially exacerbated by chest compressions performed during cpr. bedside echocardiography in the icu demonstrated severely underfilled left and right ventricles with both impella devices only able to generate flows of approximately 2.0 liters per minute. the impella cp appeared positioned deep and was repositioned; however, low flow persisted due to inadequate preload. extensive volume and blood product resuscitation was administered, including multiple units of packed red blood cells, platelets, cryoprecipitate, plasma, sodium bicarbonate boluses and infusion, albumin, lactated ringer¿s solution, and initiation of autotransfusion via chest tubes. following aggressive resuscitative efforts, device flows improved to approximately 3.0 liters per minute. the patient remained temporarily hemodynamically stable with continuous high-volume replacement and maximal vasoactive support. despite ongoing efforts to control postoperative bleeding and maintain circulatory support, the patient remained critically ill. after approximately 23 hours of biventricular impella support and continued high-dose vasoactive therapy, the family elected to withdraw care. the patient subsequently expired. the impella cp and impella rp flex are being conservatively reported as associated with a death outcome. however, the patient¿s death was most consistent with refractory cardiogenic and hemorrhagic shock following protamine administration and separation from cardiopulmonary bypass, compounded by postoperative bleeding and hemodynamic collapse requiring acls and are considered the primary factors to the patient's demise. correction. complaint coding was updated accordingly based on medical safety/clinical review to removed cardiac failure (e0611). and h6. health effect - clinical and impact coding that represents reported event has been repopulated. note: there is no impact to the investigation finding for this event as no finding/cause was identified. cardiac failure event is not associated with the impella device. h6. medical device problem, component and investigation type/finding/conclusion as previously reported remain unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3404 | 1220648-2025-48382 | 23564082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23564082&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/15/2025 |  |  | 10/17/2025 | 11/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented for a planned percutaneous coronary intervention(pci) with impella cp. left femoral artery (lfa) was accessed, however upon access the physician noted the initial access was "crunchy" once access achieved. during perclose the vessel would not seal with perclose after 4 attempts. the calcium was noted to be thick, lfa site changed to access site for pci and right femoral artery accessed for impella. angio done of the groin after accessed achieved. significant tortuosity and long impella sheath chosen. the pump was passed through the peel away- once past through the peel away the clinician met resistance. multiple attempts of pulling back to the access site and re "viper slide" and angle changes. due to the peripheral anatomy impella was aborted. decision to proceed with pci. there was no harm reported. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had significant tortuosity and peripheral anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3405 | 1220648-2025-48383 | 23564083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23564083&pc=OZD | 10/17/2025 | 11/15/2025 |  | 11/15/2025 |  |  | 10/17/2025 | 11/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734175 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented with chest pains, shortness of breath and chest ex-ray revealed bilateral pulmonary edema. echocardiogram revealed multivessel coronary artery disease and an ejection fraction of 25%. decision for high-risk surgery with elective 5.5 placement was made. however, it was noted that patient's cardiac anatomy made proper positioning difficult. aborted after 3 attempts. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. d9 updated with product received date. no issues found on the returned product. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult cardiac anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
3406 | 1220648-2025-48384 | 23564084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23564084&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/15/2025 |  |  | 10/17/2025 | 11/15/2025 |  |  |  |  | 11/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679106 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was arrhythmia, hypotension, major bleeding, anemia and a positioning issue during impella 5.5 patient support. while on support, the patient had an increase in ectopy. a point of care ultrasound showed the pump was close to the septal wall. the impella was repositioned; it was pulled back from 6.1 to 5.3 cm. the patient had a hypotensive episode while ambulating. one unit of packed red blood cells (prbcs) was administered. the patient received one unit of packed red blood cells for a drop in hemoglobin/hematocrit. hemoglobin was trending downwards and a chest, abdomen, and pelvis cat scan was done. it showed contrast extending along the drive line anteriorly in the axillary artery. no hematoma was noted. the patient was scheduled for an esophagogastroduodenoscopy (egd)and three units of prbcs were administered. the pump measured at 6.5cm and was pulled back to 5.7 cm. the egd was completed. the patient was positive for a gastrointestinal bleed from either the stomach or small bowel. the patient had a mesenteric artery embolization and one unit of prbc. three additional units were administered. the patient had two black stools. another unit of prbc was administered for a hemoglobin of 7. | [5 empty] | Correction: b1, e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the cause of the injury was unable to be determined due to insufficient clinical details. anemia, hypotension: the cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the bleed was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3407 | 1220648-2025-48386 | 23564085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23564085&pc=OZD | 10/17/2025 |  |  | 11/15/2025 |  |  | 10/17/2025 | 11/15/2025 |  |  |  |  | 11/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026639379 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella rp the patient placement signal issue on rp, audio off. impella rp currently running at p9 with intermittent ps waveforms. overnight impella rp was alarming ¿placement signal unreliable¿ with placement signal intermittently not reading. the physician would like to disable placement signal audio. impella rp weaned to p0 and explanted without issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the minor bleed/acess site adverse event: oozing noted at rp site yesterday during dressing change. dressing redone with improvement. heparin had been off for trialysis line placement and act was 197. the cause of the access site adverse event was use related to high act values. placement signal issue: alarm reported on second day of support. data logs confirm the psnr alarms and show alarms #1053, 1054, and 1055 indicating an optical issue. at this time cvp is negative with a drop in snr. there are also suction alarms. dp placement signal and mc are also variable at this time. these events are occurring at p-9. all parameters and alarms improve when the pump is turned down to p-6, p-5, and p-4, indicating that the patient could not handle higher p-levels. however without additional clinical details such as the patient's volume status, any concomitant devices, or the pump's positioning, the issue cannot be investigated further to identify the cause of the suction condition. the metrics improve by afternoon of oct-18. clinical states the patient was believed to be hypovolemic and was receiving blood volume while on support. the cause of the placement signal issue was determined to be patient condition related per data log review showing suction conditions, negative cvp values, and metrics/alarms improving by reducing p-level and giving blood volume. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3408 | 1220648-2025-48346 | 23551899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23551899&pc=OZD | 10/17/2025 | 11/12/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603443 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 was placed for patient support during percutaneous coronary intervention. during impella 5.5 support, it was noted that epinephrine was being given due to low heart rate. the pump was noted to be shallow and was awaiting pump repositioning. the patient remained on support and was successfully weaned and explanted. additional information was received. the low heart rate was related to the medication given for sedation. the medication was turned off and epinephrine was chosen as a pressor to assist in bringing the heart rate back up versus the need for true inotropic support. additionally, the impella 5.5 was inserted by the surgeon for high-risk percutaneous coronary intervention to be done in the catheterization lab. upon arrival to the unit post placement, initial echocardiogram showed the pump was shallow. the pump was repositioned by the interventional cardiologist within the hour the patient got down to the catheterization lab. the positioning was confirmed under echocardiogram guidance. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the bradycardia: the cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3409 | 1220648-2025-48347 | 23551924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23551924&pc=OZD | 10/17/2025 | 11/12/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741082 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon insertion of the impella cp, they struggled to get the repo sheath and sterile sleeve pushed back in order to allow operational length of catheter to be inserted. it took more than expected force to push back. once the impella cp was placed adequately into the left ventricle, they experienced difficulty removing the impella wire. it required more than expected force to remove and upon removal a portion of the wire broke off and lodged partially in the peel away sheath. they were able to successfully retrieve the remaining part of the wire. upon initiation of support, there were flow issues with the cp. the decision was made to explant. the device was replaced with a new impella cp and insertion and function were as expected. the procedure was completed successfully. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D9 updated to device received date. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. no physical damage was noted on the returned product. the pump was run as expected during testing in a bucket of water. data log was not provided or could not be retrieved from the remote server. low or blocked pump flow: no issue was found on the returned product. the cause of the issue could not be determined without sufficient clinical details, including data logs. device history review: the device passed all post sterile inspection checks.
3410 | 1220648-2025-48348 | 23551952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23551952&pc=PYX | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  | 11/30/2025 |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025442264 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient had successful percutaneous coronary intervention to the left main/left anterior descending artery. however during the course of the procedure, it was identified that there was a need for right ventricle support. the decision for an impella rp flex was made. however, they were unable to place the impella rp flex due to anatomy concerns. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller section: potential adverse events (united states), ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings, ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
3411 | 1220648-2025-48349 | 23551974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23551974&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713206 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 68 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a delivery issue during impella 5.5 implant. the doctor was unable to pass the innominate artery as the vessel was calcified. the decision was made to explant and attempt impella cp insertion. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. | Medical safety reviewed the event and noted the following: the patient had congenital heart disease and was receiving extracorporeal membrane oxygenation (ecmo) support. physicians attempted to insert an impella 5.5 device; however, due to vessel size and calcification, the device could not be advanced. this was classified as a non-delivery event attributable to the patient¿s known anatomy. subsequently, an impella cp device was successfully placed. while on impella cp support, the patient experienced a hypotensive episode during continuous renal replacement therapy (crrt). the hypotension was associated with excessive fluid removal and resolved after intervention. the patient¿s condition continued to deteriorate, with severe biventricular failure noted. the family elected to withdraw care, and the patient expired. device involvement and reporting: impella 5.5: non-delivery event due to anatomical limitations. impella cp: hypotensive episode conservatively reported as device-associated, though root cause was excessive fluid removal during crrt. patient death conservatively reported while on device support; however, the device is unlikely to have contributed to the outcome, which was attributed to the patient¿s underlying condition. conclusion: the events were reported in accordance with regulatory requirements. no device malfunction was identified. the inability to insert impella 5.5 was due to patient anatomy, and subsequent complications were related to clinical management and disease progression rather than device performance. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, unable to deliver or advance (delivery issue), the cause of the deliver issue was determined to be patient condition as the patient had innominate and calcified vessel preventing successful delivery of impella. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome.
3412 | 1220648-2025-48350 | 23551990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23551990&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026742086 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 60 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was thrombocytopenia, major bleeding and a stroke during impella 5.5 patient support. while on support, the patient¿s platelet count was low so heparin was held. the patient had a cerebrovascular accident. there was seizure activity so they were intubated to protect the airway. after a computed tomography, a small bleed was found and anticoagulation was held. the complication was successfully managed. the patient remained on support until successfully explanation and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: no product was returned for investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. stroke: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the stroke was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. major bleed: there was seizure activity so they were intubated to protect the airway. after a computed tomography, a small bleed was found and anticoagulation was held. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
3413 | 1220648-2025-48351 | 23552053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23552053&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705264 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was placed due to st-elevation myocardial infarction team activated due to ventricular fibrillation arrest and hypotension. the patient had begun to stabilize and was being weaned slowly from levophed, and showed signs of recovery, however a fever developed. no intervention was reported. the impella cp was successfully weaned and explanted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. fever : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3414 | 1220648-2025-48352 | 23552079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23552079&pc=OZD | 10/17/2025 | 12/30/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708980 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was successfully inserted via the left femoral artery. the pump was connected. the cannula was across the aortic valve. however, when the physician attempted to advance the repositioning sheath up the catheter length, the sheath did not move. the pump was aborted and a secondary impella cp was implanted for patient support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3415 | 1220648-2025-48325 | 23553982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23553982&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712122 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, the patient lost placement signal. an echocardiogram confirmed position. the patient was awaiting an orthotopic heart transplantation. placement signal was not reliable and motor current was fine. the patient may be moved up to tier one for transplant. the pump was removed and the patient received a transplant. | [5 empty] | The logs were returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3416 | 1220648-2025-48354 | 23554144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23554144&pc=OZD | 10/17/2025 | 11/12/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026759078 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient was resting. his left leg was showing signs of ischemia, and no pulses were detected. it is highly likely that the pump was occlusive to the right femoral artery causing a cold leg. the doctor was made aware and discussed options with the family and the patient. the physician was unhappy about the patient not receiving extended support. the doctor will pull this pump prematurely to restore flow to the lower leg. the patient was taken down to the cardiac catheterization lab for impella removal and replacement for an intra-aortic balloon pump. the site was closed with one perclose. hemostasis was achieved. an angiogram was done after the site was closed to confirm flow was restored. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3417 | 1220648-2025-48293 | 23554262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23554262&pc=OZD | 09/18/2025 |  |  | 11/14/2025 |  |  | 09/18/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662645 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 55 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient encountered ischemia during impella cp support. during impella cp insertion, the patient's pedal pulses became unattainable and it was treated by applying a femorofemoral bypass. however, after three days of support, care was withdrawn and the patient ultimately expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3418 | 1220648-2025-48327 | 23554434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23554434&pc=OZD | 10/17/2025 | 12/29/2025 |  | 11/14/2025 |  |  | 11/04/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710804 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was placed onto impella cp support for high-risk percutaneous coronary intervention. during the intervention, the patient experienced significant hemodynamic instability and was transferred to the intensive care unit with this system for further stabilization. gradual stabilization of the patient but was already accompanied by clinical malperfusion of the left leg. after removal of the impella cp, significant ischemia on the right side was noted requiring intervention (percutaneous transluminal angioplasty). | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6 impact code f2303 has been added.
3419 | 1220648-2025-48353 | 23554782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23554782&pc=OZD | 10/17/2025 | 12/31/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736344 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed onto impella support for coronary artery bypass graft. while on support, the patient was having occasional runs of supraventricular tachycardia. the patient also had premature atrial contractions. an ultrasound showed a large left pleural effusion. the patient went to the operating room for a scheduled left video-assisted thoracic surgery hemothorax evacuation. multiple clots were evacuated from the left chest with cauterization of small bleeders in the intercostal space. a transesophageal echocardiogram was obtained and the impella was shallow, measuring at 2.4 centimeters. the doctor advanced the impella but the patient had significant ventricular ectopy, so the impella was pulled back to a measurement of 3.4 centimeters and left there. minimal pulsatility was noted and there was left ventricle separation. blood products were given per mass transfusion protocol until bleeding was controlled. the patient was rushed to interventional radiology for computed tomography angiography to confirm intercostal bleeding so the patient was taken across the hall a coiling of three tortuous intercostal arteries until the blushing stopped. immediate impella waveform improvement was noted as hemostasis was obtained. the patient was placed onto continuous renal replacement therapy. the patient developed a gastrointestinal bleed with bloody stool overnight. heparin was discontinued and the patient was placed on protonix. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Investigation summary: tachycardia/ arrhythmia/ thrombosis: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. renal failure : the cause of the injury was not determined due to insufficient clinical details. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot : 1969454. device history batch: sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3420 | 1220648-2025-48355 | 23555115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23555115&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740227 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient had hemolysis. blood products were given and good pump position was confirmed via imaging during the hemolysis. there was a tinged foley in urine. the pump was removed due to hemolysis. the patient landed in the intensive care unit in stable condition. once the field representative arrived at the patient¿s bedside, the impella was alarming placement signal low, and the impella appeared to have migrated into the ventricle. the impella cp was repositioned successfully with final measurement of 4 centimeters from the aortic valve annulus with inlet mid left ventricle space. the touhy was tightened. the pump was removed due to hemolysis. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. device in wrong position: the clinical states that the impella was alarming of low ps and it was confirmed with echo that the pump migrated into the ventricle. the impella was successfully repositioned and there was no more alarm. due to a lack of sufficient clinical information, the root cause of the how the device migrated into an improper position, cannot be determined. hemolysis: the clinical details state that there was hemolysis noted with the impella cp and earlier in the day the pump was in an improper position but that was corrected. there was no other information regarding the hemolysis. the data logs show that there were confirmed positioning alarms around the time of the positioning event, but the motor current and flows were stable throughout the case run. due to a lack of sufficient clinical information regarding if the hemolysis resolved with repositioning, and no product returned, the root cause of the mechanical interaction with blood cannot be established.
3421 | 1220648-2025-48356 | 23555172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23555172&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213831 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they attempted to prime the impella cp for acute myocardial infarction/cardiogenic shock st-elevation myocardial infarction patient. the automated impella controller (aic) was continuously in prime screen. the team took the cassette out and put back it in without resolution and the aic was not priming the solution. the patient was switched to a new aic with successful priming and set up. there was no patient harm. | [5 empty] | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 medical device problem code a1102 was erroneously entered on the initial manufacturer device report. h5 and h8 were erroneously entered on the initial manufacturer device report.
3422 | 1220648-2025-48357 | 23555355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23555355&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676290 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support the patient admitted in with a plan for high risk coronary artery bypass graft in the operating room. the implant was attempted several times with suboptimal positioning of the pump. it was reimplanted per the medical doctor with improved pump position but inflow was still directed towards inferoseptal wall. multiple attempts were made to improve the pump position. inflow to the aortic valve measurement was around four centimeters with the doctor electing to leave the pump in place despite suboptimal position. the pump was later explanted successfully and patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. difficult to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had occlusion of vessel preventing successful delivery of impella. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3423 | 1220648-2025-48358 | 23555430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23555430&pc=OZD | 10/15/2025 | 12/31/2025 |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758735 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 80 KG | Hispanic |  | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a patient admitted in without a medical history and who was found to be suffering from st-elevation myocardial infarction. the impella cp was placed and allowed for the team to assess the coronary artery disease and perform percutaneous coronary intervention. the impella cp was removed after two days due to poor positioning and hemolysis. the team escalated to an impella 5.5 and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3424 | 1220648-2025-48359 | 23555694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23555694&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727192 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was a delivery issue during impella cp implant. there was a decision to abort the impella insertion due to complicated aortic stenosis and improved patient hemodynamics. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had complicated aortic stenosis preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
3425 | 1220648-2025-48360 | 23555892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23555892&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662799 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had an impella cp being delivered via the v18 wire (a non-abiomed wire) to the left ventricle for the support of a patient admitted in with acute myocardial infarction/cardiogenic shock. the patient has known systemic comorbidities inclusive of renal failure on continuous renal replacement therapy and pulmonary edema. the pump delivery failed as the wire had kinked and the entire system of wire and pump was replaced. additional information was received. the pump was discarded in the procedure room as the failure was felt to be from the v18 wire. a new impella was placed as a second pump in the same procedure. the patient was transferred to the intensive care unit. the patient was escalated to extracorporeal membrane oxygenation (ecmo) the following day due to remaining in profound cardiogenic shock. the patient had insufficient flows with ecmo and continuing of care was felt to be futile and therefore, care was withdrawn. they do not allege the delivery failure contributed to the outcome. this complaint is for first pump (serial number (b)(6)) and an additional report will be sent for the second pump as the patient expired (serial number (b)(6)). | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had abnormal anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: "dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve." section: warnings & cautions: warnings: "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance." | Medical safety/clinical review was completed. the patient presented with a one-day history of shortness of breath, progressing to worsening dyspnea and lethargy. emergency medical services were contacted, and the patient developed respiratory failure requiring intubation. upon arrival to the intensive care unit, the patient experienced a tachycardia arrest requiring cardiopulmonary resuscitation and defibrillation with return of spontaneous circulation. the patient demonstrated worsening lactic acidosis, an echocardiogram showing severe pulmonary edema with a left ventricular ejection fraction of 25¿30%, and a cardiac index of 0.9. troponins were elevated with global wall motion abnormalities and ecg changes concerning for myocardial infarction. continuous renal replacement therapy was initiated due to worsening renal function and rising potassium levels. the patient was taken to the catheterization laboratory for left heart catheterization and possible impella support. coronary angiography noted "clean" coronaries; however, due to worsening lactates, a decision was made to implant an impella cp device for mechanical circulatory support (mcs). during impella cp placement, the physician noted concern for abnormal cardiac anatomy and suspected the pigtail may have been positioned near the papillary muscle. repositioning was attempted; however, the device crossed the aortic valve and was removed. re-access of the left ventricle was attempted using a v.18 guidewire. advancement of the impella was unsuccessful due to buckling at the aortic arch/ascending aorta junction, and a kink was noted in the v.18 guidewire. the device was removed, and a new impella cp catheter was successfully placed. bleeding and oozing were noted at the access site, and a femstop device was applied. it was reported from the physician that a deeper tract was made to ensure perclose sutures were deployed into the artery rather than subcutaneous tissue. it was noted that one of the 6 fr dilators had been advanced relatively deep into the artery likely distorting the arteriotomy and contributing to bleeding. the dissected tract was not repaired prior to or during impella support. following transfer to the surgical intensive care unit, the patient continued to have worsened lactic acidosis. the patient was escalated to venoarterial extracorporeal membrane oxygenation (ecmo) with the impella used as left ventricular venting on 10/16/2025, which was the next day. despite this support, it deemed futile, and the patient subsequently expired after withdrawal of care. the first impella cp implant that was aborted due to failure to advance and appears to have been related to the patient's anatomy based on information at hand. this device was successfully replaced and is not believed to be related to the demise outcome. the first impella cp is a malfunction type of reportable event. the second impella cp is related to access site bleeding. per the physician this was a user-related access site injury with bleeding complication. the patient's demise is most likely due to the underlying condition cardiogenic shock stage e despite the patient's escalation to mcs and elcs (impella cp and ecmo). care was withdrawn leading to the demise outcome. however, the second impella cp will be reported for the death; the device was in use at the time of expiration. correction: fields b6 tests/lab data including dates and b7 medical history updated with complete, correct details respectively. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3426 | 1220648-2025-48361 | 23556001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556001&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734850 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Fever | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient found down and unresponsive in the community by the family and was transported in cardiogenic shock/acute myocardial infarction. the pump was placed and supported for two days despite being in poor position, hypotensive and febrile with possible mixed shock/pneumonia. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported event has been completed. no product was returned. no product was returned. fever, hypotension: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details.
3427 | 1220648-2025-48362 | 23556042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556042&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026760547 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient¿s urine was pink tinged now. sites are hemostatic. the team was trying to assess what the next steps are, but there is no clear exit strategy. the patient was weaned down to p4 yesterday and is only on dobutamine now. there are no concerns or issues and they plan to explant the pump. the site was within normal limits and ejection fraction was up to 35%. platelets were down to 24. heparin-induced thrombocytopenia panel was pending. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3428 | 1220648-2025-48363 | 23556098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556098&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726772 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the patient encountered arrhythmia and hematuria during impella cp support. it was reported that on the day of explant, patient had very red urine for pfhb (plasma free hemoglobin) that was 225. additionally, the next day patient had torsade's and sustained ventricular tachycardia (vt's) over night. amiodarone was given to the patient as a result. a few days later, patient experience ventricular ectopy's with loss of consciousness and amiodarone was increased subsequently. | [5 empty] | H6: added codes based on information in the complaint file. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood (hematuria / hemolysis): after implant, pfhb 225 with very red urine. positioning was unable to be fully determined due to difficulty obtaining a transthoracic echocardiogram. p level turned down and issue resolved within 24 hours and remained resolved for the remainder of the case despite increasing the p-level. data logs were not recovered. the cause of the mechanical interaction with blood was not determined due to no product returned, no data logs recovered, and insufficient clinical details. patient codes (arrhythmia) - the root cause of the injury was unable to be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3429 | 1220648-2025-48364 | 23556190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556190&pc=OZD | 10/17/2025 |  |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674794 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with left main (lm) and diffuse left anterior descending (lad) artery disease presented for a successful primary percutaneous coronary intervention of lm and lad, however it was from the case onset utilizing low dose inotropic and pressor support for hypotension which continued intermittently throughout the procedure. hemodynamics were improved as revascularization occurred and the impella was successfully weaned and explanted without complications or concerns. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hypotension: the cause of the injury was not determined due to insufficient clinical details.
3430 | 1220648-2025-48365 | 23556304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556304&pc=OZD | 10/16/2025 |  |  | 11/14/2025 |  |  | 10/16/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703459 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention (hrpci) was to be implanted with an impella cp for mechanical circulatory support. the patient was brought to the cardiac catheterization laboratory for a hrpci. the pump was opened, being prepared for insertion, and the peel away sheath was inserted into the patient when the physician had a failure with the pre-close technique which caused a massive hematoma at the access site. due to the large hematoma at the impella access site, the implantation was aborted for patient safety and the procedure was rescheduled for after the hematoma and arterial complication was resolved. additional information was received. the hematoma was caused by the perclose before the impella sheath was inserted. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the injury is most likely use related since pre-close failure caused hematoma. device history review: the device passed all post sterile inspection checks.
3431 | 1220648-2025-48366 | 23556405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556405&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702580 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. due to anatomical issues and calcification of the patient¿s vessels, delivery of the impella to the ventricle was not possible and the implant was aborted. the impella cp was explanted, and the patient's outcome at the time of explant was survived. | [5 empty] | Product complaint # (b)(4). investigation summary no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and calcification preventing successful delivery of impella. device history lot device lot: 1947068 device history batch subcomponent lot: n/a. device history review device (sn: (b)(6) passed all post sterile inspection checks. | Section a: patient demographics are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3432 | 1220648-2025-48367 | 23556489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556489&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691169 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.5 for mechanical circulatory support. the patient was in fulminant cardiogenic shock with a severely reduced ejection fraction of 5%. postpartum cardiomyopathy was suspected as the underlying cause. the patient was initially supported with extracorporeal life support via peripheral veno-arterial extracorporeal membrane oxygenation (va ecmo) and the clinical team made the decision to escalate the patient to ecpella support with an impella 5.5. the impella 5.5 could not be delivered to the left ventricle due to the patient¿s vascular anatomy. after several attempts, the procedure was aborted, and the decision was made to implant the impella cp. the impella cp was successfully implanted via the right axillary artery. on day three of impella cp support, the patient was noted to be in very critical condition and was entirely dependent on mechanical circulatory support therapy. due to severe ischemia in the patient¿s right leg from a complication caused by va ecmo, the medical team made the difficult decision to amputate the patient's right lower leg. additionally, it was noted that the pump was running well despite several suction alarms, which the medical staff attributed to supporting the patient with both va ecmo and impella cp simultaneously. the following day, the patient expired in the setting of septic shock and multiple organ failure. the physician stated that the impella functioned without problems until the time of death, and the impella cp was not the cause of death. this complaint involves two automated impella devices: pump 1: impella 5.5, with serial number (b)(6) pump 2: impella cp, with serial number 595450 this report specifically addresses pump 1: impella 5.5, with serial number 600942. a separate report will be submitted for the other device associated with this complaint. refer to section h10 for the related report number. | [5 empty] | Product complaint #(b)(4). investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history lot: device lot: 1932182. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. a 35-year-old female with an unknown prior medical history presented in fulminant cardiogenic shock with an estimated ejection fraction of approximately 5%. postpartum cardiomyopathy was suspected. primary extracorporeal life support (ecls) was initiated with peripheral veno-arterial extracorporeal membrane oxygenation (va-ecmo). due to persistent severe cardiogenic shock, a decision was made to escalate to combined ecpella support. impella 5.5 was initially selected; however, advancement into the left ventricle was unsuccessful due to patient vascular anatomy. after multiple attempts, the procedure was aborted, and a right axillary artery impella cp was successfully implanted instead. two days after initiation of impella cp in combination with va-ecmo, multiple suction alarms and low pump flow events were observed which can be expected with combined therapies and patients that are in persistent cgs. the patient remained in critical condition and was noted to be 100% dependent on mechanical circulatory support. on the same day, right lower extremity ischemia was identified and attributed to ecls cannulation. the clinical team proceeded with right lower leg amputation. subsequently, the patient developed complete multiorgan failure (mof). given continued clinical deterioration and poor prognosis, therapy was limited. the patient expired. the unsuccessful attempt to place the impella 5.5 due to vascular anatomy will be reported as a malfunction type of reportable event (inability to advance). the death will be conservatively reported against the impella cp. the patient underlying condition (fulminant cardiogenic shock, suspected secondary to postpartum cardiomyopathy) contributed most to the demise outcome complicated the limb ischemia related to va-ecmo cannulation, and subsequent multiorgan failure. per the physician, the demise was not considered related to impella device function. note: h6. component code and investigation type/findings/conclusion coding remain unchanged as previously reported. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | A4, a5, a6 are unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3433 | 1220648-2025-48369 | 23556721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23556721&pc=OZD | 10/14/2025 | 01/06/2026 |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719156 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient presented with st elevation myocardial infarction and severe coronary artery disease. on impella cp support day three, it was reported that the previous day the patient experienced multiple ventricular tachycardias (vt) that required electrical shock and subsequent intubation. the medical team suspects the vt was caused by levosimendan, which was discontinued, and the patient's condition has since improved. the impella was noted to have been running well at p7, which was adjusted down due to suction issues. the patients¿ left ventricular ejection fraction (lvef) was noted as very poor but showed a slight improvement with impella support. escalation to an impella 5.5 was being considered, and the medical team planned to discuss this option further with cardiovascular surgery. on impella cp support day five, the patient developed a cold leg below the impella insertion site. it was noted that the patient was on a high dose of catecholamine. the medical team decided to escalate the patient to an impella 5.5. the impella 5.5 was successfully implanted, and the impella cp was explanted with no complications reported. after the impella cp was removed, the mal perfusion of the right leg was resolved. it was noted that the patient remained in critical condition. the patient remained on impella 5.5 support for 11 days awaiting surgery for a durable left ventricular assist device with no complications reported. the device was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. | Investigation summary impella cp sn 607503 + pc returned for evaluation. no issues were found on the returned product. arrhythmia, ischemia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1968141. device history batch: subcomponent lot: n/a. device history review: device (sn: 607503a) passed all post sterile inspection checks.
3434 | 1220648-2025-48370 | 23557017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23557017&pc=OZD | 10/18/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758707 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision was made for impella cp when a patient's electrocardiogram showed st segment elevation myocardial infarction. the impella was placed and initially blood pressure and flows were low however eventually improved with epinephrine boluses. while on support it was noted that the patient had a gastrointestinal bleed in addition to bleeding from other areas, hemoglobin dropped and 3 units of bleed were given. the patient remained on support for a total of 41.80 hours and was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b13 and b24 h11: added the results of the investigation. investigation summary clinical symptoms (hemorrhage/bleeding, low blood pressure/hypotension): the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow: the pump was not returned for investigation. data log shows the pump flow was within specification range throughout the case except at the time of the reported suction issue, which was related to pump position. pump flow had some fluctuation at the start of the case, with a decreasing trend that correlated with the placement signal trend. according to the clinical details, flows and blood pressure were low but improved with epinephrine boluses. ejection fraction (ef) was reported at 10%, with bleeding episodes and blood transfusions given. the cause of the reported low pump flow issue was most likely due to patient condition, based on the provided clinical details and data log review. device history review ==> the pump (b)(6) passed all post sterile inspection checks.
3435 | 1220648-2025-48371 | 23557107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23557107&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731527 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | Difficult to Remove | Detachment of Device or Device Component |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered a high purge pressure during impella 5.5 support. the doctor reported decreased purge fluid (pf) at 3.3 milliliters per hour (ml/h). the bag and system were changed, and the pf was subsequently at 3.9 ml/h. further steps and the problem were discussed, including purge lysis, and contact with the hotline was recommended for a change. later, an alarm "pp high" was present, where purge pressure (pp) was over 1100 mmhg and purge fluid was less than 2ml/h. lysis protocol was sent out with recommendations for application. recommendations were also given to change purge fluid and cassette to prevent future complications. lysis was done by the medical team. there were no further issues with the purge system. the pump ran smoothly and motor current was in range. the patient was in stable condition. additionally, during explant of the pump, strength and force had to be applied in order to remove the pump. during this procedure, the cannula detached from the catheter and was hanging by a wire. the complete system could be removed without any harm of the patient. | [5 empty] | The product without event logs was received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. high purge pressure: the clinical details state that there was a high purge pressure alarm with lower purge flows. altepase was put into the purge fluid. it is unknown if/when the alarms resolved. there are no data logs returned. the returned product had biomaterial noted in the inlet and outlet cage that was unable to be removed after soaking. fluid was not able to be pushed through the purge gap as the catheter was separated near the motor housing and the purge line was impacted. however, the pump was purged with the cassette returned and there were no issues noted with the side arm an fluid expelled from the catheter at the point of the separation. due to not being able to test the device as a whole and no data logs returned as well as clinical details confirming resolution of the alarms, the root cause of the high purge pressure cannot be established. difficult/unable to remove. the clinical details stated that strong force was required to remove the pump and the catheter ended up breaking and hanging by a wire. the returned product confirmed this as the catheter was broken near the motor housing and the nitinol was uncoiled. there was significant stretch seen on the catheter jacket but the cannula remained intact. there was significant clot in the inlet and outlet cage of the pump. due to the lack of clinical details on why a strong force was required to remove the pump, the root cause of the difficulty/unable to remove cannot be established. pd-catheter-(separation, break): the clinical details stated that strong force was required to remove the pump and the catheter ended up breaking and hanging by a wire. the returned product confirmed this as the catheter was broken near the motor housing and the nitinol was uncoiled. there was significant stretch seen on the catheter jacket but the cannula remained intact. there was significant clot in the inlet and outlet cage of the pump. based on the clinical details and returned product, the root cause of the product damage is most likely due to material fracture due to use as strong force was applied to the catheter on explant, causing it to separate. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
3436 | 1220648-2025-48372 | 23557544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23557544&pc=OZD | 10/17/2025 | 12/01/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401379 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during case startup, aic showed "purge system open" alarm. staff attempted to remove and reinsert the purge cassette. the purge section of the aic screen (bottom center) remained blank and the alarm state remained. the staff swapped out the aic and was able to successfully to complete the case. the staff mentioned that they may have had the purge cassette plugged in prior to aic startup. there was no patient harm reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was received for evaluation. upon completion of the investigation, a supplemental will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. data analysis: a review of the logs confirmed the reported failure mode. device analysis: the console (ic6973) was tested. when the test cassette and pump were run under the intended settings and configurations, the reported failure mode - ¿purge system open¿ alarm - could not be reproduced. the ¿purge system open¿ alarm was forcefully reproduced by connecting the test cassette and disc without purge solution. in this condition, the purge pressure dropped below 100 mmhg, which triggered the alarm. root cause: the root cause of the ¿purge system open¿ alarm could not be determined, as the issue could not be reproduced under the intended settings and configurations. no evidence was found in the logs to confirm the reported product experience code - ¿purge disk not detected.¿ it is likely that the ¿purge system open¿ alarm was reported as ¿purge disk not detected.¿
3437 | 1220648-2025-48374 | 23558208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23558208&pc=OZD | 10/15/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712138 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 113 KG | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had a 22 day support with an impella 5.5 pump after being admitted in with cardiac co-morbidities post covid. the patient has cardiomyopathy and is on united network for organ sharing transplant list. during the 22 days of support, the pump had suction and optical signal issues. the pump was explanted on day 22 and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: based on the data logs and clinical information the root cause of the optical sensor issue is due to material wear at 17 days which is within the intended design life. pump suction: the root cause of the pump suction reported on is due to material wear of the optical sensor. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3438 | 1220648-2025-48373 | 23558683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23558683&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 56 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was placed onto impella support for high-risk percutaneous coronary intervention. patient experienced some bradycardia and aberrant rhythms post-impella placement. two perclose and angioseal were required to achieve hemostasis upon removal. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia the cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3439 | 1220648-2025-48376 | 23559001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23559001&pc=OZD | 04/19/2025 | 12/31/2025 |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608691 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had a month long of impella 5.5 pump support for cardiogenic shock and cardiomyopathy. the patient was explanted and transplanted and survived. the long-term outcome and quality indicator (loqi) impella registry reported months later reported upon the adverse event of the ventricular tachycardia treated with medical therapy. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. ventricular tachycardia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1885263 device history batch: subcomponent lot: n/a device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3440 | 1220648-2025-48375 | 23559485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23559485&pc=OZD | 10/16/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/16/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635416 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in st elevation myocardial infarction (stemi)/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a witnessed ventricular fibrillation arrest while bike riding, receiving immediate bystander cardiopulmonary resuscitation (cpr). he presented to the hospital with an anterior stemi, undergoing a total of 45 minutes of cpr and 17 defibrillations before return of spontaneous circulation was achieved in the cardiac catheterization lab (ccl). an impella cp was immediately placed via the right femoral artery, and single-access percutaneous coronary intervention (pci) of the proximal left anterior descending coronary artery was performed. it was noted in the ccl that the pump very ¿positional¿ and had intermittent suction. the pump had been placed shallow to achieve optimum flows and had occasional "in aorta" alarms, but motor current remained pulsatile. the patient was stable post pci and was transferred to another hospital. upon arrival at the hospital¿s intensive care unit, the patient had pulsatile bleeding from the impella access site. the physician noted that the impella had not been sutured at the previous hospital after insertion. forward tension sutures were placed, and the site was redressed with angle matching and surgical was deployed to achieve hemostasis. the patient was successfully weaned off of impella cp, had begun weaning pressors, and was extubated. the patient was noted to show significant improvement and the patient¿s blood pressure stabilized once sedation was discontinued, and the patient is now only on low-dose levophed and vasopressin. the patient is completely intact neurologically, has good intrinsic cardiac function, and is currently visiting with family. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3441 | 1220648-2025-48377 | 23559560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23559560&pc=OZD | 10/15/2025 |  |  | 11/14/2025 |  |  | 10/15/2025 | 11/14/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722001 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had an impella cp placed for the support of a patient admitted in with post ventricular fibrillation arrest and emergency medical services support on way to hospital. the impella cp was placed and allowed for the multivessel disease, coronary artery disease, and percutaneous coronary intervention to occur but, the team did note optical signal loss from the point of insertion, possibly from sensor damage at insertion. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella cp and usb data stick were returned for evaluation. section d9 was updated. the cause of optical sensor issue was no defect found per product and log analysis. the device passed all post sterile inspection checks.
3442 | 1220648-2025-48379 | 23559857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23559857&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026685596 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 51 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a cp inserted to support the patient admitted in actively coding and in need of care escalation from the intra-aortic balloon pump. the cp was placed but explanted after just 19 minutes with saturated flows. the team replaced the cp with another and on that 2nd pump the patient expired. | [5 empty] | Correction: e4 and h6 medical device problem code 140508 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. saturated flows: the data logs were reviewed and full case showed no events of pump having saturated flows. there was no defect found with relative to saturated flows. low pump flow/pump suction: the data logs were reviewed and motor current (mc) spike greater than1400 ma observed soon after implant. the cause of low pump flow/pump suction was due to biomaterial ingestion based on log and clinical analysis. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. an 80-year-old female with a significant medical history of coronary artery bypass grafting (cabg) performed on (b)(6) 2025 was admitted postoperatively to the intensive care unit (icu). on postoperative day one, the patient experienced cardiac arrest in the icu. an intra-aortic balloon pump (iabp) was placed for hemodynamic support. the patient experienced a subsequent cardiac arrest while in the icu and was transferred emergently to the cardiovascular laboratory for placement of an impella cp and performance of left heart catheterization (lhc). the impella cp device (pump 1) was successfully placed during active cardiopulmonary resuscitation. following placement, elevated motor current and low flow alarms were observed, consistent with suspected thrombus ingestion. due to concern for pump thrombosis and impaired device function, impella cp pump 1 was removed and exchanged for a second impella cp device (pump 2). left heart catheterization demonstrated occlusion of all coronary artery bypass grafts. percutaneous coronary intervention (pci) was performed on the native left anterior descending (lad) artery and the right coronary artery (rca). despite revascularization and mechanical circulatory support, the patient experienced multiple recurrent cardiac arrests. return of spontaneous circulation (rosc) could not be achieved. the patient expired on support. impella cp pump 1 (sn(b)(6)) is being reported as associated with the death outcome because it was utilized during the resuscitative efforts in the setting of cardiogenic shock and cardiac arrest and was exchanged due to suspected thrombus ingestion and device dysfunction. however, the patient¿s death is considered primarily attributable to severe multivessel coronary artery disease, acute graft failure, refractory cardiogenic shock, and repeated cardiac arrest unresponsive to advanced resuscitative measures. note: h6. investigation type/findings/conclusion remain unchanged as previously reported.
3443 | 1220648-2025-48381 | 23559976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23559976&pc=PYX | 10/16/2025 |  |  | 11/14/2025 |  |  | 10/16/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026652264 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was a positioning issue during impella rp flex patient support. while on support, a chest x-ray showed the impella rp flex needed adjustment. an abiomed representative noted the impella rp flex needed adjustment due to cardiopulmonary resuscitation. the pump was advanced by 2 cm. care was withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was most likely pulled out of position by the user during cpr. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review. medical safety/clinical reviewed the event. the patient had a recent st-elevation myocardial infarction involving the right coronary artery and underwent implantation of an impella cp device for mechanical circulatory support. while on impella cp support, the patient demonstrated signs consistent with right ventricular failure, including elevated central venous pressure, rising lactic acid levels, frequent suction alarms, and reduced impella cp flow. the patient subsequently experienced a cardiac arrest. one round of cardiopulmonary resuscitation was performed, with return of spontaneous circulation achieved. escalation of support was required, and the decision was made to implant an impella rp flex device for right-sided mechanical circulatory support. approximately two days later, the impella cp device was explanted. during continued impella rp flex support, a positioning issue was identified. chest radiography demonstrated that the impella rp flex required adjustment. actions taken were unnoted. the following day, the patient expired after the family elected to withdraw care. the event story involves two product complaints: impella cp - mdr 1220648-2025-48234: serious injury. impella rp flex - mdr 1220648-2025-48381: death. related medical device reports: this impella rp flex device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
3444 | 1220648-2025-48380 | 23560083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23560083&pc=OZD | 10/17/2025 | 11/14/2025 |  | 11/14/2025 |  |  | 10/17/2025 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757676 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp placed for the support of a patient admitted in with acute myocardial infarction complicated by cardiogenic shock. the pump supported despite signs of hemolysis until care was withdrawn and patient expired on day 2 of the support. | [5 empty] | H6: omit code a150202, the physician confirmed pump was in good position. added code b13. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. data review: data logs showed pump ran without issue during support. there were no alarms triggered during the case. there were no mc spikes, abnormal ps or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely patient condition related since the physician confirmed hemolysis was due to svr (systemic vascular resistance) being so high. device history review: pump sn: (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3445 | 1220648-2025-31927 | 23560178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23560178&pc=OZD | 01/31/2024 | 11/14/2025 |  | 11/14/2025 |  |  | 01/31/2024 | 11/14/2025 |  |  | 05/31/2024 |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023211916 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 for mechanical circulatory support. the sensor failed with lost placement signal on day three with no issues with the patient¿s support. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. placement signal issue: the pump was returned for investigation, however, only the part from the red handle to the cannula was returned. the pump was connected to a spare connector patient cable from the lab to connect to the console and to the electrical testing module for further testing. a placement signal not reliable (psnr) alarm was observed when the pump was connected to the console. the pump passed electrical testing. there were minor scratches seen on the dp sensor silicone under microscopic guidance. no data logs were returned for analysis. the root cause of the placement signal issue was not determined due to lack of sufficient clinical details, partially returned product, and unreturned data logs. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3446 | 1220648-2025-31928 | 23560203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23560203&pc=OZD | 08/18/2023 | 11/14/2025 |  | 11/14/2025 |  |  | 08/19/2023 | 11/14/2025 |  |  |  |  | 11/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the complainant reported that the ¿placement signal unknown¿ alarm was present. the waveform aortic and motor current pulsatile. echocardiography was completed and position correct, but unaware of left ventricle measurement. awaiting the final read of echocardiography. the clinician was informed on how to disable placement signal audio. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs (obtained from the cloud) are consistent with complaint and show placement signal not reliable alarm (#130, due to low snr), preceded by multiple events 130, indicating console's inability to reliably calculate placement due to poor quality of optical signal, manifested by low snr. ltlog and rtlog data indicates that initially optical signal had good quality, but gradually worsened over time, triggering psnr condition (130) within 5 hours from pump start. device analysis: console (b)(6) was not returned for the initial investigation (b)(4) after investigation (b)(4) against pump (b)(4) identified that the console might have contributed to the issue, therefore root cause of psnr condition could not be established at that time. the console was eventually returned for (b)(4), where complaint (b)(4) alleged psnr alarms with pump (b)(4). the investigation concluded that degradation of optical signal was caused by malfunction of the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3447 | 23542445 | 23542445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23542445&pc=OZD | 09/21/2025 | 10/22/2025 | 10/22/2025 | 11/13/2025 | 10/22/2025 | 11/13/2025 |  | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | No answer provided |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Nurse | [6 empty] | 74 YR |  | Male |  |  |  | Injury | Yes | Yes | Break | Migration | [5 empty] | Foreign Body In Patient | [12 empty] | Required Intervention | Hospitalization | [23 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During impella explant procedure, the pump broke off inside the patient and started migrating. the surgeon tried to chase the pump. he was able to control the blood flow to stop the pump for migrating further. the patient needed to have a sternotomy to retrieve the impella pump after locating it with fluoroscopy.
3448 | 1220648-2025-48329 | 23542499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23542499&pc=OZD | 10/14/2025 |  |  | 11/13/2025 |  |  | 10/14/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690360 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day 2 of support, the patient¿s impella access site was noted to be oozing and the pump was reported to have suction alarms. the patient was transfused with two units of packed red blood cells and the dressing was changed at the impella access site. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. visual inspection found no issues on the pump. major bleed: the cause of major bleed was most likely use related since bleeding resolved with dressing change and angle matching. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B7 added information that was omitted from the initial report that was submitted. d9 has been updated to reflect product was returned for investigation. h6 type of investigation, investigation findings, and investigation conclusions codes and h11 have been updated to reflect product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3449 | 1220648-2025-48330 | 23542650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23542650&pc=OZD | 10/14/2025 | 12/29/2025 |  | 11/13/2025 |  |  | 10/14/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692687 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male | 100 KG |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported on the first day of impella 5.5 support the pump experienced high purge pressures with pressures >1000 mmhg and purge flows < 1 ml/h. the physician requested tissue plasminogen activator protocol. as a result of the unresolved high purge pressure, the impella 5.5 was replaced with a new impella 5.5 pump the following day. the pump was successfully exchanged. at the time of writing this report, the patient continues to be supported by the replacement impella 5.5 with no complications reported. | [5 empty] | Investigation summary: no product or data logs were returned for evaluation. high purge pressure: clinical details noted that high purge pressure was noted on first day of support. pump was exchanged the following day. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation. device history lot: device lot: 1932200. device history batch: subcomponent lot: n/a. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3450 | 1220648-2025-48331 | 23542752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23542752&pc=OZD | 10/14/2025 | 02/26/2026 |  | 11/13/2025 |  |  | 10/14/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603902 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male | 95 KG |  | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Anemia | Arrhythmia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient presented to the hospital with a late-diagnosed st elevation myocardial infarction due to an occluded left anterior descending (lad) artery. the patient was transferred to an outside facility with stage c cardiogenic shock and has a low ejection fraction of 19% noted on echocardiogram. there were no additional complications such as mitral regurgitation, left ventricular thrombus, or mechanical abnormalities noted at that time, and was currently receiving inotropic support with milrinone and a nipride infusion. the medical team decided to implant an impella 5.5 device with initial stabilization with an intra-aortic balloon pump during the procedure. the patient was stable post-procedure and is currently in the unit for further management, with a decision regarding long-term options on day 16 of support, the patient underwent a successful coronary artery bypass graft (cabg) procedure. post-operatively, a missed bleeder caused a blood loss of approximately 400 cc which was not related to the impella 5.5 support. on day 21 of support the patient was scheduled to be extubated, however this was interrupted. the patient experienced a run of ventricular tachycardia that degenerated into torsades de pointes the patient was paced to manage the rhythm disturbance. the following day the patient was noted to be septic due to hospital acquired pneumonia. the patient was placed on a broad-spectrum antibiotic and was noted to not be related to impella use. the patient continued to show no heart recovery but was stable over the next several days. on day 34 of support, the patient was noted to be alert and oriented and continued to remain on impella support. on day 39 of support, the patient started physical therapy. it was noted that the patient¿s hemoglobin and hematocrit dropped, and the plan was to transfuse the patient to address the anemia. the patient was noted to be stable. at the time of writing this report, the patient continues to be supported by impella 5.5. | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient presented to the hospital with a late-diagnosed st elevation myocardial infarction due to an occluded left anterior descending (lad) artery. the patient was transferred to an outside facility with stage c cardiogenic shock and has a low ejection fraction of 19% noted on echocardiogram. there were no additional complications such as mitral regurgitation, left ventricular thrombus, or mechanical abnormalities noted at that time, and was currently receiving inotropic support with milrinone and a nipride infusion. the medical team decided to implant an impella 5.5 device with initial stabilization with an intra-aortic balloon pump during the procedure. the patient was stable post-procedure and is currently in the unit for further management, with a decision regarding long-term options on day 16 of support, the patient underwent a successful coronary artery bypass graft (cabg) procedure. post-operatively, a missed bleeder caused a blood loss of approximately 400 cc which was not related to the impella 5.5 support. on day 21 of support the patient was scheduled to be extubated, however this was interrupted. the patient experienced a run of ventricular tachycardia that degenerated into torsades de pointes the patient was paced to manage the rhythm disturbance. the following day the patient was noted to be septic due to hospital acquired pneumonia. the patient was placed on a broad-spectrum antibiotic and was noted to not be related to impella use. the patient continued to show no heart recovery but was stable over the next several days. on day 34 of support, the patient was noted to be alert and oriented and continued to remain on impella support. on day 39 of support, the patient started physical therapy. it was noted that the patient¿s hemoglobin and hematocrit dropped, and the plan was to transfuse the patient to address the anemia. the patient was noted to be stable. at the time of writing this report, the patient continues to be supported by impella 5.5. additional information received on 23-nov-2025. it was reported that on support day 41, there were impella alarms for high purge pressure and it was noted that the purge flow dropped to 1.5-3.0 cc/hour and the purge pressure increased to 1070 mmhg. tissue plasminogen activator (tpa) was added to the impella purge solution, and resolved the high purge pressure and low purge flow issues. following the tpa protocol, it was noted that the purge pressure dropped to 575 mmhg, and the purge flow increased to 4 cc/hour. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This follow-up report is being submitted to update b2, b5, and h6 codes for new information that was received on november 23, 2025, and for investigation conclusions. b2 is being updated to include product problem b5 is being updated to include new information that was not included in the initial report. the revised b5 provides a complete event narrative. d6b has been updated to reflect date of explantation. h6. health effect - impact code code for medication required has been added h6 medical device problem code increase in pressure has been added h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product or data logs were returned for evaluation. anemia: clinical details noted that patient had a drop of hemoglobin and hct and received transfusion. the cause of the injury was not determined due to insufficient clinical details. arrythmia: clinical details noted that patient had runs of vtach. the root cause of the injury was unable to be determined due to insufficient clinical details. high purge pressure: clinical details noted that high purge pressure alarms with purge floes of 1.5-3.0ml/hr and purge pressure of 900 that increased to 1070mmhg. tpa infusion resolved issue. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation.
3451 | 1220648-2025-48332 | 23542929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23542929&pc=OZD | 10/15/2025 | 11/13/2025 |  | 11/13/2025 |  |  | 10/15/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722329 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male | 104 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Arrhythmia | Renal Failure | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was escalated from the impella cp to bipella support of the impella 5.5 and impella rp flex. while on support, patient was placed onto dialysis. the patient has been going in & out of a wide complex quick response sequence tachycardia for the past few days, went back into that rhythm overnight requiring increased presser support. tachycardia recurred two days later and patient was cardioverted. pump presented false impella in aorta alarm; placement was verified with point-of-care ultrasound. this complaint is for the impella 5.5 additional reports will be sent for the impella cp and impella rp flex. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 codes updated.
3452 | 1220648-2025-48333 | 23543206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543206&pc=OZD | 10/15/2025 | 11/13/2025 |  | 11/13/2025 |  |  | 10/15/2025 | 11/13/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597704 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk off-pump triple coronary artery bypass graft (opcab) was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was successfully placed without complications via the aorta and the opcab was completed successfully. the patient was transferred to the intensive care unit to rest and recover. on day 2 of support, it was reported that the patient had a fever, with a temperature of 102.4 degrees fahrenheit. blood cultures were drawn and it was noted that anticoagulation was withheld. the following day, the patient¿s temperature dropped to 99.1 degrees fahrenheit and the patient was sitting up in bed and alert. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary no product was returned for investigation. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot device lot: 1872080. device history batch subcomponent lot: n/a. device history review device with sn: (b)(6) passed all post sterile inspection checks.
3453 | 1220648-2025-48334 | 23543286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543286&pc=OZD | 10/16/2025 | 11/13/2025 |  | 11/13/2025 |  |  | 10/16/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756681 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | Renal Failure | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient having intermittent ventricular tachycardia. point of care ultrasonography done at bedside, physician happy with impella position. per bedside nurse, patient is hemolyzing with red urine. initial plasma free that was drawn last night was elevated into the 80s and creatinine has risen over 3. the patient is on continuous veno-venous hemofiltration and peripheral va ecmo. p level decreased and fluids given. plan to recheck plasma free later today. dropped to p5 and fluids given. the patient was weaned off ecmo and impella. the patient survived. | [5 empty] | Correction h6 health effect - clinical code 060109 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3454 | 1220648-2025-48335 | 23543472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543472&pc=OZD | 10/16/2025 | 11/13/2025 |  | 11/13/2025 |  |  | 10/16/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531322 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 52 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp on the patient was experiencing bleeding from right femoral artery impella access site. the patient¿s anti coagulation status was found to be super-therapeutic (ptt>200). care team administered kcentra for anticoagulation reversal. the bleeding has stabilized. care team gave kcentra, manual compression of impella access site, and surgical stitch around access site. two units of blood product was given for diastolic suction. impella flows p-6 (2.3lpm). the patient had intermittent sinus rhythm to junctional rhythm. reviewed by cardiology, levo/vaso trending down, aware of constant suction alarms, which occur during junctional rhythm and alleviate during sinus episodes. the patient explanted in operating room with vascular surgery utilizing vascular cut-down. patient successfully bridged to heart recovery. | [5 empty] | The investigation for the reported event has been completed. no product was returned. major bleed: the cause of the bleed was most likely use issue related since anti coagulation status was found to be super-therapeutic. pump suction: patient had constant suction alarms. after review of logs, multiple suction alarms were observed. no sudden motor current spike or placement signal trend was observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3455 | 1220648-2025-48324 | 23543496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543496&pc=OZD | 10/16/2025 | 11/12/2025 |  | 11/13/2025 |  |  | 10/16/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720165 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 57 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Tachycardia | Ventricular Fibrillation | Pericardial Effusion | Hemodynamic instability | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while impella cp support, the patient developed tachycardia, diaphoresis, and narrowed pulse pressure. subsequent suction events occurred on the impella following a drop in mean arterial pressure bedside echocardiography revealed a large pericardial effusion, with the impella in adequate position but an underfilled left ventricle. a pericardiocentesis was performed at the bedside, and a massive transfusion protocol was initiated with packed red blood cells, platelets, and cryoprecipitate. the patient then experienced ventricular fibrillation arrest, requiring advanced cardiac life support drugs, two rounds of cardiopulmonary resuscitation and defibrillation. the patient was deemed not a candidate for extracorporeal membrane oxygenation or further surgical intervention, and after family discussion, transitioned to comfort measures and expired in the intensive care unit. | [5 empty] | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned for evaluation. data logs were returned for analysis. log analysis confirmed low flows and suction alarms. a trending placement signal was observed towards the end of the case. no motor current spikes were seen outside of p-level changes. clinical details revealed that the patient became tachycardic, diaphoretic, and narrowed pulse pressure in the intensive care unit. subsequent suction events occurred on impella following drop in mean arterial pressures. bedside echocardiogram revealed large pericardial effusion, impella in adequate position however underfilled left ventricle. a pericardiocentesis was performed at bedside. mass transfusion protocol was initiated with packed red blood cells, platelets, and cryoprecipitate. the patient expired while on support. tachycardia, ventricular fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injuries was unable to be determined due to insufficient clinical details. pericardial effusion, cardiac arrest, hemodynamic instability: the cause of the injuries was not determined due to insufficient clinical details. pump suction: the cause of the pump suction was most likely patient condition related based on clinical details and data log analysis. device history review: the pump passed all post-sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3456 | 1220648-2025-48336 | 23543634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543634&pc=OZD | 10/16/2025 | 11/13/2025 |  | 11/13/2025 |  |  | 10/16/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741037 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that due to the patient¿s anatomy¿specifically a short aorta with a steep turn angle to the left ventricle¿after impella cp insertion, the cannula was observed to be kinked. this resulted in restricted flow, with the device unable to achieve more than 1 l/min. the physician decide to replace to a new device. there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. the returned product was investigated. a cannula kink was found under microscopic guidance. no biomaterial or foreign material was found around the inflow and outflow cages, or impeller blades. no other abnormalities were seen. the pump was successfully primed and run in the lab. no low flows were reproduced, and the pump performed as per specifications. data logs were returned for analysis. data log analysis confirmed the low flow rate. suction alarm was also seen, as were the impella position in ventricle alarms. no other abnormalities were seen. clinical details mention that a cannula kink was noticed after impella insertion and flows were unable to get higher than 1 l/min due to the patient's tortuous anatomy. the patient had a short aorta and a steep turn angle to the left ventricle. low or blocked pump flow: the root cause of the low or blocked pump flow was most likely cannula kink based on the clinical details provided and data log analysis. the root cause of the cannula kink was most likely due to the patient's tortuous anatomy based on the provided clinical details. patient anatomy or condition prior to therapy: the cause of the patient's tortuous anatomy was patient condition related based on provided clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3457 | 1220648-2025-48337 | 23543741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543741&pc=OZD | 10/14/2025 |  |  | 11/13/2025 |  |  | 10/14/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758539 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Cardiac Arrest | Hemolysis | Renal Failure | Respiratory Failure | Hematuria | Multiple Organ Dysfunction Syndrome | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient arrested when laid on the table in the catheter lab. they were intubated and ventilated emergently, and the impella cp was implanted via the right femoral artery. attempted to cardiovert the patient unsuccessfully and the heart rate remained at 150s. the patient was sent to the medical intensive care unit (micu) for further care. an arterial line was placed and the patient's rhythm changed to pulseless ventricular tachycardia (vt). defibrillated and returned to normal sinus rhythm, then converted back to pulseless vt. advanced cardiac life support was initiated and cardiopulmonary resuscitation (cpr) was ongoing with mean arterial pressures in the 80s throughout. the patient converted between vt and pulseless electrical activity for approximately 30 minutes. code drugs were administered for multiple rounds and return of spontaneous circulation (rosc) was achieved. echocardiogram demonstrated cardiac function. the aortic signal was flat, apulsatile, and there were no alarms. extracorporeal membrane oxygenation (ecmo) placement requested due to persistent vt. now on peripheral veno-arterial extracorporeal membrane oxygenation (va ecmo). echocardiogram at bedside demonstrated the inlet was 2.8 cm to the aortic valve area and was advanced to 3.2 cm. heparin was restarted after holding for ecmo placement. urine output was pink red. the patient arrived at the other hospital and lidocaine and amiodarone infusions for vt storm. minimal pulsatility noted on the echocardiogram. lactate dehydrogenase (ldh) this morning was 7340 and urine was red, indicating hemolysis. no blood products were required, the pump was repositioned, and p-level decreased. plan to optimize medication and place a pulmonary artery catheter. systemic anticoagulation infusing with argatroban. liver function tests (lft) down trending and kidney function remains elevated, but the patient was making small amounts of urine. discussions to escalate to an impella 5.5 was on hold as the patient was too unstable at this time. there was an impella in ventricle alarm. the patient started decompensating, was hypoxic, and on 3 pressors. the catheter was repositioned under echocardiogram guidance, and the marking was moved from 86 to 90 cm. the patient remained intubated, sedated on propofol, and continuous renal replacement therapy (crrt). va ecmo oxygen saturation (sp02) on monitor off due to poor perfusion. the patient remains on support. | [5 empty] | Explant date provided d6b updated. correction e4. data logs were returned for analysis. data log analysis confirmed the suction alarms at the beginning of the case. pump suction: the root cause of the pump suction was most likely patient condition related based on provided clinical details and data log analysis. mechanical interaction with blood: the root cause of the hemolysis was not determined due to inconclusive evidence from the provided clinical details and data log analysis. hematuria, organ failure, renal failure, respiratory failure, cardiac arrest: the cause of the injuries was not determined due to insufficient clinical details. tachycardia: the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3458 | 1220648-2025-48338 | 23543848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23543848&pc=OZD | 10/15/2025 | 12/30/2025 |  | 11/13/2025 |  |  | 10/15/2025 | 11/13/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740418 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during a coronary artery bypass graft (cabg) procedure supported by impella cp, the pump was inadvertently pulled the impella back and the pump alarmed for impella in aorta. the patient was noted to be stable, and the medical plan had been to remove the impella after anticoagulation therapy was discontinued. the medical team was instructed to increasing the pump from p2 from surgical mode, and placement signal monitoring was disabled. the cabg was successfully completed, and the impella cp was removed after the procedure was completed. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was accidently pulled out of position by the user. device history lot: device lot: 1970213. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
3459 | 1220648-2025-48339 | 23544107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23544107&pc=OZD | 10/16/2025 | 12/12/2025 |  | 11/13/2025 |  |  | 10/16/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026717000 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 59 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | Unspecified Infection | Hypovolemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with a history of previous takotsubo syndrome that required impella support, presented and was believed to be having another episode of takotsubo. decision for impella cp was made for unloading and stabilization of her condition. however, the impella was getting suction in auto and unable to flow above p-5 due to her having a large sigmoid septum due to her proximal septal hypertrophy. the patient was left on p5. during support it was noted of possible hypovolemia and albumin was given. in addition to concerns for infectious source for takotsubo, with high white blood cells and fever broke. the cp was successfully weaned and explanted. | [5 empty] | The investigation for the reported pump suction and hypovolemia has been completed. no product was returned for evaluation. data logs were returned for analysis. data log analysis confirmed suction alarms occurring around the morning and noon. isolated placement signal spikes with no known pattern were observed through the case. no motor current spikes were seen outside of p-level changes. suction alarms resolved after dropping p-level to p-5. no other abnormalities were seen.clinical details mention the pump was experiencing suction due to the patient's large sigmoid septum, and the doctor suspected that the sigmoid septal was pressing against the cannula of the pump. patient experienced hypovolemia and albumin was administered. high white blood cells count was observed in the patient along with a fever. the root cause of the pump suction was most likely patient condition based on the clinical details provided and data log analysis. the root cause of the hypovolemia was not determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3460 | 1220648-2025-31925 | 23544804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23544804&pc=OZD | 05/07/2021 | 11/13/2025 |  | 11/13/2025 |  |  | 05/12/2021 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1538036 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support for high-risk percutaneous coronary intervention. while on support, an automated impella controller (aic) purge issue with no resolution occurred. optical signal issues have also been reported. no patient harm has been reported. | [5 empty] | The investigation has been completed. clinical review: clinical staff reported placement signal issues with console ic8396. it was visible during the beginning of the case but became invalid while the pump was being inserted. this complaint was related to a clinical procedure, and there were no adverse events related to this product malfunction. data analysis: imc revealed there was a placement signal not reliable alarm (event set #1036 ¿ opm signal invalid) on the complaint date which occurred one minute and 30 seconds before the pump was started. rt logs showed the calculated placement signal went to the max value. device analysis: the console was tested in (b)(6) field service (fs). 5s parameters were within specification according to the functional check procedure (b)(4). gain: 1,38. contrast: 55,6%. snr: 4905. power: 4,01 w running an optical pump did not reproduce any issues related to the placement signal going invalid. the root cause of optical placement signal issues could not be determined because the failure mode could not be reproduced.
3461 | 1220648-2025-31926 | 23544845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23544845&pc=OZD | 11/03/2021 | 11/13/2025 |  | 11/13/2025 |  |  | 11/03/2021 | 11/13/2025 |  |  | 04/30/2023 |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070415 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported there were placement signal unreliable alarms and that there was no patient harm. the signal went from normal signals to dashed out aorta / left ventricle placement signal with active alarm placement signal unreliable. staff de-aired the purge system, then alarm activated, then resolved, then after a few minutes, waveforms slowly re-appeared although they were erratic, ultimately resolving to normal signals with alarm extinguishing. the white cable was pushed in all the way to the automated impella controller. | [5 empty] | The investigation has been completed. the cause of the optical placement signal alarm issue was most likely use related since support continued for 3 days after troubleshooting which included pushing the white cable in all the way.
3462 | 1220648-2025-48341 | 23545395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23545395&pc=OZD | 10/14/2025 | 12/31/2025 |  | 11/13/2025 |  |  | 10/14/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the emergency room (er) with chest pain, shortness of breath, and elevated troponin levels. she suffered two cardiac arrests in the er and catheterization laboratory. following successful resuscitation and achievement of return of spontaneous circulation, a st elevation myocardial infarction code was called. a left heart catheterization revealed left main and right coronary artery disease. given the presentation of cardiogenic shock, an impella cp was successfully placed via the right femoral artery for mechanical circulatory support. following the initiation of impella flow, vasopressor support was successfully weaned. however, the patient did not regain consciousness or exhibit a positive neurological response after sedation was discontinued, indicating a poor prognosis likely due to the duration of cardiac arrests. the family designated the patient as do not resuscitate. the patient subsequently developed distal limb ischemia with an absent distal pulse. due to the poor overall prognosis and questionable neurological status, the medical team decided against escalating care to an impella 5.5 device. the patient continued to deteriorate despite increasing pressor support. the family ultimately decided to withdraw care and transition to comfort measures only. patient care was withdrawn, and the patient expired. | [5 empty] | Investigation summary: no product was returned for investigation. ischemia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1940955. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. as per qn (b)(4), for pumps sn: (b)(6), dhr documentation missing. the dhr (delivery record) for the lot was found at the sv (supply and distribution) department. after the dhr review dated (b)(6) 2025, it was returned to production to correct a missing material consumption posting for 0048-3106 au, cannula, pump 350. the dhr fa (delivery record) documented that 6 units from lot 1939189 and 4 units from lot 1942102 were used; these quantities are not reflected in the documented goods movements. this explains why no dhr report is available for the psi (process safety inspection) check, as it could not be generated due to another error. this confirms the proper functioning of the procedure according to sop-qa118 rev. b, chapter 6.2. consequently, the reported nonconformity cannot be confirmed, and the qn (quality notification) can be closed with acc (acceptable). since the cannula posting issue has already been processed for an excessive amount of time, it will be addressed and investigated in a new f3 qn. the lot will remain in the mrb area via (b)(4) until its final disposal. | A4 is unknown. although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3463 | 1220648-2025-48342 | 23546860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23546860&pc=OZD | 10/17/2025 | 11/13/2025 |  | 11/13/2025 |  |  | 10/17/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749884 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 102 KG | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Fever | Pulmonary Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed for acute myocardial infarction/ cardiogenic shock. during support it was reported that escalation was on hold due to fever and concerns for pneumonia. in addition to the patient remained intubated due to concerns of pulmonary edema prior to impella insertion however anticoagulation was on hold. ultimately the pump stopped with peaked motor current. decision was made to explant the cp and place intra-aortic balloon pump as bridge to 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b1, e4 the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. fever/pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. pump stop: the returned pump was inspected and an abnormally large purge gap was observed indicative of bearing wear. clinical reported cp device stopped at day 10 of usage bedside with peaked motor current. incomplete data logs were returned. the cause of the pump stop was determined to be due to a bearing wear as evidenced by returned product showing abnormally large purge gap device history review: the device passed all post sterile inspection checks.
3464 | 1220648-2025-48343 | 23546963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23546963&pc=OZD | 10/17/2025 | 01/06/2026 |  | 11/13/2025 |  |  | 10/17/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661007 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Renal Failure | Tachycardia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella 5, the patient was started on continuous renal replacement therapy (crrt) this morning, sedation vacation showed good neuro responses. the patient remains incredibly debilitated and weakness was able to stand one time however had ventricular tachycardia requiring patient to sit down. currently in atrial fibrillation ( afib) rvr rates in the 140s. note from critical care and cardiology discussion to be had with dr. (b)(6) regarding steps moving forward. continue dialysis. continue therapy. does not appear to have signs of kidney recovery. unsure of long-term destination plans. informed patient passed. | [5 empty] | H6 updated. investigation summary: no product returned. peri-procedural adverse event - the causes of these injuries was not determined due to insufficient clinical details. patient codes - renal failure: the cause of the renal failure was not determined due to insufficient clinical details. tachycardia, paroxysmal atrial fibrillation: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1919463. device history batch: subcomponent lot: n/a. device history review: qn (b)(4) was opened for batch 1919463 related to legibility of labels. mrb reviewed the impacted batch and confirmed that the labels are not legible or complete and must undergo rework. all reworks dispositioned as accept except the following s/n's: (b)(6). therefore, pump s/n (b)(6) passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3465 | 1220648-2025-48344 | 23547083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23547083&pc=OZD | 10/17/2025 | 11/12/2025 |  | 11/13/2025 |  |  | 10/17/2025 | 11/13/2025 |  |  |  |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026738993 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Fever | Hemolysis | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was on impella 5.5 as a bridge to transplant. during support it was noted that the patient was transferred to bed and had ventricular tachycardia(vt) with positioning alarms, point-of-care ultrasound done and impella was pulled back. vt and alarms stopped. in addition noted was hemolysis labs down trending-monitoring since repositioning. the patient also had intermittent fevers; transitioned from milrinone to dobutamine, now weaned off dobut off all vasoactive. the patient continued with fevers despite blood cultures being negative, the source of fever was unidentified. ultimately the 5.5 was weaned and explanted as bridge to transplant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and h6 codes were updated. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ventricular fibrillation/fever: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted.
3466 | 1220648-2025-48345 | 23547469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23547469&pc=OZD | 10/17/2025 |  |  | 11/13/2025 |  |  | 10/17/2025 | 11/13/2025 |  |  | 02/28/2026 |  | 11/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517683 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp flex was placed without complication after the patient had low flows on only the impella cp. considering starting inotropes for pulmonary artery pressures and reduced pulse volume recording. intermittently pacing with external pacemaker and receiving blood transfusions with hemodynamic improvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injury was not determined due to insufficient clinical details.
3467 | 1220648-2025-48277 | 23531169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531169&pc=OZD | 10/16/2025 |  |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755249 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Heart Block | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was pump suction, positioning issues, and a heart block during impella cp patient support. while on support, there were reports of suction alarms and the device¿s performance level was reduced to 4. the impella appeared to be shallowed based on a chest x-ray. the patient developed third degree heart block. a low dose of levophed was started. the patient was later successfully explanted and survived. | Additional information provided that the impella was repositioned in the operating room under transesophageal echocardiogram guidance during the coronary artery bypass graft procedure which resolved the suction alarms. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Additional information was received therefore b5 and b7 were updated to reflect that information. h6 impact code was updated to reflect the additional information that was reported. the investigation was completed and no product was returned. heart block: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: no product was returned. patient had suction alarms overnight and p level was reduced. after review of logs, multiple suction alarms were observed. no sudden motor current spike or placement signal trend was observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review was completed and passed all post sterile inspection checks.
3468 | 1220648-2025-48278 | 23531202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531202&pc=OZD | 10/16/2025 |  |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739059 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was ventricular fibrillation during impella cp patient support. while on support, additional shocks and cardiopulmonary resuscitation were needed for continued ventricular fibrillation after impella cp insertion. the patient remained on support and was later successfully explanted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3469 | 1220648-2025-48280 | 23531252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531252&pc=OZD | 10/16/2025 | 11/11/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755467 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 67 KG | Non Hispanic | White | Death | Yes | Yes | Difficult to Advance | [6 empty] | Arrhythmia | Hematoma | Low Blood Pressure/ Hypotension | Unspecified Infection | Cardiogenic Shock | [8 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was difficulty to advance, arrhythmia, hematoma, hypotension, infection, and cardiogenic shock during impella cp patient support. there was severe aortic stenosis that caused some issues crossing with the impella, but the doctor was able to torque the impella to get it to cross the aortic valve. while on support, there was concern for a small hematoma proximal to the impella. a femostop was placed. an echocardiogram was performed and showed atrial flutter. one milligram per minute of amiodarone was started. the patient experienced refractory vasoplegia and was pan cultured. the patient was started on broad spectrum antibiotics. the clinical team suspected mixed distributive cardiogenic shock empirically covered with vancomycin which was transitioned to linezolid and piperacillin-tazobactam. care was withdrawn and the patient expired. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. it was difficult to deliver since patient had severe stenosis along with difficult anatomy due to which the introducer bent and kinked most likely. the cause of the hematoma was unable to be determined due to insufficient clinical details. the cause of the injury was unable to be determined due to insufficient clinical details. arrhythmia, infection: the cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this is one of two related reports. this report represents the pump. another report was submitted to represent the introducer. refer to section h10 for the related report number.
3470 | 1220648-2025-48281 | 23531346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531346&pc=OZD | 10/16/2025 |  |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700963 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Ventricular Fibrillation | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was cardiac arrest, arrhythmia and ventricular fibrillation during impella cp patient support. while on support, the patient required intubation for coding and cardioversion for ventricular tachycardia. the patient achieved return of spontaneous circulation. later, the patient went into ventricular fibrillation. the patient¿s family made him ¿do not resuscitate¿. the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia/ ventricular fibrillation: the cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3471 | 1220648-2025-48282 | 23531485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531485&pc=OZD | 10/16/2025 | 12/31/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727804 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient was sent to the intensive care unit on bipella (impella cp and impella rp flex) support. the patient¿s groin was oozing from the left impella cp femoral arterial site, the dressing was taken down, forward tension sutures were re-done, dressing was re-done following best practices per bedside registered nurse (rn). the patient received two units of packed red blood cells (prbcs). no oozing was noted. repositioning the impella cp using echocardiogram guidance as well as due to placement signal low alarms was possible. the patient no longer was on heparin drops due to site oozing overnight. dressings were re-done by the bedside rn. impella cp was successfully weaned and explanted. the patient was stable on current support. the plan was to wean and explant the impella rp flex. | [5 empty] | Investigation summary: major bleed: the cause of the injury is most likely use related since forward tension sutures re-done, dressing re-done following best practices and no oozing noted since per bedside rn. device history lot: device lot: 1960366 device history batch: sub-component lot: n/a device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
3472 | 1220648-2025-48283 | 23531538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531538&pc=OZD | 10/14/2025 |  |  | 11/12/2025 |  |  | 10/14/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730529 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp had decoupled waveforms and lower than expected flows. transvenous pacer wires were placed, but the patient continued to have decreased flows and left ventricle tracing. the impella rp flex was placed without complication. inotropes were considered to be started for pulmonary artery pressures and reduced pulse volume recording. a left ventricle signal message to adjust was noted and the customer support center was contacted for assistance. assistance was provided and all questions were answered. the impella cp position was deep and was repositioned. some suction occurred when coughing that self-resolved. intermittently pacing with external pacemaker and receiving blood transfusions with hemodynamic improvement were done. the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation was completed and no product was returned. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. low pump flow: clinical reported decoupling waveforms and low flows at case start. it was also stated in clinical details that the patient had biventricular failure, leading to the placement of an rp flex for bipella support. once the rp flex was placed, flows reportedly improved significantly. the product was not returned for investigation and data logs were not provided. without sufficient product or data, it cannot be confirmed if low flow occurred or make conclusions about what cause it. the root cause of the low pump flow could not be determined due to lack of data or product returned. device history review was completed and passed all post sterile inspection criteria.
3473 | 1220648-2025-48284 | 23531637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531637&pc=OZD | 10/16/2025 | 11/11/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734888 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Asystole | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The use facility reported a patient was placed onto impella support for coronary artery bypass graft. while on support, the patient went into ventricular tachycardia rhythm and became asystole with flat motor current requiring cardiopulmonary resuscitation. the pump was repositioned in the catheterization lab. this resolved the premature ventricular contractions/ventricular tachycardia episodes. a temporary pacemaker was put in at the rate of 80. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia/asystole/arrhythmia : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3474 | 1220648-2025-48285 | 23531662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531662&pc=OZD | 10/16/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026718385 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 81 KG | Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received stating that the patient was explanted. | The user facility reported that multiple clots were observed near the catheter of the impella 5.5 at the time of heart transplant. a brachial thrombectomy was required. |  |  |  |  | B5 additional information was received and d6 updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e1, e4. d9 updated. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. thrombosis: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3475 | 1220648-2025-48286 | 23531775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531775&pc=OZD | 10/16/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667198 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was placed onto impella support for high-risk percutaneous coronary intervention. while on support, the medical doctor (md) accidentally advanced the impella forward which made the waveform look ventricular and gave the alarm of in the ventricle, but then converted to low placement signal. the md pulled the impella back while under fluoroscopy and the waveforms returned to normal. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3476 | 1220648-2025-48287 | 23531884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531884&pc=OZD | 10/16/2025 | 11/11/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026678488 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was placed onto impella support for coronary artery bypass graft. while on support, the patient required blood products for an internal bleeding. the chest was re-opened and the source of bleeding could not be identified. the patient went into ventricular tachycardia while coughing. the pump appeared to be at 4.0 centimeters from the aortic valve, and potentially interfering with the papillary muscle. no resolution was specified for pump position. additional information was received. the pump is not available for return. the bleeding was not related to the impella. the care team had no bleeding concerns from a device standpoint. in addition to the device being in the papillary muscle, the care team did not believe there were intracardiac contact from a device standpoint other than when the patient dramatically coughed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleed was not determined due to insufficient clinical details. tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3477 | 1220648-2025-48289 | 23531960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23531960&pc=OZD | 10/16/2025 | 11/19/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697566 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Insufficient Information | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient was placed onto impella cp support due to acute myocardial infarction complicated by cardiogenic shock. while on support, a suction alarm was noted and low flow and minimal pulsitility on motor current. echocardiogram at bedside was done and position was confirmed by the medical doctor and they were still unable to resolve the alarms. the team ultimately decided to withdraw care. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Correction h6 health effect - clinical code 0602 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. suction/low flow: data log review confirmed the rapid decline/trend in the placement signal (ps) while motor current (mc) and pump flows remained stable at p-9. the pump was turned down to a lower p-level, and during this time, pump flows remained in the expected range for each respective p-level until a sudden mc spike with ms deviation. returned product was investigated and there were no defects of the product noted (impeller damage, cannula kink). there was, however, biomaterial in the inflow cage of the pump. the pump was run in a bucket of water with the biomaterial still present, and low pump flows reproduced. the biomaterial was removed and pump flows were back within the expected range. based on the clinical details provided, signature of biomaterial ingestion followed by low pump flows on data logs, and biomaterial on returned product reproducing low flows, the cause of the suction/low pump flows was determined to be biomaterial ingestion. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3478 | 1220648-2025-48290 | 23532017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23532017&pc=OZD | 10/16/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026756743 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient was placed onto impella support for high-risk percutaneous coronary intervention. while on support, pump presented placement signal not reliable alarm. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: data log review confirms placement signal losing waveforms and triggering placement signal not reliable (psnr). at the time, snr drops to zero, contrast steps down, contrast and gain decrease, and lamp increases slightly. based on clinical details of the event, the optical sensor likely became damaged by pulses from the shockwave. the distance between the impella and shockwave is unknown. the root cause of the optical sensor issue was most likely due to a broken sensor face that resulted from the shockwave device mentioned in clinical details. device history review: the device passed all post sterile inspection checks.
3479 | 1220648-2025-31923 | 23532367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23532367&pc=OZD | 09/07/2023 | 11/12/2025 |  | 11/12/2025 |  |  | 08/20/2023 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1587053 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported there was placement signal issues with the impella 5.5. upon investigation, the engineer determined that the automated impella controller (aic) optical bench might not be operating within specification. care was withdrawn, and the patient eventually expired. the death is captured in the sister records of this record. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: from log analysis rlm reset itself, no issues found in usb, no indication from partition switching, restarting aic+rlm, with ic-test again, now there¿s usb communication, and data streaming is working. device analysis: from data logs lot of contamination on optical fiber (maybe even some damage, but can¿t confirm until cleaned) causing degradation of optical signal (and likely air gap) maybe contaminates inside the receptacle. root cause: signal issues were caused by degradation of optical placement data due to damage and contamination of optical pump connector, and noisy signal from optical bench. all pertinent information available to abiomed has been submitted at this time.
3480 | 1220648-2025-31924 | 23532413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23532413&pc=OZD | 02/23/2023 | 11/12/2025 |  | 11/12/2025 |  |  | 03/08/2023 | 11/12/2025 |  |  | 06/30/2024 |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023233372 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) failed to display a placement signal and began to alarm with a placement signal not reliable alarm. the aic was swapped, however the placement signal still failed to repopulate. patient support continued with no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump not returned for review. field lt logs were reviewed and psnr was confirmed on 2/23. snr increased and lamp increased at the time of the psnr but the other 5s parameters remained stable. see 23-13671-1 for aic investigation which includes imc log review. console logs from the reported day of event are consistent with complaint and show that after pump (b)(6) was connected to console ic2343, the pump eeprom usage data area 1 and 2 could not be correctly read (due to crc errors). this resulted in the console¿s software not recognizing the pump as previously used, which caused missing prompt to resume pump at previous speed (with the same p-level as the last p-level on prior console, ic2345). because entire usage data was not read correctly, the "erroneous" value of 0 was used for calibrated g0, which resulted in placement signal out of range, due to incorrect scaling. the pump was unplugged and re-connected in an attempt to resolve the issue, but prior to that, the usage data area 1 and 2 sections were overwritten by aic software with erroneous value of g0 but correct (matching) values of crc, so that after the pump was reconnected, there were no issues with reading its eeeprom. nevertheless, the placement signal was still out of range (as indicated by alarm #1048 ¿placement signal not reliable¿ that followed), due to pump being recognized as ¿previously initialized¿ and having calibrated value of g0=0. the case continued, with uninterrupted hemodynamic support, although without placement signal, until its completion on (b)(6) 2023. log analysis indicated that during eeprom reading issues, consistently repeating mismatch between read and calculated values of crc indicates that incorrect values were not the result of communication issues during eeprom reading but were recorded as such in the eeprom. therefore, this issue might be related to the prior console, ic2345, on which the pump¿s eeprom had been written to. analysis of logs from ic2345 showed pump (b)(6) connected for the first time on (b)(6) 2023, and its eeprom read correctly. in the process of optical sensor calibration, a new value of g0 was determined, and placement signal was within range. before starting, however, the pump had been disconnected (presumably to uncoil the catheter) and reconnected again, where its eeprom was read correctly, and the correct value of calibrated g0 was used. placement signal was present, and support continued without issues until on (b)(6) 2023, when decision was made to transfer the pump to ic2343, to facilitate return of ic2345 in relation to complaint (b)(4) (inaudible alarms). at this point, the pump had been running for 12 days 16 hours and 55 minutes (18295 minutes) and was at p-4 when unplugged. additional data, obtained from the cloud, contained in several consecutive copies of pump¿s eeprom (as read/interpreted by aic, sent to cloud after each pump connection) was analyzed and confirmed that either ic2345 recorded already corrupt data in the pump¿s eeprom, or the eeprom got corrupted during (or after) writing. the exact cause of the issue (and whether it is due to pump or console), although known to be related to the mechanism of writing to pump¿s eeprom (inside the red handle), could not be determined because the issue could not be reproduced, and pump (b)(6) was not returned for analysis. see (b)(4) for aic investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3481 | 1220648-2025-48294 | 23532959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23532959&pc=OZD | 10/13/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/13/2025 | 11/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647287 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with occluded proximal right coronary artery. drug-eluting stent was placed. during post-dil inflations blood pressure declined and impella cp was placed. it was noted that there was access complication of bleeding and a hematoma at insertin site following knee bending. one unit of packed red blood cell were given in addition to placing femstop. the patient continued on support for a total of 81.17 hours and was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The returned product was investigated. dried blood was seen on the suture hub, sterile sleeve, and inflow and outflow cages. no other abnormalities were seen. access site adverse event (major bleed, hematoma): clinical details mention that a large hematoma was seen at the insertion site after the patient bent their knee. femstop was used to manage the complication, and 1 unit prbc was given to counter the drop in hemoglobin. the root cause of the access site adverse event was most likely use related owing to patient movement as the patient bent their knee after being agitated in the icu based on the provided clinical details.
3482 | 1220648-2025-48295 | 23533070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533070&pc=OZD | 09/12/2025 | 11/03/2025 |  | 11/12/2025 |  |  | 09/12/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725150 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was escalated from an impella cp to an impella 5.5 with smartassist. the pink-red urine cleared with impella 5.5 escalation. an episode of sustained ventricular tachycardia occurred requiring reintubation and pressors were restarted. extracorporeal membrane oxygenation had to be added. the patient ultimately expressed that he did not want to go through with the transplant anymore and chose to have care withdrawn. the patient ultimately expired. this is one of two related reports. this report represents the impella 5.5 another report was submitted to represent the impella cp. refer to section h10 for the related report number. | [5 empty] | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review: medical safety/clinical reviewed the event: a 59-year-old male with a medical history of coronary artery disease, prior coronary artery bypass grafting, mitral regurgitation, hyperlipidemia, and hypertension was brought to the cardiac catheterization laboratory for mechanical circulatory support. the patient was initially supported with an impella cp. during impella cp support, red-colored urine (hematuria) was observed, and the patient required vasopressor support. the patient was considered a potential candidate for advanced heart failure therapies. a decision was made to escalate support to an impella 5.5 to allow continued evaluation. following escalation to the impella 5.5, the hematuria resolved. two days after initiation of impella 5.5 support, the patient experienced an episode of sustained ventricular tachycardia, requiring reintubation and reinitiation of vasopressor therapy. due to ongoing hemodynamic instability, extracorporeal membrane oxygenation (ecmo) was initiated. subsequently, the patient expressed a desire not to pursue heart transplantation or further advanced therapies and elected to withdraw care. the patient subsequently expired. this event story include two product complaints. this event story include two product complaints. impella cp - mdr 1220648-2025-47734: serious injury. impella 5.5 - mdr 1220648-2025-48295: death. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3483 | 1220648-2025-48296 | 23533184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533184&pc=OZD | 09/15/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 09/15/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686564 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient had an impella 5.5 placed for support when signal was lost. the loss of the signal prompted the team to replace the pump. there was no harm from the malfunction. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3484 | 1220648-2025-48297 | 23533554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533554&pc=PYX | 09/14/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 09/15/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529489 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 99 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Inflammation | Renal Failure | Multiple Organ Dysfunction Syndrome | [6 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was placed on an impella pr flex that had intermittent pump suction due to fluid volume overload. the patient is on a low dose of heparin and blood platelets were transfused. right hemothorax bleeding was noted that required three units of blood products. bedside thoracostomy was performed with chest tube insertion dramatically improved suction events. bipella support was maintained and suction coinciding with rhythms. pseudomonas in sputum were noted. high bilirubin and organ failure. care was withdrawn and the patient expired. there are two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. | [5 empty] | The impella device was not received from the customer. the event logs were returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Medical safety/clinical reviewed the event. a 70-year-old female with a medical history of persistent atrial flutter, paroxysmal atrial fibrillation, mixed hyperlipidemia, hypertension, and morbid obesity presented to the emergency department with several weeks of fatigue, dizziness, palpitations, and associated shortness of breath. the patient was newly diagnosed with heart failure with reduced ejection fraction (hfref) with a severely reduced ejection fraction of 10¿15% and severe global hypokinesis. the patient underwent left heart catheterization for ischemic evaluation, which demonstrated two-vessel coronary artery disease involving the mid left anterior descending artery and right coronary artery. right heart catheterization demonstrated worsening lactic acidosis, with lactate levels trending up to 12.9 mmol/l. the advanced heart failure team was consulted, and the patient was initiated on left-sided impella support, which was subsequently escalated to biventricular (bipella) support with impella devices: impella 5.5 and rp flex. during impella cp support, the patient developed acute renal failure requiring continuous renal replacement therapy and was treated with multiple medications. paroxysmal junctional rhythm was also reported. the patient was escalated to an impella 5.5, and bipella support was initiated with the addition of an impella rp flex. while on bipella support, the patient experienced multiple adverse clinical events. additionally, the automated impella controller (aic) connected to the impella rp flex displayed a frozen screen with a persistent ¿complete the procedure¿ alarm. replacement of the purge cassette did not resolve the issue. the aic was exchanged with no noted consequence to the patient, and the patient continued to receive support without further interruption. after approximately eight days of mechanical circulatory support, care was withdrawn, and the patient expired. h6. investigation type/findings/conclusion remain unchanged. related medical device reports: this is one of four devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | B3 revised date of event as it was entered incorrectly on the initial report that was submitted. g3 date of awareness should of reflected 15-sep-2025 om the initial report submitted. h10 this is the third pump used on the patient. the related 2 pumps report numbers were attached. investigation summary: arrythmia/infection: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension/renal failure/major bleed/organ failure/inflammation: the cause of the injury was not determined due to insufficient clinical details. data logs returned were missing about 20 minutes of support during automated impella controller swap. impella rp flex support was from (b)(6) 2025. logs were returned from (b)(6) 2025 at 11:47am till (b)(6) 2025 at 2:56pm and from (b)(6) 2025 at 3:21pm till (b)(6) 2025 at 7:58am. low or blocked pump flow: impella rp flex flows were reported low on (b)(6) 2025 with flows only at 1.4 liters per minute on p6. the medical team increased the activated clotting time goal to over 200. flows were improved. the data logs showed a sudden shift up in placement signal and a sudden drop in pump flows and purge flows with no p-level changes. there was a minor motor current shift with an effect on purge pressure and purge flow characteristic of biomaterial ingestion. after the shift, placement signal and pump flows trended inversely (from 80mmhg to 110mmhg and from approximately 2.0 liters per minute to 1.3 liters per minute respectively) for about eight minutes before flows returned to expected range for matching p-level. the cause of the issue was determined to be due to biomaterial ingestion as evidenced by log pattern and clinical information of low flows improving after increasing activated clotting time values. pump suction: clinical details reported few suction events on (b)(6) 2025 with suction coinciding with junctional rhythm change. position reported to be appropriate. no suction events reported after occurrence. lower blood pressure reported when in junctional rhythm. complications were managed with lidocaine and levophed increase when not in normal sinus rhythm. logs confirm suction and show no motor current spikes prior to suction outside p-level changes. no changes to motor current, pump flows or placement signal. the cause of the issue was determined to be patient condition related as evidenced by clinical details indicating when suction alarms occurred and resolved with no log abnormalities observed. device history review: the pump passed all post sterile inspection checks
3485 | 1220648-2025-48006 | 23533584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533584&pc=OZD | 10/03/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/03/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691900 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Death | Yes | No | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp exhibited neurologic shaking. a head computed tomography showed cerebral edema. the pump generated placement signals low alarms. an echocardiogram showed the pump in a deep position. bivalirudin was withheld in response to bleeding from the access site. pressure was held. the patient was looking much worse and was septic with neurological deficits. care was withdrawn and the patient expired. | [5 empty] | The cause of the major bleed was not determined due to insufficient clinical details. the root cause of the sepsis and stroke was unable to be determined due to insufficient clinical details. the cause of the optical sensor issue was not established as limited clinical information was provided and no product was returned for analysis. the cause of the positioning issue was not established as limited clinical information was provided. the cause of the access site bleeding was not determined as limited clinical information was provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3486 | 1220648-2025-48298 | 23533597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533597&pc=OZD | 09/15/2025 |  |  | 11/12/2025 |  |  | 09/15/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026717815 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 76 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Ventricular Fibrillation | Hematuria | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hematuria. it was suspected that blood was due to a traumatic foley insertion. an echocardiogram was performed to confirm the pump's position. the patient also experienced ventricular fibrillation requiring cardioversion, and two units of packed red blood cells were transfused in response to low hemoglobin. later, care was withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of the anemia and hematuria was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3487 | 1220648-2025-48299 | 23533643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533643&pc=OZD | 09/17/2025 |  |  | 11/12/2025 |  |  | 09/17/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702599 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male | 100 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced retroperitoneal bleeding from the right groin. three units of red blood cells were transfused and the bleeding was resolved surgically. the pump was explanted. | [5 empty] | No product was returned for investigation. the cause of the bleeding is determined to be patient condition because the reported bleeding is at non-impella site. the device passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3488 | 1220648-2025-48302 | 23533720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533720&pc=OZD | 10/14/2025 |  |  | 11/12/2025 |  |  | 10/14/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697508 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 107 KG |  | White | Injury | Yes | Yes | Restricted Flow rate | Use of Device Problem | [5 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for the patient admitted in with multiple cardiac comorbidities. it was reported that the repo sheath had not been advanced or covered. educated team on concern of sterility and impella securement. plan to place stat lock and redress. attempted to increase p level but continues to run into suction alarms. no pulses noted in bilateral lower extremities, although adequate color and temperature, no plan for vascular consult at this time. the patient did receive 4 packed red blood cells and 3 albumin overnight per extracorporeal membrane oxygenation protocol. labs ok this morning. no signs of active bleeding. the patient sent for ct overnight, which did reveal a stroke, although unsure if new or old. patient is off sedation continuing to be tachypneic and no purposeful movement. pt also had left sided chest tube placed with around 500 cc. the patient had difficulties yesterday afternoon with atrial fibrillation with rapid ventricular response and was placed on multiple drips which are all off except for a low dose of vaso. patient successfully weaned from impella and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the pump suction: the pump was not returned for investigation. data log was not retrieved from the remote server or provided for review. the cause of the pump suction issue was not determined without returned product investigation and sufficient clinical information. arrhythmia, stroke/cva): the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3489 | 1220648-2025-48300 | 23533808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533808&pc=OZD | 10/01/2025 |  |  | 11/12/2025 |  |  | 10/01/2025 | 11/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610457 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | False Alarm | Adverse Event Without Identified Device or Use Problem | [5 empty] | Gastrointestinal Hemorrhage | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient was placed on cardiac bypass for triple coronary artery bypass grafting with atriclip and following the procedure the patient was implanted with an impella 5.5 via direct aortic insertion for weaning off bypass. on day 2 of impella 5.5 support, the patient had an increased amount of drainage from the chest drain and required two units of red blood cells. on day 4 of support, the patient experienced a possible gastrointestinal bleed. the patient was transfused with one unit of red blood cells and the patient¿s bivalirudin was withheld. the same day, the medical team had concerns for an impella in aorta alarm. an echocardiogram was performed and clearly showed the impella was across the valve. the medical team was instructed to disable the alarm. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary. optical sensor issue/false alarm: clinical details report position in aorta alarms on 5/oct. position was confirmed to be good via echo. data logs were reviewed and the alarms were confirmed on 5/oct. around midday; while running at a constant p-2, a concentration of the alarms began to trigger. they were all short durations (<1 min) suggesting that the pump was on the borderline of the condition to trigger the alarm. all optical signal metrics were stable and within expected range. an rt log of one of the periods of alarms confirmed that the mc and ps waveforms were in proper phase. however, it was noted that there was a slight drop in mcmax values whenever the alarm would trigger. the mc values were within the expected range for p-2, but the delta (mcmax-mcmin) would drop from ~60 ma to ~40 ma. when mc modulation would increase, it was noted that there were 500 rpm shifts up in the ms. the pump had been at p-2 for almost 24 hours before the alarms began to trigger. an early p-2 rt log was reviewed, and the same alternation between larger/smaller mc modulation was noted, however, it appeared that the frequency of the larger amplitude mc was more consistent, therefore, reducing the number of times the modulation dropped below the threshold, so the alarm didn't trigger. this supports the idea that the pump was at the borderline condition of triggering the alarm. there is potential that the alarms were false triggers based on the fact that clinical details reported that the pump was correctly positioned and that the alarm triggering criteria can be more sensitive at a low p-level (p-2). however, limited clinical information provided that would explain why the shifts in mc/ms were occurring. product was not returned for investigation. since data log review was inconclusive based on the clinical details provided and product wasn't returned for investigation, the cause of the false positioning alarms could not be determined. major bleed: clinical details report that 1 unit of blood was given overnight due to a possible gi bleed on 5/oct. the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1885302. device history batch: subcomponent: na. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
3490 | 1220648-2025-48301 | 23533809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23533809&pc=OZD | 09/17/2025 |  |  | 11/12/2025 |  |  | 09/17/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722316 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Fever | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5 the patient went into sustained unstable ventricular tachycardia. the patient was placed on peripheral veno-arterial extracorporeal membrane oxygenation and an implantable cardioverter-defibrillator was also placed. heparin was withheld as a gastrointestinal bleed was suspected and the patient was positive for heparin induced thrombocytopenia. there was concern for shock as the patient had a fever and hemolysis with unknown causes, and the patient was transfused blood products. later, the family opted to withdraw care and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The product was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the tachycardia and fever was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the pump passed all post sterile inspection checks.
3491 | 1220648-2025-48303 | 23534040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534040&pc=OZD | 09/18/2025 |  |  | 11/12/2025 |  |  | 09/18/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026727083 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [28 empty] | Initial submission | [9 empty] | The user facility reported that there was arrhythmia during impella 5.5 patient support. while on support, the patient experienced ectopy. the team discussed starting amiodarone. | [5 empty] | The impella device is not available for return. therefore, investigation of the device is not possible. should any new information be received, a supplemental report will be filed.
3492 | 1220648-2025-48304 | 23534215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534215&pc=OZD | 09/18/2025 | 11/06/2025 |  | 11/12/2025 |  |  | 09/18/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679124 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 67 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Asystole | Heart Block | [10 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | The user facility reported that there was asystole, heart block, arrhythmia, and positioning issues during impella 5.5 patient support. while on support, the patient went into asystole and then heart block. amiodarone was discontinued and atropine was started. the patient experienced a ventricular tachycardia episode and a shock was administered. the patient was converted back to normal sinus rhythm. an echocardiogram was completed and found the impella was too deep. it was pulled back from 42 centimeters to 41 centimeters. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned. asystole, heart block: the root cause of the injury was unable to be determined due to insufficient clinical details. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3493 | 1220648-2025-48307 | 23534352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534352&pc=OZD | 10/15/2025 |  |  | 11/12/2025 |  |  | 10/15/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025486982 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella 5.5 on the patient impella is currently running at p5 and still experiencing frequent suction alarms. team is positive, it is related to right ventricular failure (rvf). but the patient deemed not extracorporeal membrane oxygenation candidates per cts. plan to treat rvf with dobutamine and veletri and optimize impella flows. later the patient was still experiencing intermittent and frequent suction alarms. however, left ventricle is severely underfilled and right ventricle is akinetic. urine output is decreased and pink. creatinine slightly elevated today. ct output is minimal. team decided given severity of rvf revealed on tee plan to send patient to ccl for rp flex. tee bedside this morning revealed ef of 45% with mild to moderate rv dysfunction. creatinine improved today with crrt. the patient has tolerated uf of 100 cc/hr. the patient down 1.5 l in last 24 hrs. inr 5 this morning which was treated with vitamin k. lfts ok. the patient received 2 packed red blood cells overnight. explant performed independently with local support enroute and available for phone support. impella 5.5 explanted without complication per the doctor. successful first bipella support utilizing impella 5.5 and impella rp. the patient was transferred back to icu with plans to extubate and continue crrt. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d3 the investigation of the reported failure mode has been completed. no product was received for evaluation. data log shows several suction alarms during the first half of the case. the pump was unable to achieve the desired flow during the suction events. there were no reported optical signal issues, motor current spikes, or positioning issues on the data log. hemolysis: the cause of the hemolysis was most likely related to patient condition based on clinical details and data log review. pump suction: based on data log review and provided clinical details, the cause of the pump suction issue was most likely related to patient volume status and right ventricular dysfunction, as indicated by the clinical findings and response to interventions. clinical symptoms (renal failure): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3494 | 1220648-2025-47999 | 23534439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534439&pc=OZD | 09/27/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 09/27/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647249 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had undergone an intervention to their left anterior descending coronary artery on (b)(6) 2025 and decompensated in the unit. a left heart catheterization and right heart catheterization were scheduled for (b)(6) 2025 and an impella cp was implanted. after the impella cp was implanted, it was noted that the patient¿s leg did not have deplorable pulses. the decision was made to escalate the patient to an impella 5.5. on day 6 of impella 5.5 support, the patient was taken to the operating room for a mitral valve repair and coronary artery bypass graft. the patient developed postoperative right ventricular failure, and the decision was made to implant and impella rp flex. on day 3 of bipella support with the impella rp flex and impella 5.5, the patient spiked a fever. the medical team decided to remove the swan and central line. post removal of the swan, the placement signal numbers spiked on the impella rp flex, and the flows dropped. the decision was made to remove the impella rp flex and escalate the patient to veno-arterial extracorporeal membrane oxygenation (va ecmo). the patient remained on impella 5.5 and va ecmo support. the following day, the patient¿s family made the decision to withdraw care of the patient. the patient¿s care was withdrawn, and the patient expired. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6) pump 3: impella rp flex with serial number (b)(6). this mdr specifically addresses pump 1: impella cp, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event and noted the following: the patient presented with cardiogenic shock (cgs) following intervention to the left anterior descending (lad) artery. an impella cp was implanted for hemodynamic support. post-implant, distal limb ischemia was observed at the impella cp access site, prompting escalation to impella 5.5. six days after impella 5.5 implantation, the patient underwent coronary artery bypass grafting (cabg) and mitral valve replacement (mvr). postoperatively, the patient developed right ventricular (rv) failure, and an rp flex device was implanted. subsequently, the patient developed a fever, and the clinical team elected to remove the swan-ganz catheter and central line. upon removal, the impella rp flex placement signal spiked and flows dropped. it is unknown whether the recommended practice of reducing impella rp flex flows prior to line removal was followed. the impella rp flex was removed, and veno-arterial extracorporeal membrane oxygenation (va ecmo) was initiated. the patient remained on ecpella support until the following day, when the family decided to withdraw care. the patient expired. assessment outcome is the following the limb ischemia on impella cp considered device-related due to access site involvement. fever on impella 5.5 and rp flex likely secondary to swan/central line infection but conservatively attributed to devices. the impella rp flex device malfunction reported due to abnormal placement signal and flow drop; root cause undetermined (possible procedural deviation by not reducing flows upon removal of lines). and the death will be reported for impella 5.5 as a conservative measure; clinical deterioration most likely contributing factor. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex and the impella 5.5, which is now associated with the demise outcome.
3495 | 1220648-2025-48308 | 23534467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534467&pc=OZD | 09/26/2025 |  |  | 11/12/2025 |  |  | 09/26/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714723 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 141 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 support experienced atrial fibrillation and ventricular tachycardia. it was confirmed that the pump was in the proper position. the patient was cardioverted. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The root cause of the arrhythmia and need for resuscitation was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3496 | 1220648-2025-48306 | 23534590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534590&pc=OZD | 10/16/2025 | 12/31/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725302 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 55 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient developed a fever overnight. the patient was given two units of packed red blood cells due to drop in hemoglobin. also, the patient had several bloody bowel movements. an abdominal computed tomography did not show a gastrointestinal (gi) bleed. the gi bleeding has been resolved. the patient was brought to the cardiovascular operating room for impella removal. the doctor exposed the axillary pocket and removed the impella without issue with surgical closure. the patient was transferred back in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. major bleed: pt had several bloody bowel movements overnight. abdominal ct did not show gi bleed. fever: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1960408 device history batch: subcomponent lot: n/a device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3497 | 1220648-2025-48313 | 23534653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534653&pc=OZD | 10/14/2025 |  |  | 11/12/2025 |  |  | 10/14/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741352 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that post left main stent placement, the patient decompensated, despite increasing pharmaceutical support and maximum impella support. the patient arrested and required cardiopulmonary resuscitation (cpr). return of spontaneous circulation was obtained and the patient was taken to the intensive care unit (icu). despite all efforts, the patient repeatedly arrested, requiring cpr two more times in the icu before the family decided to withdraw care. impella support was discontinued at bedside by the physician. with care withdrawn, the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3498 | 1220648-2025-48314 | 23534691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534691&pc=OZD | 10/12/2025 |  |  | 11/12/2025 |  |  | 10/13/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717814 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was escalated to an impella 5.5 from an impella cp device for increased mechanical circulatory support. the patient was initially transferred from an outside hospital in cardiogenic shock. the patient was placed on veno-arterial extracorporeal membrane oxygenation (va-ecmo) support. following successful decannulation from va-ecmo earlier the same week, the patient experienced runs of arrhythmia and weakening pressures, including episodes of complete heart block that required the placement of a temporary permanent pacemaker. due to persistent instability and an ongoing requirement for mechanical circulatory support, an impella cp was implanted via the right femoral artery. during impella cp support, the pump alarmed for suction and as a result of the alarms the patient received intravenous fluids. the medical team evaluated the patient¿s right iliac artery for a potential occlusion. computed tomography (ct) scan revealed that the patient had a right iliac occlusion, and to resolve the ischemia the patient had a femoral-to-femoral bypass. it was noted that the patient¿s lactate values had increased due to the prior decompensation event. the patient remained unstable on the impella cp and the care team subsequently escalated to an impella 5.5 to provide an increased level of support. the patient was brought to the cardiac catheterization laboratory, the impella 5.5 was implanted and the impella cp was explanted with no complications. on day four of impella 5.5 support, it was noted that the patient underwent a thrombectomy over the weekend for a stroke, which yielded negative results. anticoagulation therapy was withheld. the following day, the patient remained off of heparin due to heparin induced thrombocytopenia. ct scans were conducted with no changes in image findings. on day seven of impella 5.5 support, anticoagulation therapy was restarted, and the patient was given bivalirudin due to the patient¿s history of heparin induced thrombocytopenia. the patient remained on support with the care plan to rest and recover and bridge to a durable left ventricular assist device. on day 13 of impella 5.5 support, the impella experienced suction alarms and the patient was administered one liter of isolyte to manage volume status. the patient began working with physical therapy and occupational therapy in a chair. there was no change in the impella care plan, and no issues or concerns related to the device were noted by staff. at the time of writing this report, the patient continues to be supported by impella 5.5 while awaiting surgery for a durable left ventricular assist device. this complaint involves two automated impella devices: pump 1: impella cp, with serial number (b)(6) pump 2: impella 5.5, with serial number (b)(6) this report specifically addresses pump 2: impella 5.5, with serial number (b)(6). a separate report will be submitted for the other device associated with this complaint. | The device was explanted successfully after weaning during surgical implant of a durable left assist device. the explant was not considered premature and was unrelated to the reported complication. patient outcome is unknown as outcome information was not available in the complaint record accessible for mdr assessment at time of reporting. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5: additional information added for explantation, patient outcome is not available. d6b: explant date has been added. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: clinical symptoms (stroke/cva): the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. data log was not provided for the first 6.5 hours of the case run on 10-oct, and the suction issue was reported on 22-oct. the cause of the suction issue was not determined without returned product investigation and sufficient clinical details, including data logs. device history lot: device batch: 1955097. device history review: the pump (b)(6) passed all post sterile inspection checks.
3499 | 1220648-2025-48309 | 23534766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534766&pc=OZD | 09/26/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 09/26/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650457 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a subdural hematoma and hypotension. the patient was treated with levophed and systemic anticoagulation was withheld. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the hematoma and bleeding was not determined since product was not returned and insufficient clinical details provided. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3500 | 1220648-2025-48311 | 23534876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23534876&pc=OZD | 10/13/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/13/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648048 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the emergency department on (b)(6) 2025 with shortness of breath, atrial flutter, and a congestive heart failure exacerbation. the patient initially improved in the cardiac intensive care unit (cicu) but decompensated over the weekend, experiencing a ventricular tachycardia (vt) episode for which his automatic implantable cardioverter-defibrillator delivered a shock. the patient required intubation and vasopressor support (phenylephrine). a decision was made to place a heart pump to stabilize the patient for transfer to an outside hospital for advanced therapy. the impella cp heart pump was successfully placed via the right femoral artery. the patient was successfully transferred to the hospital and arrived in stable condition. however, the following day on support day two, the patient¿s plasma free hemoglobin, lactate dehydrogenase, and lactate levels all increased. the impella was reduced to p4, and vasopressors were held. a transthoracic echocardiogram (tte) showed good pump position but raised suspicion for a mobile mass on the pigtail catheter and potential pneumonia. the following day, the patient experienced multiple atrial fibrillation events requiring cardioversion and increased medication requirements. hemolysis was noted to have been resolved. the patient developed right lower lobe pneumonia on computed tomography scan. on day four of impella cp support, the patient continued to decline, and lactic acid results trended up. the clinical team determined the patient was a "bridge to nowhere" due to poor prognosis but later reversed the decision to escalate care. the medical team decided to escalate the patient to an impella 5.5. the impella 5.5 was implanted and the impella cp was explanted with no complications reported. on day two of impella 5.5 support, the patient remained critically stable, however the medical team was hopeful for recovery. on day six of impella 5.5 support, it was noted that the patient was making an incredible recovery with an improved ejection fraction of 30% on bedside echocardiogram, successful extubation, and intact neurological function. the following several days the patient continued to recover and the patient began physical therapy and was stable on impella p6 and dobutamine. the team determined the patient was not a candidate for immediate advanced therapies due to extensive ischemic coronary artery disease not amenable to intervention, suggesting destination therapies might be required if improvement continued. on day 11 of impella 5.5 support, the patient's condition worsened significantly after developing sepsis secondary to an ileus. the patient expired while on support. | [5 empty] | Medical safety reviewed the event and noted the following: a patient with exacerbated congestive heart failure, atrial fibrillation, and ventricular tachycardia underwent successful implantation of an impella cp device for hemodynamic support. the patient was subsequently transferred to a hub hospital. approximately 48 hours post-implant, laboratory findings revealed elevated ldh, plasma-free hemoglobin, and lactate levels, raising suspicion of hemolysis. a transthoracic echocardiogram confirmed appropriate pump positioning but noted a possible mobile mass on the pigtail and potential pneumonia. in response, device support was reduced to p-4, after which hemolysis resolved. during this period, the patient experienced multiple atrial fibrillation episodes requiring pharmacologic intervention and cardioversion. on day 4, due to continued clinical decline, support was escalated to an impella 5.5 device. the patient initially demonstrated improvement; however, on day 11, the patient developed sepsis secondary to ileus, resulting in significant clinical deterioration and subsequent death. the occurrence of hemolysis during impella cp support is considered a serious injury, despite resolution following standard troubleshooting measures. the patient¿s death is conservatively attributed to the impella 5.5 device, as the patient was on support at the time of death. however, sepsis secondary to ileus is assessed as the most likely contributing factor. device involvement: impella cp associated with hemolysis (resolved). and the impella 5.5: patient was on support at time of death; no direct device malfunction identified. iin conclusion: serious injury (hemolysis) and death reported. death likely related to underlying clinical condition (sepsis from ileus) rather the impella pump. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, clinical symptom, arrhythmia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. regarding, hemolysis, the pump was not returned for investigation. data log shows three motor current spikes, two on 12-oct around 10:00 pm and another on 14-oct around 9:30 am, while running on a steady p-level. the motor current spikes were indicative of biomaterial ingestion events during the case. the cause of the hemolysis issue was most likely biomaterial ingestion, based on data log review. pertaining to device malposition, device in the wrong position, there was no definitive indication that the pump was in a misposition at the start of the case. the cause of the pump position issue was not determined without sufficient clinical information. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | H6. health effect - clinical code for arrhythmia and health effect - impact code for resuscitation were not included in the initial mdr and is now correctly added.
3501 | 1220648-2025-48317 | 23535228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535228&pc=OZD | 10/16/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520528 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 68 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented with chest pain and dyspnea. the patient was admitted for acute hypo six respiratory failure, in setting lactic acidosis, and non-st-elevation myocardial infarction. cardiogenic shock was identified on the floor. the patient was not taken to the catheterization lab due to poor renal function. the patient¿s renal function improved the next morning, and the patient was taken to the catheterization lab for an impella cp. the impella cp was successfully placed. the patient was flown to a different facility for advance therapies. during support it was noted that the care team decided to start the weaning process due to suspicious of hemolysis. the next day it was noted that there were no signs of hemolysis however the patient coded in ventricular tachycardia and required chest compressions. ultimately the patient expired on support. | [5 empty] | Correction: d3 and h6 medical device problem code 01 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. hemolysis: there were no data logs returned. due to a lack of sufficient clinical information and no product or data logs returned, the cause of the hemolysis cannot be established. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3502 | 1220648-2025-48305 | 23535243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535243&pc=OZD | 09/18/2025 | 11/06/2025 |  | 11/12/2025 |  |  | 09/18/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719641 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Pain | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient's pump mispositioned due to the touhy not tightened and the groin was oozing. the groin was oozing from either the vein or the impella site. the medical doctor placed a fem stop for a few hours in the intensive care unit to slow the ooze while the act came down. the patient's pump had to be repositioned yesterday under echocardiogram and had placement signal low. the pump was sitting midventricular free of all structures. atrial fibrillation and rapid ventricular response were on the monitor. the patient was started on amiodarone and was given intravenous pyelogram digoxin. the decision was made for synchronized cardioversion for atrial fibrillation with rapid ventricular response. the patient was awake, alert, and oriented with complaints of mild back discomfort. | [5 empty] | Correction h6 health effect - clinical codes 0601, 0506 and medical device problem code 150202 were inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. back pain: the cause of the clinical symptom cannot be determined due to insufficient clinical details. atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. access site adverse event: due to lack of clinical information provided on what the activated clotting time (act) values were and if that resolved the oozing, the root cause cannot be established. device in wrong position/difficult to close/open cath anchor or tuohy based on the clinical details, the root cause of the device in wrong position/difficult to close/open cath anchor or tuohy is due to user related error. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
3503 | 1220648-2025-48319 | 23535326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535326&pc=OZD | 09/18/2025 |  |  | 11/12/2025 |  |  | 09/18/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697750 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient encountered hemolysis during impella cp support. it was noted that the patient had elevated levels of plasma free hemoglobin which is indicative of hemolysis. no further intervention was mentioned. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3504 | 1220648-2025-48312 | 23535619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535619&pc=OZD | 09/26/2025 |  |  | 11/12/2025 |  |  | 09/25/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725383 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | No | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced impella in aorta alarms that occurred under p-2 level. after changing the p level, the alarm immediately resolved. the pump position was described as good by the physician according to the sonography and that the impella was working well. later, the patient expired and per the physician, the death was not related to impella use. | [5 empty] | The cause of the change in the ps and "ps low" alarms could not be determined. based on the clinical details provided and data log review, the cause of the false "position in aorta" alarms was patient condition related. the device passed all post-sterile inspection checks. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3505 | 1220648-2025-48315 | 23535657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535657&pc=OZD | 10/16/2025 |  |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026744764 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 61 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella cp support, the patient had the left femoral artery occluded from the impella sheath. a bilateral antegrade sheath was placed for distal perfusion and the patient was placed on venoarterial extracorporeal membrane oxygenation, and the impella cp was turned down to p2. temp pacer was also removed. six units of packed red blood cells and four units of plasma were given. following initiation of extracorporeal membrane oxygenation and blood transfusion, the patient was stabilized and all pressor support was able to be weaned off. left ventricular ejection fraction appeared to have returned to her baseline. the vascular team was consulted to expedite decannulation and explant. a fixed dose of heparin was given and now on the patient is on standby, pending the operating room. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information was received and d9 was updated with this information | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary the device was returned for evaluation. visual inspection found no issues on the pump. ischemia: in order to make a cause determination, all clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. vessel perforation: the cause of the injury was not determined due to insufficient clinical details & was not related to impella since it occurred before placing impella. device history review ==> pump sn: (b)(6) passed all post sterile inspection checks.
3506 | 1220648-2025-48316 | 23535731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535731&pc=OZD | 09/25/2025 | 12/12/2025 |  | 11/12/2025 |  |  | 09/25/2025 | 11/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740854 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp required a temporary pacer due to low heart rate. the patient displayed hematuria. the pump was repositioned to resolved the hematuria. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3507 | 1220648-2025-48321 | 23535775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23535775&pc=OZD | 10/13/2025 |  |  | 11/12/2025 |  |  | 10/13/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662222 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the patient admitted in with acute myocardial infarction/cardiogenic shock. it was reported that the patient had urine output from clear yellow to red after patient sat up to vomit. an echocardiogram showed the pump was deep. pump pulled back. the patient also became hypotensive with that event. started low dose levo and milrinone that they plan to wean. the patient had significant chest output. continue to replace blood products. the impella unable to run above p2 without suction. bedside echo imaging not optimal after cv surgery. limited windows showed pump has moved and is a little deep with pigtail on pap. interventional fellow repositioned several times. suction resolved temporarily but pump is positional seems to dive back to pap. physician uncomfortable pulling back more with sub par transthoracic echocardiogram (tte). requested tee but intensivist hesitant due to bleeding after issues passing probe in operating room. team currently wants to leave pump as is despite suction and sub optimal flows. later the pump was repositioned due to impella being in the aorta. reportedly advanced 8 cm with bedside echo. nurse was unsure what measurement was from inlet to annulus. it was further reported that the pump had suction. the patient was given a unit of red packed blood cells, fresh frozen plasma, and albumin. suction has resolved. later amio was increased due to rapid atrial fibrillation. the family decided to withdraw care and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data logs were returned for analysis. data log analysis showed suction alarms and positioning alarms. of the positioning alarms, there were impella position in aorta alarms and impella position unknown alarms placement signal trends were also seen during the case. no motor current spikes or trends were seen outside of p-level changes. mechanical interaction with blood: the root cause of the hemolysis was most likely patient condition related based on clinical details and data log analysis. device in wrong position: the root cause of the initial positioning issue on the day of implant was most likely use related based on clinical details. however, the positioning issues faced later during the case where the pump seemed to move after repositioning was most likely patient condition related based on clinical details. pump suction: the root cause of the suction alarms was most likely patient condition related based on the clinical details provided. arrythmia: the root cause of the arrythmia was unable to be determined due to insufficient clinical details. major bleed: the cause of the major bleed was not determined due to insufficient clinical details. hypotension: the cause of the hypotension was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3508 | 1220648-2025-48310 | 23536031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23536031&pc=OZD | 10/13/2025 |  |  | 11/12/2025 |  |  | 10/13/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758905 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient returned to the emergency department on (b)(6) 2025 with complaints of shortness of breath, chest pain, weakness, and was status post st-elevation myocardial infarction with a percutaneous coronary intervention performed two weeks prior. the patient was admitted with an elevated troponin level in non-st-elevation myocardial infarction. shortly after admission, the patient experienced a cardiac arrest and was transferred to the intensive care unit (icu) for monitoring. the following day, the patient was urgently taken to the cardiac catheterization lab (ccl) for impella cp placement to facilitate stabilization. during transfer to the ccl table, the patient coded twice more. an impella cp was successfully implanted utilizing best practices, and a temporary pacemaker was placed and set at 80 bpm for rhythm stabilization. after the patient was stabilized, a repeat evaluation of the coronary arteries showed no changes from the studies two weeks prior. two perclose devices were deployed for access site closure, and the patient was transferred back to the icu in critical condition. the patient remained in a guarded state, and the physician care team immediately discussed goals of care with the family. following these discussions and multiple cardiac arrests over the preceding 24 hours, the family made the decision to withdraw care. the patient¿s care was withdrawn, and the patient expired. | [5 empty] | Medical safety reviewed the event and noted the following: the patient experienced cardiogenic shock following a cardiac arrest and underwent emergent placement of an impella cp device for hemodynamic support. approximately six hours after initiation of impella support, the patient¿s family elected to withdraw care. the patient expired while the impella cp remained in place. based on available clinical information, there is no indication that the impella cp device contributed to the patient¿s death. the decision to withdraw care was made by the family in the context of the patient¿s critical condition, including multiple cardiac arrests within the preceding 24 hours. the event is being conservatively reported as occurring while on impella cp support, in accordance with regulatory requirements, despite the absence of evidence of device malfunction or contribution to the fatal outcome. reported due to patient death while on impella cp support, per regulatory requirements. no evidence of device contribution to the event. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the cardiac arrest, the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | A4 is unknown. the event was reviewed. the event describes a patient who presented with chest pain and shortness of breath two weeks status post st-segment elevation myocardial infarction with one vessel percutaneous coronary intervention (surgical intervention with coronary artery bypass graft was declined). troponin levels were elevated. the patient was admitted as an non-st-segment elevation myocardial infarction. the patient coded shortly thereafter and was monitored in the intensive care unit overnight. later in the morning, the patient was urgently transported to the catheterization lab for impella cp placement for stabilization. upon transfer to the catheterization lab table the patient coded twice. the impella was successfully placed, and a temporary pacer was placed to help with rhythm stabilization. the view of coronary arteries results was unchanged from two weeks ago. the patient was transferred back to the unit on impella mechanical circulatory support in critical condition, guarded state. a discussion was made with the family. it was noted after family discussion and multiple codes over the last 24 hours, the decision was made to withdraw care, which led to the patient's demise (after approximately six hours of impella mechanical circulatory support). the impella cp did not malfunction (and was not associated with an adverse event), was in use during patient expiration, but does not appear it would have materially contributed to the patient's demise given the patient's underlying condition (cardiogenic shock extremis). however, conservatively, the death is being reported as the device was in use at the time of expiration. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3509 | 1220648-2025-48320 | 23536392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23536392&pc=OZD | 09/24/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 09/24/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735644 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 could not be completed due to the patient's anatomy. no patient harm was reported. | [5 empty] | No product was returned. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3510 | 1220648-2025-48322 | 23536715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23536715&pc=OZD | 10/16/2025 |  |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750041 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 62 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient¿s left femoral arterial impella site with bleeding and hematoma. the hematoma was marked and monitored, imaged access site and pressure held. angle match dressing applied. impella successfully explanted due to groin concerns. groin site appeared improved and softer the following morning. | [5 empty] | The cause of the bleeding and hematoma was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3511 | 1220648-2025-48323 | 23536815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23536815&pc=OZD | 10/16/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/16/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026739032 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Device Sensing Problem | [5 empty] | Renal Failure | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support for intra-aortic balloon pump placement. while on support, patient was sent to operating room for ventricular fibrillation (vf). placement signal was not correlating with arterial line readings. patient presented again with vf, was shocked 3x and given amio. pt was placed onto dialysis. during removal, there was resistance when pulling impella. fluoroscopy revealed cannula ¿stuck¿ in the subclavian artery. doctor decision to re-open sternum to investigate further. manipulation to aorta to attempt to straighten the path to graft. pump successfully removed but a biomaterial possible endothelial found on explant on the motor housing. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the thrombus/ventricular fibrillation. the clinical noted the patient had vf throughout support and that thrombus was on the motor during explant. the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the clinical noted that the patient was on dialysis on support. the cause of the injury was not determined due to insufficient clinical details. optical sensor issue: about two days into support, the clinical stated that the ps was not correlating to the arterial pressure line. the map was 62 while the ps was around 30. the imc logs show at case start the calibrated g0 was about 137 lower than the factory g0 value. sea level pressure during the time of calibration was about 30.04 inches. the lt logs show that at the beginning of starting the case the ps was around 60 but then it remained variable during the case in between the with systolic values in the 30's and 70's. during this time the 5s parameters were stable as they were not variable or noisy. there are two instances of psnr alarms and a change in 5s parameters but that occurs when the pump is shut off and unplugged but the parameters recover about 50 hours later. specifically, an rt log was observed on 10/16 during the time of the reported complaint with the ps appearing aortic with systolic values of 70 and diastolic of 25. there were no changes in mc or pump flows. another rt log was observed on the 27th before explant, and systolic values were about 60 with diastolic values of about 20. 5s were stable during this time. there is no confirmation of any abnormalities with the placement signal. due to a lack of product returned and clinical details regarding the a-line and what was occurring with the patient during this time, the root cause of the optical sensor issue cannot be determined. difficult/unable to remove: the clinical details state that there was resistance when attempting to explant the pump. it revealed that the cannula appeared to be stuck in the subclavian artery. the physicians re-opened the sternum and manipulated the aorta to straighten the path of the graft. the pump was successfully removed but biomaterial was found on the motor housing upon explant. it was stated that the motor housing may have scrapped the inside of the artery due to the sharp angle of the subclavian artery and aorta bifurcation. based on the clinical details, the root cause of the difficulty removing the device is most likely due to the patients condition as there was a sharp angle of the subclavian artery and aorta bifurcation causing resistance and a difficult angle to explant. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3512 | 1220648-2025-48326 | 23536940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23536940&pc=OZD | 10/17/2025 | 11/12/2025 |  | 11/12/2025 |  |  | 10/17/2025 | 11/12/2025 |  |  |  |  | 11/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708776 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp, the patient had cold leg or limb ischemia. the patient remains highly acidotic ph 7.1, troponin greater than 25k. lactic not clearing, alanine transaminase and aspartate transaminase all extremely high. the leg was cold and pulses, plan to upgrade to 5.5 shortly. the pump was explanted. attempted escalation to impella 5.5, but axillary cp placed instead due to small vasculature. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3513 | 1220648-2025-48250 | 23520953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23520953&pc=OZD | 10/13/2025 |  |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697933 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 69 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient had been in and out of atrial fibrillation. four bolus of amiodarone have been given. dobutamine went from seven to five. the doctor spoke with the family regarding plans for recovery. an echocardiogram was done at bedside and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3514 | 1220648-2025-48251 | 23520972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23520972&pc=OZD | 10/11/2025 |  |  | 11/11/2025 |  |  | 10/14/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691681 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male | 83 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered pump suction and positioning issues during impella cp support. due to low volume, the impella encountered a suction alarm. the doctor accidently pulled the impella out of the ventricle. as a result, the decision was made to exchange the impella cp with a new one. the pump was exchanged successfully. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product was returned for investigation. data logs were returned for analysis. pump suction: clinical details mention that the pump got a suction alarm due to less volume. data log analysis confirmed the suction alarm occurring towards the end of the case. motor current high alarms were also seen. the root cause of the pump suction was not determined due to lack of sufficient clinical details and unreturned product. device in wrong position: clinical details mention that the doctor accidently pulled the impella out of the ventricle and was not in a position to replace it, so a new pump was implanted. the root cause of the device being in the wrong position was user related based on the provided clinical details. device history review pump passed all post-sterile inspection checks.
3515 | 1220648-2025-48252 | 23521029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521029&pc=OZD | 10/14/2025 |  |  | 11/11/2025 |  |  | 10/14/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758804 |  | 000000000001000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 80 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient supported with an impella cp experienced episodes of arrhythmias: supraventricular tachycardia and frequent premature ventricular contractions, requiring medication. the patient had intermittent hypotension. the device position was confirmed via echocardiogram and remained stable. the patient was eventually weaned off support and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3516 | 1220648-2025-48253 | 23521103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521103&pc=OZD | 10/14/2025 | 12/31/2025 |  | 11/11/2025 |  |  | 10/14/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697717 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient¿s right leg lost pulse and became cold and mottled. the decision was made to wean and remove the device. the patient required levophed at p2. the patient was taken to the operating room. a cutdown of the common femoral artery was done and the device was removed. clots were pulled from the iliac. surgical closure was preformed and positive pulses in both prothrombin time and diastolic pressure were noted. | [5 empty] | Investigation summary: no product was returned for investigation. ischemia/ thrombosis: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1947607 device history batch: subcomponent lot: n/a device history review: device with sn: 610796 passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.".
3517 | 1220648-2025-48254 | 23521150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521150&pc=OZD | 10/14/2025 | 12/31/2025 |  | 11/11/2025 |  |  | 10/14/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692751 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, the patient was moved out of bed to a chair but had sustained ventricular tachycardia. levophed was increased and the patient was taken back to bed. the patient likely will need a longer duration of impella 5.5 support per the medical doctor. additional information was received from the area clinical consultant. the device is not available for return and the ventricular tachycardia had nothing to do with the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1932234. device history batch: subcomponent lot: n/a. device history review: device with sn: 603340 passed all post sterile inspection checks.
3518 | 1220648-2025-48255 | 23521259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521259&pc=OZD | 10/14/2025 | 12/23/2025 |  | 11/11/2025 |  |  | 10/14/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026728919 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 54 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was attempted to be escalated from an impella cp to an impella 5.5 for support. however, the impella 5.5 was unable to pass the thoracic inlet. the physician opted instead to place another impella cp. | [5 empty] | The investigation for the reported delivery issue has been completed. the root cause of the delivery issue was determined to be patient condition as the patient had small anatomy at thoracic inlet preventing successful delivery of impella. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3519 | 1220648-2025-48256 | 23521452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521452&pc=PYX | 10/18/2025 |  |  | 11/11/2025 |  |  | 10/18/2025 | 11/11/2025 |  |  | 01/31/2026 |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025513851 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 132 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp flex was positioned at 58 centimeters and was operating at p9, yet the patient continued to experience suction alarms and the left-sided impella could not advance beyond p4. a significant increase in the patient¿s white blood cell count from 43 to 60, likely due to ischemic bowel and sepsis. device placement was confirmed with echocardiogram. ultimately care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. clinical details revealed that suction alarms were noted on 18 oct 2025, making it difficult to run the left-sided support pump above p4 despite the rp running at p9. the patient was also suspected to be suffering from ischemic bowel causing sepsis. peri-procedural adverse event (ischemia, sepsis): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
3520 | 1220648-2025-48258 | 23521520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521520&pc=OZD | 10/21/2025 | 01/06/2026 |  | 11/11/2025 |  |  | 10/21/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026718295 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male | 102 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during the impella cp implantation, it was difficult to place the peel-away-sheath. after that correct position was noted according to fluoroscopy and contrast agent administration. when placing the impella, the pump could not be fully placed. it got stuck in the sheath and could not be advanced any further. a kink in the sheath was now visible in fluoroscopy. removal of the pump was only possible by peeling the sheath. a new pump was implanted with new peel-away-sheath. there was no harm to the patient. there are two related reports. this report represents the pump and another report was submitted to represent the introducer. refer to section h10 for the related report number. | [5 empty] | Investigation summary: unable to place or position: impella cp (b)(6), two 14 fr x 25 cm introducers, and a purge cassette were returned for investigation. one introducer was in a peel-away state and had several kinks along its length, but no tip damage. the second introducer also had several kinks, with kinked tip. no damage was found on the returned pump or purge cassette. the pump cannula was also kinked lengthwise on the returned product. no other physical damage was noted on the returned pump. the pump was tested in a bucket of water and no abnormalities were observed. the cause of the delivery issue was related to kinks on the introducer and the pump¿s cannula. device history lot device batch: 1955208. device history batch subcomponent lot: na. device history review the pump (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated.
3521 | 1220648-2025-48259 | 23521652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521652&pc=OZD | 10/13/2025 |  |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707242 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Hematuria | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on an impella 5.5 for cardiac support. during impella 5.5 support, there was a purge blocked alarm noted and purge flows were low and purge pressure was high. the physician ordered tissue plasminogen activator to run as purge. it was noted that the patient was having a lot of clots. a slight pink red tint was noted to the foley catheter output. the impella site had a small amount of bloody drainage and no hematoma was noted. however, there was bleeding from the swan ganz site followed by hypotension. one liter of lactated ringer's was administered in addition to one unit of packed red blood cell. heparin was placed on hold. the patient was successfully weaned off impella 5.5 support and survived. | The patient was noted to have been experiencing atrial fibrillation during support. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Correction: b5 and h6 additional failure mode was inadvertently omitted from initial report. the investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia/thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria/major bleed/hypotension: the cause of the injury was not determined due to insufficient clinical details. purge system blocked: the real time (rt) log shows a jagged spike in purge pressure from 600mmhg to over 1100mmhg in about a minute. high purge pressure is sustained for about 40 minutes after which there is a sharp drop before spiking back up to over 1000mmhg. during this second sustained event, there is a slight shift up in motor current and pump flow values. after 20 minutes the purge pressure values sharply drop to normal values and there are no other occurrences. the log pattern observed is characteristic of a kink in the purge lumen. purge flow values also increase. the cause of the purge system blocked was determined to be due to a kink of the purge lumen, though cause of kink can not be established since no product was returned and insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks.
3522 | 1220648-2025-48260 | 23521656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521656&pc=OZD | 10/22/2025 | 11/11/2025 |  | 11/11/2025 |  |  | 10/22/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026741746 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Noise, Audible | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was escalated to an impella 5.5 from an impella cp for mechanical circulatory support. the patient presented to the emergency department with atrial fibrillation with rapid ventricular response, and the patient was taken to the cardiac catheterization lab (ccl). in the ccl, the patient underwent a percutaneous coronary intervention (pci) to the left anterior descending (lad) artery, involving the placement of four drug-eluting stents. during the procedure, the patient required hemodynamic support with levophed. due to the patient's condition, a decision was made to place both an impella cp for mechanical circulatory support and a swan-ganz catheter for advanced hemodynamic monitoring. the impella cp was inserted without issue and support was initiated. due to the patient's large pannus, the peel-away sheath was removed slowly in incremental segments. during this process, the clear plastic retention ring from the 14 fr dilator became lodged around the heart pump catheter. during the removal of the peel-away sheath, a placement signal low alarm occurred, shortly followed by the pump migrating fully into the ventricle. fluoroscopy confirmed the pump had curled within the ventricle. the physician attempted to slowly pull the catheter back, but the pump completely exited the ventricle into the aorta. efforts to reposition the pump failed as the physician was unable to recross the valve. the pump was then fully removed, flushed with heparinized saline, and reinserted using a short sheath. the second insertion attempt was successful and completed without further issue. a reperfusion sheath was placed, and the device was locked in position at 92 cm depth and set to performance level p6. the patient was transferred to the cardiac unit for rest and recovery. there were slight oozing and hematoma noted in the right groin access site, however no additional oozing was noted at the impella access site after dressing was removed. manual pressure was applied to the groin, and the hematoma was reduced. it was noted that the physician was at the bedside to address the patient¿s low cardiac index of 1.4 and low mixed venous oxygenation of 45%. the physician ordered the patient to remain off anticoagulation, and the impella purge fluid was changed to 5% dextrose in water with bicarbonate. the following day the patient¿s lactic acid increased from 2.1 mmol/l to 4 mmol/l with an acute kidney injury noted. the decision was made to escalate the patient¿s therapy to an impella 5.5. the patient was brought to the operating room where the impella 5.5 was successfully placed, and the impella cp was explanted without issue. on support day 10 with impella 5.5, the patient experienced episodes of atrial fibrillation. the patient had been extubated the day prior but the patient¿s requirement for levophed increased. the medical team reported motor noise from the impella, however the impella¿s function was not reported to be affected by the noise. the impella 5.5 was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complications. the patient's outcome at the time of explant was survived. this complaint involves three impella devices: pump 1: impella cp with serial number: (b)(6). 14fr introducer with serial number: (b)(6), used with impella cp. pump 2: impella 5.5 with serial number: (b)(6). this mdr specifically addresses pump 2: impella 5.5 with serial number: (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: the device was not returned for evaluation. arrhythmia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. noisy device: clinical reported noise from impella motor. the data logs show pump was running at p7 on day of complaint. logs show no motor current spikes outside p-level changes. the cause of the issue could not be established as no product was returned. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
3523 | 1220648-2025-48249 | 23521837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23521837&pc=OZD | 10/13/2025 |  |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620601 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. when attempting to remove the peel away sheath at bedside, the device was pulled back into the aorta. the patient was taken emergently to cardiac catheterization laboratory to replace the device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled back into aorta when attempting to remove peel away sheath at the bedside. device history review: the device passed all post sterile inspection checks.
3524 | 1220648-2025-31922 | 23522051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23522051&pc=OZD | 10/22/2024 | 11/11/2025 |  | 11/11/2025 |  |  | 10/22/2024 | 11/11/2025 |  |  | 02/28/2026 |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562869 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support, at the beginning of support, there was an optical sensor failure and the issue persisted until the end of the case. no report of patient harm or injury. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation for the optical signal issue has been completed. no data logs were returned. pump returned and psnr was reproduced. pump passed laser continuity test. uson test showed psnr alarm and raw optical reading in the -250 mmhg range during testing. the 5s parameters were all within normal ranges though contrast was slightly low. sensor face was imaged with no abnormalities. cavity length was within normal ranges. without the field data logs and imc logs, the failure cannot be investigated further to determine the issue in the field. the g0 is unable to be checked without the imc logs. the root cause of the optical signal issue was not determined since field data logs were not returned for investigation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3525 | 1220648-2025-48261 | 23522056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23522056&pc=OZD | 09/12/2025 | 11/11/2025 |  | 11/11/2025 |  |  | 10/29/2025 | 11/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1683279 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported after investigation of the pump, the investigation team suspected an issue with the automated impella controller (aic). investigation team suspected that the low snr (signal to noise ratio) was related to the aic. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | E4 was inadvertently omitted on the initial manufacturer device report. h6 medical device problem code a0709 was erroneously entered on the initial manufacturer device report.
3526 | 1220648-2025-48262 | 23522062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23522062&pc=OZD | 08/06/2025 | 12/12/2025 |  | 11/11/2025 |  |  | 08/06/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672079 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 92 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was escalated from an impella cp to an impella 5.5 for a case of acute myocardial infarction/cardiogenic shock. while on impella 5.5 support, the patient received two units of blood after the operating room and was placed onto continuous renal replacement therapy. additionally, a chest tube was placed for pleural effusion. the patient ultimately expired. | [5 empty] | The investigation for the reported renal failure/ pulmonary edema has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure/ pulmonary edema could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3527 | 1220648-2025-47908 | 23522187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23522187&pc=OZD | 08/18/2025 |  |  | 11/11/2025 |  |  | 08/18/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712087 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered placement signal issues, false alarm and the patient had ventricular tachycardia (vt) during impella 5.5 support. the patient had continuous vt during support which required shock and an amiodarone intravenous drip. additionally, it was reported that the patient had persistent suction alarms and placement signal (ps) low alarms. manual cuff pressures consistently read 77/62/66. impella ps had been decreasing over time and then read as low as 32/27/29. the pump had been in for 41 days. no change in hemodynamics and the team further discussed inaccurate aorta sensor. ao placement was off and the team believed there was a sensor drift and inaccurate ao sensor. an echocardiogram showed the impella 5 centimeters in mid ventricular space. support continued and the patient remained stable. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Clinical details report that there were persistent suction and ps low alarms. data logs were reviewed and the reports were confirmed. the product was not returned for investigation. based on the patterns noted in data log review, the cause of the optical signal issue and suction alarms was determined to be optical sensor wear. clinical details report intermittent but reoccurring ventricular tachycardia throughout support. the injury was treated with shocks, medication adjustment, and pump repositioning. the root cause of the injury was unable to be determined due to insufficient clinical details. the device passed all post-sterile inspection checks.
3528 | 1220648-2025-48264 | 23522452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23522452&pc=OZD | 10/13/2025 |  |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697126 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 86 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a stemi patient presented with occluded left anterior descending. during impella cp support the patient had pink-red urine and labs were hemolyzed. the p level was turned down and the urine output was slowly improving. in addition continuous renal replacement therapy was started due to low creatine and was creatine slightly improved. the patient continued on support and was eventually escalated to a impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction h6 health effect - clinical code 1302 and medical device problem code 01 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. renal failure/hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3529 | 1220648-2025-48265 | 23522619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23522619&pc=OZD | 09/02/2025 | 11/03/2025 |  | 11/11/2025 |  |  | 09/03/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714485 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Asystole | [8 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was escalated from an impella cp smartassist to an impella 5.5 smartassist. the axillary site was oozy post implant. the patient received multiple units of packed red blood cells and a sandbag was placed over the impella site. the patient is on anticoagulants, antiplatelets, and medication. the patient became bradycardic after suctioned by respiratory and went into ventricular standstill, with only p-waves on electrocardiogram. same sequence occurred again, ventricular tachycardia requiring defibrillation three times before return of spontaneous circulation. another episode of bradycardia occurred requiring atropine and transcutaneous pacing pad placement. the patient remained in asystole under pacing pads before rhythm was regained. an echocardiogram showed the impella position was deep. the pump was repositioned with waveform improvement. additional units of packed red blood cells were required due to low hemoglobin and hematocrit. bleeding from the mouth and retroperitoneal bleed were noted. suction events requiring p-level decrease due to ectopy. a jackson pratt drain was placed at the access site. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. major bleed: clinical details note bleeding from the mouth throughout the case. the cause of the major bleed was not determined since insufficient clinical details were provided. retroperitoneal hemorrhage: the cause of the injury was not determined due to insufficient clinical details. bradycardia/asystole/tachycardia/arrythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: clinical details note suction on the morning of (b)(6) 2025. it is unknown if there were positioning issues at this time. data log review shows some suction alarms while changing p-levels that all resolved quickly. there were no motor current spikes near this time and there was no clear trend in ps. the cause of the suction was not determined since product was not returned and insufficient clinical details were provided.
3530 | 1220648-2025-48267 | 23523089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23523089&pc=OZD | 10/14/2025 |  |  | 11/11/2025 |  |  | 10/14/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737979 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient with heart failure was escalated from impella cp to impella 5.5 and ecmo was weaned. during impella 5.5 support the patient had a subarachnoid hemorrhage. it was further reported that care was withdrawn and the patient expired. | Additional information was received. the subarachnoid hemorrhage was an event that occurred prior to any impella implants. the patient presented with this neurological condition alongside the heart failure. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and b5 was updated. | Medical safety/clinical reviewed the event. it was initially reported a patient with heart failure was escalated from impella cp to impella 5.5 and ecmo was weaned. during impella 5.5 support the patient had a subarachnoid hemorrhage. care was withdrawn and the patient expired. additional information was received and noted the subarachnoid hemorrhage was an event that occurred prior to any impella implants. the patient presented with this neurological condition alongside the heart failure. upon further review, no allegation or reportable device-related event was identified. this follow-up 2 report is being submitted to redact this previously reported event. no further updates will follow.
3531 | 1220648-2025-48269 | 23523099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23523099&pc=OZD | 10/15/2025 |  |  | 11/11/2025 |  |  | 10/15/2025 | 11/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726781 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was ischemia during impella cp patient support. while on support, there was a lack of a pedal pulse and the impella was removed. the patient survived. | [5 empty] | The investigation for the reported event, ischemia, has been completed. no product was returned for investigation. the root cause of the injury was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3532 | 1220648-2025-48268 | 23523282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23523282&pc=OZD | 10/13/2025 | 02/24/2026 |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682500 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information has been provided upon request: "no, the device is not available for return." | It was reported that a patient presented for coronary artery bypass graft and exchange of impella 5.5. the procedures were successfully. during 5.5 support it was noted that amio bolts were given for atrial fibrillation. the patient remained on support for 305.22 hours and was successfully weaned and explanted. however, after explant a clot was noted on the cannula. |  |  |  |  | B5 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia / thrombosis : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3533 | 1220648-2025-48270 | 23523504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23523504&pc=OZD | 10/13/2025 | 11/11/2025 |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725112 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 96 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Pulmonary Edema | Hypovolemia | Respiratory Failure | [9 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted was escalated to an impella 5.5 from an impella cp device for mechanical circulatory support. the patient underwent a transcatheter aortic valve replacement (tvar) procedure on (b)(6) 2025, which was immediately complicated by a st-elevation myocardial infarction due to a left anterior descending (lad) artery occlusion. the patient received lad stenting and was cannulated for extracorporeal membrane oxygenation (ecmo) support, then successfully decannulated on (b)(6) 2025, the patient experienced pulseless electrical activity arrest and was recannulated for ecmo and then brought to the cardiac catheterization laboratory for impella cp placement for left ventricular unloading. on day 2 of impella cp and emco support, the patient was given volume, a unit of blood, and albumin. it was noted that the patient had hematuria, however the patient¿s bladder had a known bleed and there was no concern from the physician for hemolysis. the following day it was noted that the patient continued to bleed, and required blood products. on day 4 of impella cp and emco support left heart catheterization was performed and found a severe 70% mid-distal lad lesion, and two severe blockages in the right coronary artery: a 90% blockage and a 50% blockage in the proximal and mid sections. it was noted that the previously placed stents in the left main (lm) to lad and lm to left circumflex (lcx) arteries were patent. there was no immediate interventions planned for the patient. on day 5 of impella cp and emco support a transesophageal echocardiogram (tee) was to assess for thrombus on the tavr valve leaflets before considering an escalation to the impella 5.5. the tee revealed no thrombus on the aortic valve leaflets but indicated poor intrinsic heart function and restricted leaflet opening. a thrombus was noted in the left atrial appendage (laa). intermittent impella suction alarms occurred when the patient coughed. the decision was made to escalate the patient to an impella 5.5 for advanced support. the impella 5.5 was successfully placed, the impella cp was weaned and explanted. the patient remained on ecmo support with the impella 5.5. of note, 3 units of packed red blood cells were transfused to the patient during the implantation due to blood loss from the right axillary arterial line which was removed prior to the case. on support day 1 2 of impella 5.5 and ecmo, large left pleural effusion was identified with a chest x-ray. a pigtail catheter was placed for drainage. the patient required emergency reintubation this morning in the intensive care unit (icu) due to acute respiratory distress and hypoxia, occurring prior to a planned procedure in the operating room. there were reported impella suction alarms, which the medical team suspected to be related to hypovolemia. the impella performance was consequently reduced to p4 with flows of 1.6 l/min to because of the alarms. on support day 2 of impella 5.5 and ecmo, the patient¿s arterial blood gas results indicated the patient had hypoxia and hypercarbia upon admission from the operating room. ecmo flow rates were increased and the impella p level was increased to p3. overnight, the patient developed unequal pupils, prompting a head computed tomography (ct) and the temporary holding of anticoagulation therapy. the ct revealed no acute changes but noted findings consistent with "cortical laminar necrosis." this morning, the patient is stable and following commands in all extremities. the patient continued continuous renal replacement therapy (crrt). on support day 9 of impella 5.5 and ecmo, the patient experienced an acute failure of the ecmo system when the oxygenator clotted off. during the subsequent procedure to exchange the oxygenator and the entire circuit, the impella heart pump performance level was increased to p8. shortly after the increase, the impella experienced suction events and low flow. an echocardiogram performed at the bedside showed the left ventricle to be completely decompressed with no volume, and failed ramp tests suggested acute right ventricular failure. the patient was successfully placed back onto the ecmo circuit, achieving stable flow at 5 l/min, which stabilized the patient's condition. post-stabilization, the impella was stable at performance level p3. there were no immediate concerns or questions reported regarding the impella device function itself after the circuit exchange was completed. on support day 16 of impella 5.5 and ecmo, the patient¿s family decided to withdraw the care of the patient. the patient¿s care was withdrawn and the patient expired. this complaint involves two automated impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event and noted the following: the patient experienced pulseless electrical activity (pea) and was placed on extracorporeal membrane oxygenation (ecmo) support. impella cp was implanted for left ventricular unloading. during support, hematuria was observed, attributed to a pre-existing bladder bleed, and one unit of blood was administered. the patient was subsequently escalated to impella 5.5 with continued ecmo support. during impella 5.5 implantation, significant bleeding occurred at the axillary access site, requiring additional blood products. the patient later developed acute respiratory distress secondary to a left pleural effusion, which was drained. after 16 days of impella support, the patient¿s family elected to withdraw care, and the patient expired. adverse events impella cp: bleeding event (hematuria requiring transfusion); of note, a thrombus was noted in the left atrial appendage (laa). intermittent impella suction alarms occurred when the patient coughed. thrombosis is a high risk with ecmo use. impella 5.5: bleeding event (axillary site bleeding requiring transfusion) and hypoxia/respiratory failure (pleural effusion). the death is conservatively reported on impella 5.5; unlikely device-related, attributed to patient¿s underlying acuity. correction: concomitant device ecmo inadvertently omitted from initial report was added to section d10 accordingly. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the cause was unable to be determined due to insufficient clinical details. additionally, regarding the pulmonary edema/respiratory failure/hypovolaemia, the cause of these adverse events was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
3534 | 1220648-2025-48271 | 23524387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23524387&pc=OZD | 07/14/2022 |  |  | 11/11/2025 |  |  | 07/14/2022 | 11/11/2025 |  |  | 09/30/2023 |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136217 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported placing a second impella cp smartassist after the first impella cp immediately had suction and low flows. insertion of the second pump was successful but yielded the same results - immediate suction and very low flows. an echocardiogram was completed and showed no blood flow to the left ventricle, compression, and potential bleeding into the lungs. the medical team was going to discuss with the patient's family but the patient passed. | [5 empty] | The investigation was completed. major bleed: clinical reported that the lv became compressed and no volume was flowing, and it was attributed to a bleed into the lungs. the product was returned for investigation but no abnormalities were observed. the cause of the bleed could not be determined due to insufficient clinical details. pump suction/low pump flow: clinical reported that the pump was getting low flows and suction immediately after starting the pump. this was the second pump attempted in the case. the product was returned for investigation and no kinks or abnormalities were observed. the returned pump was not run in the lab to reproduce the low flow. data log review confirms low flows and suction during the case, which was only about 15 minutes long. at p-9, flows were less than 3.0l/min. placement signal was very low from the start and lost pulsatility over time. after ~4 minutes of running, the pump was turned off by the user for ~6 minutes. after restarting the pump, placement signal pulsatility and magnitude were lower than before and trended down for the remainder of the case. data logs from the prior pump showed a decrease in placement signal and low pulsatility as well. it was mentioned in clinical details that the patients there was no blood volume in or entering the lv and it became compressed due to a bleed into the lungs. according to clinical details and investigation results, the root cause of the low flow and suction was due to the patients condition, as the low placement signal aligns with clinical stating there was no volume in the lv. device history review was completed and passed all post sterile inspection criteria. | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3535 | 1220648-2025-48273 | 23524600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23524600&pc=OZD | 10/16/2025 | 11/11/2025 |  | 11/11/2025 |  |  | 10/16/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639716 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 83 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was hemolysis and a positioning issue during impella cp patient support. while on support, it was noted that the staff was unable to get a full set of labs due to hemolysis. the nurse called about probable hemolysis. they stated that the pump was under the papillary muscle. the fellow did not want to reposition the pump and wanted to know if reducing the pump¿s performance level would help. the company clinical support center consultant advised an echocardiogram and positioning the pump. the pump was repositioned at the bedside. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Correction: h6 medical device problem code a01 was inadvertently omitted. the investigation of the reported failure mode has been completed. no product was received for evaluation. hemolysis: unable to get full set of land due to hemolysis. after review of logs, no suction alarms or sudden motor current spike was observed. positioning issues was observed in the logs. the cause of hemolysis was not determined since insufficient clinical details were provided. device in wrong position (positioning issue): the cause of positioning issue was not determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks.
3536 | 1220648-2025-48266 | 23524821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23524821&pc=OZD | 09/08/2025 |  |  | 11/11/2025 |  |  | 09/08/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 71 KG |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support there were had high purge pressure alarms. a nurse called in stating that they have attempted changing the purge cassette, tubing and traced for kinks. no visible obstruction or resolution in the alarms. purge pressure high alarm that could not be resolved through normal troubleshooting means were noted. the pump was removed, and a clot was found in the pump per surgeon and anesthesiologist. the patient was not on heparin due to an arterial perforation during percutaneous coronary intervention. the surgical wound was repaired under conscious sedation. the patient remained stable post procedure. | [5 empty] | Investigation summary: no product was returned. high purge pressure: clinical details state that there were high purge pressure alarms during the case that could not be resolved with standard troubleshooting. the pump was removed and a clot was found in the pump. the patient was not on heparin due to arterial perforation during percutaneous intervention. data log review showed a slight motor current and speed disturbance shortly before the gradual rise in purge pressure lasting about an hour. the purge pressure was not resolved. the cause of the high purge pressure was biomaterial buildup as data log review showed a gradual purge pressure rise and clinical details mentioned a lack of anticoagulation. thrombosis: in order to make a root cause determination, all of the clinical details would have to be provided. the cause of the injuries was unable to be determined due to insufficient clinical details.¿ peripheral artery dissection: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided.¿ device history review: the pump passed all post-sterile inspection checks. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿.
3537 | 1220648-2025-48263 | 23525152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23525152&pc=OZD | 08/29/2025 | 12/09/2025 |  | 11/11/2025 |  |  | 08/29/2025 | 11/11/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495044 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient on impella cp support showed bleeding on a computed tomography scan of the abdomen. there was also bleeding and clots in the lungs so anticoagulation was held. daily bronchoscopies were done to clear the clots. the pump was escalated from an impella cp to an impella 5.5. | [5 empty] | Additional information was received in the form of medical safety review and noted the following: medical safety reviewed the event and noted there is not enough information to exclude impella 5.5 pump 3 as an associated factor in the patient's outcome. impella cp pump 1 and impella 5.5 pump 2 events are serious injury. related medical device reports. this impella cp device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5 pump 2 and impella 5.5 pump 3, which is associated with the demise outcome. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H.6 code f19 was reported in error on the initial report that was submitted. a new code was added. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed, the cause was not determined since product was not returned and insufficient clinical details provided. thrombosis, in order to make a cause determination, all of the clinical details would have to be provided. the cause of the thrombosis was unable to be determined as well due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3538 | 1220648-2025-48274 | 23525236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23525236&pc=OZD | 10/12/2025 |  |  | 11/11/2025 |  |  | 10/13/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541271 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day seven of impella support, the automated impella controller battery level dropped to 51% after showing full charge and was unplugged. within a minute of being unplugged, the battery dropped to 50% with an alarm for battery level low. the automated impella controller remained plugged in to resume alternating current power. there was no loss of impella 5.5 support, and no patient harm reported. the impella 5.5 was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A4 is unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data analysis: imc logs from the complaint date confirmed the reported ¿battery level low¿ alarm (event set #23), which was triggered approximately two hours after ac power was disconnected (event set #109) was observed. the alarm cleared automatically once ac power was restored. refer to (b)(4) for the relevant imc logs. long-term (lt) power log data showed the battery¿s absolute charge dropping to 3500 and then increasing as ac power was restored. see attachments for relevant lt powerdata (ltpowerdata_(b)(6) png. device analysis the console (b)(6) was returned to fs for investigation. the reported issue was successfully reproduced by disconnecting the ac power and operating the aic solely on battery for an extended duration without reconnecting to the ac power supply. as the batteries discharged, the charge level dropped to 50%, which triggered the battery level low alarm. battery testing revealed no abnormalities; both batteries charged according to specifications, and no cell imbalances were detected. see attachments for batttest.png. communication between the pbm and the batteries was monitored throughout the tests and showed no irregularities. additionally, a thorough visual inspection of the power-related circuitry was performed, and no defects were observed. root cause: the ¿battery level low¿ alarm was triggered because the aic remained disconnected from ac power for an extended period, causing the batteries to discharge. no issues within the aic were identified. d9 updated.
3539 | 1220648-2025-48275 | 23525957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23525957&pc=OZD | 09/11/2025 | 11/11/2025 |  | 11/11/2025 |  |  | 09/12/2025 | 11/11/2025 |  |  |  |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697598 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 105 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during support of impella cp, that the patient was found to have a high plasma free hgb of 459, and urine had turned dark red. echo was done at bedside and found inlet of impella was very close to mitral valve. mds adjusted impella successfully away. measuring deep at 5.3cm. awaiting dr (b)(6) to reposition. educated bedside staff on impella management. assisted registered nurse (rn) in changing purge fluid to d5w with sodium bicarbonate. plan to rest heart and discuss with care team whether to transfer patient out for extracorporeal membrane oxygenation. impella repositioned by dr (b)(6) utilizing a bedside ultrasound. medical doctor (md) satisfied with position. the patient was transferred from (b)(6) last night, on arrival plasma free hgb was 459, urine was red, and echo showed inlet on mitral. mds were able to manipulate impella to get it into a good position last night and this morning plasma free clearing now 159. pt will be started on continuous veno-venous hemofiltration in the interim as cr has started rising. rn will recheck another plasma free hgb in an hour and make sure it continues to trend down. pt is still requiring 3 lower dose pressors so keeping impella would be beneficial. pt is being cooled and remains intubated and sedated. no alarms with impella overnight, recheck of plasma free several hours later showed large decrease in plasma free hgb. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3540 | 1220648-2025-48276 | 23526225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23526225&pc=OZD | 09/12/2025 | 11/11/2025 |  | 11/11/2025 |  |  | 09/12/2025 | 11/11/2025 |  |  | 02/28/2026 |  | 11/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473835 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Inflammation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered a placement signal issue, pump suction and that the patient had access site bleeding and right arm swelling during impella 5.5 support. it was reported that the ¿placement signal not reliable¿ alarm alarmed three times on night shift per rn (registered nurse) and that it was confirmed via impella connect. later, patient had swelling in right arm so plans were made for duplex today to assess for clot, but no further mention of clot was made. access site bleeding was also noted and a few days later oozing was noted again. the dressing was changed. oozing was also noted from incision site where graft was hanging out of body. no hematoma. finally, patient encountered suction alarms but after 2 units of packed red blood cells, the pump was able to go back up to p7. | [5 empty] | The pump was not returned for investigation. the cause of the inflammation was not determined due to insufficient clinical details. the cause of the access site bleeding issue was not determined due to lack of returned product for investigation and insufficient clinical information. the cause of the suction issue was most likely related to patient condition, based on clinical details and returned log analysis. the cause of the optical signal issue was not determined since the product was not returned. the pump passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3541 | 1220648-2025-48199 | 23508667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23508667&pc=OZD | 10/29/2025 |  |  | 11/10/2025 |  |  | 10/29/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1847282 |  | 0042-0040-US |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 125 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) displayed a ¿purge disc not detected¿ alarm, which appeared spontaneously and resolved without any intervention. the purge system parameters remained within normal limits and at baseline throughout the event. there were no reports of patient harm associated with the alarm. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h10 for the related report number.
3542 | 1220648-2025-48200 | 23508727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23508727&pc=PYX | 10/11/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026655497 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 84 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | Increase in Pressure |  |  |  |  | Pulmonary Edema | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp flex exhibited elevated purge pressure and reduced purge flow overnight, prompting administration of tissue plasminogen activator purge solution to mitigate potential thrombus formation. no alarms were triggered. additionally, the rp flex was providing higher flow than the impella cp, raising the doctor's concern for pulmonary edema. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. high purge pressure: the cause of high purge pressure was not determined due to product/data logs not being returned. saturated or high pump flow: the cause of saturated pump flow was not determined due to product/data logs not being returned. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3543 | 1220648-2025-48201 | 23508769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23508769&pc=OZD | 09/10/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673758 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 75 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock. while on support and following transfer to another facility, the patient arrested shortly after arriving to intensive care unit. the patient was resuscitated for a few minutes before arresting again and was not able to be revived after approximately 30 minutes of cardiopulmonary resuscitation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | After further review, it was determined (b)(4) is a duplicate report of (b)(4). please see (b)(4) for any additional information.
3544 | 1220648-2025-48202 | 23508879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23508879&pc=OZD | 10/11/2025 |  |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632516 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 132 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Renal Failure | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a hematoma and oozing at the impella insertion site. laboratory testing showed a partial thromboplastin time greater than 100, and the intensivist opted to continue heparin therapy with incremental dose reductions after scheduled draws. the patient¿s mean arterial pressure was labile, fluctuating between the 60's and over 100 mmhg. imaging confirmed the impella pump position and a hematoma was suspected to have affected device positioning. the physician performed bedside intervention, including repositioning of the impella, opening and cleaning the pocket, cauterizing, and applying surgicel. hemoglobin was low, prompting transfusion of two units of packed red blood cells. creatinine increased to 2.4 milligrams per deciliter, and continuous renal replacement therapy was considered if renal function did not improve. the patient was implanted with right-sided impella rp flex and expired while on this new pump. this report is one of two reports. this report represent the impella 5.5. another report will be submitted to represent the impella rp flex and the patient's death will be captured in that report. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. hemodynamic instability/ renal failure: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3545 | 1220648-2025-48204 | 23509016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509016&pc=OZD | 10/12/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during implant of impella cp the physician obtained access in right common femoral artery, peripheral angiogram showed vessels unable to accommodate impella cp. aborted placement due to patient anatomy. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. difficult to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3546 | 1220648-2025-48206 | 23509039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509039&pc=OZD | 10/12/2025 |  |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712355 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 97 KG | Hispanic |  | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Atrial Fibrillation | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp, the patient¿s labs hemolyzing overnight, down titration of p-level, urine red and output low - pump waveforms and metrics satisfactory. neurology status was improving, mixed dosing of vasopressors slowly being down titrated as well. recommendations were given to assess pump position under echocardiogram versus weaning and benefits or risks. the pump access site was stable, and the pulmonary artery catheter site was bleeding. the pump was withdrawn by one centimeter and waveforms and metrics were appropriate. during the visit, the patient went into atrial fibrillation (afib) with rapid ventricular response (rvr) and was successfully cardioverted electrically. an echocardiogram was to be done later that day. the day after, the patient continued to go in and out of afib with rvr, sometimes requiring cardioversion. severe hemodynamic instability with these runs was noted. cardiology and nephrology teams were at the patient's bedside at time of visit. the family made the decision to refuse hemodialysis and ultimately switched to comfort care. hospital staff discontinued life support devices and the patient passed peacefully thereafter. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the root cause of the injury was not determined due to insufficient clinical details; however, data log shows a decreasing placement signal trend at start with fluctuations throughout the case run. pump flow was low at the start, with instances of suction events. no other abnormalities were reported related to motor current spike or positioning issues during the case run. no conclusive evidence from the data log was found related to the hemolysis issue. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3547 | 1220648-2025-48207 | 23509130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509130&pc=OZD | 10/10/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026678030 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during support of impella 5.5 there was an episode of ventricular tachycardia (vt) and the patient was cardioverted twice and amiodarone was given to convert to normal sinus rhythm. the patient was intubated and now is extubated. an echocardiogram confirmed the impella was in stable position. the team believes ventricular tachycardia (vt) was related to ischemia not impella position. the impella was now up to p-8. lactate dehydrogenase was stable with no suction. the patient was being worked up for a left ventricular assist device (lvad). a lung module was seen on computed tomography scan. a biopsy was needed but the plan is unclear for holding ac (bivalirudin and cangrelor). vt episodes were noted during support. a large clot was noted at the inlet at time of explant to the lvad. no hemolysis or suction. the patient was well supported prior to explant and had been on cangrelor and bivalirudin. notably, pf had been trending down but was still within acceptable range and aorta ps had upward drift/was no longer correcting with non-invasive blood pressure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. impella 5.5 cannula and motor assembly returned for evaluation. the returned product was investigated. a lot of dried blood was seen on the cannula. the inflow cage was full of biomaterial. biomaterial was also seen inside the cannula around the impeller blades. since only the cannula and motor assembly was returned, further investigation of the rest of the pump could not be performed. no other abnormalities were seen. tachycardia, ischemia, thrombosis: clinical details mention that the patient suffered from ventricular tachycardia and ischemia during impella support, and a large clot was noted at the inlet during explant. the cause of the injury was unable to be determined due to insufficient clinical details. | Investigation summary: the investigation has been completed. the impella 5.5 cannula was returned for evaluation. the device was received after analysis was completed as no product returned or photo/video analysis. tachycardia/ischemia/thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
3548 | 1220648-2025-48208 | 23509149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509149&pc=OZD | 10/12/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717744 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Infection | Pain | Renal Failure | Thrombosis/Thrombus | Gastrointestinal Hemorrhage | Hemodynamic instability | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella 5.5 experienced high purge pressure alarms, which were resolved with administration of tpa; possible thrombus formation in the purge system. the patient underwent attempted femoral-popliteal bypass surgery, which was aborted due to hemodynamic instability. the patient had gastrointestinal bleeding as well as pain and mobility limitations. renal failure was noted, requiring dialysis. blood infection was documented at the femoral access site, necessitating antibiotics and surgical wound care. patient had necrotic femoral artery and iliac artery, which was repaired and the area was infected. multiple medications were administered | [5 empty] | Correction: impella access site was right axillary. the adverse events occurred at the extracorporeal membrane oxygenation (ecmo) access site - femoral artery. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3549 | 1220648-2025-48209 | 23509181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509181&pc=OZD | 10/11/2025 |  |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026718420 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Hematuria | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while on impella 5.5 support, the patient experienced ongoing bleeding, evidenced by red output from chest tubes and hematuria. the patient also had multiple runs of ventricular tachycardia and required additional hemodynamic support for blood pressure management. during the night, the patient received packed red blood cells (prbcs), cryoprecipitate, and fresh frozen plasma, with plans to transfuse additional prbcs. lidocaine bolus was administered for arrhythmia control. the patient was eventually weaned off support and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported issue has been completed. no product was returned for investigation. major bleed: patient continues to have red output from chest tubes and red in urine. the cause of the bleeding is determined to be patient condition because red output is noted from multiple sites. arrhythmia/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria/ hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details.
3550 | 1220648-2025-48205 | 23509219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509219&pc=OZD | 10/12/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734903 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 74 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella 5.5 experienced pain and hematoma at the access site and required washout. due to ongoing placement signal not reliable alarms, the doctor replaced the impella 5.5 pump | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
3551 | 1220648-2025-48211 | 23509335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509335&pc=OZD | 10/29/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/29/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 74 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) did not register the purge disc. cassettes were changed multiple times without resolve. the console was switched out. there were no reports of patient harm associated with the alarm. this report represents the first automated impella controller used. separate report(s) will be submitted for related device(s). | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3552 | 1220648-2025-48212 | 23509343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509343&pc=OZD | 10/13/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623806 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Detachment of Device or Device Component | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during removal of the impella 5.5 through the graft, the cannula snapped off just above the motor housing. the remaining portion of the cannula was visualized and successfully retrieved without any patient harm. no adverse clinical outcomes were reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. product damage (cannula detached from motor housing) the clinical details state that the cannula detached from the motor housing in the graft upon explant. it was noted that there were challenges with the removal as it seemed like the pump became caught on something but it was unable to tell where it was catching. with mild maneuvering, no excessive pulling it was able to break free. this catch was on the initial removal process but the cannula did not officially break until it was in the graft. there were no indications of clamps or anything used to cause the detachment. the returned product confirmed that the cannula was detached from the motor housing at the location of the outlet. there were no deformities in the cannula as in any clamp marks or scratches seen. the delo band remained on the cannula and part of it was torn off but did not remain on the pump housing. residual glue marks are noted on the inside of the cannula and the indentation of it being bonded to the motor housing can been seen, but there appears to be blood in between the delo band and the cannula. on the surface of the lbb (motor housing) some delo also remains as evidenced by the darker material on the surface, but it is relatively clean. there were no kinks seen on the catheter. the root cause of the product damage (cannula detached from the motor housing) is due to a material fracture but the cause cannot be further established. major bleed: the clinical stated that there was bleeding on october 9th. no further details were provided regarding the bleeding. the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3553 | 1220648-2025-48203 | 23509517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509517&pc=OZD | 10/12/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717608 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Thrombocytopenia | Hemodynamic instability | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella 5.5 support experienced an gastrointestinal bleed which required heparin to be withheld. the patient had ongoing thrombocytopenia previous to the implant which continued while on support. additionally, the impella 5.5 triggered a suction alarm and low placement signal during an episode of orthostatic hypotension while the patient was ambulating. point-of-care ultrasound confirmed stable device position. no repositioning was required. | [5 empty] | Correction: g4. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed/hypotension/hemodynamic instability/thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. optical sensor issue: the cause of the issue was not established as limited clinical information was provided, no data log abnormalities were observed and no product was returned. false alarm: the cause of the issue was not established as limited clinical information was provided, no data log abnormalities were observed and no product was returned. device history review: the device passed all the post sterile inspection checks. | Additional information added to d6b: explantation day, month and year. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3554 | 1220648-2025-48214 | 23509599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509599&pc=OZD | 10/13/2025 | 11/07/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720041 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Unspecified Infection | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, patient experienced non-sustained ventricular tachycardia, fever, and infection. broad-spectrum antibiotics were given and patient was ultimately revised to impella cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4, h6 health effect - clinical code the investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia/ infection: the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3555 | 1220648-2025-48210 | 23509617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509617&pc=OZD | 10/13/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756583 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 80 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient experienced a right common femoral artery (rcfa) hematoma while supported with impella cp. the device was explanted, and the patient required placement of a covered stent to achieve closure and hemostasis. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed / hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3556 | 1220648-2025-48215 | 23509944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23509944&pc=OZD | 10/11/2025 |  |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739735 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 105 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient presented to an outside hospital) complaining of shortness of breath and a left heart catheterization was performed which revealed significant multivessel disease. the patient was transferred to the complainant facility for an impella supported primary percutaneous coronary intervention (ppci). the cardiothoracic surgeon subsequently declined surgical intervention due to the patient's low ejection fraction. the patient was brought to the cardiac and vascular lab for an impella cp supported ppci. the patient¿s right radial artery was accessed to perform an angiography. the procedure revealed a chronic total occlusion (cto) in the right coronary artery (rca). upon attempting to engage the left coronary artery (lca), the patient's blood pressure dropped significantly. the impella cp was then inserted via the right femoral artery for mechanical support. the impella cp was placed without issue, and the ppci was completed successfully, and the patient was transferred to the intensive care unit. on support day 2, the patient was noted to have pink-red urine output and additionally it was noted that the patient was hemolyzing. a plasma free hemoglobin sample was drawn, and it was noted the results were elevated. the medical team declined to reposition the impella cp, and a formal echocardiogram was scheduled for later in the day. of note, the impella was functioning as expected with no reported alarms. the patient¿s impella access site was bleeding, and the site was re-dressed. on support day 3, the impella access site continued to bleed. heparin was decreased, and the site was again redressed with angle matching. the patient¿s urine had cleared significantly, however still had a slight pink tinge. the following day it was noted that the bleeding had resolved and the impella access site was dry, clean, and intact. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemorrhage/bleeding: bleeding was reported from the impella site during the first two days and was resolved after redressing, angle matching, and decreasing heparin levels. the cause of the bleeding issue was most likely user-related, based on the given clinical details. mechanical interaction with blood (hemolysis, hematuria): the pump was not returned for investigation. clinical details confirm hemolysis based on increased plasma free hemoglobin levels on the first day of implant, with resolution noted by the third day of use. data log shows no abnormalities related to suction, motor current spike, or positioning issue during the case run. some fluctuation in placement signal and pump flow was reported while running on a steady p-level; however, this does not indicate any specific patient condition. the cause of the hemolysis issue was not determined without returned product and sufficient clinical information. device history review: the device passed all post sterile inspection checks
3557 | 1220648-2025-48216 | 23510033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23510033&pc=OZD | 10/11/2025 |  |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758581 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support experienced two sustained runs of ventricular tachycardia overnight, one of which required a single shock for rhythm conversion. during the episode, the patient maintained a pulse but became hemodynamically unstable. lidocaine and amiodarone infusions were initiated for arrhythmia management. a transthoracic echocardiogram was performed the following morning revealed a large left ventricular thrombus, prompting continued weaning of impella support and planned device removal. dual antiplatelet therapy was continued, and a heparin infusion was started. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Investigation summary: the investigation has been completed. arrhythmia/ thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
3558 | 1220648-2025-48213 | 23510038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23510038&pc=OZD | 10/13/2025 | 12/12/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750250 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 95 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient supported with impella cp experienced intravenous thrombus formation noted around the impella catheter, starting from the graft insertion point and extending to mid-shaft, requiring systemic heparin and consideration of tpa therapy. the patient received multiple blood transfusions (2 units packed red blood cells overnight and additional units later) and bicarbonate infusion for metabolic management. episodes of hypotension were documented, requiring norepinephrine and vasopressin support, which were later weaned. patient had pink urine (hematuria) and oral bleeding. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombosis, hematuria, and major bleed has been completed. the root cause of the thrombosis and hematuria was not determined due to insufficient clinical details.the root cause of the major bleed is most likely patient condition related since excessive oral bleeding noted.
3559 | 1220648-2025-48217 | 23510171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23510171&pc=OZD | 10/11/2025 |  |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701199 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 47 KG | Non Hispanic | White | Death | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Ischemia | Ventricular Fibrillation | Thrombocytopenia | [7 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support was extremely hypovolemic with thick left ventricle walls and could not achieve flows above 2.7 liters per minute (l/min) initially, later dropping to 1.5 l/min despite support. the patient experienced progressive deterioration, including limb ischemia, acidosis, significant bleeding requiring multiple transfusions and blood products, and a drop in hemoglobin to 6 and platelets to 34. the patient developed idiopathic ventricular arrhythmias and experienced four episodes of ventricular fibrillation requiring defibrillation. despite interventions with fluids, vasopressors, amiodarone, and lidocaine, the patient expired while on impella support. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleed was unable to be determined due to insufficient clinical details. thrombocytopenia, arrhythmia, ventricular fibrillation, ischemia: the cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest, resuscitation: the cause of the injury was unable to be determined due to insufficient clinical details. patient anatomy or condition prior to therapy: the cause of the patient anatomy or condition prior to therapy was most likely patient condition related. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3560 | 1220648-2025-48218 | 23510334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23510334&pc=OZD | 10/12/2025 | 11/26/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026699490 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was a delivery issue during impella 5.5 implant. during implant, there was difficultly crossing the left ventricle. despite multiple attempts, the impella 5.5 would not cross the axillary into the aorta. the decision was made to abort the implant and proceed with the impella cp. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the unable to deliver or advance (delivery issue) cause was determined to be patient condition as the patient had difficult anatomy and resistance at axillary artery preventing successful delivery of impella. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3561 | 1220648-2025-48219 | 23510448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23510448&pc=OZD | 10/12/2025 |  |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632720 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, a hematoma developed at the femoral access site. the patient became vasoplegic, pulsatility loss, and pulseless electrical activity arrest. return of spontaneous circulation was achieved. the left ventricle collapsed and the patient was brought to the unit and support was withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety/clinical reviewed the event. a 75-year-old female with a medical history significant for peritoneal dialysis¿dependent end-stage renal disease (esrd), hypertension, diabetes mellitus, congestive heart failure, atrial fibrillation, peripheral arterial disease (pad), coronary artery disease (cad), prior nstemi, tobacco use, and asthma presented with chest pain and was diagnosed with non-st elevation myocardial infarction (nstemi) complicated by cardiogenic shock. the patient was taken to the cardiac catheterization laboratory, where significant peripheral arterial disease was noted. an intra-aortic balloon pump (iabp) was placed, and cardiothoracic surgery was consulted. the patient subsequently underwent coronary artery bypass grafting x2vessels. postoperatively in the intensive care unit (icu), the patient developed post-cardiotomy cardiogenic shock (pccs). a decision was made to proceed with axillary impella 5.5 placement for additional mechanical circulatory support. during insertion of the impella 5.5 via the axillary approach, the was an axillary artery tear. the patient required blood transfusion and surgical repair of the vessel. following repair, an impella cp was successfully implanted via the axillary artery, and the iabp was removed. a couple of hours later, a hematoma developed at the femoral access site. the patient subsequently became vasoplegic with loss of pulsatility and developed pulseless electrical activity (pea) arrest. return of spontaneous circulation (rosc) was achieved; however, left ventricular collapse was noted. the patient was transferred back to the unit, and the care was withdrawn. the patient expired. correction. complaint/patient coding was updated/corrected following medical safety/clinical review as follows: removed arrhythmia (e0601) and inserted hypotension (e2321) and cardiac arrest (e0602). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3562 | 1220648-2025-48220 | 23510598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23510598&pc=OZD | 10/12/2025 |  |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520815 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the care team stated the patient sustained a massive stroke. the patient was taken to the operating room for vascular removal of impella cp. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3563 | 1220648-2025-48221 | 23511116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511116&pc=OZD | 10/12/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730516 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was pump suction and low pump flow during impella cp patient support. while on support, the device was unable to flow above performance level 1 without suction. fluoroscopy and transesophageal echocardiogram confirmed optimal placement. volume did not resolve the suction event and the device was unable to flow above 0.5 liters per minute. the doctor decided to abort and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. low or blocked pump flow/pump suction: the logs confirm suction alarms a review of the logs confirm suction triggers are correct suggesting patient is experiencing suction at both phases of the cardiac cycle with low motor current (mc) max values and low flows during systole. logs also showed pump flows were averaging 0.35 l/min (spec: 0.0 - 0.9l/min) when at p1 and averaging 0.7l/min. the cause of the issue was unable to be determined as no product was returned for analysis and no conclusions could be drawn from clinical information and log analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Medical safety/clinical reviewed the event. a 71-year-old female with unknown prior medical history presented for a scheduled three-vessel coronary artery bypass graft (3v cabg) procedure. upon completion of the cabg and during transfer from the operating room table to the bed, the patient experienced cardiac arrest. cardiopulmonary resuscitation (cpr) was initiated. the chest was reopened, and the right coronary artery (rca) graft was found to be occluded. the graft was revised. due to ongoing hemodynamic instability, a decision was made to proceed with impella 5.5 insertion via the axillary approach. however, there was difficulty advancing the device across the aortic valve into the left ventricle. despite multiple attempts, the impella 5.5 could not be advanced from the axillary artery into the aorta. the procedure was aborted. the physician then proceeded with placement of an impella cp via right femoral artery (rfa) access, which was successfully implanted. support was initiated at p-2; however, the device was unable to provide flows above p-1 without suction alarms. fluoroscopy and transesophageal echocardiography (tee) confirmed appropriate device positioning. the patient required multiple blood product transfusions throughout the day. volume administration did not resolve the suction events, and pump flow remained limited to approximately 0.5 l/min or less. given persistent hemodynamic instability and inability to achieve adequate flow despite confirmed positioning and volume resuscitation, the physician decided to discontinue mechanical support. the patient subsequently expired. the unsuccessful attempt to implant the impella 5.5 due to inability to advance will be reported as a malfunction-type reportable event. the death will be reported against the impella cp. however, the patient¿s death is most likely attributable to the postoperative cardiogenic shock following cabg complicated by intraoperative graft failure and cardiac arrest, with ongoing hemodynamic instability despite surgical revision, transfusion support, and mechanical circulatory support. coding. patient coding was updated following medical safety and clinical review as follows: removed: device explantation (f1903) and no clinical signs, symptoms or conditions (e2403) as the patient was hemodynamically unstable and required blood transfusion. added: hemodynamic instability (e2343) and blood transfusion (f2302). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the final coding that accurately reflects the reported event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3564 | 1220648-2025-31919 | 23511263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511263&pc=OZD | 04/09/2021 | 11/10/2025 |  | 11/10/2025 |  |  | 04/09/2021 | 11/10/2025 |  |  | 02/28/2023 |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022052281 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Problem | [6 empty] | Tachycardia | Asystole | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complainant reported the patient had a coronary artery bypass graft, then decompensated. impella cp was placed for hemodynamic support. while on support impella support the patient developed ventricular tachycardia and required defibrillation, progressed into asystole requiring chest compressions. during chest compressions the impella cp became out of position in left ventricle. the impella was repositioned correctly by physician at bedside under echo guidance. the patient remained on support for three more days before the family made the decision made to withdraw care and the patient expired. it was unknown if the impella device contributed to the patient outcome. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
3565 | 1220648-2025-31920 | 23511283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511283&pc=OZD | 11/09/2022 | 11/10/2025 |  | 11/10/2025 |  |  | 11/09/2022 | 11/10/2025 |  |  | 07/31/2024 |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257388 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | While the patient was under impella cp support, the patient¿s leg was noted to have diminished and then absent pulses. it was cool/cold to the touch. the patient was escalated to the impella 5.5. the impella cp was removed from the left ventricle and left in thoracic/descending aorta until vascular surgery was completed. the vascular surgery was to continue and the return of the flow in the leg was to be monitored. | [5 empty] | The investigation of the reported failure mode has been completed. no device was received for investigation. limb ischemia: to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
3566 | 1220648-2025-48222 | 23511502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511502&pc=OZD | 10/12/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736995 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient was placed onto impella support for cardiomyopathy. while on support, slight oozing was noted at the incision site. the catheter appeared to be pulling due to being snagged. oozing began after this event. oozing occurs only with excessive coughing, otherwise, the site is dry. the patient had left hemiparesis and was found to have right middle cerebral artery infarct. the patient was taken for a thrombectomy and heparin was withheld. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Explant date added d6b updated.
3567 | 1220648-2025-48224 | 23511525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511525&pc=OZD | 10/12/2025 |  |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725121 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 91 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 experienced suction alarms associated with hypotension and aggressive crrt fluid removal. point-of-care ultrasound confirmed device position after suction events. the patient experienced a gastrointestinal bleed, which required blood transfusion and surgical intervention. medication adjustments were necessary while on support. arrhythmia (ectopy) was noted during support. access site complications included hematoma. patient was eventually weaned off support and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed/hematoma/hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: the data logs confirm continuous suction alarms #14 characteristic curve diastolic suction and real time (rt) log analysis shows alarm triggered correctly. there were no position related alarm abnormalities. the cause of the pump suction was determined to be patient condition related as evidenced by clinical detail of issue resolution by pausing continuous renal replacement therapy (crrt) volume removal and administering volume to the patient. access site adverse event: clinical reported hematoma developed at access site. the cause of the issue was not established as limited clinical information was provided. device history review: the device passed all post sterile inspection checks.
3568 | 1220648-2025-48226 | 23511643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511643&pc=OZD | 10/12/2025 |  |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026676295 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | (b)(4): (b)(6) 2025 - patient had episode of hypotension early this morning, now on epi at 0.1 with maps around 70. systemic heparin started. diuresis overnight, large bumex bolus. | [5 empty] | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received: d6b explant date. corrections: d1, d4, e1, e3, e4, g1. the investigation of the reported failure mode has been completed. no product was received for evaluation. hypotension: the cause of the injury was not determined due to insufficient clinical details. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of bleed was most likely use issue related since withholding the systemic heparin, site no longer appeared to be oozing. device history review: the device passed all the post sterile inspection checks.
3569 | 1220648-2025-48227 | 23511759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511759&pc=OZD | 10/13/2025 |  |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Tamponade | Pericardial Effusion | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was arrhythmia, cardiac tamponade, and pericardial effusion during impella cp patient support. while on support, atrial fibrillation with a rapid ventricular rate was noted. amiodarone was started. additionally, there was a large apical effusion. a tamponade was suspected. 600cc of blood came out of the chest tubes within 30 minutes. the patient went back into surgery to evacuate. two units of red blood cells were administered. the impella cp was exchanged for intra-aortic balloon pump and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no device was returned. arrhythmia : the cause of the injury was unable to be determined due to insufficient clinical details. cardiac tamponade/ pericardial effusion : the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
3570 | 1220648-2025-48225 | 23511775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511775&pc=OZD | 10/11/2025 |  |  | 11/10/2025 |  |  | 10/11/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on bipella support required initiation of dobutamine for bradycardia and low cardiac output shortly after intensive care unit arrival. despite impella cp (left side) support, the patient demonstrated signs of inadequate perfusion, including low urine output and rising creatinine. additionally, the patient experienced a run of ventricular tachycardia requiring one shock for conversion. amiodarone was administered to treat the arrythmia. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." this file is one of two related reports. this report represents the impella cp. another report was submitted to represent the impella rp flex. refer to h10 for the related report number. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary no product was returned for investigation. renal failure: the cause of the injury was not determined due to insufficient clinical details. arrhythmia/ tachycardia: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review device with sn: (b)(6) passed all post sterile inspection checks.
3571 | 1220648-2025-48228 | 23511818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511818&pc=OZD | 10/12/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026758533 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 77 YR |  | Female |  | Hispanic | White | Death | Yes | No | Use of Device Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Impella placed independently. solution used to prime device noted to be d5w + ns when local team arrived to support. total run time on solution was 1 hour. patient expired. | [5 empty] | Corrections: d1, d3, d4, e3, e4, g1, g4, g6. the investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. use against labeling: the cause of the use against labeling was a misuse issue since the incorrect purge solution was used. device history review: the device passed all the post sterile inspection checks. | Product complaint # (b)(4). the impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3572 | 1220648-2025-48230 | 23511885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511885&pc=OZD | 10/12/2025 | 12/29/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025592345 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Failure to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon insertion of axillary impella 5.5, the axillary artery tore. the patient needed a blood transfusion and surgical repair of the vessel. the axillary artery tore just distal to anastomosis during insertion. attempt to advance was done and the axillary artery was dissected. the decision to abort the impella 5.5 placement, repair the vessel and implant an impella cp via the axillary artery was made. additional information was received that the bleeding was from a skin tract ooze as identified by surgery. the pump is not available for return. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the reported vascular dissection and delivery issue has been completed. the root cause of the deliver issue was determined to be patient condition as the patient had resistance at anastomosis preventing successful delivery of impella.the root cause of vascular dissection cannot be determined since insufficient clinical details were provided. | Medical safety/clinical reviewed the event. a 75-year-old female with a medical history significant for peritoneal dialysis¿dependent end-stage renal disease (esrd), hypertension, diabetes mellitus, congestive heart failure, atrial fibrillation, peripheral arterial disease (pad), coronary artery disease (cad), prior nstemi, tobacco use, and asthma presented with chest pain and was diagnosed with non-st elevation myocardial infarction (nstemi) complicated by cardiogenic shock. the patient was taken to the cardiac catheterization laboratory, where significant peripheral arterial disease was noted. an intra-aortic balloon pump (iabp) was placed, and cardiothoracic surgery was consulted. the patient subsequently underwent coronary artery bypass grafting x2vessels. postoperatively in the intensive care unit (icu), the patient developed post-cardiotomy cardiogenic shock (pccs). a decision was made to proceed with axillary impella 5.5 placement for additional mechanical circulatory support. during insertion of the impella 5.5 via the axillary approach, the was an axillary artery tear. the patient required blood transfusion and surgical repair of the vessel. following repair, an impella cp was successfully implanted via the axillary artery, and the iabp was removed. a couple of hours later, a hematoma developed at the femoral access site. the patient subsequently became vasoplegic with loss of pulsatility and developed pulseless electrical activity (pea) arrest. return of spontaneous circulation (rosc) was achieved; however, left ventricular collapse was noted. the patient was transferred back to the unit, and the care was withdrawn. the patient expired. note: h6. component code and investigation type/findings/conclusion coding remain unchanged as previously reported. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3573 | 1220648-2025-48229 | 23511999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23511999&pc=OZD | 10/13/2025 | 12/31/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725988 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male | 117 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella 5.5 the patient had short runs of ventricular tachycardia, lido increased. no additional information. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1960432. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3574 | 1220648-2025-48232 | 23512007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512007&pc=OZD | 10/13/2025 | 01/06/2026 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697852 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient had a small hematoma at the access site. manual pressure and tightening of perclose were done to manage the complication. a few hours later the pump was explanted and the patient survived. | [5 empty] | Investigation summary: hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot : 1951035. device history batch: sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3575 | 1220648-2025-48233 | 23512266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512266&pc=OZD | 10/13/2025 |  |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761334 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 78 YR |  | Female | 117 KG | Non Hispanic | Black Or African American | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | Unspecified Infection | Sepsis | Hematuria | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was arrhythmia, sepsis, infection, hemolysis and hematuria during impella cp patient support. while on support, the patient experienced atrial fibrillation. there were concerns for hemolysis verbalized. the patient¿s urine was red. the pump¿s position was confirmed by echocardiogram. the patient had pneumonia and was likely more septic related than cardiogenic. the patient expired on support after being septic for several days. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered near the end of the case. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. arrhythmia/sepsis/infection: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3576 | 1220648-2025-48236 | 23512350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512350&pc=OZD | 10/12/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755470 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. pump was implanted, but excessive bleeding occurred upon removal of the introducer. the pump was removed but could not be loaded back onto the guidewire (wire could not get through pigtail). a new impella was opened and implanted in the patient. | [5 empty] | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the delivery issue was unable to be determined as insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3577 | 1220648-2025-48237 | 23512391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512391&pc=OZD | 10/13/2025 |  |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646615 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Decrease in Pressure | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was placed onto impella cp support due to heart failure/cardiogenic shock associated with heart transplant rejection. while on support, the patient was placed onto dialysis. the impella cp presented ongoing purge pressure low alarm, which was noted briefly prior to scheduled explant. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. purge pressure low: data log review confirms purge pressure low alarm. purge pressure dropped from around 400 mmhg to ~200 mmhg over approximately 20 minutes with purge flow increasing from about 15 to 30 ml/hr. the cause of the low purge pressure was not determined since product was not returned. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3578 | 1220648-2025-48234 | 23512435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512435&pc=OZD | 10/14/2025 | 01/02/2026 |  | 11/10/2025 |  |  | 10/14/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741824 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was cardiac arrest during impella cp patient support. return of spontaneous circulation was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned for investigation. cardiac arrest: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot device lot: 1977171. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3579 | 1220648-2025-48239 | 23512505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512505&pc=OZD | 10/14/2025 |  |  | 11/10/2025 |  |  | 10/14/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686527 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 98 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was hematoma, major bleed, and cardiac arrest during impella 5.5 patient support. while on support, the patient was developing a hematoma on the right chest wall and right side of the abdomen. there was a ct scan overnight. two units of packed red blood cells were given and there was minimal oozing at the impella site. systemic heparin was put on hold. the patient coded and passed away. | [5 empty] | The cause of the bleeding and hematoma was not determined since product was not returned and insufficient clinical details provided. the cause of the cardiac arrest was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3580 | 1220648-2025-48238 | 23512508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512508&pc=OZD | 10/13/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698712 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 85 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient¿s groin site was bleeding despite angle match and appropriate dressing. the patient was transferred back to a room in the coronary care unit. urine was clear yellow. they were waiting for the fellow to do a bedside echocardiogram and orders were placed. a full panel of labs was pending. the groin site was bleeding mildly. the medical doctor came to the intensive care unit and perclose was tightened. bleeding was resolved. the patient was successfully weaned from support and the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury is most likely use related since md tightened perclose and bleeding resolved. device history review: the device passed all post sterile inspection checks.
3581 | 1220648-2025-48240 | 23512623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512623&pc=OZD | 10/13/2025 |  |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support experienced tachy-arrhythmia, and an attempt was made to correct the rhythm using electrical cardioversion. the cardioversion was unsuccessful. the device was replaced with impella 5.5. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3582 | 1220648-2025-48241 | 23512735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512735&pc=OZD | 10/13/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698923 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 54 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed onto impella cp support for high-risk percutaneous coronary intervention. while implanting the impella cp and during high-risk percutaneous coronary intervention, the patient went into pulseless electrical activity arrest and required an increase in dobutamine drops to 15 and a dose of epinephrine. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. minor bleed/ hematoma: after transitioning from peel-away to repositioning sheath, site bleeding was observed followed by small hematoma formation. hematoma managed in lab. noted that activated clotting time (act) was 344 when in icu. the root cause of the bleeding and hematoma was a use related issue as patient act was still high after implant when in icu. cardiac arrest: clinical details noted that patient went into cardiac arrest. the root cause of the injury was unable to be determined due to insufficient clinical details.
3583 | 1220648-2025-48242 | 23512864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23512864&pc=OZD | 10/13/2025 |  |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694698 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was supported with an impella cp and experienced absent palpable pulses in both legs, requiring doppler assessment, which was unsuccessful. the first set of laboratory results was hemolyzed and had to be redrawn. a bicarbonate drip was ordered. the patient was eventually weaned off support and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3584 | 1220648-2025-48243 | 23513250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23513250&pc=OZD | 10/13/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026760781 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the decision for impella cp was made on a patient with multiple vessel disease. the cp was successfully placed and started in auto however flows were low with suction alarm. decision was made to explant and place a new cp. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: data logs confirmed low pump flow when pump started and remained to the rest of the case. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. visual inspection found a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. low or blocked pump flow: the cause of low or blocked pump flow was most likely kinked cannula as per clinical information and observed on the returned product. product damage (cannula kink): the cause of the product damage (cannula kink) cannot be determined as insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d3, e4, h6. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated
3585 | 1220648-2025-48244 | 23513425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23513425&pc=OZD | 10/13/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731303 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 90 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to cardiomyopathy. while on support, patient may have experienced a stroke but she continued to follow commands. patient intubated out of concern for protecting her airway. doctor noted possible transient ischemic attack | [5 empty] | Correction: e4, h6 health effect - clinical code e0509 was inadvertently omitted. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia/ stroke - the cause of the stroke was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3586 | 1220648-2025-48245 | 23518679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23518679&pc=OZD | 10/13/2025 | 11/10/2025 |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a large hematoma developed at the right groin access site after impella removal. the physician initially had difficulty achieving adequate compression, but the patient remained hemodynamically stable. serial hemoglobin and hematocrit checks did not indicate a significant drop, and no transfusion or surgical intervention was required. the hematoma was managed with manual compression. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the bleeding and hematoma were unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
3587 | 1220648-2025-48246 | 23518682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23518682&pc=OZD | 10/12/2025 |  |  | 11/10/2025 |  |  | 10/12/2025 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527694 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was escalated to an impella 5.5 from an impella cp for mechanical circulatory support. the patient presented to the emergency department with atrial fibrillation with rapid ventricular response, and the patient was taken to the cardiac catheterization lab (ccl). in the ccl, the patient underwent a percutaneous coronary intervention (pci) to the left anterior descending (lad) artery, involving the placement of four drug-eluting stents. during the procedure, the patient required hemodynamic support with levophed. due to the patient's condition, a decision was made to place both an impella cp for mechanical circulatory support and a swan-ganz catheter for advanced hemodynamic monitoring. the impella cp was inserted without issue and support was initiated. due to the patient's large pannus, the peel-away sheath was removed slowly in incremental segments. during this process, the clear plastic retention ring from the 14 fr dilator became lodged around the heart pump catheter. during the removal of the peel-away sheath, a placement signal low alarm occurred, shortly followed by the pump migrating fully into the ventricle. fluoroscopy confirmed the pump had curled within the ventricle. the physician attempted to slowly pull the catheter back, but the pump completely exited the ventricle into the aorta. efforts to reposition the pump failed as the physician was unable to recross the valve. the pump was then fully removed, flushed with heparinized saline, and reinserted using a short sheath. the second insertion attempt was successful and completed without further issue. a reperfusion sheath was placed, and the device was locked in position at 92 cm depth and set to performance level p6. the patient was transferred to the cardiac unit for rest and recovery. there were slight oozing and hematoma noted in the right groin access site, however no additional oozing was noted at the impella access site after dressing was removed. manual pressure was applied to the groin, and the hematoma was reduced. it was noted that the physician was at the bedside to address the patient¿s low cardiac index of 1.4 and low mixed venous oxygenation of 45%. the physician ordered the patient to remain off anticoagulation, and the impella purge fluid was changed to 5% dextrose in water with bicarbonate. the following day the patient¿s lactic acid increased from 2.1 mmol/l to 4 mmol/l with an acute kidney injury noted. the decision was made to escalate the patient¿s therapy to an impella 5.5. the patient was brought to the operating room where the impella 5.5 was successfully placed, and the impella cp was explanted without issue. on support day 10 with impella 5.5, the patient experienced episodes of atrial fibrillation. the patient had been extubated the day prior but the patient¿s requirement for levophed increased. the medical team reported motor noise from the impella; however, the impella¿s function was not reported to be affected by the noise. the impella 5.5 was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complications. the patient's outcome at the time of explant was survived. this complaint involves three impella devices: pump 1: impella cp with serial number (b)(6), 14fr introducer with serial number (b)(6), used with impella cp, pump 2: impella 5.5 with serial number (b)(6). this mdr specifically addresses pump 1: impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: h6 health effect - clinical code 0506 was inadvertently omitted on the initial manufacturer device report.
3588 | 1220648-2025-48248 | 23518707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23518707&pc=OZD | 10/14/2024 |  |  | 11/10/2025 |  |  | 10/13/2025 | 11/10/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025506819 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient was placed onto impella support due to post-cardiotomy cardiogenic shock. while on support, patient experienced hemolysis. 1 packed red blood cells was given. | [5 empty] | Correction: e4; h6 medical device problem code 01 was inadvertently omitted on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. mechanical interaction with blood: after 7 days on support, labs show free hgb at 380. however, lab noted issue with the sample and requested a redraw. 2 days later, hgb noted to be less than 30, as was shown by every sample taken from the patient throughout support. there was no mechanical interaction with blood due to the field reporting a false positive in the sample that had high free hgb.
3589 | 1220648-2025-31921 | 23518753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23518753&pc=OZD | 10/26/2023 | 11/10/2025 |  | 11/10/2025 |  |  | 10/26/2023 | 11/10/2025 |  |  |  |  | 11/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1719164 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was supported by the impella 5.5 and while on support, there were placement signal issues. the placement signal monitoring was disabled due to issues with the fiber optic cable. support continued. the investigation team determined that the optical signal issue was likely due to the automated impella controller (aic). no known patient harm or injury has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the data logs were reviewed and a temporary placement signal not reliable (psnr) with intermediate spike in placement signal, gain and contrast was observed. based on the rt log, the first psnr was observed at 02:12 pm on (b)(6) 2023 and other 9 psnr alarms were noted in the span of 1 hour and 12 minutes. further analysis of the rt logs showed that the snr value do not vary even though there is a variation in the raw optical sensor reading. snr and contrast is all over the place with snr avg/min between 2500/0 and contrast following similar pattern to snr. no other significant issues with gain/lamp/light. the first ps adjust (-150 mmhg) was done when ps was close to 300 mm hg. post the first ps adjust ps was dropped back to less than 100 mmhg and slowly started to rise later and lost completely maxing to 3000 mmhg. later again ps adjust of -125 mmhg was done and corrected back values below 100 mhg and started to drift downwards later. furthermore, the observed trend of the 5s parameters - snr, light, gain, lamp and contrast do not correlate with any of the failure modes from the optical failures. ic10472 was not returned for further investigation and another complaint investigation (b)(4) was opened for a similar placement signal issue. device analysis: product was not returned. product damage (sterile sleeve) . per the clinical, the back of the anti-contamination sterile sleeve compromised and the hospital staff placed tape on it; to hold it. there was no mention of how this damage happened. thus, the root cause of the product damage (sterile sleeve) is unable to be determined as no product was returned and provided clinical details are insufficient. root cause: thus, the root cause of the optical signal issue is unable to be determined as no product was returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3590 | 1220648-2025-31917 | 23495431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495431&pc=OZD | 08/10/2023 | 11/07/2025 |  | 11/07/2025 |  |  | 08/10/2023 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1151537 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that once support was initiated a placement signal not reliable alarm occurred. the console was exchanged for a different one and the alarm was no longer present. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show large number of error messages related to invalid placement signal samples. console inspection also revealed evidences of corrupt microsd card. placement signal not reliable issues were caused by a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3591 | 1220648-2025-48103 | 23495677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495677&pc=OZD | 10/10/2025 | 11/05/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734943 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Pain | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced arm pain and ventricular tachycardia. the patient was treated with dilaudid, antitachycardia pacing, amio, and lido. impella support continues. | Additional information was received. the registered nurse called to report that a purge system blocked alarm and purged pressure high alarm were displayed on 17-jan-2026. five percent dextrose in water with 25 meq/liter was purged at that time. it was recommended to first switch to (b)(4) units per milliliter heparin purge, which was done. they intend to switch to alteplase purge if the heparin was not successful. they also are considering p-level alteration. it was recommended to closely monitor for rising motor current as a sign of pump failure. a plan will be made in case of pump stop. |  |  |  |  | The cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 added selection due to new information received. b5 additional information was added. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. h6 medical device problem code was changed due to new information received. h10 related report numbers were added. this is one of three related reports. this report represents the pump and other reports were submitted for the related automated impella controllers.
3592 | 1220648-2025-48105 | 23495680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495680&pc=OZD | 10/15/2025 | 11/05/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 112 KG | Non Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a notification on the supporting automated impella controller indicating battery life was lower than expected. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The console was returned and the reported issue could not be reporduced. data logs were reviewed and confirm the complaint. the reported battery issue was likely due to a momentary communication glitch between the power battery manager and the batteries.
3593 | 1220648-2025-48173 | 23495708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495708&pc=OZD | 10/09/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742626 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male | 109 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced the pump pulled back into the aorta during patient movement. the pump was surgiacally re-advanced into the left ventricle. later, the pump was explanted and exchanged for an impella 5.5. | [5 empty] | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position into aorta by the patient during large bath turn. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3594 | 1220648-2025-48174 | 23495720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495720&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731316 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 80 KG | Hispanic |  | Injury | Yes | No | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure and had a thrombus in the left ventricle. tissue plasminogen activator was added to the purge with little resolution. the patient underwent a washout and the pump was explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3595 | 1220648-2025-48178 | 23495724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495724&pc=OZD | 10/09/2025 |  |  | 11/07/2025 |  |  | 10/11/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743286 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 89 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Valvular Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, the patient had mitral and tricuspid regurgitation. pump was turned down. patient was also placed onto dialysis. the patient was later explanted and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. heart valve incompetence: the cause of the injury could not be determined due to insufficient clinical details and no product returned. renal failure: the cause of the injury could not be determined due to insufficient clinical details and no product returned. device history review: the device passed all post sterile inspection checks.
3596 | 1220648-2025-48176 | 23495731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495731&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690030 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Great Vessel Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after explantation of an impella cp a hematoma developed in the groin and perforation of the superficial femoral artery. a balloon tamponade was attempted several times without success. the artery was repaired surgically; two units of red blood cells were transfused and levophed was increased. the patient survived. | [5 empty] | No product was returned for investigation. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of vessel perforation was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3597 | 1220648-2025-48177 | 23495735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495735&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739649 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | [6 empty] | Anemia | Arrhythmia | Ventricular Fibrillation | Hypovolemia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an unreliable placement signal. the patient was hypovolemic and had ventricular tachycardia and was treated with magnesium. later, the patient also experienced atrial fibrillation. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. optical sensor issue: the clinical details state that a patient was requiring volume resuscitation as they were hypovolemic and the ps was reading half of the arterial line even though it was appropriately zeroed at case start. the data logs show that upon initial plug in of the pump the mean placement signal was around 70mmhg. shortly after the pump started the values became lower with systolic values of 55mmhg and diastolic of about 30mmhg. there were many instances of low-pressure (ps) low alarms triggered during this time and then again on 10/12. stable around 8000, with contrast about 25% throughout the whole case. there were no instances of a hard failure. due to a lack of product returned, the root cause of the optical sensor issue cannot be established. anemia/hypovolemia: the cause of the injury was not determined due to insufficient clinical details. ventricle fibrillation/arrythmia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: impella passed all post sterile inspection checks.
3598 | 1220648-2025-48180 | 23495785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495785&pc=OZD | 10/22/2025 | 12/30/2025 |  | 11/07/2025 |  |  | 10/22/2025 | 11/07/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599934 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the impella cp had a leak on the purge filter. the team found its own solution to maintain the purge flow. the patient is still on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary. purge sidearm returned for investigation. the returned sidearm was visually inspected and the tubing off the filter was observed to be cut. possible purge sidearm bypass performed in the field to resolve the leak. the joint between the filter and the reservoir was observed to be fractured. the cause of the purge leak was a damaged filter to reservoir joint likely due to an unintended use error. device history review: pump passed all post sterile inspection criteria. | D6b added. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3599 | 1220648-2025-48181 | 23495805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495805&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  | 11/30/2025 |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423421 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 100 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the emergency department with complaints of chest pain and a history of drug abuse. the decision was made to implant an impella cp and transfer the patient to another hospital for advanced therapies. prior to impella placement, the physician elected to place a femoral-to-femoral bypass due to the patient¿s vessel size. the impella was implanted, and support was initiated with no complications. the patient was successfully transferred to another hospital and upon arrival the patient was cannulated with veno-arterial extracorporeal membrane oxygenation (va ecmo). on day 2 of impella cp support, the patient experienced a hematoma in the right groin where the impella access site was located. the patient was transfused with 1 unit of packed red blood cells, cryoprecipitate, a unit of platelets and 6 units of fresh frozen plasma. a femstop was placed to control the bleeding and heparin was withheld, and hemostasis was achieved. on day 5 of impella cp support, the medical team decided to escalate the patient to an impella 5.5 as a bridge to wean the patient from va ecmo. prior to the impella 5.5 escalation, the patient was cardioverted from atrial fibrillation with rapid ventricular response to normal sinus rhythm. the impella 5.5 was implanted without complications and the impella cp was explanted. the patient was then weaned and decannulated successfully from va ecmo. on day 4 of impella 5.5 support, the patient experienced multiple bradycardic episodes overnight which was treated with atropine. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: pump 1: impella cp, with serial number 472213 pump 2: impella 5.5, with serial number 599111 this mdr specifically addresses pump 1: impella cp, with serial number 472213, which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> no product was returned for investigation. hematoma: patient had a hematoma in the right groin. the cause of hematoma was determined to improper anticoagulation management due to high act values after transferring to icu. issue was managed by stopping heparin and placing femstop at the site. device history lot ==> device lot: 1752998 device history review ==> device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3600 | 1220648-2025-48182 | 23495808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495808&pc=OZD | 10/22/2025 | 01/06/2026 |  | 11/07/2025 |  |  | 10/22/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542869 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was switched to surgical mode to perform a bypass operation. immediately after the aorta was clamped, the impella displayed purge system blocked alarm. after a successful bypass operation, the pump was returned to the p-level. the purge system blocked alarm remained present, and no meaningful flow through the purge system was observed. due to the risk of an impending pump stop and because the patient appeared hemodynamically stable after successful revascularization, the decision was made to explant the pump in the operating room. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3601 | 1220648-2025-48183 | 23495916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23495916&pc=OZD | 10/11/2025 |  |  | 11/07/2025 |  |  | 10/11/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729348 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Sepsis | Respiratory Failure | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient had a mitral valve replacement (b)(6) 2025 during which the patient had a dissection related to one of the cardiac bypass cannulas. the patient required a large amount of blood products and was cannulated for central veno-arterial extracorporeal membrane oxygenation (va ecmo). the following day the patient was taken back to the operating room for chest closure and decannulation of va ecmo. the patient continued to decline hemodynamically, requiring multiple high doses of inotropes and vasopressors. the following day, the patient was taken to the operating room for impella 5.5 placement. during preparation of the device for implant, it was reported that the automated impella controller (aic) displayed an error message stating the console was not configured to work with the optical sensor. the aic was power cycled; however, the aic would not reboot and displayed a system self-check failed message. the aic would then not turn off and made a continuous noise. the battery was turned off to shut the system down. the aic was exchanged for a back up aic and the 5.5 was successfully implanted and support initiated. on day 3 of impella 5.5 support, the patient became hypoxic during a bronchoscopy. methylene blue was administered, and vasopressor dosage was increased. the impella was unable to flow higher than p7 without suction alarms. the impella position was verified, and the patient¿s right heart was suggested to be investigated for dysfunction. the following day, the patient¿s code status was changed to ¿do not resuscitate¿. the patient was determined to not be a candidate for advanced therapies. the patient was noted to be showing signs of sepsis. the impella was unable to flow above p4 and the patient had poor oxygenation. the medical team declined investigating the patient¿s right heart function as they believed the patient¿s outcome would be unchanged. the patient continued to decline and expired while on support. this complaint involves two impella devices: impella 5.5 with serial number (b)(6) automated impella controller with serial number (b)(6) this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary : respiratory failure/sepsis: the cause of the injury was not determined due to insufficient clinical details. low pump flow: clinical reported inability to flow at higher p-levels. the product was not returned for investigation. data log review confirms suction and flow levels slightly below spec at higher p-levels (3.7l/min at p-7). placement signal is seen trending down gradually throughout the case. on the last day, placement signal began to decrease rapidly from 70mmhg to <30mmhg. on this day, clinical details stated the patient's condition was declining. the cause of the low flows was most likely due to the patients condition, as placement signal was trending down throughout the case and clinical details implied that the patient was decompensating. device history lot : device lot number: 1962574. device history review : pump sn (b)(6) passed all post sterile inspection criteria. | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3602 | 1220648-2025-31918 | 23496363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23496363&pc=OZD | 11/03/2024 | 11/07/2025 |  | 11/07/2025 |  |  | 11/12/2024 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124582 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a 5.5 pump being used to support a patient admitted in with a cardiac history and having a bentall procedure. the pump was placed but lost reliable placement signal (ps). the pump remained on for support despite the lost ps. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause for the psnr alarm due to oscillating contrast was most likely due to the defective lamp. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3603 | 1220648-2025-48184 | 23496382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23496382&pc=OZD | 10/16/2025 |  |  | 11/07/2025 |  |  | 10/16/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682536 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 100 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with an impella cp was escalated to an impella 5.5 device for mechanical circulatory support. the patient presented to the emergency department with complaints of chest pain and a history of drug abuse. the decision was made to implant an impella cp and transfer the patient to another hospital for advanced therapies. prior to impella placement, the physician elected to place a femoral-to-femoral bypass due to the patient¿s vessel size. the impella was implanted, and support was initiated with no complications. the patient was successfully transferred to another hospital and upon arrival the patient was cannulated with veno-arterial extracorporeal membrane oxygenation (va ecmo). on day 2 of impella cp support, the patient experienced a hematoma in the right groin where the impella access site was located. the patient was transfused with 1 unit of packed red blood cells, cryoprecipitate, a unit of platelets and 6 units of fresh frozen plasma. a femstop was placed to control the bleeding and heparin was withheld, and hemostasis was achieved. on day 5 of impella cp support, the medical team decided to escalate the patient to an impella 5.5 as a bridge to wean the patient from va ecmo. prior to the impella 5.5 escalation, the patient was cardioverted from atrial fibrillation with rapid ventricular response to normal sinus rhythm. the impella 5.5 was implanted without complications and the impella cp was explanted. the patient was then weaned and decannulated successfully from va ecmo. on day 4 of impella 5.5 support, the patient experienced multiple bradycardic episodes overnight which was treated with atropine. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6) pump 2: impella 5.5, with serial number (b)(6) this mdr specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned for investigation. bradycardia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the bradycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1929987. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3604 | 1220648-2025-48185 | 23497427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23497427&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731183 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to heart transplant. the patient presented to an outside hospital on (b)(6) 2025 with complaints of shortness of breath and had a pleural effusion. the patient was then transferred to the complainant hospital, where the patient went into atrial fibrillation with rapid ventricular response, was cardioverted but continued to decompensate despite multiple vasopressors. a femoral intra-aortic balloon pump (iabp) was placed. on (b)(6) 2025, the iabp was moved to an axillary insertion due to positioning issues but was later moved back to the femoral placement. the medical team decided to escalate the patient to an impella 5.5 for increased support. the impella 5.5 was implanted without issues, and the iabp was weaned and explanted. the patient was transferred to the intensive care unit following the implant of impella 5.5 in stable condition. on support day 2, the patient developed a hematoma post operatively. a pressure dressing was applied, and the hematoma was monitored. the following day, bivalirudin was withheld due to the hematoma. on support day 6, it was noted that the patient required cardioversion overnight. on support day 9, during the patient¿s daily walk the impella moved out of the left ventricle and into the aorta. the patient was brought to the cardiovascular operating room to replace the impella 5.5 with a new impella 5.5. the new impella 5.5 was inserted without issues and following support and site closure, the patient was transferred to the icu in stable condition. at the time of writing this report, the patient continues to be supported by the replacement impella 5.5 while awaiting a donor heart for transplantation | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. no product was returned. hematoma: the cause of the hematoma was unable to be determined due to insufficient clinical details. paroxysmal atrial fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3605 | 1220648-2025-48187 | 23498258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23498258&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740520 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient was admitted the previous day after experiencing a right coronary artery st-elevation myocardial infarction. further examination revealed a chronic total occlusion in the left anterior descending (lad) artery. the medical team decided to bring the patient to the cardiac catheterization lab for a primary percutaneous coronary intervention (ppci) to address the blocked lad artery with impella cp support. the impella cp was successfully inserted via the right femoral artery. following the successful procedure, the impella cp was weaned and removed without complications. the physician utilized a dry closure technique; however, during a subsequent balloon inflation above the peel-away sheath, an iliac artery perforation was noted. a covered stent was placed to repair the perforation, but the stent was not adequately sized and began to migrate. vascular surgery was immediately contacted for assistance. the patient's condition rapidly deteriorated, requiring immediate administration of multiple vasopressors (neo-synephrine, dopamine, and levophed) as mean arterial pressures dropped into the 30s mmhg. medical staff placed another covered stent, after which the patient's condition stabilized. the peel-away sheath remained in place during the stabilization efforts. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product returned. post procedure adverse event (vessel perforation, hypotension) - after impella removal, md closure with ballooning above peel away sheath, but iliac perforated. pt decompensated but eventually stabilized with covered stent. the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1970208. device history batch: subcomponent lot: n/a. device history review: pump s/n (b)(6) passed all post sterile inspection checks.
3606 | 1220648-2025-48186 | 23498276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23498276&pc=OZD | 10/09/2025 | 11/07/2025 |  | 11/07/2025 |  |  | 10/09/2025 | 11/07/2025 |  |  | 10/31/2025 |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406460 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 11 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Plan to give platelets now to mitigate surgical oozing, start bival overnight. | The us complainant had placed impella 5.5 to support the pediatric patient admitted in with extensive known cardiac and surgical history. the pump was placed and there was surgical bleed that prompted the team to infuse platelets and heparin was on/off throughout the support. the pump remained on for 15 days until explanted at the time of heart transplant. |  |  |  |  | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b5, d1, d3, d4, e4, g1. the investigation of the reported failure mode has been completed. no product was received for evaluation. no signs, symptoms or patient involvement: plan to give platelets now to mitigate surgical oozing, start bival overnight. the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3607 | 1220648-2025-48188 | 23498352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23498352&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701186 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk for surgical intervention was implanted with an impella 5.5 pump for mechanical circulatory support. the patient presented with a st-elevation myocardial infarction. a cardiac catheterization revealed severe blockages in the right coronary artery, left main, left anterior descending, and circumflex arteries, prompting an emergent quadruple coronary artery bypass graft procedure. following surgery, the patient was successfully weaned from cardiopulmonary bypass (cpb) but became hemodynamically unstable approximately 20 minutes later, experiencing a sudden drop in blood pressure. the surgical team resumed cpb and successfully placed an impella 5.5 for mechanical circulatory support. at the time of implant, the patient¿s ejection fraction was less than 20% and the patient¿s pressures were managed with epinephrine and levophed. on support day 4, the patient experienced atrial fibrillation and two episodes of ventricular tachycardia within one hour, requiring defibrillation with one shock per episode. the patient was started on lidocaine and amiodarone infusions. the clinical team noted a prolonged qt interval and initiated a cooling protocol. the following day it was noted that the patient¿s care was withdrawn and the patient expired. | [5 empty] | Investigation summary : arrhythmia/ tachycardia : the cause of the injury was unable to be determined due to insufficient clinical details. device history lot : device lot : 1937708. device history review : the pump s/n: (b)(6) passed all the post sterile inspection checks. | Given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3608 | 1220648-2025-48189 | 23498432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23498432&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 1172893 |  | 0042-0000-US |  | No |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Stemi nurse called at 0600 to inform me a ¿controller error¿ alarm appeared. instructed to perform an automated impella controller (aic) to aic transfer. | [5 empty] | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3609 | 1220648-2025-48190 | 23498777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23498777&pc=OZD | 10/11/2025 |  |  | 11/07/2025 |  |  | 10/11/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026700861 |  | 0048-0047 |  | No |  |  | No |  |  |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 55 YR |  | Male | 120 KG |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Impairment | [11 empty] | Required Intervention | Life Threatening | [27 empty] | Initial submission | Followup | [8 empty] | Icu decided impella too deep and causing arythmias / pulled back and got red alarm impella in the aorta. doctors on call repositioned under toe & resolved red alarm / pump secured & stable. | [5 empty] | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. arrhythmia: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. acute kidney failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
3610 | 1220648-2025-48192 | 23498984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23498984&pc=OZD | 10/10/2025 | 11/07/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702558 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 88 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was being escalated from a impella cp to a 5.5. the escalation was successful; however, it was noted that post impella 5.5 placement and cp removal, the decision was made to check the femoral artery. small micro puncture access was obtained and images were taken. a large clot in the superficial femoral artery/profunda bifurcation was noted. an open thrombectomy procedure was completed and the groin site was closed surgically. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4 . the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3611 | 1220648-2025-48193 | 23499095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23499095&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725262 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 139 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient presenting with worsening end organ function for an impella 5.5 as a bridge to a durable left ventricle assist device or for donor heart transplant. the impella 5.5 was implanted via the right axillary artery and support initiated without issue. on day 3 of support, the patient reported arm pain. the following day the patient experienced episodes of ventricular tachycardia (vt) overnight. on day 6 of support, the patient experienced additional episodes of vt which required treatment with a bolus and continuous infusion of lidocaine. on day 20 of support, the patient experienced further vt and treatment with amiodarone was initiated. at the time of writing this report, the patient continues to be supported by impella 5.5 while awaiting a donor heart for transplantation. | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient presenting with worsening end organ function for an impella 5.5 as a bridge to a durable left ventricle assist device or for donor heart transplant. the impella 5.5 was implanted via the right axillary artery and support initiated without issue. on day 3 of support, the patient reported arm pain. the following day the patient experienced episodes of ventricular tachycardia (vt) overnight. on day 6 of support, the patient experienced additional episodes of vt which required treatment with a bolus and continuous infusion of lidocaine. on day 20 of support, the patient experienced further vt and treatment with amiodarone was initiated. at the time of writing this report, the patient continues to be supported by impella 5.5 while awaiting a donor heart for transplantation updated (b)(6) 2025: the patient was successfully bridged to a heart transplant. the impella 5.5 was weaned and explanted without complication, and the patient survived. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This follow-up report is being submitted to update date of explant, patient outcome and to report investigation findings. b5 is being updated to include new information that was not included in the initial report. the revised b5 provides a complete event narrative. d6b: explant date added. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. pain: the cause of the clinical symptom was not determined due to insufficient clinical details. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1960411. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3612 | 1220648-2025-48191 | 23499103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23499103&pc=OZD | 10/10/2025 | 11/07/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758126 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision was made for high risk percutaneous coronary intervention(pci) with impella cp after diagnostic workup revealed multi vessel disease on a patient. the impella was successfully placed. during the pci it was noted that they post dilated the mid right coronary artery stent and patient had st changes as well as became hypotensive. levophed was started and echocardiogram was done to check for effusion. no effusion was seen. slow flow on angiogram resolved with ic nitro and adenosine. patients electrocardiogram normalized and levophed was weaned off. the physician stating it was a reperfusion arrythmia. impella was weaned and explanted without issue. | [5 empty] | Correction e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the cause of the injury was unable to be determined due to insufficient clinical details. hypotension : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3613 | 1220648-2025-31872 | 23500194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23500194&pc=OZD | 11/07/2024 | 11/07/2025 |  | 11/07/2025 |  |  | 11/14/2024 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815750 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella support, and, during support, there were placement signal not reliable alarms. support continued. the investigating engineer requested that the automated impella controller be looked at for the placement signal issue. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the case before the complaint, pump (b)(6) was run from (b)(6) with intermittent contrast spikes above 100%. pump (b)(6) was implanted on (b)(6) and run until (b)(6) with the same intermittent but more frequent spiking contrast with snr plummeting and placement signal rocketing to 3002 mmhg during these times. the extended high placement signal is what triggered the placement signal not reliable alarms from the complaint as contrast is not visible to the user. rt logs during this case confirm the lt logs. the following case had the same spiking contrast from the previous cases but did not have the unreliable placement signal from the previous case. device analysis: the console was eventually returned for investigation. the oscillating contrast was not reproduced while running a test pump, the console performed as expected. root cause: the root cause of the intermittent unreliable placement signal and oscillating contrast could not be determined as the failure was not reproduced during testing. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3614 | 1220648-2025-48196 | 23500634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23500634&pc=OZD | 10/11/2025 | 12/09/2025 |  | 11/07/2025 |  |  | 10/11/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026683680 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | (b)(4): (b)(6) 2025 - pt will go for a ct abd/pelvis to check for a bleed. systemic heparin is paused due to concerns for bleeding. (b)(4): (b)(6) 2025 - surgery was consulted yesterday for gi bleed but pt ended up sealing the bleed off on its own. pt did receive 5 units prbc total yesterday. | [5 empty] | The investigation for the reported major bleed and renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed and renal failure could not be determined. | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3615 | 1220648-2025-48104 | 23500684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23500684&pc=OZD | 10/15/2025 | 11/05/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1752054 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 112 KG | Non Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a notification on the supporting automated impella controller indicating battery life was lower that expected. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | B3 revised date as it was reported incorrectly on the initial report that was submitted. g3 date should of reflected 15-oct-2025 on the initial report that was submitted. h10 add related report numbers. this is one of three related reports. this report represents an automated impella controller. other reports were submitted to represent the pump and the other automated impella controller. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: data review: console logs were consistent with what was reported in the complaint. according to the imc logs: [15/10/25 18:18:51 imcgui: event set: #109 from 17 (0)] the console was disconnected from ac power [15/10/25 18:24:03 imcgui: event clear: #109 from 17 (0)] console was connected back to ac power about 6 minutes later. [15/10/25 18:24:05 imcgui: event set: #23 from 17 (0)] a battery level low alarm was issued about 2 seconds after console was reconnected to ac power [15/10/25 18:24:12 imcgui: event clear: #23 from 17 (0)] the battery level low alarm quickly resolved about 7 seconds after it was issued battery data embedded in the ltlogs revealed that ¿battfullcap 2¿ momentarily spiked to 65000mah around the time the battery level low alarm occurred but recovered back to a around 6000mah shortly after. device analysis: console was tested after arriving in field service. the reported battery issue was unable to be reproduced. testing was done by disconnecting the console from alternating current power and reconnecting it shortly after but no battery alarms occurred. no evident anomalies with the consoles batteries were observed in the batttest results. conclusion: the reported battery issue was likely due to a momentary communication glitch between the power board module and the batteries where a single sample from battery data (in this case value of ¿battfullcap 2¿) was corrupt. the exact component responsible could not be conclusively determined. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of console (b)(4).
3616 | 1220648-2025-48197 | 23500838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23500838&pc=OZD | 10/11/2025 | 11/07/2025 |  | 11/07/2025 |  |  | 10/11/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026687345 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | (b)(4): (b)(6) 2025 - having a lot of vt runs per rn. md at bedside talking about potential goals of care discussion. (b)(4): (b)(6) 2025 - received platelets and blood overnight. right arm/right side of chest swollen. bival was cut in half. | [5 empty] | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections: b1, d1, d3, d4, d7a, e4, g1, g2, h8. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleed was not determined due to insufficient clinical details. arrhythmia, tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3617 | 1220648-2025-48194 | 23500861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23500861&pc=OZD | 10/10/2025 | 11/07/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699723 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Tachycardia | Thrombocytopenia | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility patient was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, patient was placed on continuous renal replacement therapy. pt had ventricular tachycardia and was given amio and lido. heparin was turned off due to 50% drop in platelets. care was ultimately withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4; h6 health effect - clinical code 060109 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. no product was received for evaluation. renal failure/thrombocytopenia: the cause of the clinical symptom was not determined due to insufficient clinical details. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3618 | 1220648-2025-48175 | 23500872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23500872&pc=OZD | 10/10/2025 | 11/07/2025 |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736475 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 74 KG | Non Hispanic | Black Or African American | Injury | Yes | No | False Alarm | Patient-Device Incompatibility | [5 empty] | Atrial Fibrillation | Bradycardia | Thrombosis/Thrombus | Heart Block | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 exhibited a low heart rate so a transvenous pacemaker inserted. the pump generated impella in aorta alarms that self resolved. the patient continued to experience atrial fibrillation and a complete heart block. the patient developed deep vein thrombosis at the right internal jugular transvenous pacemaker site and an atrioventricular node ablation. impella support continues. | [5 empty] | Since product wasn't returned and data logs were inconclusive, the cause of the false alarms couldn't be determined. the cause of the heart block was not determined due to insufficient clinical details. the root cause of the bradycardia, atrial fibrillation, and thrombosis was unable to be determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3619 | 1220648-2025-48198 | 23501053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23501053&pc=OZD | 10/10/2025 |  |  | 11/07/2025 |  |  | 10/10/2025 | 11/07/2025 |  |  |  |  | 11/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678477 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller had a power cord completely cut off. there was no harm reported and there was no patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3620 | 1220648-2025-48114 | 23483434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483434&pc=OZD | 10/07/2025 |  |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623354 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing from a cardiac rupture. a thoracotomy was performed and a hematoma was removed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the cardiac perforation was not determined due to insufficient clinical details. the cause of the hematoma was not determined due to insufficient clinical details. the cause of the pump suction was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3621 | 1220648-2025-48115 | 23483435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483435&pc=OZD | 10/07/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105809 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced an automated impella controller failure. the controller was replaced to support the patient. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b, explantation date, was added. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3622 | 1220648-2025-48116 | 23483438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483438&pc=OZD | 08/22/2025 |  |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1120454 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump support ongoing for weeks for the patient admitted in with known cardiac history and awaiting ventricular assist device (vad) device for care escalation. the pump lost optical signal but, there was no explant as the team awaited vad device. the signal loss was investigated in the pump, and the team concluded the automated impella controller needed to be investigated as well. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3623 | 1220648-2025-48117 | 23483441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483441&pc=OZD | 10/07/2025 |  |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026734065 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male | 100 KG | Non Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Respiratory Failure | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support due to acute myocardial infarction / cardiogenic shock. while on support, the pump experienced placement signal issues with no specified resolution. additionally, the patient required tracheostomy. no additional information provided. the patient outcome was noted as survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding respiratory failure, the cause was not determined due to insufficient clinical details. optical sensor issue: clinical details state the patient's mean arterial pressure (map) was 134 but the blood pressure map was 60. the data logs show the placement signal was variable between 60-70mmhg. the p-level was increased from p-6 up to p-9 over the course of a few hours and caused many fluctuations in the placement signal of it increasing up to systolic values of about 100mmhg and then crashing with diastolic values of about 40mmhg. during this time 5s parameters were stable with signal not reliable (snr) about 2500 and contrast of about 25% but the placement signal did not appear to be ventricular. placement signal became less noisy and then stabilized. during the time of the reported issue, there was a rapid increase in ps from about 55mmhg to about 110mmhg over the span of 7 minutes. during this time there was a shift increase in mean motor current values as well as pump flows without a change in p-level. contrast increased, snr increased, while gain and light remained constant. the placement signal still did not appear ventricular. the placement signal then continued to increase and would go to values as high as 250mmhg. it continued to remain elevated throughout the case with instances of the placement signal increasing enough to trigger 1048 placement signal not reliable alarms as well as false positioning alarms, all with fluctuations/shifts up in motor current. based on the data logs provided, the root cause of the optical sensor issue is most likely due to biological material ingestion that is impacting the optical sensor. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3624 | 1220648-2025-48118 | 23483443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483443&pc=OZD | 10/09/2025 |  |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712981 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp. the patient had a history of peripheral artery disease in the same limb where the pump was implanted. the pump was operating with slow flow so a perfusion fem-fem was placed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the injury was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | Impella cp was returned for evaluation. peri-procedural adverse event (ischemia): the returned device was investigated and no product damage was noted. there was some dried blood near the suture hub area but no other abnormalities were found. the cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3625 | 1220648-2025-48119 | 23483446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483446&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026715836 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered high purge pressure during impella 5.5 support. high purge pressure was found and after the discussion with the team, the decision was made to wean the patient, without giving the thrombolysis. pump would not be returned due to infection. no other information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported events have been completed. purge pressure high - after 5 days on support, hpp occurred. the team decided to explant without troubleshooting. data logs were not recovered. the cause of the hpp was not determined as no product was returned, no data logs recovered, and insufficient clinical details. the cause of the infection was unable to be determined due to insufficient clinical details.
3626 | 1220648-2025-48120 | 23483448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23483448&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026733801 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was fever and infection during impella 5.5 patient support. while on support, the patient had a low grade fever and the patient¿s white count was slightly elevated. the patient¿s fever resolved. after a ct scan, bowel ischemia was found. it completely resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. infection: the cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3627 | 1220648-2025-48121 | 23484771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23484771&pc=OZD | 10/08/2025 |  |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749138 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered positioning issues during impella cp support. it was noted that the impella pump was very difficult to adjust. however, bedside echocardiogram with cardiologist was performed, and the impella position was corrected. a positioning issue was encountered a second time and the cardiologist corrected the pump position again. after repositioning of the pump, the pump ran well the whole time of support. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the difficult to place or position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump position was corrected again the second time by an experienced cardiologist on ward and was stable then the whole time of support most likely incorrect repositioning the first time. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3628 | 1220648-2025-48122 | 23484913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23484913&pc=PYX | 10/08/2025 |  |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025484859 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was thrombosis during impella rp flex support. the device was removed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3629 | 1220648-2025-48123 | 23485027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485027&pc=OZD | 10/08/2025 |  |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658655 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was escalated from an impella cp to an impella 5.5 and had suction alarms and ectopy after repositioning of the pump occurred. the pump was unable to flow up without suction events. an echocardiogram revealed a thrombus was interacting with the inlet. later, care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review: medical safety/clinical reviewed the event. the patient was undergoing treatment for high-risk percutaneous coronary intervention (pci). the patient initially presented to an outside hospital emergency department with chest pain and was transferred for further management. the patient was diagnosed with non¿st-elevation myocardial infarction (nstemi) on (b)(6) 2025. cardiac catheterization performed on (b)(6) 2025 demonstrated a chronic total occlusion (cto) of the right coronary artery (rca). the patient¿s medical history was significant for prior coronary artery bypass grafting (cabg), congestive heart failure, hypertension, implanted cardioverter-defibrillator (icd), and prior cerebrovascular accident (cva). a high-risk pci of the native rca was planned. on (b)(6) 2025, the patient underwent successful pci to the saphenous vein graft (svg) to posterior descending artery (pda) and the native rca with impella cp support placed pre-pci. the patient remained hemodynamically stable during the procedure, although arterial pulsatility was noted to be reduced. right ventricular function was discussed prior to the case and was noted to be moderately reduced. the patient was started on low-dose dobutamine and given intravenous volume. the impella cp provided hemodynamic support throughout the procedure. following the procedure, the patient was transferred to the unit. the impella cp access site was reported as clean, dry, and intact, with appropriate waveforms at p-7 performance level. it was reported that the 7 french companion sheath from the insertion kit was leaking around the catheter guide at the diaphragm. on (b)(6) 2025, the patient was noted to have severe left ventricular dysfunction, which had been present prior to the high-risk pci. overnight, the patient clinically decompensated and required multiple inotropes and vasoactive medications. the clinical team determined that escalation of mechanical circulatory support was indicated. echocardiography confirmed adequate right ventricular function and severely hypokinetic left ventricular function with an estimated ejection fraction of 15%. an old, encapsulated left ventricular apical thrombus was identified. the physician acknowledged the contraindication and associated risk, but determined that the potential benefit outweighed the risk, and proceeded with impella 5.5 implantation. the impella 5.5 device was successfully implanted with flows at p-8. volume was administered, the impella cp was removed, and hemostasis was achieved. overnight, the impella 5.5 was unable to increase flow without suction events. echocardiography demonstrated that the apical thrombus was interacting with the device inlet, and the aortic and left ventricular waveforms were uncoupled. the most recent measurement showed the device inlet positioned approximately 4 cm, and ectopy was noted. recommendations included optimization of device position and volume resuscitation. the team discussed options for device removal. at a bedside evaluation, the physician did not visualize thrombus within the inlet and believed it remained in the left ventricular apex; however, the risk of thrombus ingestion and waveform abnormalities was discussed. ultimately, the family elected to withdraw care, and the patient expired. the event story involves three product complaints. impella 7fr sheath - mdr 1220648-2025-48098: malfunction. impella cp pump - mdr 1220648-2025-48100: serious injury. impella 5.5 - mdr 1220648-2025-48123: death. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp and impella 7fr sheath. | The pump was not returned for investigation. the data log confirms the complaint. the cause of the pump suction issue was not determined without returned product investigation and sufficient clinical details. the cause of the pump positioning issue was not determined without returned product investigation and sufficient clinical details. the root cause of the arrhythmia, thrombosis, and thrombus was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks.
3630 | 1220648-2025-48125 | 23485193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485193&pc=OZD | 10/07/2025 |  |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1297944 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | During preparation of an automated impella controller, there were multiple shutdown alarms that occurred. the controller was replaced and no patient harm was reported. | [5 empty] | The device was returned for investigation. section d9 has been updated. | Patient information was not provided. therefore, section a was left blank. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3631 | 1220648-2025-48126 | 23485213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485213&pc=OZD | 10/07/2025 |  |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis. the pump was explanted and replaced with an impella 5.5. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the device passed all post sterile inspection checks.
3632 | 1220648-2025-48127 | 23485240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485240&pc=OZD | 10/01/2025 |  |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722929 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 suffered a left sided hemorrhagic stroke. the pump was explanted and no clot was observed on the pump. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
3633 | 1220648-2025-48128 | 23485266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485266&pc=OZD | 10/07/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761702 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hematuria | Thrombosis/Thrombus | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and swollen lower extremities. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Correction: b1 and h6 health effect - impact codes 2301and 1903 was inadvertently omitted on the initial manufacturer device report. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data was not returned for review. pump was returned for analysis. visual inspection found no issues on the pump. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no issue was found on the pump. minor bleed/access site adverse event: oozing from either impella or ecmo cannulation site. the cause of the injury was unable to be determined as limited clinical details were provided and no issue was found on the pump. ischemia/thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3634 | 1220648-2025-48130 | 23485394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485394&pc=OZD | 10/10/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026735778 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella 5.5 had purge pressure high and purge system blocked alarms. tissue plasminogen activator was administered but the purge system blocked alarm was not resolved. the impella 5.5 was exchanged for a new one. | [5 empty] | The impella device is not available for return from the customer. therefore, investigation of the device is not possible. should any new information be received, a supplemental report will be filed.
3635 | 1220648-2025-48124 | 23485604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485604&pc=OZD | 10/09/2025 |  |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763071 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 124 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | Aortic Valve Insufficiency/ Regurgitation | Hemodynamic instability | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp and experienced low blood pressure. the physician suspected the impella was keeping the leaflet open, causing aortic insufficiency post balloon aortic valvuloplasty. the physician made the decision to explant the pump. during explanting of the pump, the patient went into ventricular tachycardia/fibrillation, coded, and was defibrillated. however, the patient expired. | [5 empty] | No product was returned for investigation. the cause of the aortic valve insufficiency, regurgitation, and hemodynamic instability was not determined due to insufficient clinical details. the root cause of the tachycardia and ventricular fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta. ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
3636 | 1220648-2025-48131 | 23485686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485686&pc=OZD | 10/07/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730512 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp and extracorporeal membrane oxygenation supporting the patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed femorally but there was a bleed observed at the site and manual pressure was not enough to manage the bleed. femstop placed. perclose strings tightened and hemostasis achieved. the following day patient had swelling in abdomen, could not maintain any pressure. the patient expired on support. the care team believes patient died to bleeding and multi organ failure. the prompted intervention and questions of retroperitoneal bleed due to abdominal swelling. the patient expired on the support after just 9 days. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. retroperitoneal hemorrhage/ major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks.
3637 | 1220648-2025-48133 | 23485741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485741&pc=OZD | 10/09/2025 |  |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728838 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was major bleeding and thrombocytopenia during impella cp patient support. while on support, the patient went back to the cardiac catheterization laboratory for bleeding at the impella site. it is unknown how the bleeding was resolved. the patient¿s hemoglobin and platelet count dropped so one unit of blood was given. there were suspicions of heparin induced thrombocytopenia, so the team switched to heparin. it was noted the patient had tarry stool. the patient was successfully weaned and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombocytopenia: the cause of the clinical symptom was not determined due to insufficient clinical details. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
3638 | 1220648-2025-48132 | 23485748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485748&pc=OZD | 10/09/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722848 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 105 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient was identified as not a surgical candidate, and due to their severe peripheral artery disease, the physician determined that the left axillary artery was the only appropriate conduit for impella implantation. the impella 5.5 was implanted, and position was confirmed, however the impella had a low placement signal alarm. the physician determined that the pump was in an ideal position and that the alarm was a false alarm. the placement monitoring was disabled. on support day 3, the patient experienced bleeding from the nose and mouth. the patient was transfused with two units of blood. additionally, it was noted that the patient had a left ventricle thrombus. it was noted that the patient had been heparinized, however it was the physician¿s discretion to outweigh risk versus benefit of stopping heparin for the bleeding or to continue heparin for infusion while patient was supported by impella. the same day, the device was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3639 | 1220648-2025-48135 | 23485752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485752&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659103 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 85 KG |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered high purge pressure during impella 5.5 support. the purge system had already been replaced, without success. lysis was administered and doctor on duty was instructed to monitor purge system and motor current very closely. purge flow then became ~5ml/h pp 570mmhg. lysis had worked and pump was now working as intended. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary ==> no product or data logs were returned for evaluation. high purge pressure: clinical details noted that pump had high purge pressure & purge system blocked alarms on third day of support. tpa was added to purge and was able to resolve high purge pressure alarms, with purge flows and pressures within normal range. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation. device history lot ==> device lot: 1911471 device history batch ==> subcomponent lot: n/a device history review ==> pump sn (b)(6) passed all post-sterile inspection checks.
3640 | 1220648-2025-48136 | 23485754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485754&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the patient flight transport, the impella cp side arm retainer was damaged, and purge fluid started leaking from sidearm. no purge pressure maintaining. side arm retainer bypass with 20g needle was done. the impella was actively being weaned for removal due to the side arm retainer damage. the impella was removed, and the patient was stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. product damage (purge sidearm crack) - the cause of the product damage (purge sidearm crack) was not determined due to no product returned and insufficient clinical details. purge leak - (pump) - the cause of the purge leak - (pump) was not determined due to no product returned and insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3641 | 1220648-2025-48134 | 23485755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485755&pc=OZD | 10/11/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737626 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 65 KG | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the impella 5.5 was to be removed in the cardiovascular operating room and was swapped out for another impella 5.5. the team became aware of hemolysis and completed an echocardiogram where the device was noted to be shallow. the patient is to have the impella 5.5 redo using the right axillary arterial site given hemolysis and a suboptimal depth placement of 4.0 centimeters in the left ventricle. the patient was to remain on veno-arterial extracorporeal membrane oxygenation during the process. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks.
3642 | 1220648-2025-48137 | 23485756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485756&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587967 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient came into the hospital in cardiogenic shock and impella cp was urgently implanted. no echocardiogram was performed. impella cannula kinked in the left ventricle and pump stopped. after repositioning and straightening the cannula, impella was unable to restart. the impella was exchanged. the patient was stable during intervention. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported events has been completed. the pump was not returned for investigation. as per the data log, the pump ran for 10 minutes before encountering a pump stop with high motor current up to 1400 and alarm 13. no purge-related issues were recorded on the data log. there was one attempt reported to restart; however, a similar pump stop occurred with another motor current spike. according to the clinical details, the cannula was kinked during the pump stop, and the pump was unable to restart, so a new pump was used. the cause of the pump stop and cannula kink was not determined without returned product investigation and sufficient clinical information.
3643 | 1220648-2025-48138 | 23485766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485766&pc=OZD | 10/09/2025 |  |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1384825 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller was being used for training intensive care unit staff and when the device was unplugged from the wall outlet, it began beeping. the beeping did not stop until it was plugged back into the wall power source. there was a reported that there was a ¿battery failure¿ alarm. the device was tagged to not be used with a patient. there was no patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information noted the automated impella controller device was returned, and an evaluation/analysis has been completed. correspondingly, sections d9: device received by manufacturer with receipt date was updated accordingly. the automated impella controller (aic) with data logs were received and the investigation was completed. service history review was completed, and there is no quality notification associated with the complaint device¿s most recent service report. additionally, there are no issues related to the complaint discovered when reviewing the product history of console. data logs were analyzed and noted the following: battery failure (transport connection blown/missing) alarms were observed in the imc logs. regarding the device analysis, the console was tested after arriving in field service. the reported battery failure could not be reproduced during testing. no evident anomalies were observed in the batttest results. the console operated without any issues while testing with a known good pump and purge cassette kit on an engineering lab bench. in conclusion, the cause of the reported battery failure is suspected to be use related. the console was likely booted up with the battery transport connection switch disengaged while the console was connected to ac-alternating current power. the battery failure alarm occurred on the first boot after field-service preventive maintenance on 02-dec-2024. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3644 | 1220648-2025-48139 | 23485774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485774&pc=OZD | 10/10/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722279 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support received one unit of blood due to drop in hemoglobin which was less than 8. additionally, the patient was also having orthostatic hypotension when standing up but quickly recovered. nurse stated that the patient was on multiple anti-hypertension medications and adjustments were being made to the regiment. the patient continued with impella support until heart transplantation. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 added zip code extension as it was omitted from the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. hypotension/anemia: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
3645 | 1220648-2025-48141 | 23485861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485861&pc=OZD | 10/11/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714496 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 59 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient was oozing at the impella site. it was soft without a hematoma. no additional information was provided. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
3646 | 1220648-2025-48140 | 23485898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485898&pc=OZD | 10/09/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026762919 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient presented to the emergency department with complaints of chest pain and an electrocardiogram revealed the patient was in st elevation myocardial infarction. the patient received tenecteplase and then experienced cardiac arrest. cardiopulmonary resuscitation was performed and return of spontaneous circulation was achieved. the patient was brought to the cardiac catheterization laboratory for left heart catheterization (lhc) and percutaneous coronary intervention (pci). lhc revealed a blockage of the left anterior descending coronary artery and a stent was placed successfully. the patient¿s lactic acid results were 10 mmol/l and the medical team decided to implant the impella cp. on day 1 of support, the patient experienced oozing from multiple sites, which the medical team attributed to the patient receiving tenecteplase in the emergency department. a femstop was applied to the impella site to resolve the bleeding. the following day, it was noted that a hematoma at the right groin of the impella access site had resolved. the patient had received albumin overnight to manage fluid balance following diuresis and bleeding. of note, there were no impella related issues reported. on day 5 of support, it was noted that the patient was transfused with one unit of packed red blood cells. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. major bleed: oozing from multiple site since she received thrombolytic in er. femstop applied at impella site. the root cause of the bleeding is determined to be patient condition because of the bleeding from multiple sites. hematoma: hematoma of the right groin resolved. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history lot: device lot: 1992905. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3647 | 1220648-2025-48143 | 23485903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485903&pc=OZD | 10/10/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026697657 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after inserting the impella cp and removing guidewire, cp was started in auto mode however unable to achieve flows >0.5lpm. md attempted to reposition without any increase in flow. the md decided to remove the impella, no thrombus noted. device was flushed with heparinized saline however purge solution was not visualized by motor housing. the medical doctor opened a new impella. there was harm to the patient. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. low or blocked pump flow - data log review showed one suction alarm and no purge related alarms. flows low after less than 1 minute of support. the returned pump was found to have no physical abnormalities and was able to be run in a bucket and was able to be run at p-4, p-9, and auto without recreating low flows. the cause of the low pump flows was not determined as the issue was unable to be recreated during investigation of the returned pump and insufficient clinical details. purge fluid not detected - after pump was removed, device was flushed with hep saline however purge solution was not visualized by motor housing. data log review showed no purge related alarms or abnormalities. the returned pump was able to purge without issue. the cause of the purge fluid not detected was an unintended error caused by the user due to the issue no being able to be recreated and no abnormalities found in the data logs. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3648 | 1220648-2025-48144 | 23485908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485908&pc=OZD | 10/10/2025 | 11/19/2025 |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756405 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 105 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that impella cp was implanted, and low flows were noted. after troubleshooting, it was found that the cannula was kinked in aorta. un able to un-kink the cannula the impella cp was exchanged. there was no harm to the patient. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: impella cp sn (b)(6) returned for evaluation. data was not returned. visual inspection found a kink on cannula about 15 mm from outflow cage & cannula bonding site. no other issues were found on the rest of the pump. low or blocked pump flow: the cause of low or blocked pump flow was most likely kinked cannula as per clinical information and observed on the returned product. product damage (cannula kink): the cause of the product damage (cannula kink) cannot be determined as insufficient clinical details were provided. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
3649 | 1220648-2025-48129 | 23485933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23485933&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688939 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Hemodynamic instability | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented to the emergency room (er) after coding in the ambulance. cardiopulmonary resuscitation (cpr) was given and return of spontaneous circulation was achieved. in the er, st-elevation myocardial infarction was activated and another two rounds of cpr were given to the patient. the decision was made for impella cp insertion and possible percutaneous coronary intervention (pci). the impella cp was inserted and pci was performed. however, the patient remained unstable. after stents were placed in the left anterior descending artery (lad), no reflow became a problem. medications were given into the coronary, with minimum resolution. the patient became unstable and developed acidosis. an x-ray was taken to determine pulmonary edema. the patients mean arterial pressure started decreasing and ventricular tachycardia started. multiple shocks were given, to return to normal sinus rhythm. the patient never came back and expired on support. | [5 empty] | No product was returned for investigation. the cause of the hemodynamic instability was not determined due to insufficient clinical details. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 and a5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3650 | 1220648-2025-48142 | 23486037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486037&pc=OZD | 10/11/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728380 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 74 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the patient¿s foot was cold without pulse and thigh appeared slightly mottled. the physicians are aware and they did a vascular ultrasound that showed slow flow. they will reevaluate when the pressors are lowered. pressors were actively weaning. the patient had poor prognosis and was not trending in the positive direction. the family made the decision to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3651 | 1220648-2025-48145 | 23486100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486100&pc=OZD | 10/09/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/09/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761431 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention (pci) was being implanted with an impella cp device for mechanical circulatory support. the patient was noted to have severe peripheral artery disease, which prevented adequate vascular access for the impella cp placement. the pci proceeded without impella cp support, where the patient developed an acute stent thrombosis involving the stents placed in the left anterior descending coronary artery and left main coronary artery. the patient went into cardiac arrest, and despite 40 to 45 minutes of advanced cardiac left support and multiple attempts to restore coronary flow, there was no return of spontaneous circulation and the patient expired. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise as the pump was intended for use but could not be placed due to severe peripheral vascular disease. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical symptoms (cardiac arrest): the cause of the injury was not determined due to insufficient clinical details. patient anatomy or condition prior to therapy: the patient's anatomical condition prior to pump use was most likely related to a pre-existing patient condition.
3652 | 1220648-2025-48146 | 23486144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486144&pc=OZD | 10/10/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729973 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp, the patient taken to ct scan and found to have right iliac occlusion. the patient has had fem to fem surgical bypass. device was removed prior to observation of limb injury ischemia. impella 5.5 inserted afterwards. the device was replaced. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 added zip code and zip code extension as they were omitted from the initial report that was submitted. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. ischemia: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
3653 | 1220648-2025-48147 | 23486152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486152&pc=OZD | 10/11/2025 | 12/09/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026760665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 62 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella cp, the patient¿s right groin was soft with continued pressure being held to the right groin hematoma. the doctor at bedside gave the order to give brilinta and then stop cangrelor drop. the next day, the right groin site with bruising but no more oozing overnight or today was noted. cangrelor drop was stopped yesterday and a low dose heparin was started today. per the registered nurse, epinephrine/lidocaine were injected to the right groin insertion site and pressured dressing was applied. hemolysis and hematoma. were noted. heparin was stopped. plasma free hemoglobin was sent. the patient was on one of levophed and one of dobutamine at p8. echocardiogram showed the impella cp was at 3.7 centimeters and clear of anterior mitral valve leaflet but was lying more towards the septum. the patient was running a fever. hematuria was improving. lactate dehydrogenase was down trending. one unit of packed red blood cells was given. no signs or symptoms of bleeding. were noted the pump was removed and the impella 5.5 was implanted. | [5 empty] | The data logs were received for review, however there was no change to the investigation results previously reported in mfr report # 1220648-2025-48147-1. data review: data logs showed pump ran without issue during support. there were suction alarms triggered in the end of the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported mechanical interaction with blood and hematoma/major bleed has been completed. no product was returned for review. the data logs showed pump ran without issue during support. there were suction alarms triggered in the end of the case but resolved quickly.there were no mc spikes, abnormal ps or purge issues observed during support. the root cause of mechanical interaction with blood and hematoma/major was unable to be determined as limited clinical details were provided and no product was returned for review.
3654 | 1220648-2025-48148 | 23486383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486383&pc=OZD | 10/11/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709975 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient was listed status 2 for a transplant and coded. the patient went into ventricular tachycardia and one round of compressions and shock were given before achieving return of spontaneous circulation. the impella was pulled back to two centimeters after the code. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3655 | 1220648-2025-48149 | 23486545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486545&pc=OZD | 10/10/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733015 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 49 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient¿s groin oozing but patient is coagulopathic. correcting coags. the day after, it was oozing from groin still. bival held this am, partial thromboplastin time improving, another stitch placed, epi at site, blood given. vascular consulted for further eval. impella site without bleeding or issues. later the groin looks great, vascular placed 7 sutures in groin to achieve hemostasis. the pump was weaned and explanted successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: bleeding from impella groin access site. the cause of the bleeding is determined to be use issue because the access site bleeding complication was managed by angle match and placing additional stitch the next day. device history review: the device passed all post sterile inspection checks.
3656 | 1220648-2025-48151 | 23486620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486620&pc=OZD | 10/11/2025 | 11/06/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704423 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 49 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient had worsening thrombocytopenia. platelet counts were down to 62. in addition to mechanical circulatory support, the surgical patient had some post-operative blood loss. the patient was weaned successfully. at time of explant, the patient was only requiring 2.5 mcg/kg/min of dobutamine. the patient was transported to the operating room for impella cp removal and surgical closure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3657 | 1220648-2025-48152 | 23486679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486679&pc=OZD | 10/23/2025 |  |  | 11/06/2025 |  |  | 10/23/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708507 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp red plug was leaking. the pump was successfully weaned and explanted. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3658 | 1220648-2025-48153 | 23486778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486778&pc=OZD | 10/22/2025 |  |  | 11/06/2025 |  |  | 10/22/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1284913 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella automated impella controller displayed a system review failure. the controller was exchanged for a working controller. no patient harm was reported. | [5 empty] | Patient information was not provided. therefore, section a was left blank. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3659 | 1220648-2025-48154 | 23486843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23486843&pc=OZD | 10/11/2025 | 11/05/2025 |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692737 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 101 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was arrhythmia and positioning issues during impella 5.5 support. while on support, the patient went into a ventricular tachycardic (vt) episode. the patient was cardioverted into normal sinus rhythm with one shock. the nurse reported some positioning issues with the device overnight. amiodarone was started after the vt event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the arrhythmia, in order to make a cause determination, all of the clinical details would have to be provided. the cause was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue), the cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3660 | 1220648-2025-48155 | 23487022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487022&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569916 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was hematuria, hemolysis, arrhythmia, and a positioning issue during impella cp support. while on support, pink urine was noted. the impella¿s position was checked with fluoroscopy, and it looked slightly deep. the pump was pulled back to 81 centimeters and locked into place. the patient remained on amiodarone for ongoing ectopy. the urine remained pink but lighter in color. there was a plan to start a systemic heparin infusion. a cerebellum device was placed to assess neurological function as there was possibly seizure like activity. there were concerns about hemolysis. the patient went into ventricular tachycardia and required advanced cardiovascular life support as well as two shocks. the return of spontaneous circulation was achieved. the urine was still dark. later, it had cleared, becoming yellow and adequate. there was minimal ectopy since the arrest as well as start of lidocaine. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. clinical symptoms (arrhythmia): the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (hematuria): the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood (hemolysis): data log shows no signs related to suction or motor current spike during the entire case run. before 12 pm on 11-oct, some improvement was noted in pump flow while running on a steady p-level. placement signal was around 115 mmhg at the start of the case, decreased to 75 mmhg, and remained with a fluctuating trend. the cause of the hemolysis issue was most likely related to patient condition, based on the given clinical details and data log review. device history review: the device passed all the post sterile inspection checks.
3661 | 1220648-2025-48156 | 23487110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487110&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716863 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a major bleed during impella cp support. while on support, there was coffee ground emesis and a suspected gastrointestinal (gi) bleed. the patient received blood transfusions for the gi bleed and heparin remained on hold. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: receiving blood transfusions due to gi bleed. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
3662 | 1220648-2025-48157 | 23487315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487315&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739610 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient noted as high-risk for a coronary artery bypass surgery was implanted pre-operatively with an impella 5.5 device for mechanical circulatory support. the patient had presented to the emergency room on (b)(6) 2025 with dyspnea, chest pain, nausea, dizziness, diaphoresis, and altered mental status. the patient was found to have elevated troponin levels and had an ejection fraction of 15% with mitral valve regurgitation, aortic valve stenosis, and right bundle branch block. the decision was made to bring the patient to the operating room for a high-risk surgery for aortic valve replacement, mitral valve replacement, tricuspid valve repair, triple coronary artery bypass grafting, and maze procedure. an impella 5.5 was placed pre-operatively. after successful surgery, the patient was removed from cardiopulmonary bypass and was bridged back to impella 5.5 support. the patient was released to the intensive care unit (icu) for rest and recovery. in the icu, the patient received two units of fresh frozen plasma, one unit of platelets, and 750 milliliters of albumin to replace lost volume from the surgery. it was noted that the drainage from the patient¿s chest tubes was slowing down, and when the bleeding slowed down sufficiently the impella performance power would be increased to p9. the following day, the patient began to experience impella flow issues and left ventricle waveform decoupling and required increased inotrope and pressor requirements. despite continued resuscitation, the patient continued to decompensate. an echocardiogram was performed to review the impella placement. the impella placement was viewed and determined to be adequate, however the decision was made to reposition the impella to investigate if pump flows would improve. during the repositioning, the impella 5.5 was inadvertently pulled out of the left ventricle. the physician declined to cannulate the patient for extracorporeal membrane oxygenation (ecmo) and instead controlled the patient¿s blood pressure with vasopressors. the patient was taken to the cardiac catheterization laboratory to try and advance the impella 5.5 back across the valve with a buddy wire and pacing, but attempts were unsuccessful. the patient was brought to the operating room to re-open the chest cavity, pull the impella 5.5, and re-access the left ventricle using a wire. during the process, the pump began having purge flow issues and eventually a purge block. the physician decided to exchange the impella 5.5 for a new pump. a new impella 5.5 pump was placed without issue, however the new pump continued to have lower than expected flows and displayed low left ventricular pressures. the physician again declined to cannulate the patient for ecmo due to the patient¿s high inotrope and pressor requirements and very poor arterial blood gas results. the decision was made to bring the patient back to the icu and withdraw care. the peel away sheath was removed, the patient¿s chest was closed, and the patient was returned to the icu on impella 5.5 support. the family decided to withdraw the patient¿s care due to poor prognosis. the patient¿s care was withdrawn and the patient expired. this complaint involves two impella devices: pump 1: impella 5.5, with serial number (b)(6) pump 2: impella 5.5, with serial number (b)(6) this report specifically addresses pump 2. a separate report will be submitted for the other devices associated with this complaint. | [5 empty] | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review medical safety/clinical reviewed the event. a 70-year-old male with a medical history significant for coronary artery disease (des ×3), chronic obstructive pulmonary disease, hyperlipidemia, chronic angina, prior cerebrovascular accident, neuropathy, and a 30 pack-year smoking history presented to the emergency department on (b)(6) 2025 with dyspnea, altered mental status, chest pain, nausea, dizziness, and diaphoresis. evaluation revealed elevated troponin, right bundle branch block, severe left ventricular dysfunction (ef ~15%), aortic valve stenosis, and mitral valve regurgitation. cardiothoracic surgery was consulted. the patient was taken to the operating room for aortic, mitral, and tricuspid valve replacement, coronary artery bypass grafting ×3, and maze procedure. a decision was made preoperatively to provide mechanical circulatory support (mcs) with an impella 5.5 device. following implantation, the patient developed low flow with decoupled left ventricular waveforms. echocardiography was performed to confirm device position. during attempts to reposition the device to improve flows, the impella 5.5 was accidentally withdrawn from the left ventricle. the patient was transferred to the cardiac catheterization laboratory, where attempts to re-advance the device across the aortic valve using a buddy wire and pacing were unsuccessful. the patient was then taken back to the operating room for open chest access, where the device was removed and re-access of the left ventricle was attempted using wire and catheter under mobile fluoroscopy and transesophageal echocardiography. during this process, the impella 5.5 developed purge flow issues with eventual purge blockage. the physician elected to exchange the device. after implantation of the replacement impella 5.5, the device continued to demonstrate low flow and low left ventricular pressure despite support. the patient required high-dose inotropes and vasopressors and had poor arterial blood gas results. given the patient's illness, poor prognosis, and concern for further escalation (including ecmo), the physician elected not to pursue additional mechanical support. after discussion with the family, a decision was made to withdraw care, and the patient expired after approximately two hours after start of the impella 5.5 pump mechanical support. the event story includes two product complaints: impella 5.5 pump 1 sn 513842 - mdr 1220648-2025-48069: serious injury. impella 5.5 pump 2 sn 624797 - mdr 1220648-2025-48157: death. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow: data log shows low pump flow without any suction alarms throughout the case run. no abnormalities were reported in pump position or motor current spike. the case run was around 30 minutes, so the trend in placement signal could not be adequately assessed. the cause of the low pump flow issue was most likely related to patient condition, based on the given clinical details and data log review. device history review: the device passed all post sterile inspection checks.
3663 | 1220648-2025-48150 | 23487349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487349&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662667 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 99 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella 5.5 due to the patient having hemolysis, the plan is to remove the pump this afternoon. staff are unsure of the cause of hemolysis, but plan is to move forward with removal and not to try to reposition. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the product was not returned for review. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current spikes, abnormal ps or purge issues observed during support. the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review.
3664 | 1220648-2025-48158 | 23487407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487407&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690252 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced intermittent suction alarms which were addressed by administering albumin. the impella cp was positioned deep in the ventricle and was adjusted under imaging. the patient suffered multiple organ system failures and was given a do not resuscitate status. later, the patient expired on support. | [5 empty] | The device was not returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the cause of the device positioning issue was not determined without returned product investigation and sufficient clinical details. the cause of the low pump flow issue was most likely related to patient condition, based on the given clinical details and data log review. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3665 | 1220648-2025-48159 | 23487456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487456&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 75 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation and lost pulsatility. the patient was defibrillated, but the patient expired on support. | [5 empty] | D9. in the -01 report was erroneously filled out; the product was not returned for investigation. investigation summary: pump suction: the pump was not returned for investigation. the total case run was about 18 hours, from midnight on (b)(6) to the afternoon of (b)(6). data log shows multiple suction alarms throughout the case run, with a deteriorating placement signal trend and flow drops while running on a steady p-level with corresponded with low points in the placement signal. no abnormalities were reported with improper position or motor current spike during case run. clinical details indicate that multiple suction alarms occurred on (b)(6) 2025, prompting a call to the team. the patient became hypotensive with declining maps and pulsatility, and escalation of support was discussed. echo showed the pump at 3.2 cm with normal rv function and an ef of 10%. the cause of the suction alarms was most likely related to patient condition, based on data log review and provided clinical details. clinical symptoms (ventricular fibrillation): in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (low blood pressure/hypotension): the cause of the injury was not determined due to insufficient clinical details. device history lot: device batch: 1958337. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
3666 | 1220648-2025-48160 | 23487553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487553&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026739625 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5, resistance was met at the subclavian-aortic junction due to calcium at the ostium. implantation was aborted and an impella cp was successfully implanted instead. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned. the cause of the delivery issue was determined to be patient condition. the device passed all post sterile inspection checks.
3667 | 1220648-2025-48161 | 23487699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487699&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704265 |  | 000000000001000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female | 78 KG |  | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was a major bleed, positioning issues, pump suction, thrombus and infection during impella cp support. while on support, the impella cp was having intermittent suction events and could not flow above performance level 3. hemoglobin dropped three points. a bleed was suspected and most likely pneumothorax as shown on the chest x-ray, which displayed a whiteout. the plan was the following: to determine the location of the bleed; going to the operating room for a potential chest washout; extracorporeal membrane oxygenation reconfiguration to veno-arterial-venous; abdominal exploration for potential compartment syndrome; and change impella cp to axillary impella 5.5. the patient was confirmed to have bilateral pneumothoraxes. the impella was shallow in the left ventricle cavity and the team was aware. there was no repositioning at that time. there was a large left ventricle thrombus, bacteremia, and minimal pulsatility. there was continuous suction due to the thrombus. the thrombus was successfully managed with heparin. care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed/thrombosis/infection: the cause of the injury was not determined due to insufficient clinical details. device in wrong position: the cause of the positioning issue was unable to be determined due to insufficient clinical details and no product returned. optical signal issues: data log review shows that upon starting the pump, signal not reliable (snr) decreases and contrast is noisy, with stable gain and light. during this time, placement signal appears to be ventricular, and suction alarms are accurately triggering. due to insufficient product returned and data logs, the root cause of the optical signal could not be determined. pump suction: data logs were returned but no data logs were returned for the date of the reported suction events. however, data log review confirms there were suction events that occurred. the root cause of the suction could not be determined due to insufficient product and data logs returned. device history review: the device passed all post sterile inspection checks. | B7 added information that was omitted from the initial report that was submitted. h6 added medical device problem code a0709 as it was omitted from the initial report that was submitted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3668 | 1220648-2025-48162 | 23487807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487807&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701381 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. ultimately, the patient expired on impella support. | [5 empty] | No product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | Patient information was not provided. therefore, section a was left blank. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3669 | 1220648-2025-48163 | 23487893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23487893&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737621 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was major bleeding, hemolysis, and a cardiac tamponade during impella 5.5 support. while on support, it was noted that the chest tube output would continue to be monitored and more blood products as well as albumin were given at the bedside. a cardiac tamponade was described with ten units of red blood cells given. the patient returned to the operating room (or) for a washout. after the or, the patient was given another three units of packed red blood cells. there were concerns with free hemoglobin and hemolysis. the plasma free hemoglobin measurement was 50 and the previous measurement was 40. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood (hemolysis): data log shows no abnormalities related to suction, motor current spike, or positioning issue during the time of the reported hemolysis. placement signal and pump flow were fluctuating during the case run while running on a steady p-level, without any issue on the optical signal. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. clinical symptoms (hemorrhage/bleeding, cardiac tamponade): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3670 | 1220648-2025-48164 | 23488062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23488062&pc=OZD | 10/11/2025 |  |  | 11/06/2025 |  |  | 10/11/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730309 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was getting placed onto impella cp support for high-risk percutaneous coronary intervention. while inserting the impella cp, they were unable to cross from the iliac into the aorta. several attempts were made. long sheath was not long enough to cross the lesion in the bifurcation. angiogram revealed a bifurcation lesion coming into distal aorta. the decision was made to abort the impella and continue with percutaneous coronary intervention. the patient remained stable. | [5 empty] | D9 the device was returned and date received was added. investigation summary: the investigation has been completed. unable to deliver or advance: only the pump was returned for investigation. no defect was found on the returned product. the cause of the inability to deliver or advance the impella cp was most likely related to the patient¿s vascular anatomy, specifically the presence of a bifurcation lesion in the distal aorta and inadequate sheath length to traverse the lesion. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3671 | 1220648-2025-48166 | 23488239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23488239&pc=PYX | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026623825 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 83 KG |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella rp flex experienced surgical bleeding from the patient's chest, unrelated to the impella. impella flows and motor current decreased and coded. the patient expired on impella support. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: impella rp flex sn (b)(6) + purge cassette returned. no field data logs available. return product found biomaterial around the impeller. low flow was reproduced when running in a water bucket in the lab and a large thrombus was expelled during the run. thrombosis/cardiac arrest: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. low flow: no field data logs available. clinical reports increase in mc and decrease in flows. patient has known history of dvt including existing ¿superficial¿ thrombus to the right upper extremity. the cause of the low flow was biomaterial based on return product investigation. device history lot device lot: 1885552. device history batch subcomponent lot: na. device history review the pump sn (b)(6) passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3672 | 1220648-2025-48167 | 23488302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23488302&pc=OZD | 10/10/2025 |  |  | 11/06/2025 |  |  | 10/10/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient had signs and symptoms of a gastrointestinal bleed and was scoped. heparin was stopped for the procedure and was then restarted. the patient is still on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D6b explant date provided. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks.
3673 | 1220648-2025-48165 | 23488329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23488329&pc=OZD | 10/08/2025 | 12/29/2025 |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 115 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to a durable left ventricular assist device (lvad). on day 10 of impella 5.5 support, it was reported that the automated impella controller (aic) would not respond to the operator regardless of which button was pressed. the aic indicated that the purge system needed to be replaced, but the medical team was unable to perform the operation on the aic. the aic was exchanged for a back up unit. no patient harm was reported. the following day the patient was noted to be feeling well, and the pump had continued to run without alarms. the patient remained on impella 5.5 support until successfully weaned during surgery for an lvad. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3674 | 1220648-2025-31322 | 23489212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23489212&pc=PYX | 07/24/2024 | 09/09/2025 |  | 11/06/2025 |  |  | 07/24/2024 | 11/06/2025 |  |  | 01/31/2026 |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025513849 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal failed and placement signal not reliable alarm occurred. there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. on the second day of support, it was noticed that the placement signal had been drifting for several hours and mean flows had reached 5 l/min. then, the psnr alarm triggered, placement signal and flow calculations disappeared from the aic. at the time of the complication, various hemodynamic measurements had worsened since case start, including co, ci, and papi. the pump was not removed due to the failure mode and was weaned a day later. product was not returned for investigation. data log were reviewed: for roughly 9 hours leading up to the psnr alarm, the ps was trending downward which was then causing the mean flow calculation to drift upward to 5 l/min. additionally, lt logs show cvp starting out noisy, but then stabilized during the time that ps was trending down. all the while, motor current remained stable. when the psnr alarm triggered, ps, mean flows, and cvp calculations were all disabled. the mc remained stable as well as all optical sensor parameters. in looking at the rt logs, it was confirmed that cvp remained stable based on the raw optical sensor pressure reading. the root cause of the placement signal issue was determined to be sensor drift. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3675 | 1220648-2025-31323 | 23489322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23489322&pc=PYX | 10/22/2024 | 09/09/2025 |  | 11/06/2025 |  |  | 10/22/2024 | 11/06/2025 |  |  | 01/31/2026 |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025461505 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a rp flex support ongoing for a patient admitted in post coronary artery bypass graft. the pump lost placement signal after 3 days of support. the pump remained on for support despite the loss/malfunction. no report of patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. data logs were reviewed and it was confirmed that the placement signal not reliable alarm triggered on 10/21. at this time alarm 1054 triggers and pap and cvp rail indicating an optical signal issue. as for the 5s parameters, snr drops to zero, contrast drops but the amplitude increases, lamp was already at 100%, so it remains there, light drops initially but becomes variable, and gain increases. the psnr alarm remains active through the rest of the case. the behavior of the 5s parameters show characteristics of both an air gap and fiber break, without a clear distinguishing factor that would suggest one over the other. clinical details do not provide any clear evidence for the cause of the complication and product was not returned for investigation. due to limited clinical details and data logs not showing a clear optical failure pattern, the root cause of the optical signal issues could not be determined. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3676 | 1220648-2025-48168 | 23489965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23489965&pc=OZD | 10/08/2025 |  |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729997 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump supporting a patient. the pump was supporting when there were concerns of hemolysis with plasma free increased to 28. plan to hold diuresis. later the patient was bridged to left ventricular assist device and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood (hemolysis): data log shows no abnormalities related to suction, motor current spike, or positioning issue during the time of the reported hemolysis issue. there was no trend in pump flow, motor current, or conclusive placement signal trend. the cause of the hemolysis issue was not determined without returned pump investigation and sufficient clinical details. device history review: the device passed all post sterile inspection checks.
3677 | 1220648-2025-48169 | 23490066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23490066&pc=PYX | 10/08/2025 |  |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026741924 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella rp flex placed for the patient with existing clotting issues. it was reported that labs were hemolyzing and not making urine. impella thought to be causing hemolysis with potential clot in cannula therefore decision made to explant. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: g4 the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed motor current (mc) spike seen on prior day of hemolysis event reported. there were placement signal (ps) unreliable alarms triggered in early of the case but resolved quickly. there were no other abnormalities seen further. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely biomaterial ingestion since the data logs and clinical information meet the biomaterial ingestion pattern. device history review: the device passed all post sterile inspection checks.
3678 | 1220648-2025-48170 | 23490226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23490226&pc=OZD | 05/19/2025 |  |  | 11/06/2025 |  |  | 10/07/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658586 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported to abiomed's long term outcomes program a patient on impella support had atrial fibrillation (a fib) with rapid ventricular response. patient on heparin for impella and amiodarone infusion for a fib. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. paroxysmal atrial fibrillation : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3679 | 1220648-2025-48171 | 23490491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23490491&pc=OZD | 10/08/2025 |  |  | 11/06/2025 |  |  | 10/08/2025 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752297 |  | 000000000001000851 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Pump | Pumping Stopped | Device Damaged by Another Device |  |  |  |  | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was in cardiogenic shock and received a moderate low-dose of noradrenaline iv infusion for hemodynamic support and underwent diagnostic cardiac catheterization. after a left main coronary artery stenosis was diagnosed in the diagnostic coronary angiography, the decision was made to proceed with impella support. the pump was advanced into the left ventricle. after uneventful positioning and implantation, the pump immediately stopped upon initiation of support with the error message "motor current too high." a second wire was inserted with the 14 fr. sheath for single access and to better verify the impella position. repeated attempts to restart the pump remained unsuccessful with the same error message, despite exchanging the automated impella controller. the pump was withdrawn back into the aorta. interaction of the impella cp with the j wire in the aorta prevented the removal of the j wire before removing the impella cp. the pump and j wire were pulled back to just before the 14 fr. sheath. the sheath and pump were then removed together from the site, and the puncture site was closed with a preclosure proglide system. the decision was made to proceed with the intervention without impella support due to the patient's worsening hemodynamic condition. it was reported that the patient later expired during the intervention. | [5 empty] | Product complaint # (b)(4). investigation summary: impella 5.5 (b)(6)+ purge cassete + introducer returned for evaluation. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. pump stop: clinical reported motor current high pump stop immediately upon activation. the data logs confirm pump stop alarms and show #13 motor current high alarms with failed restart attempts. there was also an alarm #115 pump stop for rpm deviations. clinical reported guidewire and red lumen were removed from pump before it was started. the pump was returned with j-wire through the cannula with the pump stuck in the 14fr sheath. there was also damage to the catheter at about 32cm mark. the lines were exposed and purge line was damaged. the pump passed electrical testing. there was small biomaterial on the impeller and it is likely some biomaterial was dislodged with the j-wire interaction. the pump was soaked and run and the pump stop did not reproduce. the cause of the pump stop was not determined as no defects were found with the pump, no conclusive causes from data log analysis were observed and limited clinical information was provided. device interaction or damage by another device: clinical reported interaction of the impella cp with the lying j wire in the aorta prevented the removal of the j wire before removing the impella cp. the pump and j wire were pulled back to just before the 14 fr. sheath. the sheath and pump were then removed together from the site. the pump was returned with j-wire through the cannula with the pump stuck in the 14fr sheath. there was also damage to the catheter at about 32cm mark. no patient anatomy issues were mentioned. the cause of the issue with the interaction of the impella cp with the lying j wire in the aorta was determined to be use-related as issue occurred after pump pulled back from ventricle with no patient anatomy issues noted or product related issues. device history lot: device lot number: 1982367. device history batch: subcomponent lot: n/a device history review: the pump (b)(6) passed all post sterile inspection checks. | A4 is unknown. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3680 | 1220648-2025-31916 | 23490577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23490577&pc=OZD | 07/06/2023 | 11/06/2025 |  | 11/06/2025 |  |  | 10/18/2023 | 11/06/2025 |  |  |  |  | 11/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the impella pump had an optical sensor failure that caused loss of placement signal. the position was monitored along with motor current and flows. the automated impella controller was also sent back for evaluation. no known patient harm of injury has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the case associated with the reported complaint was initiated on (b)(6), 2023, with the pump sn: (b)(6). on the 4th day of support (jul 05, 2023), the console displayed the alarm: ¿placement signal not reliable (opm invalid signal, optical pump connected)¿ - event # 1036. the rt log indicates that the optical parameters were abnormal, with the ps spiking to 3000, the snr dropping, and the contrast oscillating between -3276.6 and 3274.2. this confirms the reported failure mode. event # 1036 was not cleared until the pump was removed on (b)(6) 2023. note: this console was not returned for investigation of complaint (b)(4). however, it was returned for complaint (b)(4), which indicates the same failure mode: abnormal optical parameters and the alarm ¿controller error (defective optical sensor lamp)¿ ¿ event 1039, which was displayed during the next console power-on after the case with pump sn: (b)(6), which ended on july 6, 2023. device analysis: this console was not returned for investigation of complaint (b)(4); however, it was returned and investigated for complaint (b)(4), which indicates the same failure mode. root cause: the root cause of the loss of the placement signal was most likely the malfunction of the optical bench lamp. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3681 | 1220648-2025-48091 | 23473002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23473002&pc=OZD | 10/12/2025 | 11/04/2025 |  | 11/05/2025 |  |  | 10/12/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026671587 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 5 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that there were clots and access site bleeding during impella cp support and that patient encountered ectopy. a few days into support, patient encountered ectopy. as a result, adrenaline infusion was discontinued and patient was given bolus of amiodarone, potassium and magnesium. some oozing was noted this morning around groin dressing. as a result, pressure was applied and groin area was cleaned and redressed. the site was no longer oozing. additionally, patient was noted to have significant clots during implella implant and was on a number of iv anticoagulants, such as heparin, tirofiban, asprin, and clopidorel. the patient later remained on heparin only. then, several days later, the patient had some fresh oozing noted from the groin and 1 unit of packed red blood cells was given. the groin was cleaned and redressed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3682 | 1220648-2025-48092 | 23475579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23475579&pc=OZD | 10/18/2025 | 11/04/2025 |  | 11/05/2025 |  |  | 10/18/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1586834 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that they encountered a "purge not detected" alarm. it was reported that during set-up, the alarm "purge disc not detected" appeared, followed by the alarm "purge system failure". changing the purge cassette did not solve the problem. after switching to the back-up controller, set-up was performed without any problem. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 has been updated to reflect product was returned for investigation. device status. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
3683 | 1220648-2025-48094 | 23475699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23475699&pc=OZD | 10/14/2025 | 11/04/2025 |  | 11/05/2025 |  |  | 10/14/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026758517 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 66 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella cp support was taken to the cath lab for escalation of treatment to an impella 5.5. the patient experienced ventricular tachycardia and was defibrillated and placed on veno-arterial extracorporeal membrane oxygenation peri-arrest. impella cp was successfully weaned and explanted in the procedural area. patient hemodynamically stable on impella 5.5 support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4, h6 health effect - clinical code 060109 was inadvertently omitted on the initial manufacturer device report the investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3684 | 1220648-2025-48095 | 23475817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23475817&pc=OZD | 10/13/2025 | 12/01/2025 |  | 11/05/2025 |  |  | 10/13/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734226 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female | 82 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient encountered a stroke after impella 5.5 support. the patient had been on impella 5.5 support for 8 days and 2 days after pump explant, the patient experienced a stroke. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
3685 | 1220648-2025-48097 | 23476021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476021&pc=OZD | 10/07/2025 |  |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657871 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female | 65 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced sepsis and multiple organ failure. care was withdrawn and the patient expired. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the cause of the organ failure was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3686 | 1220648-2025-48099 | 23476063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476063&pc=OZD | 10/07/2025 |  |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677261 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male | 70 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implante with an impella cp experienced access site bleeding. the physician stated that the bleeding was related to the initial puncture of the femoral artery. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the major bleed was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3687 | 1220648-2025-48100 | 23476083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476083&pc=OZD | 10/07/2025 |  |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026744982 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp had left ventricle dysfunction and required multiple inotropes and vasoactives. an old, encapsulated apical thrombus was identified. the pump was explanted and the patient was escalated to an impella 5.5. | [5 empty] | No product was returned. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. the pump passed all post-sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3688 | 1220648-2025-48101 | 23476089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476089&pc=OZD | 10/09/2025 |  |  | 11/05/2025 |  |  | 10/09/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211527 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an optical sensor error on the supporting automated impella controller (aic). the system self-check failed, the aic would not turn off, and it was making a continuous noise. a different aic was used to support the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3689 | 1220648-2025-48102 | 23476125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476125&pc=OZD | 10/13/2025 | 11/04/2025 |  | 11/05/2025 |  |  | 10/10/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704878 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 75 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Sepsis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient encountered sepsis and infection during impella cp support. the bipella patient remained on cp support after rp was removed. during cp support, the doctor suspected sepsis and a few days later patient was confirmed to have sepsis. it was noted that the blood cultures revealed that they found patient to have sever candidiasis fungal infection. the medical team decided to withdraw care and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported event has been completed. no product was returned for investigation. the root cause of the injury was unable to be determined due to insufficient clinical details.
3690 | 1220648-2025-48096 | 23476154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476154&pc=OZD | 10/13/2025 | 11/04/2025 |  | 11/05/2025 |  |  | 10/13/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659742 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered pump stop and kink during impella 5.5 support. it was reported that a pump stop was encountered. attempts to start the impella at different p levels have been unsuccessful. the perfusionist also reported that there was a kink between the red plug and the pump. as a result the aic console was exchanged as well as the pump. the patient remained stable on the new impella 5.5. the first pump was collected and planned to be sent back to abiomed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: impella 5.5 sn (b)(6) + purge cassette returned. temporary pump stop/product damage (catheter kink). there are no data logs returned. the clinical stated that there was a sudden pump stop but was able to restart on another aic. due to the stoppage that occurred before, the physicians opted to replace the pump. the pump was returned and the catheter was kinked at the red impella plug. it was electrical tested and passed continuity from phases 1 to 3, but 1 to 2 and 2 to 3 would go in and out from in-spec to 0l when the kink was positioned a certain way. the pump was run and in a bucket and was able to run on p-9 without issue but when the catheter was kinked near the plug and pressed upon, the pump would stop, triggering #119 alarms. based on the returned product, the root cause of the temporary pump stop was most likely due to electrical open lines due to a user error. the cause of the product damage (catheter kink) is most likely due to use as excessive force was applied to the catheter at that location. device history review: impella 587742 passed all post sterile inspection checks.
3691 | 1220648-2025-48106 | 23476457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476457&pc=OZD | 10/09/2025 | 11/05/2025 |  | 11/05/2025 |  |  | 10/09/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667824 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. it was reported that the impella insertion was aborted due to the patient¿s vasculature. no patient harm was reported. | [5 empty] | No product was returned. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy and calcification along with resistance at aorta preventing successful delivery of impella. the device passed all post sterile inspection checks. | A2, a4 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3692 | 1220648-2025-48107 | 23476537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476537&pc=OZD | 10/08/2025 | 11/05/2025 |  | 11/05/2025 |  |  | 10/09/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689085 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient after undergoing a coronary artery bypass graft, experienced cardiac arrest after being transferred to the intensive care unit. the patient was cannulated for veno-arteria extracorporeal membrane oxygenation (va ecmo) and the decision was made to place an impella 5.5 to unload the left ventricle. the impella 5.5 was successfully placed, and the patient was sent back to the unit on impella 5.5 and va ecmo support. on day 5 of impella 5.5 and va ecmo support, the patient was brought to the operating room for a chest washout. the following day, the patient was noted to have heparin induced thrombocytopenia. heparin was discontinued and the patient was started on agatroban. on day 7 of impella 5.5 and va ecmo support, the patient experienced a subdural bleed noted which was noted on computed tomography. all anti-coagulation therapy was discontinued. neurosurgery was consulted, but the decision was the patient was not a candidate for a craniotomy due to being on va ecmo and impella 5.5 support. on day 15 of impella 5.5 and va ecmo support, the family made the decision to withdraw the patient¿s care. care was withdrawn and the patient expired. | [5 empty] | Based on the information available, it is unlikely that the stroke was related to the use of the impella device; however, a device-related contribution cannot be definitively ruled out. although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary : stroke : the cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia : the cause of the injury was unable to be determined due to insufficient clinical details. device history lot : device lot : 1932198. device history review : the pump s/n: (b)(6) passed all the post sterile inspection checks.
3693 | 1220648-2025-48108 | 23476813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23476813&pc=OZD | 10/09/2025 |  |  | 11/05/2025 |  |  | 10/09/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734165 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Inflammation | Ischemia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced swelling in their right arm and ectopy. the patient was treated with milrinone. during explantation of the pump a clot was removed from the graft. the patient survived. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the arrhythmia and thrombosis was unable to be determined due to insufficient clinical details. the cause of the inflammation was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3694 | 1220648-2025-31914 | 23478730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23478730&pc=OZD | 12/18/2024 | 11/05/2025 |  | 11/05/2025 |  |  | 12/18/2024 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1621395 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella automated impella controller had optical signal issue. the impella cp was running without any problems. there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. partial console logs (obtained from the cloud) from the reported day of event were analyzed with respect to pump performance and suggested that the issue might be caused by the console. console had not been returned for investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3695 | 1220648-2025-48109 | 23479120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23479120&pc=OZD | 09/07/2025 | 11/05/2025 |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657743 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while a patient was on impella cp support that the patient became agitated yesterday which resulted in the 5.5 becoming malpositioned and ultimately explanted. the next day patient¿s pressor/intropic support drastically increased and the decision was made to place another 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient when the patient became agitated. device history review: the device passed all post sterile inspection checks.
3696 | 1220648-2025-48110 | 23479128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23479128&pc=OZD | 10/06/2025 | 12/05/2025 |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025558755 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that after the impella cp was implanted for cardiogenic shock. the patient had to go in for vascular surgery due to peripheral complications and the impella had been removed. no other information has been provided. | [5 empty] | The impella device was discarded by the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product was returned. vessel perforation (peripheral): the cause of the injury was not determined due to insufficient clinical details. clinical review: 2025-oct-6: impella cp inserted. the physician said that they are undertaking a debrief but no further information was received. impella explant date unknown. device history review device passed all post sterile inspection checks. | H6 impact code f19 has been added.
3697 | 1220648-2025-48111 | 23479613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23479613&pc=OZD | 10/08/2025 | 01/07/2026 |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713513 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge system block and low purge flow. lytics were added to the purge to resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Investigation summary impella 5.5 sn (b)(6) + purge cassette returned for evaluation. the returned product was investigated. the cannula was received cut from the catheter, so the pump was unable to be run in the lab. the cannula was soaked in a mixture of warm water and terga zyme to get rid of dried blood and check for biomaterial. a large amount of dried blood was found in the inflow cage. biomaterial was found in the purge gap, underneath the impeller blades. no other abnormalities were seen. data logs were returned for analysis. suction alarms were seen at the beginning of the case, along with impella position unknown alarms. high purge pressure (hpp) and purge system blocked alarms were confirmed around 11pm on (b)(6) 2025, which did not resolve for the remainder of the case. drop in purge flows was confirmed during the alarms. a gradual rise in the purge pressure was seen at the onset of the hpp alarms. no motor current spikes were seen outside of p-level changes, but the gradual rise in purge pressure coincided with a gradual rise in motor current for the same p-level. no other abnormalities were seen. clinical details mention that the purge system was blocked, and purge flow dropped below 1 ml, and lysis protocol was discussed. the patient was also on ecmo support during the case. purge pressure high: biomaterial was found in the purge gap from the product evaluation. data log analysis confirmed drop in purge flows and high purge pressure alarms, with gradually rising purge pressure values. clinical details mention that the purge system was blocked and purge flows dropped below 1 ml. the root cause of the high purge pressure was most likely biomaterial buildup based on the evaluation of the returned product, data log analysis, and provided clinical details. device history lot device lot: 1946957 device history batch subcomponent lot: n/a device history review pump #(b)(6) passed all post-sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3698 | 1220648-2025-48112 | 23479818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23479818&pc=OZD | 10/06/2025 | 11/05/2025 |  | 11/05/2025 |  |  | 10/06/2025 | 11/05/2025 |  |  |  |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725419 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 73 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. the patient was transfused blood products and the site was surgically addressed. the patient also had hemolyzed labs. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A review of the data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. no product was returned for investigation. the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the device passed all post sterile inspection checks.
3699 | 1220648-2025-48113 | 23479972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23479972&pc=OZD | 09/18/2025 | 02/18/2026 |  | 11/05/2025 |  |  | 10/07/2025 | 11/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631600 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while implanted in a patient the impella cp pump lost optical signal. impella support continues. | The complainant reported additional information upon request: "to address the issue, the customer was reassured by ensuring the pump was functioning properly and providing adequate flow for the patient. i have the pump and can return it." |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. | No product or data logs were returned. the clinical details state that after impella cp and mitraclip implantation, the aortic and ventricular curves did not appear on p-8 or p-9. it is unknown if this resolves, if any alarms were triggered or when the pump was explanted. due to lack of clinical information, data logs or product, the root cause of the optical sensor issue cannot be determined. the device passed all post sterile inspection checks.
3700 | 1220648-2025-31915 | 23480672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23480672&pc=OZD | 02/24/2023 | 11/05/2025 |  | 11/05/2025 |  |  | 02/24/2023 | 11/05/2025 |  |  | 01/31/2025 |  | 11/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024296595 |  |  |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported intermittent false position alarms during patient support. there was no patient harm and support continued with the implanted pump. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for investigation. data logs were returned for analysis. data logs confirmed the impella positioning alarms. the 5s parameters were stable but snr and contrast were widespread, with snr low values dropping to 0 throughout the case. no placement signal alarms were observed, however, placement signal did spike to 3000 in the middle and towards the end of the case. no motor current trends or spikes were seen out of p-level changes. optical signal issue: the root cause of the optical signal issue was not determined due to lack of sufficient clinical details, unreturned product, and inconclusive evidence from data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3701 | 1220648-2025-48036 | 23462743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23462743&pc=OZD | 10/05/2025 | 11/04/2025 |  | 11/04/2025 |  |  | 10/05/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311806 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 103 KG | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that an automated impella controller (aic) had a self-diagnosis failure when starting up. the aic was restarted, but the same issue occurred again. another aic was used successfully. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3702 | 1220648-2025-48047 | 23463129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23463129&pc=OZD | 10/05/2025 |  |  | 11/04/2025 |  |  | 10/05/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713395 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Atrial Fibrillation | [11 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient complained of chest pain over the last two weeks and presented to the emergency department in anterior st-elevation myocardial infarction. the patient was taken for coronary intervention to the left anterior descending coronary artery, and the right coronary artery was found to be in chronic total occlusion and the circumflex artery with 95% lesion. an impella cp was successfully placed, and surgery was consulted for revascularization options. on impella cp support day 4, the patient was taken to the operating room for coronary artery bypass grafting and aortic valve replacement. after the procedure, the impella cp was explanted, and the patient was escalated to an impella 5.5. on impella 5.5 support day 3, the patient experienced an episode of atrial fibrillation (afib) with rapid ventricular response (rvr). the patient was cardioverted three times, however the patient continued to be in afib. the following morning the patient was successfully cardioverted into normal sinus rhythm. on support day 6, it was reported that the patient experienced another episode of afib with rvr. the patient was cardioverted and was cardioverted into normal sinus rhythm. on support day 10, the patient experienced low hemoglobin and hematocrit results and was transfused with one unit of red blood cells. the patient was unable to separate from mechanical circulatory support. the decision was made to transition the patient to palliative/hospice care and to withdraw care. the patient expired. | [5 empty] | Additional information added to d6b date of explant. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. no product was returned for investigation. paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. anemia: the cause of the clinical symptom was not determined due to insufficient clinical details. device history (lot) device lot: 1951070 device history review device with sn: (b)(6) passed all post sterile inspection checks. | B.2 date of death is unknown. d.6b date of explant is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3703 | 1220648-2025-48048 | 23463175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23463175&pc=OZD | 10/04/2025 | 01/08/2026 |  | 11/04/2025 |  |  | 10/05/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720033 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient noted as high-risk ventricular septal defect (vsd) repair was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a previous impella cp implanted and was additionally cannulated for extracorporeal membrane oxygenation for added support. the patient¿s pulmonary artery pressure remained high and mean atrial pressure was low, so the decision was made to escalate the patient to an impella 5.5 after two days of support on the impella cp. the vsp repair was performed successfully, however a myocardial injury occurred during manipulation of the heart for the procedure. it was noted that the impella 5.5 did not fully penetrate, it was slightly visible through the epicardium. tachoseal was applied to the injured area, and the patient was noted to be stable. the patient remained on impella support to allow the left ventricle to decompress and to await hemostasis. | [5 empty] | A4 is unknown. follow-up information regarding patient outcome and whether the impella 5.5 was explanted or remains implanted is currently unknown and not available. the patient's status currently is unknown. a supplemental report will be submitted if additional information is received. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. mechanical interaction w tissue: clinical details noted that a myocardial injury occurred during rotation of the heart in a vsp repair. the impella was not penetrating but was slightly visible through the epicardium. the cause of the mechanical interaction with tissue was a use related issue as the heart was rotated during the procedure according to clinical details. device history lot: device lot: 1932248. device history batch: subcomponent lot: n/a. device history review device with sn: (b)(6) passed all post sterile inspection checks. | D6b: explant date added.
3704 | 1220648-2025-48049 | 23463259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23463259&pc=OZD | 10/08/2025 |  |  | 11/04/2025 |  |  | 10/08/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704408 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 61 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock. during percutaneous coronary intervention, the patient had a moment of pulseless electrical activity arrest during stent placement but recovered after balloon was down. the patient had a warm foot but no pulse while the catheter was in place. pulse was present at popliteal. after pump explant, pulse still was not present. angiograph found a clot that was removed by thrombectomy, restoring flow to the foot. angioseal was used for hemostasis following impella removal. the patient survived. | [5 empty] | No product was returned for investigation. the root cause of the cardiac arrest, ischemia, and thrombosis was unable to be determined due to insufficient clinical details. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.".
3705 | 1220648-2025-48050 | 23463602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23463602&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740535 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the patient had a previous intra-aortic balloon pump placed in the cardiac catheterization laboratory after becoming hemodynamically unstable due to a pinched left circumflex artery and a pseudoaneurysm on her inferior wall. due to severe aortic calcification, despite many attempts, the impella cp would not advance past the aortic arch. the medical team decided to abort the impella cp implantation and decided to place an impella 5.5 instead. the impella 5.5 was successfully placed with no complications. the impella 5.5 was explanted successfully after weaning. the patient's outcome at the time of explant was survived. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had calcification and difficult anatomy preventing successful delivery of impella. device history lot: device lot: 1970191. device history review:device (sn: (b)(6) passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3706 | 1220648-2025-48051 | 23463663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23463663&pc=OZD | 09/18/2025 |  |  | 11/04/2025 |  |  | 09/18/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713022 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Flutter | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced patient respiratory arrest. the patient was intubated and stabilized without need for resuscitation. the next day the patient experienced bleeding from bilateral groin sites. in response, the patient was transfused blood products. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the respiratory failure and atrial flutter was not determined due to insufficient clinical details. the cause of the bleeding is determined to be patient condition because of the bleeding form multiple sites. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3707 | 1220648-2025-48053 | 23463932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23463932&pc=OZD | 09/19/2025 | 11/04/2025 |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719763 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Sepsis | Tachycardia | [10 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support (mcs) as elective placement. flows were compromised and the device was replaced with another impella 5.5 device. the patient experienced adverse events and would ultimately expire; care was withdrawn. the patient was brought in for a scheduled mini mitral valve. plans for thoracotomy approach. the patient attempted to separate from bypass and experienced signs of postcardiotomy cardiogenic shock (pccs) & low cardiac output syndrome (lcos). the operating room team and the impella team for implant of impella 5.5 device were notified. the right axillary was cut down and a 10 mm hemashield platinum graft was sewn on. introducer with graft locks in place. rosen curve and glide catch used to access left ventricle under fluoroscopy and transesophageal echocardiogram. this was exchanged for abiomed 0.018 wire. the impella 5.5 was placed in left ventricle and device started at p-2. goal flows of 2.0-2.3 lpm achieved with impella on p-6. at end of case before transitioning patient to veno-arterial-venous extracorporeal membrane oxygenation (vav ecmo) as planned, the physician to unplug white cable from impella console as it was twisted in the ecmo cannulas he was needing to access. upon removal and immediate reconnection, the impella resumed at p-6 but flows were only 0.8-0.9 lpm. flows before removal were around 2 lpm. troubleshooting was performed to access function at different performance levels. flows would not come above .9lpm. the physician ultimately decided to explant and replace with a new impella 5.5 device using same right axillary access, which was successfully completed. post replacing the impella 5.5 device, the patient was converted from veno-arterial (va) ecmo to vav ecmo without issue. the patient was transferred to the unit in stable condition (now on impella 5.5 pump 2). after arriving to the unit, the patient had two episodes of ventricular tachycardia that required multiple shocks and administration of lidocaine and amiodarone. thereafter the patient's rhythm was stable with nurse able to wean inotropes/pressors. the following day, the impella was repositioned under echocardiogram. left ventricle measurement of 4.3 cm with flows of 2.2 lpm and appropriate waveforms. the patient passed the awakening trial in the am but planned to put the patient back under sedation once the patient was more stable. now the following day, the plan to decannulate ecmo in three or four days was made. the patient was awake on vent following commands. the patient continued on support. weaning ecmo as tolerated and the patient remained hemodynamically stable in critical yet stable condition. now day 6 and the plan was to originally decannulate ecmo today but infectious disease and heart team now stating the patient was septic and will remain on support for now. antibiotic was initiated. the patient continued on support in critical condition. the patient was on multiple pressors, continuous renal replacement therapy (crrt) and ecmo. lactate continued to be elevated despite therapies. heart team felt this would be difficult to salvage. the patient still in critical condition. the plan was unclear as prognosis was "very poor." the patient eventually expired on day 11 of support. care was withdrawn. | [5 empty] | Patient-specific information a4 weight is unknown. medical safety review: medical safety reviewed the event. the event involving the first impella 5. 5 device describes failure of the pump to obtain adequate flows after the physician had the white cable from impella console be unplugged as it was twisted in the ecmo cannulas he was needing to access. troubleshooting was performed (details on actions taken unnoted) to no avail. the impella 5.5 device was explanted and replaced with a new impella 5.5 device. the physician's negligible actions of unplugging the connect cable that links the impella pump catheter to the control unit likely caused the event. the pump was unable to recover the appropriate flow level. this first pump had a device malfunction but did not contribute to the demise outcome. this is malfunction type of reportable event. the second impella 5.5 device did not malfunction. however, the patient experienced ventricular tachycardia shortly after implant arriving to the unit treated with shocks and medication to resolve the arrhythmic event. there was difficulty weaning vav ecmo and eventually was unable to be weaned noted on day 6. day 6, the patient was noted to be septic started on antibiotic. the patient remained hemodynamically stable in critical but stable condition. the patient was on multiple pressors, crrt, and ecmo. lactate continued to be elevated despite therapies. the plan was unclear as prognosis was "very poor" and the patient's condition was unsalvageable. day 11 care was withdrawn and the patient expired. the patient became septic treated with antibiotic and it is unknown how, if any, the impella 5.5 pump played a role. there is no information that indicates or reasonably suggests that the patient expired from the arrhythmic event or being septic. however, there is not enough evidence to rule out, dissociate the impella 5.5 from the patient's outcome. although this dissociation cannot be made, given the patient's condition was unsalvageable and that care was electively withdrawn, it is unlikely that the impella would be associated to the patient's cause of death. the patient's underlying condition likely contributed most to an outcome of death. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the first impella 5.5 device. | H6 revised codes to remove arrhythmia and add ventricular tachycardia. resuscitation code was also removed. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. h10 related manufacturer report number as it was omitted from the initial report that was submitted. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the pump passed all the post sterile inspection checks. regarding, sepsis and ventricular tachycardia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details.
3708 | 1220648-2025-48054 | 23464112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464112&pc=OZD | 09/19/2025 | 11/04/2025 |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700092 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp developed a hematoma. the physician explanted the pump and implanted another impella on the left side. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. the patient experienced cardiac arrest at home. upon emergency medical services (ems) arrival, the patient was found to be in ventricular fibrillation. cardiopulmonary resuscitation (cpr) was initiated by the patient¿s spouse and continued by ems. the patient was in cardiogenic shock from an acute myocardial infarction. the patient was transported emergently to the cardiac catheterization laboratory. coronary angiography revealed "severe" left main (lm) and left anterior descending (lad) coronary artery disease. plain old balloon angioplasty of the lad was attempted. during the procedure, the patient developed recurrent ventricular fibrillation requiring cpr. an impella cp device was implanted via right femoral access on to provide hemodynamic support. thereafter, the patient was transferred to a tertiary care facility for surgical evaluation; coronary artery bypass graft surgery was declined. upon arrival, a right femoral access-site hematoma was noted. the physician elected to explant the initial impella cp pump 1 (after 3 hours of mechanical circulatory support). the neurologic status remained a concern due to unknown pre-hospital downtime. the outcome of the patient was noted as expired. although a right femoral access-site hematoma was observed following the placement of the impella cp device, vascular access-site hematoma is a known potential complication of a femoral access. there is no information or evidence to suggest that the hematoma contributed directly to the patient¿s demise. however, the death will be conservatively reported. the patient's underlying condition contributed most if not solely to the demise outcome (refractory cardiogenic shock secondary to acute myocardial infarction involving severe left main (lm) and left anterior descending (lad) coronary artery disease, complicated by recurrent ventricular fibrillation and prolonged cardiac arrest). note: h6: device problem/component/investigation coding remains unchanged. | The investigation of the reported failure mode has been completed. no product was received for evaluation. asae/hematoma: the cause of the access site adverse event is use related to patient movement as hematoma was noted after hospital transfer. device history review: the device passed all post sterile inspection checks. complaint coding was updated to better reflect the reported event. consequently, section h6 impact codes and medical device problem codes were updated/corrected and repopulated accordingly. date of death was unknown.
3709 | 1220648-2025-48055 | 23464144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464144&pc=PYX | 09/19/2025 | 11/04/2025 |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026712277 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Renal Failure | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced heparin-induced thrombocytopenia and a clot in the groin. the patient was treated with argatroban and one unit of packed red blood cells. the patient was given volume in response to suction events. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the renal failure, thrombosis, and thrombocytopenia was not determined due to insufficient clinical details. the cause of the suction was most likely related to the patient's condition, as clinical reported the patient had right ventricle dysfunction and was receiving volume. the placement signal was trending down at the time of suction, indicating that the patient was likely experiencing volume issues. the device passed all post sterile inspection criteria.
3710 | 1220648-2025-48056 | 23464150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464150&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740342 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced hematuria. proper pump position was confirmed via echocardiogram. systemic vascular resistance (svr) was lowered and they added nipride to lower svr. the impella cp was successfully weaned and explanted as the patient went in for ventricular assist device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The pump was not returned for investigation. the data logs do not confirm the complaint. the cause of the hemolysis was most likely related to patient condition, based on data log review and provided clinical details. the pump passed all post sterile inspection checks.
3711 | 1220648-2025-48057 | 23464181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464181&pc=PYX | 10/09/2025 |  |  | 11/04/2025 |  |  | 10/09/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661942 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 105 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced a placement signal issue. no patient harm was reported. impella support continues. | [5 empty] | No product was returned for investigation. the cause of the placement signal issue cannot be determined without returned product. the device passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3712 | 1220648-2025-48059 | 23464235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464235&pc=OZD | 10/20/2025 |  |  | 11/04/2025 |  |  | 10/20/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716671 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 was suspected of suffering a clot. repositioning the pump was unsuccessful. the impella 5.5 was explanted and exchanged for an impella cp. the patient survived. | [5 empty] | No product was returned for investigation. the cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿.
3713 | 1220648-2025-48058 | 23464370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464370&pc=OZD | 09/19/2025 |  |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729993 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to heavy calcification of the patient's vessels. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had heavy calcification in the vessels preventing successful delivery of impella. the device passed all post sterile inspection checks.
3714 | 1220648-2025-48060 | 23464446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464446&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722940 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impella pump was implanted without any problems, and protect percutaneous coronary intervention was also performed without any problems. the pump explant went smoothly. the groin was closed with two femoseal and was unremarkable. the patient was stable and catecholamine-free after explant. during the procedure, however, it was noticed that placement signal and left ventricle were so elevated, sometimes for a screen length, that the signal was completely out of range (see screenshots). pump flow and purge system were unremarkable during this time. external monitoring also revealed no abnormalities. the pump function did not appear to be impaired by the event. | [5 empty] | The investigation for the reported event, optical sensor issue, has been completed. no product or data logs were returned. the clinical details state that during the pci procedure with the pump, the ps and lv were elevated and the signal was completely out of range. external bp monitoring revealed no abnormalities. due to a lack of sufficient information, product and data logs returned, the root cause of the optical sensor issue cannot be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3715 | 1220648-2025-48061 | 23464456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464456&pc=OZD | 09/18/2025 |  |  | 11/04/2025 |  |  | 09/18/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717720 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost pedal pulses on both feet caused by a thrombus. the patient was treated with aggrastat and brillinta. the pump was repositioned in response to hemolysis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Data logs confirmed the complaint condition. the product was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the ischemia and thrombus was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3716 | 1220648-2025-48062 | 23464547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464547&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743649 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp was initially plugged in and alarmed impella pump malfunction. the device supported the patient for approximately two and a half hours before being weaned and explanted in the procedural area. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6 medical device problem code was updated to a1301 to better reflect what was reported. investigation summary: the investigation was completed. no product was returned. pump not detected: the initial pump was plugged in and said impella pump malfunction. data log review showed no issue alarms noted. real time log showed slightly lower than expected lamp at 76. the logs showed the pump was plugged in and recognized without issue but then eeprom write blocked. logs end one hour later. there are no relevant alarms in the logs. the cause of the pump not detected was not determined due to no pump returned, inconclusive data log review, and insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
3717 | 1220648-2025-48063 | 23464568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464568&pc=OZD | 09/19/2025 |  |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725294 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support (mcs) as elective placement. flows were compromised and the device was replaced with another impella 5.5 device. the patient experienced adverse events and would ultimately expire; care was withdrawn. the patient was brought in for a scheduled mini mitral valve. plans for thoracotomy approach. the patient attempted to separate from bypass and experienced signs of postcardiotomy cardiogenic shock (pccs) & low cardiac output syndrome (lcos). the operating room team and the impella team for implant of impella 5.5 device were notified. the right axillary was cut down and a 10 mm hemashield platinum graft was sewn on. introducer with graft locks in place. rosen curve and glide catch used to access left ventricle under fluoroscopy and transesophageal echocardiogram. this was exchanged for abiomed 0.018 wire. the impella 5.5 was placed in left ventricle and device started at p-2. goal flows of 2.0-2.3 lpm achieved with impella on p-6. at end of case before transitioning patient to veno-arterial-venous extracorporeal membrane oxygenation (vav ecmo) as planned, the physician to unplug white cable from impella console as it was twisted in the ecmo cannulas he was needing to access. upon removal and immediate reconnection the impella resumed at p-6 but flows were only 0.8-0.9 lpm. flows before removal were around 2 lpm. troubleshooting was performed to access function at different performance levels. flows would not come above .9lpm. the physician ultimately decided to explant and replace with a new impella 5.5 device using same right axillary access, which was successfully completed. post replacing the impella 5.5 device, the patient was converted from veno-arterial (va) ecmo to vav ecmo without issue. the patient was transferred to the unit in stable condition (now on impella 5.5 pump 2). after arriving to the unit, the patient had two episodes of ventricular tachycardia that required multiple shocks and administration of lidocaine and amiodarone. thereafter the patient's rhythm was stable with nurse able to wean inotropes/pressors. the following day, the impella was repositioned under echocardiogram. left ventricle measurement of 4.3 cm with flows of 2.2 lpm and appropriate waveforms. the patient passed the awakening trial in the am but planned to put the patient back under sedation once the patient was more stable. now the following day, the plan to decannulate ecmo in three or four days was made. the patient was awake on vent following commands. the patient continued on support. weaning ecmo as tolerated and the patient remained hemodynamically stable in critical yet stable condition. now day 6 and the plan was to originally decannulate ecmo today but infectious disease and heart team now stating the patient was septic and will remain on support for now. antibiotic was initiated. the patient continued on support in critical condition. the patient was on multiple pressors, continuous renal replacement therapy (crrt) and ecmo. lactate continued to be elevated despite therapies. heart team felt this would be difficult to salvage. the patient still in critical condition. the plan was unclear as prognosis was "very poor." the patient eventually expired on day 11 of support. care was withdrawn. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. low pump flow: clinical reported that the pump was unplugged from the console in order to untangle the white cable from the extracorporeal membrane oxygenation (ecmo) cannulas. once the pump was restarted, flows were no longer reaching expected levels at p-6 and troubleshooting did not resolve the issue. the product was not returned for investigation. data log review confirms low flows occurring after the pump was restarted. suction alarms were triggering prior to restarting the pump due to low pulsatility, which was likely due to the patient being on veno-arterial (va) ecmo support. however, the pump was still achieving expected flows at this time. after the pump was restarted, flows become clipped. the cause of the suction was most likely patient condition because suction started once the patient was placed on ecmo, and data logs show low placement signal pulsatility when flows dropped. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Patient-specific information a4 weight is unknown. medical safety review: medical safety reviewed the event. the event involving the first impella 5. 5 device describes failure of the pump to obtain adequate flows after the physician had the white cable from impella console be unplugged as it was twisted in the ecmo cannulas he was needing to access. troubleshooting was performed (details on actions taken unnoted) to no avail. the impella 5.5 device was explanted and replaced with a new impella 5.5 device. the physician's negligible actions of unplugging the connect cable that links the impella pump catheter to the control unit likely caused the event. the pump was unable to recover the appropriate flow level. this first pump had a device malfunction but did not contribute to the demise outcome. this is malfunction type of reportable event. the second impella 5.5 device did not malfunction. however, the patient experienced ventricular tachycardia shortly after implant arriving to the unit treated with shocks and medication to resolve the arrhythmic event. there was difficulty weaning vav ecmo and eventually was unable to be weaned noted on day 6. day 6, the patient was noted to be septic started on antibiotic. the patient remained hemodynamically stable in critical but stable condition. the patient was on multiple pressors, crrt, and ecmo. lactate continued to be elevated despite therapies. the plan was unclear as prognosis was "very poor" and the patient's condition was unsalvageable. day 11 care was withdrawn and the patient expired. the patient became septic treated with antibiotic and it is unknown how, if any, the impella 5.5 pump played a role. there is no information that indicates or reasonably suggests that the patient expired from the arrhythmic event or being septic. however, there is not enough evidence to rule out, dissociate the impella 5.5 from the patient's outcome. although this dissociation cannot be made, given the patient's condition was unsalvageable and that care was electively withdrawn, it is unlikely that the impella would be associated to the patient's cause of death. the patient's underlying condition likely contributed most to an outcome of death. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the second impella 5.5 device, which is associated with the demise outcome.
3718 | 1220648-2025-48065 | 23464683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464683&pc=OZD | 09/19/2025 |  |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738945 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp accidentally pulled the pump out of the ventricle, and experienced hemolysis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data logs confirmed the complaint. the product was not returned for review. the cause of device in wrong position was most likely use related since the clinical team confirmed impella got pulled on possibly by patient who is sundowning. the device passed all post sterile inspection checks.
3719 | 1220648-2025-48064 | 23464739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464739&pc=OZD | 10/06/2025 | 02/18/2026 |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736355 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported additional information upon request: "nothing was done! patient recovered full aortic valve function by herself and her ef went back to normal! device was discarded i believe" | It was reported that a patient was on impella 5.5 support for a total of 95.26 hours without any complications and was successfully weaned and explanted. however, upon removal of the impella an echocardiogram revealed moderate aortic insufficiency that was not present prior to impella insertion. there was a small gap between non-coronary cusp and left coronary cusp. no calcification or history of endocarditis. |  |  |  |  | The investigation of the reported failure mode has been completed. no product was received for evaluation. heart valve incompetence : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | B5 updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿. section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. section: use of echocardiography for positioning of impella catheter: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿. section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿.
3720 | 1220648-2025-48067 | 23464824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464824&pc=OZD | 10/09/2025 |  |  | 11/04/2025 |  |  | 10/09/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737492 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Fever | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced a fever so antibiotics were given. the patient was suspected of having heparin-induced thrombocytopenia, and had chronic bleeding from the nose and lungs, so anticoagulation was withheld and bivalirudin was added to the purge. the patient remained in a state of atrial fibrillation as cardioversion was unsuccessful. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the bleeding is determined to be patient condition because of the bleeding from multiple locations. the cause of the fever, thrombocytopenia, and epistaxis was not determined due to insufficient clinical details. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular. perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3721 | 1220648-2025-48066 | 23464872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464872&pc=OZD | 09/19/2025 |  |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689074 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed heparin induced thrombocytopenia and a clot in the groin. one unit of blood was transfused and continuous renal therapy was started. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the thrombocytopenia and renal failure was not determined due to insufficient clinical details. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3722 | 1220648-2025-48068 | 23464984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23464984&pc=OZD | 09/19/2025 |  |  | 11/04/2025 |  |  | 09/19/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | False Alarm | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported that there was arrhythmia and a false alarm/buzzer during impella 5.5 patient support. while on support, there was a white advisory for impella position unknown. to resolve the issue, the abiomed representative showed the nurse how to adjust the left ventricle signal. additionally, the patient had atrial fibrillation and was treated with amiodarone. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3723 | 1220648-2025-48070 | 23465090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23465090&pc=OZD | 10/09/2025 |  |  | 11/04/2025 |  |  | 10/09/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642604 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Crack | Device Sensing Problem | [5 empty] | Stroke/CVA | Unspecified Infection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced an unreliable placement signal. the patient suffered a cerebrovascular accident, clots, and an infection. due to manual pressure applied to the purge sidearm, the device broke. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3724 | 1220648-2025-48069 | 23465217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23465217&pc=OZD | 10/09/2025 |  |  | 11/04/2025 |  |  | 10/09/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485980 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced flow issues and left ventricle waveform decoupling. the device malposition was confirmed with ecmo. repositioning was attempted but the pump was accidently pulled out of the left ventricle. the patient was taken to surgery to access the left ventricle when the pump began to have purge flow issues and eventually a purge block. the doctor decided to exchange the pump for a new impella 5.5. another report will be submitted to capture the patient expiring while being supported by the second pump. | [5 empty] | The pump was not returned for investigation. this complaint investigates the first utilized 5.5 s2 pump. as per the clinical details, both the current and replacement devices consistently exhibited low pump flow despite attempts to optimize support and reposition the pump. persistent flow issues occurred alongside ongoing hemodynamic instability, decoupled lv waveforms, and increasing requirements for inotropes and pressors. these findings suggest that low pump flow was most likely due to the patient¿s critical condition and poor cardiac function. the cause of the positioning issue was most likely unintentional use-related, based on the given clinical details. the cause of the hemodynamic instability was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3725 | 1220648-2025-48071 | 23465529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23465529&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722810 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support was presented with a gastrointestinal bleed that required heparin to be withheld. explanting of the device was delayed. the patient was eventually weaned and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks.
3726 | 1220648-2025-48072 | 23465692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23465692&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728961 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 117 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the impella 5.5 experienced placement signal lost (psl) alarms as well as some suction alarms. waveforms appeared as the device was potentially a little deep. they were advised to confirm the position. the psl alarm audio was disabled. there were no concerns of hemolysis. there is harm to the patient. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: placement signal low alarms reported. the data logs confirm placement signal low alarms at various points during support which coincided with the occurrence of suction alarms. the cause of the placement signal low alarms was not established as no product was returned and insufficient clinical details were provided. optical sensor issue: the logs show placement signal not reliable alarms and the 5s parameters are of the pattern identified as an oscillating contrast issue. this issue has been characterized as a console-related issue. pump suction: the data logs confirm characteristic curve continuous suction alarms which were suspended during loss of placement signal related to a console issue. there were no motor current spikes outside p-level changes and pump flows remained stable and within expected range for matching p-level. purge pressure and flow remained stable. suction was reported again later during support and clinical noted waveforms appeared as if device is potentially a little deep. device was repositioned. logs show suction alarms occur at various instances after this. no volume related issues were noted. the cause of the pump suction was not established as limited clinical information was provided, log analysis was inconclusive and no product was returned. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B7 added information that was omitted from the initial report that was submitted. d6b explant date was added. h10 this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. related report number was added.
3727 | 1220648-2025-48073 | 23465885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23465885&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679953 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 50 KG | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support experienced access site bleeding/oozing that required heparin to be withheld. additionally, the patient¿s heart rhythm was converted from very slow (bradycardia) with pauses to atrial fibrillation causing intermittent suction alarms on the impella. the patient was successfully weaned from the impella cp and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B2 revised selection as it was entered incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. major bleed/access site adverse event: clinical details reported oozing at the site 13 hours after implant. activated clotting time was not provided for the time of the bleed. the product was not returned for investigation. the cause of the bleeding could not be determined due to insufficient clinical details and no product returned. pump suction: clinical details reported that the patient was converting from bradycardia with pauses to atrial fibrillation which triggered intermittent suction. the product was not returned for investigation. data log review confirmed suction events on 7-oct-2025. placement signal was gradually trending down for the entire case. at the times when suction was triggering most frequently, placement signal briefly dropped to below 50mmhg. clinical details stated that the heart rhythm was variable, which could have led to abnormal filling and flow in the heart. the cause of the suction was most likely related to the patients condition, as placement signal was trending down throughout the case indicating some volume issues, and clinical details mentioned the patient had variable heart rhythm which could have led to inadequate filling in the heart. device history review: the pump passed all post sterile inspection criteria.
3728 | 1220648-2025-48074 | 23466377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466377&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761611 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella cp came out when the physician moved the guidewire. the device was repositioned. there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3729 | 1220648-2025-48075 | 23466517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466517&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698175 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. the patient was scheduled for an upcoming transcatheter aortic valve replacement. in preparation for the procedure, the patient was brought to the cardiac catheterization laboratory for right heart catheterization and left heart catheterization, percutaneous coronary intervention (pci) of the right coronary artery (rca) and left anterior descending (lad) artery, and balloon aortic valvuloplasty (bav). the decision was made to insert an impella cp to provide mechanical circulatory support during the procedures because the patient had complex coronary anatomy and a reduced stroke volume. an impella 14fr x 25cm sheath was placed without issue and the aortic valve was crossed with a wire. bav was performed successfully and the impella cp was loaded over wire. multiple unsuccessful attempts were made to deliver the impella cp to the left ventricle; however, resistance was met at the aortic valve. the medical team decision was ultimately made to explant the impella and perform the pci unsupported due to aortic stenosis. impella cp explant was attempted through the impella 14fr x 25cm peel away sheath. the device was pulled entirely into sheath lumen and would not exit hemostatic valve. multiple attempts were made to explant the device but all were unsuccessful. a guide wire was introduced into sheath to maintain vessel access, and the sheath was explanted along with impella. upon explant, a kink was noted in the sheath after the impella cp was explanted, the pci of the rca and lad were performed without issue. this complaint involves two impella devices: impella cp, with serial number (b)(6) impella 14fr x 25cm peel away sheath with lost number s9517744 this report specifically addresses impella cp, with serial number (b)(6), which is the subject of this report. a separate report will be submitted for the other device associated with this complaint. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the difficult/unable to remove through introducer: clinical details indicate that during explant, the impella cp was pulled entirely into the sheath lumen but could not be removed through the hemostatic valve. multiple attempts to explant the device were unsuccessful until a wire was introduced to maintain vessel access, and the sheath was explanted together with the impella. upon removal, a kink was noted in the sheath. the cause of the device removal issue was most likely related to the patient¿s condition of tortuous vessel anatomy, which caused product damage and made it difficult to remove the pump. however, product damage was not confirmed without returned product investigation. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3730 | 1220648-2025-48077 | 23466677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466677&pc=OZD | 10/09/2025 |  |  | 11/04/2025 |  |  | 10/09/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1613085 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during support of a patient an automated impella controller failed to boot up. no patient harm was reported. | [5 empty] | The device was evaluation. section d9 was updated with the date the device was returned for analysis. the reported failure mode was reproduced during testing and the complaint was confirmed. the root cause of the ¿system self check error¿ was determined to be power and battery manager (pbm) corrosion due to leakage of the electrolyte from the battery failure alarm super caps. before returning the console to the customer, the pbm was replaced and a full functional check was performed. | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3731 | 1220648-2025-48079 | 23466702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466702&pc=OZD | 10/09/2025 |  |  | 11/04/2025 |  |  | 10/10/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656243 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia requiring defibrillation and intubation. the impella was positioned too deep and against the septum so the pump was explanted. | [5 empty] | No product was returned for investigation. the cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the positioning issue cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3732 | 1220648-2025-48080 | 23466725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466725&pc=OZD | 10/21/2025 |  |  | 11/04/2025 |  |  | 10/21/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659747 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female | 55 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 could not be implanted due to the patient¿s anatomy. an impella cp was implanted instead. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy, too tight and small size preventing successful delivery of impella. the device passed all post sterile inspection checks.
3733 | 1220648-2025-48078 | 23466750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466750&pc=OZD | 10/06/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620591 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pain | Tachycardia | Heart Block | [32 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. upon arrival to the hospital, the patient was found to be hypotensive with audible crackles in bilateral lung bases. electrocardiogram results were noted to be largely unremarkable, showing no acute st changes. the patient¿s beta-natriuretic peptide was elevated with a result of 1627 pg/ml and creatinine was elevated with a result of 1.5 mg/dl. a transthoracic echocardiogram (tee) was performed to reveal a left ventricular ejection fraction (lvef) of 15% with severe left ventricular (lv) hypokinesis, moderate mitral regurgitation, a laminar thrombus at the lv apex, and grade iii diastolic dysfunction. further cardiac evaluation included a left heart catheterization with no coronary artery disease (cad) noted. right heart catheterization was performed and revealed an elevated right atrial pressure of 27 mmhg and pulmonary capillary wedge pressure of 33 mmhg, decreased pulmonary artery saturation of 34%, and a fick of 1 l/min. cardiac magnetic resonance imaging (cmri) was performed to assess the lv thrombus and ventricular function. the cmri found no indication of a lv thrombus, worsening lvef of 13% while receiving dual inotropes, decreased right ventricular ejection fraction of 28%, and a small non-cad scarring of mid inferior wall. due to the decompensated heart function, the medical team decided to implant an impella cp in the setting of society for cardiovascular angiography and interventions stage ¿c¿ decompensated cardiogenic shock with a plan to escalate the patient to an axillary impella 5.5 the next day for continued evaluation of advanced heart failure therapy. after the impella cp was placed, the position was imaged with fluoroscopy. the position of the impella did not look to be ideal, and the physician was concerned the pigtail was entangled with the mitral valve apparatus. despite numerous attempts to reposition, the decision was made to remove the impella cp device and re-insert the same pump after obtaining repeat access with a guidewire. the device was successfully placed, and good placement was noted with fluoroscopy and tee. the following day, the patient was taken to the operating room for planned escalation to impella 5.5. upon removal of the impella cp, a clot was observed to be hooked on the pigtail. it was noted that although the patient had a history of lv thrombus, no lv thrombus had been observed on the cmri prior to impella cp placement. on day 5 of impella 5.5 support, the patient¿s pulmonary artery pressures were noted to be elevated and impella support was increased from p7 to p8. it was reported that the patient was intermittently experiencing second degree heart block and briefly, complete heart block a couple of days ago. the following day the patient experienced small, nuisance nose bleeds. it was noted that the partial thromboplastin time was therapeutic. the next day it was reported that the nosebleeds continued and the patient¿s bivalrudin was started and stopped intermittently to address the bleeding. on day 8 of support, it was noted that the nosebleeds had resolved, and bivalrudin was resumed. on day 9 of support, it was noted that the patient¿s hemoglobin had dropped to 6.2 g/dl and the patient received one unit of packed red blood cells (prbcs). the repeated hemoglobin level was 7.3 g/dl post transfusion. on day 10 of support, it was reported that the previous day the patient was out of bed and that overnight the patient had increasing anxiety. the patient was tachycardic with heart rate around 130 beats per minute. the patient experienced additional minor nose bleeds and was given medication for anxiety. on day 13 of support, the patient was taken to the cardiac catheterization laboratory for cannulation of venoarterial extracorporeal membrane oxygenation (va ecmo) for ongoing progressive biventricular failure. after va ecmo support was initiated, impella 5.5 support was decreased to p6. the patient was noted to have been transfused with an additional two units of prbcs and 250 mls of albumin. overnight, it was reported that there were suction alarms and the patient experienced increased ectopy with reduced impella flows. a point of care ultrasound revealed that the pump was orientated towards the mitral valve. the physician repositioned the device, which increased the impella flows. the patient was noted to feel much better after repositioning and the patient¿s mixed venous oxygenation levels were noted to be 65%. a small hematoma was noted at the axillary site found with computed tomography and was not appreciable with visual examination. it was additionally noted that the patient was noted to have had a laminar lv thrombus which measured 1.6cm x 1cm on (b)(6) 2025 and was not appreciable on (b)(6) 2025 cmri prior to impella placement. the patient was taken to the cardiovascular operating room for escalation to central va emco placement and remove the impella 5.5 and peripheral va ecmo. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this report specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. a separate report will be submitted for the other device associated with this complaint. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. arrythmia/tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms: the cause of the injury was not determined due to insufficient clinical details. pump suction/low pump flow: clinical details reported suction and reduced flows overnight, and at the time, the pump was oriented towards the mitral valve. flows improved after repositioning the pump. the product was not returned for investigation. data log review confirmed low flows and suction alarms triggering on 10/14/25 when it was reported. at this time, flows were below spec (2.0 liters per minute at p-6), placement signal started to trend down over time and lost pulsatility. motor current pulsatility also decreased, which could imply something was inhibiting flows. clinical details mentioned that the pump was oriented towards the mitral valve at this time. after repositioning, flows and placement signal improved slightly. the cause of the low pump flow and suction was most likely improper positioning of the pump, as data log review indicates positioning issues and clinical details stated that the pump was oriented towards the mitral valve during the events. device history review: the pump passed all post sterile inspection criteria. | Section d catalog number was corrected.
3734 | 1220648-2025-48081 | 23466824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23466824&pc=OZD | 10/08/2025 |  |  | 11/04/2025 |  |  | 10/08/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276779 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during preparation for impella insertion it was reported that when the automated impella controller (aic) was started, it displayed a system error with the message that the aic must be replaced, and a replacement aic should be used. the aic was replaced. no patient harm was reported. | [5 empty] | Product complaint # (b)(4). investigation summary: data analysis: console logs from the reported day of event are consistent with complaint and show persistent communication failures of sau_powermod_1. aic reboots were not resolving the issue. device analysis: device was tested at fs, and the issue was reproduced. it¿s been observed that communication with pbm1 was not working. inspection of the console revealed corrosion near c69, c70 supercaps, due to electrolyte leak. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | A.2, a.4, a.5 and a.6 are unknown. the automated impella controller was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3735 | 1220648-2025-48082 | 23467099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23467099&pc=OZD | 09/14/2025 |  |  | 11/04/2025 |  |  | 10/30/2025 | 11/04/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075037 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) generated a placement signal not reliable alarm during patient support. the alarm self-resolved. proper pump placement was confirmed with echocardiogram. no patient harm was reported. this report represents the aic. another report was submitted to represent the pump. refer to section h.10 for the related report number. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4, a5, and a6 are unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3736 | 1220648-2025-48083 | 23472464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472464&pc=OXD | 10/04/2025 |  |  | 11/04/2025 |  |  | 10/06/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OXD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658786 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that removing the guidewire from the impella cp during positioning of the pump was difficult. however, once the pump was explanted from the patient the guidewire came out easily. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3737 | 1220648-2025-31912 | 23472646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472646&pc=OZD | 12/05/2023 | 11/04/2025 |  | 11/04/2025 |  |  | 12/05/2023 | 11/04/2025 |  |  | 05/31/2025 |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345524 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a lost placement signal during support. the patient remained stable and support continued with no adverse impact. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. investigation was done to determine the root cause of the failure mode 'optical signal issue'. upon review of the lt and rt data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3738 | 1220648-2025-31913 | 23472654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472654&pc=OZD | 06/04/2025 | 11/04/2025 |  | 11/04/2025 |  |  | 06/04/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella cp support, and during support, there were false impella in aorta alarms. position was confirmed via echocardiography. the placement signal monitoring alarm was disabled. it is unknown if the automated impella controller (aic) was related to the optical signal issue. care was eventually withdrawn, and the patient expired. the patient death is captured in this record¿s sister records. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ¿ the case associated with the reported complaint was initiated on june 4, 2025, with the pump (sn: (b)(6). ¿ after 35 seconds from the pump being started, the console displayed ¿impella position in aorta (for up to 14s) ¿ alarm,¿ event # 128. rt logs confirm that the snr was within specifications during event # 128. ¿ there were no alarms indicating any issues on the console. note: lt log shows that the snr was sometimes close to the passing threshold of 2000, and at other times below it, but it recovered automatically in the associated case (lt_log_ic4303.png), which is unrelated to the reported failure mode. device analysis: this console was tested after arriving at fs and was found to have an abnormal optical bench ccd array analog output. root cause: the root cause of the inaccurate impella position alarm was most likely a malfunction of the optical bench ccd array. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3739 | 1220648-2025-48084 | 23472971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472971&pc=OZD | 10/10/2025 |  |  | 11/04/2025 |  |  | 10/10/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736987 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 121 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia which resolved without intervention. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3740 | 1220648-2025-48085 | 23472980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472980&pc=OZD | 10/10/2025 | 02/17/2026 |  | 11/04/2025 |  |  | 10/10/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Unspecified Infection | Sepsis | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis so the pump was repositioned. the patient entered a ventricular tachycardic storm and was defibrillated multiple times. the patient developed bacteremia so the pump was explanted. care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3741 | 1220648-2025-48086 | 23472986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472986&pc=OZD | 10/10/2025 |  |  | 11/04/2025 |  |  | 10/10/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714118 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 86 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Multiple Organ Dysfunction Syndrome | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced multiple organ failure, a subdural hematoma, and a cerebrovascular accident. the patient was transfused two units of packed red blood cells. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the stroke was unable to be determined due to insufficient clinical details. the cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks
3742 | 1220648-2025-48087 | 23472988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472988&pc=OZD | 10/10/2025 |  |  | 11/04/2025 |  |  | 10/10/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724384 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia. the pump was explanted and an impella 5.5 was implanted to support the patient. the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3743 | 1220648-2025-48088 | 23472991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472991&pc=OZD | 10/11/2025 |  |  | 11/04/2025 |  |  | 10/11/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729475 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia. an echocardiogram was ordered. the patient survived. | [5 empty] | The cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3744 | 1220648-2025-48089 | 23472996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472996&pc=OZD | 10/11/2025 |  |  | 11/04/2025 |  |  | 10/11/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026755162 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing around the blue suture. the patient was transfused one unit of blood. pressure was held and a femostop was applied. the bleeding continued so a pressure dressing was applied and another unit of red blood cells was transfused. the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3745 | 1220648-2025-48090 | 23472997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23472997&pc=OZD | 10/11/2025 |  |  | 11/04/2025 |  |  | 10/11/2025 | 11/04/2025 |  |  |  |  | 11/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026757951 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced atrial fibrillation. care was withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3746 | 1220648-2025-47991 | 23449630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23449630&pc=OZD | 10/04/2025 |  |  | 11/03/2025 |  |  | 10/04/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743533 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 56 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. on support day 2, the patient experienced ectopy. an echocardiogram was performed which showed the impella slightly deep at 5.2 centimeters and appeared to be somewhat on the septum. company support recommended to the medical team to pull the impella back to 3.5 centimeters., however it was not reported if interventions were actioned. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B7 added information omitted from the initial report that was submitted. h6 code b18 should of been on the initial report instead of b17 as the device was discarded. h11 the device was discarded and should of been noted on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3747 | 1220648-2025-47993 | 23449757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23449757&pc=OZD | 10/06/2025 |  |  | 11/03/2025 |  |  | 10/06/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026731304 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Difficult or Delayed Positioning | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was being prepared for impella 5.5 support due to postcardiotomy cardiogenic shock. there was initial difficulty accessing the left ventricle, difficultly placing the placement catheter and wire due to new aortic valve and the patient being on extracorporeal membrane oxygenation and intra-aortic balloon pump. the surgeon opted to use al1 with proglide stiff access. the impella was successfully placed. the patient ultimately expired as care was withdrawn. | [5 empty] | No product was returned. the cause of the positioning issue was determined to be related to the anatomy of the patient since there was difficultly placing placement catheter and wire due to new aortic valve and patient being on ecmo and intra aortic balloon pump. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3748 | 1220648-2025-47992 | 23449802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23449802&pc=OZD | 10/06/2025 |  |  | 11/03/2025 |  |  | 10/06/2025 | 11/03/2025 |  |  | 12/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442169 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation which was treated with amiodarone. later, the patient was successfully weaned off the pump and survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the returned pump was investigated. dried blood was seen around the inflow and outflow cages, on the sterile sleeve, and around the suture hub. no other abnormalities were seen. peri-procedural adverse event (arrhythmia): patient¿s hr is currently atrial fibrillation in the 130-140s and uncontrolled, now bolusing amio and starting a drip. the cause of the injury was unable to be determined due to insufficient clinical details. | The cause of the cardiac arrhythmia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3749 | 1220648-2025-47995 | 23450002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450002&pc=OZD | 10/07/2025 |  |  | 11/03/2025 |  |  | 10/07/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740501 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female | 60 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced junctional paroxysmal ventricular tachycardia and renal failure. the patient was placed on continuous renal replacement therapy. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned for investigation. the cause of the renal failure was not determined due to insufficient clinical details. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3750 | 1220648-2025-47965 | 23450004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450004&pc=OZD | 10/03/2025 | 10/30/2025 |  | 11/03/2025 |  |  | 10/04/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740374 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp required repositing of the pump because it appeared deep. the pump then migrated into the aorta and was repositioned again. the patient had a do not resuscitate order. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The pump was returned for evaluation. no issues were found with the returned product. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 and a5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3751 | 1220648-2025-47997 | 23450200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450200&pc=OZD | 09/12/2025 |  |  | 11/03/2025 |  |  | 09/12/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583754 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 64 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, a hematoma was observed at the groin area. iit was noted that the patient was well and there were no complications at the access site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided.
3752 | 1220648-2025-31909 | 23450803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450803&pc=OZD | 03/10/2024 | 11/03/2025 |  | 11/03/2025 |  |  | 03/10/2024 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during patient care the placement signal disappeared. patient support was not impacted by the issue, and no intervention was required. no harm to the patient was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imc logs confirmed that the controller error alarm opm comm loss 1045 occurred after placement signal not reliable error. following the set of errors, the optical bench was intentionally reset, which resolved the placement signal issue. real time rt logs confirmed that immediately prior to the controller error alarm, the optical bench reported an oscillating contrast value far outside the normal range of contrast. this oscillating contrast issue has been previously observed in a few other consoles that exhibited similar errors. a root cause has not been yet identified, however it is known that an oscillating contrast is the result of a hardware fault on the optical bench. device analysis: console ic was sent back to field service (fs) for further investigation. the issue was not able to be reproduced in the lab. this issue is a low probability event and is very difficult to reproduce. the console is designed to automatically reset the optical bench when this issue emerges, which resolves the issue. before being returned to the customer, fs was instructed to replace the optical bench 0042-3015 or equivalent). root cause: the root cause was determined to be caused by a faulty optical bench which resulted in an unreliable placement signal due to an oscillating contrast value, triggering the alarm. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3753 | 1220648-2025-47996 | 23450811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450811&pc=OZD | 12/27/2024 | 11/03/2025 |  | 11/03/2025 |  |  | 12/27/2024 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558929 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 converted to atrial fibrillation with rapid ventricular response rate. epi was discontinued and amiodarone was increased. no patient harm was reported. | [5 empty] | No product was returned for investigation. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A.4. is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3754 | 1220648-2025-48000 | 23450814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450814&pc=OZD | 10/04/2025 |  |  | 11/03/2025 |  |  | 10/04/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706528 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced epistaxis and gastrointestinal bleeding, resulting in hematochezia. red blood cells were transfused to the patient. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the epistaxis was not determined due to insufficient clinical details. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3755 | 1220648-2025-48001 | 23450888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450888&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724689 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 101 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support there was slight oozing at the impella site. manual pressure held for 15 minutes. heparin was put on hold. the patient also had pink-tinged urine, hemolyzing labs, and one suction alarm. the impella was pulled back under echocardiogram guidance to approximately 3 centimeters across the aortic valve, not interfering with any structures. later, the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed/hemostasis: the cause of minor bleed/hemostasis was most likely anti-coagulation management related since act value at the time of bleeding was higher than recommended dose. device history review: the device passed all the post sterile inspection checks.
3756 | 1220648-2025-48004 | 23450981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23450981&pc=OZD | 10/06/2025 |  |  | 11/03/2025 |  |  | 10/06/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705220 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a hematoma at the axillary cutdown site. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks.
3757 | 1220648-2025-47998 | 23451035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451035&pc=PYX | 10/04/2025 |  |  | 11/03/2025 |  |  | 10/04/2025 | 11/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026705443 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Fever | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp flex for mechanical circulatory support. it was reported that the patient had undergone an intervention to their left anterior descending coronary artery on (b)(6) 2025 and decompensated in the unit. a left heart catheterization and right heart catheterization were scheduled for (b)(6) 2025 and an impella cp was implanted. after the impella cp was implanted, it was noted that the patient¿s leg did not have dopplerable pulses. the decision was made to escalate the patient to an impella 5.5. on day 6 of impella 5.5 support, the patient was taken to the operating room for a mitral valve repair and coronary artery bypass graft. the patient developed postoperative right ventricular failure, and the decision was made to implant an impella rp flex. on day 3 of bipella support with the impella rp flex and impella 5.5, the patient spiked a fever. the medical team decided to remove the swan and central line. post removal of the swan, the placement signal numbers spiked on the impella rp flex, and the flows dropped. the decision was made to remove the impella rp flex and escalate the patient to veno-arterial extracorporeal membrane oxygenation (va ecmo). the patient remained on impella 5.5 and va ecmo support. the following day, the patient¿s family made the decision to withdraw care of the patient. the patient¿s care was withdrawn, and the patient expired. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). pump 3: impella rp flex with serial number (b)(6). this report specifically addresses pump 3: impella rp flex with serial number (b)(6), which is the subject of this report. separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event and noted the following: the patient presented with cardiogenic shock (cgs) following intervention to the left anterior descending (lad) artery. an impella cp was implanted for hemodynamic support. post-implant, distal limb ischemia was observed at the impella cp access site, prompting escalation to impella 5.5. six days after impella 5.5 implantation, the patient underwent coronary artery bypass grafting (cabg) and mitral valve replacement (mvr). postoperatively, the patient developed right ventricular (rv) failure, and an rp flex device was implanted. subsequently, the patient developed a fever, and the clinical team elected to remove the swan-ganz catheter and central line. upon removal, the impella rp flex placement signal spiked and flows dropped. it is unknown whether the recommended practice of reducing impella rp flex flows prior to line removal was followed. the impella rp flex was removed, and veno-arterial extracorporeal membrane oxygenation (va ecmo) was initiated. the patient remained on ecpella support until the following day, when the family decided to withdraw care. the patient expired. assessment outcome is the following the limb ischemia on impella cp considered device-related due to access site involvement. fever on impella 5.5 and rp flex likely secondary to swan/central line infection but conservatively attributed to devices. the impella rp flex device malfunction reported due to abnormal placement signal and flow drop; root cause undetermined (possible procedural deviation by not reducing flows upon removal of lines). and the death will be reported for impella 5.5 as a conservative measure; clinical deterioration most likely contributing factor. correction: after review of the case by medical safety, it was determined the impella rp flex is a serious injury type of reportable event (rather than death. the death outcome is now associated with the impella 5.5). sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes and medical device problem code) have been updated, corrected accordingly. device status. additional information also noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this impella rp flex device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp and impella 5.5 which is associated with the demise outcome. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Impella device was returned for evaluation. returned pump had significant biomaterial around the bottom side of the impeller. slight catheter kink observed. low flow: the cause of the low flow was biomaterial ingestion based on data log and returned pump analysis. placement signal issue: the cause of the placement signal issue was biomaterial ingestion based on data log and returned pump analysis. fever: the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
3758 | 1220648-2025-48002 | 23451058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451058&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702570 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 100 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and a retroperitoneal bleed. the bleed was addressed surgically and the patient was placed on continuous renal replacement therapy. the patient then experienced cardiac arrest that required advanced life support to obtain return of spontaneous circulation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The cause of the hematuria was not determined due to insufficient clinical details. the cause of the bleeding was not determined since product was not returned and insufficient clinical details provided. the cause of the cardiac arrest was unable to be determined due to insufficient clinical details. the cause of the renal failure was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
3759 | 1220648-2025-48003 | 23451159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451159&pc=OZD | 10/06/2025 |  |  | 11/03/2025 |  |  | 10/06/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653594 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation. amiodarone was administered. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3760 | 1220648-2025-48005 | 23451206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451206&pc=OZD | 09/09/2025 |  |  | 11/03/2025 |  |  | 09/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722887 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing from the cutdown site secondary to anti-coagulation. the patient underwent a pocket revision and washout, was successfully weaned from the pump, and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3761 | 1220648-2025-48007 | 23451320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451320&pc=OZD | 10/06/2025 |  |  | 11/03/2025 |  |  | 10/06/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729943 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was difficult due to tortuous patient anatomy. instead, a new impella 5.5 was implanted on the other side. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned. the cause of the delivery issue was determined to be patient condition as the patient had tortuous anatomy with stenosis right axil preventing successful delivery of impella. the device passed all post sterile inspection checks.
3762 | 1220648-2025-47726 | 23451422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451422&pc=OZD | 09/09/2025 |  |  | 11/03/2025 |  |  | 09/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697581 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 141 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella cp was aborted due to the patient's anatomy. an impella 5.5 was implanted instead. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3763 | 1220648-2025-48008 | 23451656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451656&pc=OZD | 10/04/2025 |  |  | 11/03/2025 |  |  | 10/07/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1483134 |  | 0042-0000-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was being prepared for impella support for an unspecified indication. the power was turned on when trying to use the impella but the self-diagnosis failed and it became unusable. a backup automated impella controller was used to support the patient. | [5 empty] | Patient information was not provided. therefore, section a was left blank. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3764 | 1220648-2025-48011 | 23451691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451691&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730399 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock. while on support, the patient experienced hypotension followed by loss of placement signal. the pump presented suction issues with no specified resolution. later, the patient was successfully explanted and survived. additional information was received. the area manager does not have the pump to return. the decoupling manifested due to low cardiac output. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned. optical sensor issue/pump suction: the clinical details state that the lv signal became disconnected from the ao ps and there was suction at p-5. the data logs show that there was suction alarm #14 triggering on 10/04 and then triggering again at the time of the reported event on 10/05. during this time the lvps can be seen to begin to mildly decouple from the ao ps, as the motor current max began to trend down gradually causing a trend up in the dpmin. the placement signal remain relatively stable during this time along with snr and the rest of the 5s parameters. this issue seems to resolve with adjustments of the p-levels and suction is not triggered for the duration of the two days. there were no motor current spikes or drastic changes of flow during this time. due to lack of clinical details regarding positioning of the pump and no product returned, the root cause of the optical sensor issue and the suction issue cannot be established. hypotension: the cause of the clinical symptom cannot be determined due to insufficient clinical details. device history lot: device lot: 1960372. device history batch: subcomponent: n/a. device history review: impella 617314 passed all post sterile inspection checks.
3765 | 1220648-2025-31910 | 23451829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451829&pc=OZD | 10/25/2024 | 11/03/2025 |  | 11/03/2025 |  |  | 10/25/2024 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142852 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella support and there were placement signal not reliable alarms. no changes in patient hemodynamics or motor current. the investigating engineer found that this issue likely stemmed from the automated impella controller (aic). no known patient harm or injury has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause for psnr alarm was defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3766 | 1220648-2025-48010 | 23451982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23451982&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641434 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 73 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support the patient had a hematoma. manual pressure was held for 1.25 hours. later, the patient was successfully explanted and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3767 | 1220648-2025-48012 | 23452015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452015&pc=OZD | 09/12/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/12/2025 | 11/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721377 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | No | Fluid/Blood Leak | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 pump supporting the patient admitted in with cardiogenic shock (cgs) and cardiomyopathy. the 5.5 pump was supporting when there was a purge leak observed from the pump that did not prompt explant or intervention. there additionally was vt that prompted the team to readjust position. the pump remained on for 13 days til care was withdrawn for the patient. the impella 5.5 will be coded for ventricular tachycardia, device repositioning and conservatively for death. the ventricular tachycardia is coded because it occurred during impella 5.5 support and required clinical intervention in the form of device repositioning. | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia and had a purge leak. the pump was repositioned. later, care was withdrawn and the patient expired. |  |  |  |  | No product was returned for investigation. the cause of the purge leak was not determined since product was not returned and insufficient clinical details were provided. the cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | Event description details and complaint/patient coding was updated/corrected following further clinical review. therefore, sections b5 event description has been updated and h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. note: investigation type, findings, conclusion, and component coding remain unchanged. the update is health effect - impact coding. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3768 | 1220648-2025-48009 | 23452024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452024&pc=OZD | 09/12/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/12/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693555 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced suction events. a new impella cp was placed to resolve the issue. no patient harm was reported. | [5 empty] | A4 and a5 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3769 | 1220648-2025-48013 | 23452234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452234&pc=OZD | 09/12/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/12/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653647 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 73 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during explant of an impella 5.5 a clamp was used incorrectly causing the pump to detach at the cannula. the entire graft had to be removed with the pump. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The pump was returned for analysis. the returned product did not have any damage to the outflow cage as it remained on the cannula. towards the inlet the cannula was compressed and had clamp marks on it. the delo band was torn on the top and the inlet was seen to be completely detached from the cannula. the inlet was crushed as it had a clear clamp mark on it. the cause of the detached inflow cage is due to a material fracture of the inlet based on an unintended user error as the physician clamped on the inlet causing it to separate. the device passed all post sterile inspection checks.
3770 | 1220648-2025-48015 | 23452303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452303&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741080 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 85 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Asystole | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. while on support, the patient presented asystole, stabilized by temporary pacemaker. the next day, brain computed tomography showed severe anoxic brain injury. care was ultimately withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6: type of investigation, added code b11. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. asystole/ stroke: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1970236 device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3771 | 1220648-2025-48021 | 23452402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452402&pc=OZD | 08/26/2025 |  |  | 11/03/2025 |  |  | 08/26/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627747 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male | 70 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a gastrointestinal bleed. heparin was withheld and an exploratory gastroscopy was performed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the bleed was not determined since product was not returned and insufficient clinical details were provided. the pump passed all the post sterile inspection checks.
3772 | 1220648-2025-48019 | 23452466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452466&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667580 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support went into cardiac arrest and required advanced cardiac life support. return of spontaneous circulation was obtained. some suction alarms were observed with change in intrathoracic pressure. treatment was discussed to escalate to an impella 5.5 however patient had signs of infection. patient was on anti-arrhythmics to induce bradycardia in an effort to prevent ventricular tachycardia. the patient was later escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: clinical symptom (unspecified infection): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. clinical symptom (cardiac arrest): the cause of the injury was not determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. according to the clinical details, suction was reported with changes in intrathoracic pressure on (b)(6) 2025, and on (b)(6) 2025, suction was resolved after giving albumin and reducing the p-level. data log confirmed multiple suction alarms on the reported days, with the majority occurring at higher p-levels (7 or above). placement signal had a fluctuating trend during the entire case run. no position or motor current spike-related abnormalities were reported. the cause of the suction issue was most likely related to patient condition, based on data log review and provided clinical details. device history lot: device batch: 1920393. device history batch: subcomponent lot: na. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
3773 | 1220648-2025-48014 | 23452631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452631&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710790 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella cp support for an unspecified condition. post-explant and closure device deployment, there appeared to be a new flow-limiting lesion in the common femoral artery at the level of the access closure. contralateral access prior to explant allowed an angiogram runoff and sequential 6mm and 7mm pta catheter inflations at the level of the closure, post-closure. flow was restored and zero extravasation documented. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product was returned. peripheral artery dissection: the cause of the injury was unable to be determined due to insufficient clinical details. thrombosis: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. clinical review: implant/ explant date unknown. impella cpc9+ inserted. no filling defect prior to explant. however, post explant and closure device deployment, there appeared to be a new flow-limiting lesion in the common femoral artery at the level of the access closure. contralateral access prior to explant allowed an angiogram runoff and sequential 6mm and 7mm pta catheter inflations at the level of the closure, post-closure. flow was restored and zero extravasation documented. iliofemoral vessels were very diseased, and i want to get a runoff and make sure we did not cause any damage during impella insertion. device history review ==> device passed all post sterile inspection checks.
3774 | 1220648-2025-48017 | 23452646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452646&pc=OZD | 10/05/2025 | 12/02/2025 |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287216 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) displayed "placement signal not reliable" alarms after explanting the extracorporeal membrane oxygenation. there were no placement signals for aorta + left ventricle, but the pump continued running on p2 with purge flow and pressure within the correct range and the motor current pulsatile. there was no patient harm. the patient remains on support. | [5 empty] | H6 added code b13 as it was omitted from the initial report that was submitted. h10 this is one of two related reports. this report represents the automated impella controller. another report was submitted to represent the pump. related report number was added as this information was omitted from the initial report that was submitted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3775 | 1220648-2025-48018 | 23452795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452795&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698995 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed ischemia of the right leg. unknown intervention was done to successfully manage the condition. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary the device was returned for evaluation. ischemia: in order to make a cause determination, the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history review device passed all post sterile inspection checks.
3776 | 1220648-2025-48020 | 23452847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452847&pc=PYX | 10/22/2025 | 11/04/2025 |  | 11/03/2025 |  |  | 10/22/2025 | 11/03/2025 |  |  | 01/31/2026 |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025518027 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 95 KG |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced sepsis and low flows. a clot was suspected so the pump was explanted. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3777 | 1220648-2025-48024 | 23452901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452901&pc=OZD | 09/15/2025 |  |  | 11/03/2025 |  |  | 09/15/2025 | 11/03/2025 |  |  | 09/30/2025 |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400460 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp was placed, peel away sheath removed and repositioning sheath was advanced into arteriotomy. sutured utilizing forward tension sutures, dressed and secured. gurney was locked and secured. when getting patient ready to move over to gurney to be transferred 'impella in aorta' alarm occurred. device was noted to be pulled back into ascending aorta. ic (interventional cardiology) attempted to push device back across av (aortic valve) but was unsuccessful. decision made to replace impella cp. hematoma was also noted at insertion site pre and post impella cp replacement. manual pressure and light pressure fem stop was applied. new device was successfully placed. patient is stable post replacement. | [5 empty] | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of hematomas cannot be determined since insufficient clinical details were provided. unable to place or position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was noted to be pulled back into ascending aorta when getting patient ready to move over to gurney to be transferred to hub hospital and was replaced with another pump successfully. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
3778 | 1220648-2025-48016 | 23452903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452903&pc=OZD | 09/12/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/12/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593557 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Flutter | Fever | Sepsis | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a low-grade fever and the patient was treated for sepsis. the patient was given bival and bicarb in response to heparin induced thrombocytopenia. additionally, the patient was cardioverted in response to atrial flutter. later, the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the fever and thrombocytopenia was not determined due to insufficient clinical details. the root cause of the sepsis and atrial flutter was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3779 | 1220648-2025-48025 | 23452998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23452998&pc=PYX | 09/16/2025 |  |  | 11/03/2025 |  |  | 09/16/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600689 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 68 KG | Hispanic |  | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced a purge pressure low alarm, oozing at the insertion site, hematuria, and liver shock. the patient was treated with albumin but the patient's condition continued to deteriorate and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3780 | 1220648-2025-48023 | 23453044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23453044&pc=OZD | 09/10/2025 |  |  | 11/03/2025 |  |  | 09/10/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722134 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 99 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Unspecified Infection | Inflammation | Renal Failure | Multiple Organ Dysfunction Syndrome | [6 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced hypokinesis so an impella rp flex was implanted. the patient experienced intermittent suction due to volume overload. epinephrine was given and platelets were transfused. due to right hemothorax bleeding the patient was transfused three units of blood products and underwent a thorascostomy with chest tube insertion. the patient continued to have arrhythmia and failure of the liver and kidneys. lidocaine and levophen were given in response. the patient had a positive sputum culture. the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A review of the data logs confirmed the complaint. the product was not returned for review. the cause of low pump flow and suction was most likely patient related since the patient had right ventricle failure, fluid volume overload, and suction seemingly coincided with junctional rhythm change. the cause of the hemothorax bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the inflammation, hypotension, arrhythmia, infection, and organ failure was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. a 70-year-old female with a medical history of persistent atrial flutter, paroxysmal atrial fibrillation, mixed hyperlipidemia, hypertension, and morbid obesity presented to the emergency department with several weeks of fatigue, dizziness, palpitations, and associated shortness of breath. the patient was newly diagnosed with heart failure with reduced ejection fraction (hfref) with a severely reduced ejection fraction of 10¿15% and severe global hypokinesis. the patient underwent left heart catheterization for ischemic evaluation, which demonstrated two-vessel coronary artery disease involving the mid left anterior descending artery and right coronary artery. right heart catheterization demonstrated worsening lactic acidosis, with lactate levels trending up to 12.9 mmol/l. the advanced heart failure team was consulted, and the patient was initiated on left-sided impella support, which was subsequently escalated to biventricular (bipella) support with impella devices: impella 5.5 and rp flex. during impella cp support, the patient developed acute renal failure requiring continuous renal replacement therapy and was treated with multiple medications. paroxysmal junctional rhythm was also reported. the patient was escalated to an impella 5.5, and bipella support was initiated with the addition of an impella rp flex. while on bipella support, the patient experienced multiple adverse clinical events. additionally, the automated impella controller (aic) connected to the impella rp flex displayed a frozen screen with a persistent ¿complete the procedure¿ alarm. replacement of the purge cassette did not resolve the issue. the aic was exchanged with no noted consequence to the patient, and the patient continued to receive support without further interruption. after approximately eight days of mechanical circulatory support, care was withdrawn, and the patient expired. change in reportability. after review of the case by medical safety, it was determined the impella is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been repopulated/updated, corrected accordingly. date of death. the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this is one of four devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
3781 | 1220648-2025-48022 | 23453151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23453151&pc=OZD | 09/13/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/13/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661325 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp support had multiple runs of ventricular fibrillation requiring defibrillation. additionally, the patient was positive for heparin-induced thrombocytopenia. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, an investigation has not been done. should the device or any additional information be received, a supplemental mdr will be filed.
3782 | 1220648-2025-48026 | 23453936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23453936&pc=OZD | 09/18/2025 |  |  | 11/03/2025 |  |  | 09/18/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740469 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp on (b)(6) 2025, that overnight, the patient had bright red vomiting( emesis) requiring intubation and 3 units prbcs. gi consult with plans for egd. heart team relates bleeding to underlying cirrhosis with esophageal varicies. egd completed yesterday with 2l blood and clots. multiple blood products given. ef 40-45% per cardio fellow; plans for explanting today or tomorrow. the patient off heparin due to bleeding, act 180s. impella was removed successfully as bridge to post pci recovery, site is dry and intact, staff pleased with outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
3783 | 1220648-2025-48029 | 23454939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23454939&pc=OZD | 10/09/2025 |  |  | 11/03/2025 |  |  | 10/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717603 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 82 KG | Hispanic |  | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hematuria and positioning issues. a transthoracic echocardiogram confirmed the pump was positioned in a shallow manner. repositioning of the device were unsuccessful so the p level of the pump was kept low. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller. displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | No product was returned for investigation. the cause of hemolysis was most likely due to positioning issue but the cause for position issues is unknown. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3784 | 1220648-2025-48027 | 23454952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23454952&pc=OZD | 09/18/2025 |  |  | 11/03/2025 |  |  | 09/18/2025 | 11/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698844 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced atrial fibrillation and lost doppler pulses. the pump was explanted and the patient underwent an embolectomy and an endarterectomy. | [5 empty] | The cause of the ischemia, thrombosis, and paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | E1 added initial reporter phone number as it was omitted from the initial report that was submitted. h11 additional information was received that the pump is not available. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3785 | 1220648-2025-48030 | 23456405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23456405&pc=OZD | 09/18/2025 |  |  | 11/03/2025 |  |  | 09/18/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739166 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a gastrointestinal bleed. after a while, care withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3786 | 1220648-2025-48028 | 23456668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23456668&pc=OZD | 09/18/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/18/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672899 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that there was cardiac arrest during impella cp patient support. cardiopulmonary resuscitation was given until return of spontaneous circulation was achieved. later, care was withdrawn and the patient expired. | [5 empty] | The root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3787 | 1220648-2025-48031 | 23457711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23457711&pc=OZD | 09/18/2025 |  |  | 11/03/2025 |  |  | 09/18/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716171 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 58 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a thrombus in the left ventricle. the pump was explanted and the patient survived. | [5 empty] | No product was returned for investigation. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3788 | 1220648-2025-48032 | 23458072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23458072&pc=PYX | 09/18/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 09/18/2025 | 11/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659271 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced hematuria and hemolysis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Data logs showed that the pump ran without issue during support. there were no mc spikes, abnormal ps or purge issues observed during support. the product was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the hematuria was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3789 | 1220648-2025-48034 | 23459499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459499&pc=OZD | 10/04/2025 |  |  | 11/03/2025 |  |  | 10/04/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728958 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Other | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient had undergone an intervention to their left anterior descending coronary artery on (b)(6) 2025 and decompensated in the unit. a left heart catheterization and right heart catheterization were scheduled for (b)(6) 2025 and an impella cp was implanted. after the impella cp was implanted, it was noted that the patient¿s leg did not have dopplerable pulses. the decision was made to escalate the patient to an impella 5.5. on day 6 of impella 5.5 support, the patient was taken to the operating room for a mitral valve repair and coronary artery bypass graft. the patient developed postoperative right ventricular failure, and the decision was made to implant an impella rp flex. on day 3 of bipella support with the impella rp flex and impella 5.5, the patient spiked a fever. the medical team decided to remove the swan and central line. post removal of the swan, the placement signal numbers spiked on the impella rp flex, and the flows dropped. the decision was made to remove the impella rp flex and escalate the patient to veno-arterial extracorporeal membrane oxygenation (va ecmo). the patient remained on impella 5.5 and va ecmo support. the following day, the patient¿s family made the decision to withdraw care of the patient. the patient¿s care was withdrawn, and the patient expired. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). pump 3: impella rp flex with serial number (b)(6). this report specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event and noted the following: the patient presented with cardiogenic shock (cgs) following intervention to the left anterior descending (lad) artery. an impella cp was implanted for hemodynamic support. post-implant, distal limb ischemia was observed at the impella cp access site, prompting escalation to impella 5.5. six days after impella 5.5 implantation, the patient underwent coronary artery bypass grafting (cabg) and mitral valve replacement (mvr). postoperatively, the patient developed right ventricular (rv) failure, and an rp flex device was implanted. subsequently, the patient developed a fever, and the clinical team elected to remove the swan-ganz catheter and central line. upon removal, the impella rp flex placement signal spiked and flows dropped. it is unknown whether the recommended practice of reducing impella rp flex flows prior to line removal was followed. the impella rp flex was removed, and veno-arterial extracorporeal membrane oxygenation (va ecmo) was initiated. the patient remained on ecpella support until the following day, when the family decided to withdraw care. the patient expired. assessment outcome is the following the limb ischemia on impella cp considered device-related due to access site involvement. fever on impella 5.5 and rp flex likely secondary to swan/central line infection but conservatively attributed to devices. the impella rp flex device malfunction reported due to abnormal placement signal and flow drop; root cause undetermined (possible procedural deviation by not reducing flows upon removal of lines). and the death will be reported for impella 5.5 as a conservative measure; clinical deterioration most likely contributing factor. correction: after review of the case by medical safety, it was determined the impella 5.5 is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the fever, the cause of the clinical symptom was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp and impella rp flex. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3790 | 1220648-2025-48035 | 23459507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459507&pc=OZD | 10/05/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682541 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 123 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp the patient with gastrointestinal bleed were given several blood products. no other information was available. the patient was doing well. pump was explanted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3791 | 1220648-2025-48037 | 23459511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459511&pc=OZD | 10/05/2025 |  |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733487 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Low Blood Pressure/ Hypotension | Sepsis | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received. the patient was not improving per the medical doctor. the decision was made to withdraw care. the patient expired shortly after the impella was turned off surrounded by family. | A notification received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry noted the patient also experienced hemolysis. both the thrombocytopenia and hemolysis have been reported to have had a possible relationship to the impella. | The complainant reported that during support of impella 5.5 thrombocytopenia continues with heparin-induced thrombocytopenia panel being sent, though platelets given and improved some. the patient has carbapenem-resistant organisms bacteria - multi drug resistant. the patient is back up on vasoactive medications. stopped pulling fluid due to lower blood pressure issues. placing a new arterial line in the axillary. unable to evaluate advanced therapies at this time due to ongoing acute critical medical needs. there was questioning new sepsis noted however no other information provided. the patient remains on support. |  |  |  | B5 additional information was received. d6b explant date was added. h1 revised due to new information that was received. h6 added code b13 as it was omitted from the initial report that was submitted. added health effect - impact code due to new information that was received. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: table 3.2 impella catheter components : ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ investigation summary: the investigation has been completed. no product was returned for investigation. thrombocytopenia: the cause of the clinical symptom was not determined due to insufficient clinical details. bacterial infection/sepsis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | B5 updated; new information has been received. the investigation is still ongoing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3792 | 1220648-2025-48038 | 23459512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459512&pc=OZD | 10/05/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668229 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after 48 hours of impella cp placement, high purge pressure and system blocked alarm occurred. there was no kink in the purge line. the patient was correctly anticoagulated and the impella position was correct. troubleshooting was done and the issue was finally resolved by decreasing the dextrose concentration from 5% to 2.5%. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3793 | 1220648-2025-48043 | 23459519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459519&pc=OZD | 10/09/2025 |  |  | 11/03/2025 |  |  | 10/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662654 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 89 KG | Hispanic |  | Injury | Yes | No | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing and hematuria. the patient was treated with diuretics and fluid, and the p level of the pump was reduced to resolve the hemolysis. suction alarms occurred and there was concern for a clot. upon explant, a clot was visualized on the device. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 has been updated.
3794 | 1220648-2025-48042 | 23459520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459520&pc=OZD | 10/08/2025 |  |  | 11/03/2025 |  |  | 10/08/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636251 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 102 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Pumping Stopped | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Arrhythmia | Hematoma | Ventricular Fibrillation | Heart Block | Hemodynamic instability | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma at the impella access site. the patient experienced heart block and was in constant ventricular fibrillation and arrhythmia that did not respond to defibrillation. the impella also had suction events. impella support continues. | [5 empty] | The cause of the hematoma, heart block, and hemodynamic instability was not determined due to insufficient clinical details. the root cause of the ventricular fibrillation and arrythmia was unable to be determined due to insufficient clinical details. the cause of the pump suction was not established as limited clinical information was provided and insufficient data logs were returned. the cause of the temporary pump stop was not established as no product and insufficient data logs were returned. the cause of the hematoma was not established as limited clinical information was provided to confirm how hematoma developed. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3795 | 1220648-2025-48044 | 23459526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459526&pc=OZD | 09/05/2025 |  |  | 11/03/2025 |  |  | 10/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687514 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had runs of ventricular tachycardia. the patient underwent one round of advanced cardiovascular life support to achieve return of spontaneous circulation. at that time, the patient was cannulated for va ecmo. per doctor¿s comment the vt was not attributed to the impella. the patient was successfully explanted with heart recovery. | [5 empty] | The cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3796 | 1220648-2025-48045 | 23459572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459572&pc=OZD | 07/31/2025 |  |  | 11/03/2025 |  |  | 10/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On october 9, 2025 a publication entitled "empowering early recovery: the role of impella 5.5 in takotsubo cardiomyopathy complicated by cardiogenic shock" reported that a female patient in her 70s implanted with an impella cp experienced an access site bleed that was treated with transfusion of two units of red blood cells. the patient survived. | [5 empty] | The device was not returned for analysis. the cause of the bleed was not determined since product was not returned and insufficient clinical details provided. | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3797 | 1220648-2025-48046 | 23459585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459585&pc=OZD | 10/09/2025 |  |  | 11/03/2025 |  |  | 10/09/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A publication entitled: "extended use of axillary impella 5.5 as a bridge to post-infarct ventricular septal defect repair" stated an inability to get the p levels of an impella 5.5 higher than 4 and reported the patient experienced renal failure requiring renal replacement therapy. | [5 empty] | The cause of the injury was not determined due to insufficient clinical details. the cause of the low pump flow issue was not determined without returned product and sufficient clinical information. | Patient information was not provided. section a was left blank. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3798 | 1220648-2025-48039 | 23459702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23459702&pc=OZD | 10/05/2025 | 11/03/2025 |  | 11/03/2025 |  |  | 10/05/2025 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735720 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Backflow | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | Great Vessel Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, the patient's blood pressure dropped after percutaneous coronary intervention. external iliac artery was perforated along with a retroperitoneal hematoma. antegrade angiography did not initially reveal any vessel damage due to retrograde pump flow. extracorporeal membrane oxygenation was added, impella catheter removed, and an 8mm catheter was placed through the introducer at the ruptured site, achieving hemostasis. vessel repaired by covered stents and blood replacement required. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3799 | 1220648-2025-31911 | 23460376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23460376&pc=OZD | 08/27/2024 | 11/03/2025 |  | 11/03/2025 |  |  | 08/26/2024 | 11/03/2025 |  |  |  |  | 11/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308461 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during impella support, an error code in the pump occurred shortly after placement. the patient was in poor condition. the pump was not replaced due to the patient¿s condition. the patient ultimately expired due to subdural hematoma. the event of death captured under the sister records for this record. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: review both the lt and rt log. utilize the lt log to visualize the time of the failure and to see if the placement signal (ps) issue resolved. utilize the rt log to visualize the time the first placement signal issue occurs. include snapshots of both below. this snapshot should show relevant 5s parameters (snr, raw optical sensor signal, calculated placement signal, contrast %, lamp %, gain). device analysis: the console ((b)(6)) was investigated in (b)(4), which occurred after (b)(4). the root cause of the failure was a defective optical bench, specifically a localized defect of its ccd, which led to the optical bench passing the 5s testing but failing during pump use. complaint (b)(4), console sn (b)(6), pump sn (b)(6). root cause: the root cause of the failure was a defective optical bench, specifically a localized defect of its ccd. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3800 | 1220648-2025-47967 | 23434645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434645&pc=OZD | 10/02/2025 |  |  | 10/31/2025 |  |  | 10/02/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714026 |  | 000000000001000100 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 57 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day 2 of support, there were reported purge pressure high alarms. the medical team first verified there were no kinks or blockages in the purge tubing and then decided to add tissue plasminogen activator (tpa) to the purge solution. additionally, it was noted that a thrombus was present. the following day, there was still high purge pressure alarms, and it was noted that there was still about a half a bag of purge solution with tpa to be used. the tpa was used through the next day, on support day 4, however no further alarms were reported. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. purge pressure high: the data logs confirmed high purge pressure alarms and showed a gradual rise in purge pressure for about four hours before an alarm triggered. before purge pressure rise, flows were at p6 became saturated with increasing pump flows were outside the expected range of 3.4 - 4.1 liters per minute and motor current was trending up. there was no motor current spike before the rise. purge flow also was decreasing. there were two bag changes following high purge pressure trigger, both within standard time. clinical details mentioned tissue plasminogen activator was added to the purge. elevated purge pressure was sustained through till end of logs. the cause of the purge pressure high was determined to be due to biomaterial build up as evidenced by log pattern showing gradual purge pressure increase with no motor current spike prior. device history review: the pump passed all post sterile inspection checks.
3801 | 1220648-2025-47961 | 23434677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434677&pc=OZD | 10/04/2025 | 10/30/2025 |  | 10/31/2025 |  |  | 10/04/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672878 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp support had atrial fibrillation and an amiodarone bolus and magnesium/potassium were replaced. additionally, the impella was alarming low flows. albumin was running. the patient continued with impella support and was explanted as planned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: the cause of the low pump flow could not be determined due to insufficient clinical details, and no product or data returned. device history review: the device passed all post sterile inspection checks.
3802 | 1220648-2025-47968 | 23434794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434794&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712868 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 110 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient came to the emergency room with worsening chest pain. an impella cp was placed and percutaneous coronary intervention was completed. soon after arrival to another facility, the patient was turned to the side. the left ventricle (lv) placement signal became severely negative and could not be visualized on the placement screen, followed by an immediate suction alarm. flows decreased to p4, at which point an lv placement signal was visible but completely decoupled from the placement signal, with suction observed again. at this point, no lv placement signal values were provided on the automated impella controller. an echocardiogram (echo) was requested for repositioning. bedside echo by the fellow failed to visualize the impella behind the valve due to a poor imaging window. suction was noted at each p-level. motor current with each suction suggested that suction could be resolved by dropping one p-level. however, decreasing the p-level triggered a new suction alarm indicating that decreasing would cause detachment, which could not be resolved. the impella was pulled back by 4 centimeters to the 91 centimeters mark, but the issue persisted. on further review, it appeared the pump may have ingested a clot at the outflow/motor housing junction, causing suction at every p-level and dampened motor current. the patient was not deemed a candidate for escalation of care and expired on impella support. | [5 empty] | No product was returned. the root cause root cause of the optical sensor issue/suction cannot be established. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3803 | 1220648-2025-47969 | 23434894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434894&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620891 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced suction. the pump's p level was reduced without resolution. the patient was given volume and the pump was repositioned many times without resolution. the pump was removed and inspected for a clot, but none was observed. no other impella cp pumps were available for use so the same impella cp was readvanced. upon restart, the suction events occurred at all p levels. an ultrasound showed that the patient's mitral valve leaflet was near the inlet chamber, possibly resulting in suction events. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The pump was not returned for investigation. the data logs confirmed the complaint. the cause of the positioning issue and pump suction could not be determined due to insufficient clinical details. the root cause of the suction was most likely related to the patient's condition, as the patient had mitral regurgitation and the mitral valve leaflets also may have been interfering with the pump. the device passed all post sterile inspection criteria.
3804 | 1220648-2025-47970 | 23434923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434923&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729654 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp after exchanging the peel away sheath for the reposition sheath, the site began to bleed and a hematoma developed. manual pressure was held but failed to stop the bleed. pressors were started and the impella was explanted. the patient survived. | [5 empty] | No product was returned for investigation. the root cause of the bleeding and hematoma were unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3805 | 1220648-2025-47971 | 23434950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434950&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713145 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 116 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 anticontamination sleeve became disconnected after repositioning. a tegaderm was placed. the patient was weaned and explanted as planned. | [5 empty] | The cause of the sterile sleeve damage issue was not determined without returned product investigation and sufficient clinical details. the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3806 | 1220648-2025-47972 | 23434994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23434994&pc=OZD | 10/17/2025 |  |  | 10/31/2025 |  |  | 10/17/2025 | 10/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685110 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Multiple Organ Dysfunction Syndrome | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp encountered pump suction and developed ischemia. the ischemia in the patient's right leg was caused by extracorporeal life support cannulation. the medical team decided to amputate the lower leg. later, the patient developed complete multi organ failure and therapy was limited. this was noted to not be impella related and the patient expired. | [5 empty] | Complaint/patient coding was updated/corrected following medical safety/clinical review as follows: removed: ischemia (e0509) and amputation (f1902) as right lower extremity ischemia was identified and attributed to ecls cannulation resulting in amputation. note: the impella was placed via the axillary artery. inserted: multi-organ failure/multiple organ dysfunction syndrome (e2342). therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the final coding that accurately reflects the reported event. | Medical safety/clinical reviewed the event. a 35-year-old female with an unknown prior medical history presented in fulminant cardiogenic shock with an estimated ejection fraction of approximately 5%. postpartum cardiomyopathy was suspected. primary extracorporeal life support (ecls) was initiated with peripheral veno-arterial extracorporeal membrane oxygenation (va-ecmo). due to persistent severe cardiogenic shock, a decision was made to escalate to combined ecpella support. impella 5.5 was initially selected; however, advancement into the left ventricle was unsuccessful due to patient vascular anatomy. after multiple attempts, the procedure was aborted, and a right axillary artery impella cp was successfully implanted instead. two days after initiation of impella cp in combination with va-ecmo, multiple suction alarms and low pump flow events were observed which can be expected with combined therapies and patients that are in persistent cgs. the patient remained in critical condition and was noted to be 100% dependent on mechanical circulatory support. on the same day, right lower extremity ischemia was identified and attributed to ecls cannulation. the clinical team proceeded with right lower leg amputation. subsequently, the patient developed complete multiorgan failure (mof). given continued clinical deterioration and poor prognosis, therapy was limited. the patient expired. the unsuccessful attempt to place the impella 5.5 due to vascular anatomy will be reported as a malfunction type of reportable event (inability to advance). the death will be conservatively reported against the impella cp. the patient underlying condition (fulminant cardiogenic shock, suspected secondary to postpartum cardiomyopathy) contributed most to the demise outcome complicated the limb ischemia related to va-ecmo cannulation, and subsequent multiorgan failure. per the physician, the demise was not considered related to impella device function. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3807 | 1220648-2025-47973 | 23435173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23435173&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739060 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp on (b)(6) 2025, that the patient was on rp and was then put on cp as well. the cp was then explanted at bedside with manual pressure hold. unable to gain hemostasis, therefore patient taken to ccl and h a covered stent was needed to achieve hemostasis after cp explant. | [5 empty] | A4 is unknown. the impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. no root cause was identified for the bleeding. the device passed all post sterile inspection checks.
3808 | 1220648-2025-47974 | 23435884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23435884&pc=PYX | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/03/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575359 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 79 KG |  |  | Death | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | Tachycardia | Hematuria | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella rp flex had a run of ventricular tachycardia and had to get shocked. additionally the patient had confirmed hemolysis and hematuria. during support the impella rp flex had possible sensor drift with negative waveforms with pulmonary artery pulsatility index intermittently dashing out. trend screen shows central venous pressure with a negative value. a decision was made by the family to withdraw care, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria: the cause of the injury was not determined due to insufficient clinical details. ventricular tachycardia/arrythmia: the cause of the injury was unable to be determined due to insufficient clinical details. mechanical interaction with blood/hemolysis: the cause of the mechanical interaction with blood was most likely related to patient condition based on clinical details that patient had hemolysis with their previous rvad, data logs confirm suction events throughout support, and the patients condition was deteriorating. optical signal issue: the cause of the optical signal issue was determined to be patient condition related based on data log review showing negative cvp during the drift down in pa waveforms and clinical detail that patient was deteriorating while on support. device history review: the device passed all post sterile inspection checks.
3809 | 1220648-2025-47975 | 23435927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23435927&pc=OZD | 10/03/2025 |  |  | 10/31/2025 |  |  | 10/03/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728401 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 93 KG |  | Black Or African American | Death | Yes | No | False Alarm | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Low Blood Pressure/ Hypotension | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced intermittent position unknown alarms. the abiomed representative explained the meaning of the alarm and actions required. the patient was also hypotensive. later, care was withdrawn and the patient expired. | [5 empty] | Neither data logs nor product were returned for investigation. the cause of the optical signal issue could not be determined. the cause of the suction could not be determined. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3810 | 1220648-2025-47977 | 23436055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23436055&pc=OZD | 10/04/2025 |  |  | 10/31/2025 |  |  | 10/04/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723536 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support that heparin was stopped due to a concern for gastrointestinal (gi) bleed. gi consult was being placed and vascular surgery consults were made for eventual removal. the pump was removed the next day. the patient survived. | [5 empty] | The investigation for the reported issue has been completed. no product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3811 | 1220648-2025-47976 | 23436074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23436074&pc=OZD | 10/04/2025 |  |  | 10/31/2025 |  |  | 10/04/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667199 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 114 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp that the patient had oozing from intra-aortic balloon pump (iabp) site, nose and mouth from intubation, and oozing from impella site in left groin. heparin assay supratherapeutic. site soft with no signs or symptoms of hematoma. hgb within normal range. not starting systemic anticoagulation due to lab results. bicarb purge fluid has been ordered but not hung, d5 still infusing through purge. assisted bedside nurse with re-dressing site and angle matching. iabp removed and care team having difficult time achieving hemostasis to site. one unit of blood was ordered. groin site bleeding stopped but no alt or ptt drawn since yesterday. on 10/6 care team stated the patient¿s urine began to turn pink this morning. concerns for hemolysis were conveyed. provider stated he believed that the patient had hemolysis occurring. echo ordered to verify placement and impella was weaned down. the patient successfully weaned from impella cp and did not require further use of vasopressors. hemostasis achieved successfully with perclose x2, pulses present. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of major bleed was most likely patient condition related since the bleeding occurred at multiple sites. device history review: the device passed all post sterile inspection checks.
3812 | 1220648-2025-47978 | 23436183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23436183&pc=OZD | 10/04/2025 |  |  | 10/31/2025 |  |  | 10/04/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624540 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 138 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support the physician made several attempts to place an impella 5.5 in proper position across the left ventricle. the patient is morbidly obese and the axillary artery was very deep and hard to get to. however, due to vascular anatomy and the pressure used to try and cross the motor into the aorta using many different techniques, the catheter looked to be kinked, bent and damaged from the pressure that was used to try and advance the pump. a new impella 5.5 was successfully placed. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
3813 | 1220648-2025-47979 | 23436258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23436258&pc=OZD | 10/17/2025 |  |  | 10/31/2025 |  |  | 10/17/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727347 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female | 70 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced cardiac tamponade and bleeding. the patient was transfused blood products and underwent surgery to resolve the conditions. the patient survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned for investigation. the cause of the cardiac tamponade was not determined due to insufficient clinical details. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
3814 | 1220648-2025-47981 | 23436564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23436564&pc=OZD | 10/02/2025 |  |  | 10/31/2025 |  |  | 10/02/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723580 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 98 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. upon arrival to the hospital, the patient was found to be hypotensive with audible crackles in bilateral lung bases. electrocardiogram results were noted to be largely unremarkable, showing no acute st changes. the patient¿s beta-natriuretic peptide was elevated at 1627 pg/ml and creatinine was elevated at 1.5 mg/dl. a transthoracic echocardiogram (tee) was performed to reveal a left ventricular ejection fraction (lvef) of 15% with severe left ventricular (lv) hypokinesis, moderate mitral regurgitation, a laminar thrombus at the lv apex, and grade iii diastolic dysfunction. further cardiac evaluation included a left heart catheterization with no coronary artery disease (cad) noted. right heart catheterization was performed and revealed an elevated right atrial pressure of 27 mmhg and pulmonary capillary wedge pressure of 33 mmhg, decreased pulmonary artery saturation of 34%, and a fick of 1 l/min. cardiac magnetic resonance imaging (cmri) was performed to assess the lv thrombus and ventricular function. the cmri found no indication of a lv thrombus, worsening lvef of 13% while receiving dual inotropes, decreased right ventricular ejection fraction of 28%, and a small non-cad scarring of mid inferior wall. due to the decompensated heart function, the medical team decided to implant an impella cp in the setting of society for cardiovascular angiography and interventions stage ¿c¿ decompensated cardiogenic shock with a plan to escalate the patient to an axillary impella 5.5 the next day for continued evaluation of advanced heart failure therapy. after the impella cp was placed, the position was imaged with fluoroscopy. the position of the impella did not look to be ideal, and the physician was concerned the pigtail was entangled with the mitral valve apparatus. despite numerous attempts to reposition, the decision was made to remove the impella cp device and re-insert the same pump after obtaining repeat access with a guidewire. the device was successfully placed, and good placement was noted with fluoroscopy and tee. the following day, the patient was taken to the operating room for planned escalation to impella 5.5. upon removal of the impella cp, a clot was observed to be hooked on the pigtail. it was noted that although the patient had a history of lv thrombus, no lv thrombus had been observed on the cmri prior to impella cp placement. on day 5 of impella 5.5 support, the patient¿s pulmonary artery pressures were noted to be elevated and impella support was increased from p7 to p8. it was reported that the patient was intermittently experiencing 2nd degree heart block and briefly, complete heart block a couple of days ago. the following day the patient experienced small, nuisance nose bleeds. it was noted that the partial thromboplastin time was therapeutic. the next day it was reported that the nosebleeds continued and the patient¿s bivalrudin was started and stopped intermittently to address the bleeding. on day 8 of support, it was noted that the nosebleeds had resolved, and bivalrudin was resumed. on day 9 of support, it was noted that the patient¿s hemoglobin had dropped to 6.2 g/dl and the patient received 1 unit of packed red blood cells (prbcs). the repeated hemoglobin level was 7.3 g/dl post transfusion. on day 10 of support, it was reported that the previous day the patient was out of bed and that overnight the patient had increasing anxiety. the patient was tachycardic with heart rate around 130 beats per minute. the patient experienced additional minor nose bleeds and was given medication for anxiety. on day 13 of support, the patient was taken to the cardiac catheterization laboratory for cannulation of venoarterial extracorporeal membrane oxygenation (va ecmo) for ongoing progressive biventricular failure. after va ecmo support was initiated, impella 5.5 support was decreased to p6. the patient was noted to have been transfused with an additional 2 units of prbcs and 250mls of albumin. overnight, it was reported that there were suction alarms and the patient experienced increased ectopy with reduced impella flows. a point of care ultrasound revealed that the pump was orientated towards the mitral valve. the physician repositioned the device, which increased the impella flows. the patient was noted to feel much better after repositioning and the patient¿s mixed venous oxygenation levels were noted to be 65%. a small hematoma was noted at the axillary site found with computed tomography and was not appreciable with visual examination. it was additionally noted that the patient was noted to have had a laminar lv thrombus which measured 1.6cm x 1cm on (b)(6) 2025 and was not appreciable on (b)(6) 2025 cmri prior to impella placement. the patient was taken to the cardiovascular operating room for escalation to central va emco placement and remove the impella cp and peripheral va ecmo. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves two impella devices: pump 1: impella cp, with serial number 620601 pump 2: impella 5.5, with serial number 574972 this mdr specifically addresses pump 1: impella cp, with serial number 620601, which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. thrombosis: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot: device lot: 1967150 device history review: device (sn: (b)(6) passed all post sterile inspection checks.
3815 | 1220648-2025-47980 | 23436616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23436616&pc=OZD | 10/04/2025 |  |  | 10/31/2025 |  |  | 10/04/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740978 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 79 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support a patient had a foley placed when the patient got to the intensive care unit. hematuria was noted. question traumatic foley insertion. on echocardiogram, the pump appeared slightly deep. the team decided to reposition the pump. under echocardiogram guidance, the pump was repositioned to 3.2 centimeters mid inlet to valve. urine cleared shortly after repositioning. bleeding was noted from the access site and manual pressure was held. a pressure dressing was applied. the patient weaned to p3. the patient was taken to the cardiac catheterization lab for explant. hemostasis was achieved with perclose x 2. | [5 empty] | The investigation was completed and no product was returned for investigation. hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3816 | 1220648-2025-47982 | 23437165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23437165&pc=OZD | 10/04/2025 |  |  | 10/31/2025 |  |  | 10/04/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | P140003 |  |  | ABIOMED, INC. |  |  | US |  | 2026714006 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 57 KG | Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Anemia | Arrhythmia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was placed onto impella 5.5 support due to cardiomyopathy. after insertion, an echocardiogram showed the impella pointed towards the mitral valve, likely not able to advance further unless position was able to be changed more apical as the inlet would likely interact with the mitral valve. the next day, the patient received one unit of blood for low hemoglobin and hematocrit. the following night and into the next morning, the patient experienced ectopy. the impella was still oriented toward the mitral valve. the impella 5.5 was repositioned twice, but the patient continued to present supraventricular tachycardia. the physicians decided not to reposition or medicate supraventricular tachycardia. the patient was bridged to a transplant and survived. | [5 empty] | The investigation was completed and no product was returned. anemia: the cause of the injury was not determined due to insufficient clinical details. arrhythmia, tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3817 | 1220648-2025-31905 | 23437290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23437290&pc=OZD | 07/14/2024 | 10/31/2025 |  | 10/31/2025 |  |  | 07/14/2024 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276242 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an impella 5.5 implant, and, after the implant, the optical sensor failed and there were placement signal not reliable alarms. support was continued and staff monitored. it was found that this is likely an issue with the automated impella controller (aic). no known patient harm or injury has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that soon after pump started, a placement signal not reliable alarm (#1036) presented and psnr condition did not resolve for the duration of support (35 days). there were also 4 isolated occurrences of a controller error alarm (#1045 due to opm communication invalid crc), which coincided with psnr alarms (#1036), however any direct relation (causality) could not be established: (b)(6) 2024 02:50:43 ossiopsens set invalid_bad_pressure_sample_mask, (b)(6) 2024 02:50:43 processsampleforopmdatapacketcrcerror: calculator placement signal 1045, (b)(6) 2024 02:50:43 impellaconnect_bridge: slay usb - serusb write error 26: 9 bad file descriptor, (b)(6) 2024 02:50:43 imcgui: event set: #1045 from 9 (1) (b)(6) 2024 02:50:43 processsampleforopminvalidsignalerror: calculator placement signal 1036 (b)(6) 2024 02:50:43 imcgui: event remove: #1036 from 9 (1) (b)(6) 2024 02:50:43 imcgui: event set: #1036 from 9 (1) (b)(6) 2024 02:50:43 imcgui: event remove: #1045 from 9 (1) ltlog and rtlog data confirmed that during the support placement signal values were invalid, snr had very low level, and contrast had unrealistic values between -3000 and +3000. after the pump was unplugged, the console presented with a controller error alarm (#1039 - defective optical sensor lamp). console was received in (b)(6), in connection with complaint (b)(4) (shared failure mode), and tested - the issue was reproduced: there was no light visible in the optical pump connector, due to optical bench lamp assembly failure. with a temporary replacement lamp (light bulb), the optical bench operated within specification: cavity length: 14335 nm. snr: (b)(4). contrast: 63.5%. lamp: 100.0%. gain: 1.44. mano: 764.5 mmhg. light: 2.70 v. during testing of the console, a corruption of rlm som microsd card was also observed (unrelated to this complaint). a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3818 | 1220648-2025-47983 | 23437649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23437649&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696289 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 175 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Death | [5 empty] | KIT, 14FR INTRODUCER, 13CM&25CM, STERILE. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hypotension. the patient recovered by adding epi-cal. later, care was withdrawn and the patient expired. | [5 empty] | Product complaint # (b)(4). the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event and noted the following: the patient underwent a protected pci with impella cp support. following the intervention, the impella cp was explanted; however, the patient subsequently decompensated, and a second impella cp was implanted. post-implant, the abiomed representative observed that the access site had not been secured and dressed per abiomed¿s recommendations, resulting in access site bleeding. the device was re-secured and dressed, femstop was applied, and one unit of prbc was transfused. the patient later became hypotensive during an intubation attempt and remained in critical condition. the family elected to withdraw care, and the patient expired approximately 12 hours after the second impella cp implant. the access site bleeding constitutes a reportable bleeding adverse event for the second impella cp catheter and hypotension is a noted event, but the likely cause is the result of intubation not the second impella cp device. the death will be conservatively reported on the second impella cp as the patient was on support at the time of death; however, the device is not considered a probable contributing factor, as the death is attributed to the patient¿s deteriorating condition. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the hypotension, the cause was not determined due to insufficient clinical details provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. correction. concomitant device(s) initially inadvertently omitted was added to section d10 accordingly. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella introducer.
3819 | 1220648-2025-47986 | 23439034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23439034&pc=OZD | 10/01/2025 |  |  | 10/31/2025 |  |  | 10/01/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658065 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Heart Failure/Congestive Heart Failure | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had been admitted to the hospital for pneumonia recently and was discharged after two days. the patient then presented to the hospital again in cardiogenic shock, fluid overloaded with complaints of shortness of breath, in atrial fibrillation with rapid ventricular response, with an atrial thrombus and a severely reduced ejection fraction of 10%. the patient was taken to the cardiac catheterization laboratory and was noted to be hypotensive with an elevated pulmonary wedge pressure of 40 mmhg. the decision was made to place an impella cp for mechanical support. on day 2 of support, it was reported that the patient experienced hemolysis due to a malposition of the pump that occurred overnight. the pump was repositioned to resolve the hemolysis. it was also noted that the patient was somewhat hemodynamically stable and had stabilized overnight. however, due to worsening laboratory results, sub-optimal cardiac output and cardiac index, and decreased urine output, the decision was made to escalate the patient to an impella 5.5. the patient was successfully escalated to the impella 5.5 and the impella cp was explanted. the patient was noted to have been hemodynamically stable post impella 5.5 support, vasopressors were weaned as tolerated, and continuous renal replacement therapy was to be started. the impella 5.5 device was explanted successfully after weaning following 10 days of support with no reported complications. the patient's outcome at the time of explant was survived. | [5 empty] | H6 medical device problem code patient device interaction problem was not included in the initial mdr and is now correctly added. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> no product was returned. cardiac failure, acute kidney failure: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was most likely use issue related since pump position was deep and left in ventricle for a while and later repositioning resolved positioning and hemolysis. it was observed that there was some patient agitation or something which led to pump position being deep. and they were looking for knee immobilizer to stop any patient movements. hemolysis: patient had hemolysis. after review of logs, multiple impella in ventricle alarms were observed. no suction alarms or motor current alarms were observed. the cause of hemolysis was most likely use issue related since hemolysis was resolved by repositioning of pump per clinical. device history lot ==> device lot: 584113 device history review ==> device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3820 | 1220648-2025-47985 | 23439042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23439042&pc=OZD | 10/17/2024 |  |  | 10/31/2025 |  |  | 10/17/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025514297 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted wtih an impella 5.5 had a cerebral hemorrhage. according to the reported the bleeding was not related to the impella. however, there is not enough information to rule out impella contribution to event. | [5 empty] | The cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3821 | 1220648-2025-47987 | 23439189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23439189&pc=OZD | 10/18/2025 |  |  | 10/31/2025 |  |  | 10/18/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670295 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation so catecholamines were increased. later, the patient was successfully weaned from the pump and survived. in addition, it was reported the purge cassette was leaking. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product returned. peri-procedure adverse event (arrhythmia, tachycardia) - in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the arrhythmia, tachycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1931432. device history batch: subcomponent lot: n/a. device history review: pump s/n (b)(6) passed all post sterile inspection checks.
3822 | 1220648-2025-47988 | 23439258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23439258&pc=OZD | 10/01/2025 | 01/07/2026 |  | 10/31/2025 |  |  | 10/01/2025 | 10/31/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662074 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented to the emergency department with tachycardia and became hypotensive and hypoxic. it was noted that the patient¿s troponin results were elevated. the patient was intubated and brought to the cardiac catheterization laboratory for intervention. an impella cp was placed, and support was initiated, however the impella immediately alarmed ¿impella stopped¿ and ¿impella stopped motor current high¿. the impella was attempted to be restarted twice but was not successful. due to the patient¿s status, the decision was made to remove the first impella cp and place a new device. the new impella cp was delivered successfully, and support was initiated. a cardiac output response to stress test revealed that the right coronary artery was 100% occluded and 99% occluded left circumflex coronary artery. intravascular lithotripsy and balloon angioplasty were used to address the occluded left circumflex coronary artery, and a stent was placed. the patient then became hemodynamically unstable, and the decision was made to stop the procedure, admit the patient to the intensive care unit and transfer to an outside hospital for advanced therapy. it was noted that the patient was transferred and there were no distal pulses able to be dopplered. on day two of support with the second impella cp pump, the patient was noted to have a fever, and the patient was treated with vancomycin and piperacillin. blood cultures were drawn, and the medical team was concerned that the patient was experiencing a mixed sepsis and cardiogenic shock. additionally, the automated impella controller (aic) had a ¿controller error¿. to resolve the issue, the aic was replaced with a back up unit. on day three of support, the medical team believed that the patient was experiencing distributive shock. the patient experienced a blood stool overnight, which the medical team believed was caused by low perfusion upon arrival. the patient¿s distal pulses were noted to be dopplerable. due to the blood stool, heparin was withheld. the medical team attempted to wean the patient down to p4 on the impella; however, the patient decompensated and went into supraventricular tachycardia. the patient was shocked and the patient was converted into ventricular tachycardia. the patient was shocked again and converted to sinus tachycardia. fluoroscopy was performed to confirm impella placement. transesophageal echocardiogram was recommended due to the patient being intubated by company support, however the medical team declined. on day four of support, the patient¿s platelets were noted to be decreased at 30,000 platelets per microliter and the medical team was concerned for heparin induced thrombocytopenia (hit). heparin continued to be withheld. the following day, the cardiology team wanted to wean the impella as soon as possible in accordance with the patient¿s hemodynamic stability as the patient¿s platelets continued to decline to 22,000 platelets per microliter. on day six of support, the patient¿s blood cultures were positive for yeast growth. the medical team suspects the patient is in septic shock in the setting of fungemia. hit test results came back negative, and the patient¿s white blood count was increased with a result of 22,400 per microliter of blood. the patient remained on antibiotics and antifungal medications. the second impella cp device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella cp, with serial number (b)(6). automated impella controller with serial number (b)(6). this report specifically addresses pump 1: impella cp, with serial number 590316, which is the subject of this report. separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | Investigation summary: impella cp sn (b)(6) returned for evaluation. the returned pump was investigated. the pump was soaked in warm water to get rid of any dried blood and check for biomaterial. no biomaterial was found in or around the impeller cage and cannula. the impeller cage was found to be slightly deformed upon closer inspection. before running the pump, the pump was successfully primed. the pump was then run in the lab and produced high motor current pump stops, unable to start. the pump was subject to electrical testing to check for any open electrical lines, and the resistance values were all within spec; no evidence of an open electrical line was found. it was observed that the impeller blades were not spinning, and the pump was subject to a teardown to look for further evidence. after cutting the cage struts off, it was seen that the impeller blades started rotating freely. data logs were returned for analysis. it was seen that there was an "impella stopped - motor current high" alarm when the pump was started. during this alarm, motor current spiked to 1471 ma, while placement signal and purge pressure remained within expected ranges. pump flow and motor speed remained at 0. "optical system error controller" and "optical system error catheter" alarms were also seen. no other abnormalities were seen. pump stop: clinical details mention that the impella pump immediately started alarming "impella stopped" "impella stopped motor current high" when the pump was started in auto mode. impella was reattempted to start 2 additional times unsuccessfully. due to the status of the patient, decision was made to open new impella cp pump. the root cause of the pump stop was most likely deformation of the outflow cage based on the evaluation of the returned product and provided clinical details. device history lot device lot #: 1917464. device history batch subcomponent lot #: n/a. device history review pump sn (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3823 | 1220648-2025-47989 | 23439840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23439840&pc=OZD | 10/05/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026763908 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 73 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp manual pressure applied for small hematoma. bicarb for purge. will start systemic heparin per protocol. the pump was later explanted successfully and the patient survived. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: manual pressure applied for small hematoma. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3824 | 1220648-2025-47990 | 23441431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23441431&pc=OZD | 10/02/2025 |  |  | 10/31/2025 |  |  | 10/05/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607182 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 88 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during support of impella cp, the nurse called today regarding alarms, low placement signal and impella in ventricle. the impella depth was assessed with bedside and echo which confirmed it was deeper than the previous echo. the physician plans to stop by for repositioning later this afternoon. later the impella is currently alarming in aorta. echo at bedside with visualization of being shallow. awaiting physician for reposition. the patient currently oozing from multiple locations. heparin stopped. per nurse, patient was turned around 1am and oozing and impella position alarms ensued. the pump was explanted the next day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3825 | 1220648-2025-31906 | 23443480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23443480&pc=OZD | 01/22/2023 | 10/31/2025 |  | 10/31/2025 |  |  | 01/22/2025 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1129885 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 lose placement signal on day two of support. there was no harm or injury from the sensor failure, and the pump remained on for support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. analysis: on (b)(6), rlfs logs show that the rt logs indicate oscillating contrast with snr dropping to near zero with unreliable placement signal. imc logs note the placement signal not reliable alarm coinciding with the optical signals going awry. at preventative maintenance on (b)(6) 2023, the lamp, 2450-0038, was replaced due to failed pm spec. root cause: the cause of the optical signal issue was most likely the lamp malfunctioning as confirmed by the component replacement at the subsequent pm. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3826 | 1220648-2025-31907 | 23443527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23443527&pc=OZD | 06/13/2024 | 10/31/2025 |  | 10/31/2025 |  |  | 06/14/2024 | 10/31/2025 |  |  |  |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1413258 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a yellow alarm after inserting the impella cp indicating placement signal was not reliable. the event was resolved successfully with the same aic. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: optical placement signal issues on ic7087 with pump (b)(6) were confirmed in device logs, but they were most likely related to the pump. a separate complaint regarding pump¿s contribution to the same issue is investigated under (b)(4). device analysis: there was minor contamination on blue receptacle fiber but would not cause sudden drop in snr. after it was cleaned, ¿5s¿ parameters were measured and were found to be within specification. minor contamination of the fiber couldn¿t have contributed to the abrupt drop of snr observed in the case data, which was most likely related to the pump (investigated under 24-33006-1). console operated within specification. please test the optical system of the console: inspect optical pump connector for any debris that may have prevented the pump connector from being pushed all the way in, inspect connector fiber for cleanliness, and measure ¿5s¿ parameters. root cause: unable to determine due to inability to reproduce issue. console operated within specification. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3827 | 1220648-2025-31908 | 23443673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23443673&pc=OZD | 07/15/2024 | 10/31/2025 |  | 10/31/2025 |  |  | 07/15/2024 | 10/31/2025 |  |  | 01/31/2026 |  | 10/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454961 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and that a coronary shockwave was used. while "shocking" the left ventricle, a placement signal alarm was triggered with the left ventricle placement signal being lost. the pump was functioning properly otherwise. the doctor decided to leave the pump as it was and monitor. no known patient harm or injury has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. clinical reported multiple balloons, wires, and shockwave were used. clinical reported impella optical sensor was pulled back from lmain prior to shockwaving however distance between shockwave and device was not specified. the pump was returned for investigation. no physical abnormalities were observed. the pump passed the laser continuity test for the optical sensor. the optical signal parameters indicate a hard failure trend on returned product. data log shows the placement signal not reliable alarm along and all 5s parameter was down which indicate hard failure of the sensor. placement signal was lost immediately after the shockwave device was used. the cause of the optical signal issue was a damaged sensor due to the shockwave device interaction. per ifu 1000349 ri use with shockwave coronary intravascular lithotripsy (ivl) catheter at a distance of less than 20 mm between the optical sensor and ivl device may interfere with or damage the impella optical sensor. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3828 | 1220648-2025-47945 | 23423438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23423438&pc=OZD | 09/30/2025 |  |  | 10/30/2025 |  |  | 09/30/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722288 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 116 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Inflammation | Nerve Damage | Tachycardia | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, patient experienced epistaxis requiring stoppage of heparin. patient experienced tingling in hand/fingers throughout most of the duration of support with swelling in right arm. suction occurred periodically with no definitive resolution. impella site had a hematoma. short runs of non-sustained ventricular tachycardia overnight. patient was put on dobutamine for low indexes. the patient was later bridged to transplant successfully and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction h6 health effect - clinical code. the investigation of the reported failure mode has been completed. no product was received for evaluation. inflammation, epistaxis, nerve compression, hypotension: the cause of the injury was not determined due to insufficient clinical details. tachycardia: the cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the cause of the hematoma was not determined due to insufficient clinical details. pump suction: after reviewing logs, multiple suction alarms were observed. no sudden motor current spike before the suction alarms was observed. the cause of low pump cannot be determined due to insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks.
3829 | 1220648-2025-31898 | 23423466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23423466&pc=OZD | 03/10/2023 | 10/30/2025 |  | 10/30/2025 |  |  | 03/10/2023 | 10/30/2025 |  |  | 11/30/2024 |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024285059 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump exhibited placement signal issues. the device was explanted and replaced, resolving the reported issue. no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the device was returned for evaluation, and the following was found: the returned pump was plugged into a console and a controller error alarm immediately triggered. no placement signal reading prompted. the light test was performed and no light emitted from the sensor. the location of the light loss was in the catheter next to the pump. the catheter was torn down and a disconnect of the optical fiber line from the motor was observed. the cause of the optical signal issue was a damaged optical fiber line due to an insufficient delo bond of the nitinol tube to the yoke. the root cause of this issue was determined to be manufacturing process. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3830 | 1220648-2025-47946 | 23423540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23423540&pc=OZD | 10/16/2025 |  |  | 10/30/2025 |  |  | 10/16/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709722 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Male | 72 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge leak. it is not yet clear how the issue was resolved as the patient remains on impella support. | [5 empty] | No product returned for investigation. the cause of the purge leak (pump) was not determined due to no product returned and insufficient clinical details. the device passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3831 | 23423606 | 23423606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23423606&pc=OZD | 09/04/2025 | 09/11/2025 | 09/11/2025 | 10/30/2025 | 09/11/2025 | 10/30/2025 |  | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | Yes | Nurse | [6 empty] | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Therapeutic or Diagnostic Output Failure | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.5 was not functioning properly for the patient. was giving motor error codes. we were advised to switch out the impella. on removal of the impella we found the motor to be full of clot causing it to not function.
3832 | 1220648-2025-47951 | 23423942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23423942&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740855 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Pulmonary Embolism | Arrhythmia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed too deep resulting in ventricular tachycardia (vt). the impella was pulled back into the correct position and the patient stabilized with occasional vt throughout the case. the physician stated the vt was due to instability and not related to the pump. it was further reported that the patient developed a tension pneumothorax overnight which was managed with placement of a chest tube. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.1 added selection that was omitted from the initial report that was submitted. b.7 added information that was omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was omitted from the initial report that was submitted. h.6 code e0601 that was on the reported on the initial report was changed to code e060109. investigation summary: no product was returned. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. pulmonary embolism: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was placed initially too deep by the user. device history review the device passed all post sterile inspection checks.
3833 | 23423964 | 23423964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23423964&pc=OZD | 10/16/2025 | 10/16/2025 | 10/16/2025 | 10/30/2025 | 10/16/2025 | 10/30/2025 |  | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  | 605319 |  | 1000100 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The surgical team implanted and explanted a cath impella smart assist because the product was defective.
3834 | 1220648-2025-47953 | 23424170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424170&pc=OZD | 10/01/2025 | 10/30/2025 |  | 10/30/2025 |  |  | 10/01/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717761 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 82 KG | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Ischemia | Renal Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 had recent suction alarms although there was adequate volume. the impella 5.5 was found to have been flipped toward the mitral valve. the device was repositioned at bedside. a few days later the patient was extubated due to signs of delirium. the patient was placed on do not intubate. patient decompensated and given milrinone to support hemodynamic stability. patient was placed on dialysis and presented with bowel ischemia. care was withdrawn due to continued decompensation. patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. pump suction: the cause of pump suction was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3835 | 1220648-2025-47950 | 23424419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424419&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714086 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during the impella cp placement while exchanging the peel away sheath for the repo sheath, the patient went into ventricular fibrillation requiring defibrillation. a small hematoma was noted at the insertion site. the hematoma resolved with mild pressure and appropriate positioning of the repo sheath. ectopy was noted and the pump was adjusted. the patient was transfused four units of packed red blood cells but the cause of the bleeding was unknown. a ct scan revealed that patient had a stroke and neuro was consulted. the patient was successfully weaned from the pump and survived. | [5 empty] | D6b. date of explant, has been added. | No product was returned for investigation. the cause of the bleed and hematoma was unable to be determined due to insufficient clinical details. the cause of the ventricular fibrillation, arrhythmia, and stroke was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3836 | 1220648-2025-47948 | 23424450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424450&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026732555 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had multiple runs of ventricular tachycardia but did not lose consciousness. the impella's position was confirmed under echo guidance. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3837 | 1220648-2025-47949 | 23424451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424451&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026737603 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced hematuria. the impella p level was decreased to resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the hematuria was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3838 | 1220648-2025-47954 | 23424574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424574&pc=OZD | 10/01/2025 |  |  | 10/30/2025 |  |  | 10/01/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730312 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp implant, as soon as the impella started the patient experienced ventricular fibrillation and required a shock of 360j to normalize rhythm. while on support the patient bradycardia ensued. the patient experienced hematuria and oozing from everywhere. blood products were given and heparin was held. the patient was eventually weaned off support and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed : the cause of the injury is most likely patient condition related since the patient experienced oozing from everywhere. hematuria : the cause of the injury was not determined due to insufficient clinical details. ventricular fibrillation/ bradycardia : the root cause of the injury was unable to be determined due to insufficient clinical details. resuscitation : the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3839 | 1220648-2025-31900 | 23424654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424654&pc=OZD | 11/09/2023 | 10/30/2025 |  | 10/30/2025 |  |  | 11/09/2023 | 10/30/2025 |  |  | 03/31/2025 |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024319072 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella 5.5 pump experienced an intermittent loss in placement signal during patient support. the alarm was subsequently disabled and support continued while the staff monitored the patient's hemodynamics, motor current, and pump flow. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were returned and there was no placement signal not reliable alarm. only snr and contrast optical parameters were affected without seeing any trend in placement signal and p-level change throughout the case run. the data matches no established consistent trends. the before and after pumps were investigated as well and no similar abnormalities were found. therefore, the root cause of the optical signal issue was not determined since product was not returned and data logs were inconclusive. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3840 | 1220648-2025-31899 | 23424741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23424741&pc=OZD | 01/10/2025 | 10/30/2025 |  | 10/30/2025 |  |  | 01/10/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107434 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal of an implanted impella 5.5 pump began to drift during patient support. the position was verified via echo and flow levels were manipulated. the placement signals were adjusted on the console and support continued uninterrupted. there was no patient harm. | [5 empty] | D4 unique identifier (udi) # added. h4 device manufacture date added. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show downward trend of placement signal, with offsets manually applied several times, until it could no longer be corrected due to signal being not reliable due to values out of range (psnr alarm #(b)(4)). ltlog and rtlog data shows that optical parameters of the raw placement signal - snr, contrast, lamp, gain - were stable during that time, therefore optical bench contribution to placement signal drift is unlikely. pump (b)(6) contribution to placement signal drift is investigated in (b)(4). console was tested in (b)(6), and although optical bench was found to be out-of-spec due to low contrast, it was unrelated to the complaint. there were no problems found with the console during the case. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3841 | 1220648-2025-31901 | 23425620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23425620&pc=OZD | 03/09/2023 | 12/01/2025 |  | 10/30/2025 |  |  | 03/09/2023 | 10/30/2025 |  |  | 10/31/2024 |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293445 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon opening impella cp¿s catheter and plugging in, an "optical sensor system error" immediately appeared. the complainant unplugged the catheter and blew air into port thinking it was dust possibly obstructing the optical sensor from connecting to the console. that did not work, so they got a connected to another automated impella controller, and the same error appeared. a new catheter had to be opened, and upon connecting the new catheter, it worked as it normally would. no patient harm has been reported. | [5 empty] | H6: corrected the investigation findings and investigation conclusions codes based on the results of the investigation. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the lt logs show an immediate controller error alarm as soon as the pump was plugged into two separate aics. the laser light test showed light exiting a broken optical fiber in the red handle on the catheter side. the cause of the optical signal issue was a damaged optical fiber line in the red handle on the catheter side. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3842 | 1220648-2025-47955 | 23426601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23426601&pc=OZD | 10/01/2025 |  |  | 10/30/2025 |  |  | 10/01/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705151 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female | 86 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Sepsis | Tachycardia | Hypovolemia | [9 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s proximal left anterior descending coronary artery was perforated during a percutaneous coronary intervention, and a covered stent was placed to seal the perforation. the stent then inadvertently blocked the flow down the ramus and circumflex coronary arteries, which led to the patient experiencing multiple episodes of ventricular tachycardia (vt) which required defibrillation and levophed was initiated to resolve the arrhythmia. the decision was then made to place an impella cp, and support was initiated. following impella cp support, the patient¿s vt resolved, and blood pressure was noted to have improved. the physician was able to address the stent and restore flow to the ramus and circumflex arteries and the patient was weaned off levophed. it was then noted that the pericardial drain began filling faster, and the patient¿s blood pressure dropped. the physician found the perforation to be actively bleeding again. more stents were placed to try to control the bleeding and a total of 4 drug-eluting stents and 5 covered stents were placed. the patient lost a total of 3.5 liters of blood and multiple blood products were given. at the end of the procedure, the patient was still bleeding a small amount, and the physician administered protamine to lower the patient¿s activated clotting time. the patient was then transferred to the intensive care unit. on day 2 of support, it was noted that the patient was being worked up for hypovolemia and sepsis and the patient¿s heart rate was increased to 140-160 beats per minute. it was noted that the patient expired during impella cp support. | [5 empty] | Given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 health effect - clinical code e60109 code was changed from e601. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. tachycardia/ sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. hypovolaemia: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot: device lot: 1945897. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
3843 | 1220648-2025-47956 | 23426710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23426710&pc=OZD | 10/01/2025 |  |  | 10/30/2025 |  |  | 10/01/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577160 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hematuria and hemolyzed labs. the pump was repositioned using an echocardiogram to resolve the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The device was not returned for review. the cause of the hemolysis and hematuria was not determined due to insufficient clinical information and inconclusive log evidence. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3844 | 1220648-2025-47958 | 23429707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23429707&pc=OZD | 10/01/2025 | 10/30/2025 |  | 10/30/2025 |  |  | 10/01/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724512 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was hematoma and minor bleeding during impella cp patient support. while on support, a minor hematoma was noted around the insertion site. there was also some oozing from the right femoral site. manual pressure and tissue displacement were performed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma/ minor bleed : the cause of the injury is most likely use related since upon return to unit after implant of cp, nursing staff discovered hematoma and some oozing from right fem site. device history review: the device passed all post sterile inspection checks.
3845 | 1220648-2025-47959 | 23429817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23429817&pc=OZD | 10/01/2025 |  |  | 10/30/2025 |  |  | 10/01/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634287 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 46 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had an episode of hematemesis with concerns of aspiration. systemic heparin was withheld. the patient also experienced ventricular fibrillation/torsades de pointes. lidocaine and levophed were administered. the patient was transfused packed red blood cells in response to depressed hemoglobin and hematocrit values. later, care was withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the ventricular fibrillation and arrhythmia was unable to be determined due to insufficient clinical details. the cause of the anemia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3846 | 1220648-2025-47960 | 23430014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430014&pc=OZD | 10/03/2025 |  |  | 10/30/2025 |  |  | 10/03/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694358 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient was febrile from pneumonia. care team states it¿s unrelated to impella. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. fever: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3847 | 1220648-2025-47962 | 23430076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430076&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient with possible gastrointestinal (gi) bleed heparin on hold for the moment gi consult on bumex drip good response to diuretics -1.6l this am. no further weaning of impella until gi bleed resolved starting on protonix. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks.
3848 | 1220648-2025-47963 | 23430077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430077&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678002 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 138 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced hematuria when the pump's p level was at 6. reducing the p level to 5 resolved the issue. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned for investigation. the data log was inconclusive without sufficient calibration information and returned product investigation to derive the cause of the hemolysis. the device passed all post sterile inspection checks.
3849 | 1220648-2025-47964 | 23430097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430097&pc=OZD | 09/28/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728382 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 58 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Pulmonary Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient was put on lasix for pulmonary edema. echo done pending results. continue impella and vasopressor support. the patient went into atrial fibrillation rapid ventricular response requiring amiodarone bolus. later the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.3 revised date of event as it was entered incorrectly on the initial report that was submitted. g.3 date should of reflected 9/28/25 on the initial report that was submitted. h.6 added code b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was received for evaluation. paroxysmal atrial fibrillation: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pulmonary edema: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3850 | 1220648-2025-31902 | 23430275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430275&pc=OZD | 05/03/2025 | 10/30/2025 |  | 10/30/2025 |  |  | 06/27/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204292 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had placed a cp to support the patient admitted in with acute myocardial infarction / cardiogenic shock. the pump supported for five days and was then explanted as a left ventricular assist device was placed instead. the pump had been running with lost placement signal, but no harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data review: console logs were consistent with what was reported in the complaint. multiple placement signal not reliable (opm invalid signal, optical pump connected]) alarms [event set #1036] were observed in the imc logs. rtlog data revealed that snr was well below spec when these alarms occurred. device analysis: console was tested after arriving in field service. during testing, it was discovered that one of the 5s parameters on this console was out of spec. conclusion: the root cause of the reported placement signal not reliable alarms was likely a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3851 | 1220648-2025-31903 | 23430318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430318&pc=OZD | 10/17/2023 | 10/30/2025 |  | 10/30/2025 |  |  | 10/05/2023 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165235 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a cp with loss of placement signal after just one day of support. the optical signal failure did not lead to any harm or injury. the pump remained on for support. the decision was made to withdraw care, and the patient expired. the event of death has been filed under the sister complaint/record. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show multiple events and alarms ¿placement signal not reliable¿ (#1036, due to opm signal invalid) during the first day of support, one isolated occurrence of momentary loss of opm data, and multiple suction alarms throughout the case. logs also confirmed instances of placement monitoring having to be disabled. ltlog and rtlog data shows very low snr of the optical signal, which was causing placement signal values to become invalid, or inaccurate (¿barcoding¿). console was tested in danvers, and its optical bench was found to be passing snr testing only marginally. analysis of the analog output of optical bench ccd detector revealed significant noise in the baseline, which resulted in inability to correctly process the fringe pattern and calculate placement. inspection of the console also revealed contamination of the optical connector fiber ferule, resulting in further degradation of the optical signal. during console inspection and data analysis it¿s been discovered that the rl05609 had been frequently disconnecting from the cloud because of rlm reboots due to video signal noise, that was traced to defective rlm and compromised backstrap cable (unrelated to the original complaint). the single controller error and momentary loss of placement signal was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the cause of the aic issue was a defective optical bench. the cause of the aic issue was a software issue. the cause of the aic issue was a defective impella connect module. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3852 | 1220648-2025-31904 | 23430345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430345&pc=OZD | 02/11/2024 | 10/30/2025 |  | 10/30/2025 |  |  | 03/28/2024 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a cp support ongoing for a cardiogenic shock / acute myocardial infarction patient when the pump lost signal on the second day of support. the optical signal loss did not cause any harm or injury. the pump remained on for support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imc logs confirmed that multiple placement signal not reliable alarm had occurred. there was also one occurrence of a controller error alarm (opm comm loss ¿ 1045). following the set of errors, the optical bench was intentionally reset, which resolved the communication-related errors. real time (rt) logs confirmed that immediately prior to the controller error alarm, the optical bench reported an oscillating contrast value far outside the normal range of contrast. device analysis: console ic5307 was sent back to fs for further investigation. the issues were not able to be reproduced in the lab when running a known-good pump and purge system. a root cause was not determined, but it is known that an oscillating contrast is the result of no light output from the lamp on the optical bench. the cause of the intermittent optical bench communication-related alarms was not determined due to the inability to reproduce. ¿placement signal not reliable¿ alarms and controller error alarm were most likely caused by malfunction of the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3853 | 1220648-2025-47966 | 23430373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23430373&pc=OZD | 10/04/2025 |  |  | 10/30/2025 |  |  | 10/04/2025 | 10/30/2025 |  |  |  |  | 10/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp was suspected of having heparin-induced thrombocytopenia. heparin was withheld. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia was unable to be determined due to insufficient clinical details. | Patient information was not provided. section a was left blank. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3854 | 1220648-2025-31890 | 23410795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23410795&pc=OZD | 02/18/2025 | 10/28/2025 |  | 10/29/2025 |  |  | 02/18/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186358 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 experienced a placement signal not reliable alarm that was resolved after purge cassette was changed. additionally, preliminary data analysis indicated that this complaint involves console failure for placement signal issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imc logs of console (b)(6) during a case with pump 568179 showed there were multiple instances of ¿processsampleforopminvalidsignalerror: calculator placement signal 1036¿ and "ossiopsens set invalid_bad_snr_sample_mask,¿ one instance of ¿ossiopsens set invalid_bad_pressure_sample_mask.¿ lt logs on the complaint date confirmed the reported failure mode, given that the contrast oscillating between -3000 and 3000, consistent with intermittent loss of light. lt logs also showed that the placement signal pattern is not realistic. device analysis: reported ¿placement signal not reliable¿ alarms on (b)(6) during a case with pump 568179 were caused by failure of optical bench. console troubleshooting also revealed persistent contamination of the optical pump connector (unrelated to complaint), corrupt microsd card on rlm causing frequent interruptions of data streaming (unrelated to complaint), as well as sharp kink on impellatronic communication cable (unrelated to complaint). a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3855 | 1220648-2025-31891 | 23411597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23411597&pc=OZD | 05/07/2025 | 10/28/2025 |  | 10/29/2025 |  |  | 08/08/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112109 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella connect showed placement signal unreliable alarm with intermittently appropriate waveforms. soon after, the alarm resolved and waveforms appeared appropriate. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: during the complaint case, while running cp pump 535622, there were numerous optical signal issues in the imc logs. the lt logs showed low signal-to-noise ratio (snr) throughout the case and intermittent low placement signal further confirmed by the rt logs which would have triggered the reported unreliable placement signal. device analysis: the failure mode was confirmed as snr was 1405 with a test cavity root cause: the cause of the unreliable placement signal was a defective optical bench with low snr. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3856 | 1220648-2025-47899 | 23412050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412050&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728733 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Cardiac Arrest | Hemolysis | Low Blood Pressure/ Hypotension | Ischemia | Renal Failure | Cardiogenic Shock | Respiratory Failure | Hematuria | Multiple Organ Dysfunction Syndrome |  |  |  |  | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented to the emergency department with a suspected st elevation myocardial infarction, which was later cancelled after a review of electrocardiogram results. the patient's condition deteriorated, which required the patient to be intubated and the initiation of norepinephrine for cardiogenic shock, indicated by a lactic acid of 11.4 milligrams per deciliter (mg/dl) and acute kidney injury with a creatinine of 2.38 mg/dl. the patient was transferred to the intensive care unit (icu) and then urgently to the catheterization lab for a coronary angiogram and potential intervention. during an attempt to engage the left main coronary artery, the patient experienced cardiac arrest with ventricular tachycardia/ventricular fibrillation that progressed to asystole. the patient received defibrillation and 5-10 minutes of cardiopulmonary resuscitation (cpr), achieving return of spontaneous circulation (rosc). following rosc, an impella cp was urgently placed via the right femoral artery. during the subsequent percutaneous coronary intervention with impella cp support, the patient experienced two additional cardiac arrests, requiring further cpr. throughout the procedure, the patient was persistently hypoxic and hypotensive for extended periods of time. following the procedure, impella cp placement was confirmed, the device was secured and the patient was transferred to the icu. after arriving to the icu, it was noted that hemolysis was suspected due to lysed laboratory results and the observation of red urine, prompting a renal consultation for possible continuous renal replacement therapy (crrt). on the second day of support, the patient's condition worsened, with severely deteriorated lab results and increased vasopressor needs. the patient¿s right foot showed signs of ischemia with a mottled appearance and an ankle-brachial index of 0.4, but the cardiology team was not inclined to intervene and would consider a surgical consult for potential femoral to femoral bypass or distal perfusion catheter procedures. it was noted that crrt was initiated overnight for renal failure, and the patient¿s hemolysis continued with noted red urine. the patient continued to deteriorate and had care withdrawn in the setting of worsening shock and multi-organ failure. the patient¿s care was withdrawn and the patient expired. | [5 empty] | H6. medical device problem code patient device interaction problem was not included in the initial mdr and is now correctly added. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms: the cause of the injury was not determined due to insufficient clinical details. hemolysis: clinical reported hemolysis occurring. the product was not returned for investigation. clinical details implied that the hemolysis may not have been entirely due to the impella, as the patient also had a traumatic foley insertion. data log review shows positioning alarms occurring on the day hemolysis was reported. placement signal was also abnormal at the time, showing a general trend down. clinical details implied that the patient was in very poor condition throughout the case, placement signal reflected that, trending down and losing pulsatility for the entire case duration. based on the downward trend in placement signal and clinical details stating the patient was deteriorating, the cause of the hemolysis was most likely related to the patients condition. device history lot: device lot number: 1960394 device history review: pump sn (b)(6) passed all post sterile inspection criteria. | Although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3857 | 1220648-2025-47900 | 23412155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412155&pc=OZD | 10/03/2025 |  |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692312 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female | 65 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ischemia of the leg. the patient was escalated to an impella 5.5. the impella cp has been discarded. | [5 empty] | A5 and a6 are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | No product was returned for investigation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3858 | 1220648-2025-47901 | 23412226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412226&pc=OZD | 10/03/2025 |  |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734863 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella 5.5, the patient had blood products being given for implant site hematoma. the plan was to start bivalirudin after pressure was held on the site. | [5 empty] | No product was returned for investigation. the cause of hematoma is determined to be use issue due to patient transport. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3859 | 1220648-2025-47902 | 23412312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412312&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708284 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pulmonary Edema | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced sepsis secondary to fungemia and was treated with antibiotics. a chest x-ray showed the patient had significant pulmonary edema, and the left lower leg was mottled. an external femoral-femoral bypass was placed without resolution so the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | No product was returned for investigation. the root cause of the ischemia and sepsis was unable to be determined due to insufficient clinical details. the cause of the pulmonary edema was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3860 | 1220648-2025-47904 | 23412315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412315&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/30/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605564 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received. there was no active ischemia with the patient but the vascular medical doctor was concerned due to the patients know pvd. the pump was removed without complication and there was no ischemia to the leg. | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had suffered cardiac arrest during a percutaneous coronary intervention procedure, and the decision was made for an impella cp placement for stabilization. after the impella cp was placed, there were loss of pulses in the left lower extremity which indicated possible acute limb ischemia, prompting vascular consultation. the impella cp was explanted due to patient improvement and concern for limb ischemia, with the patient surviving the procedure. |  |  |  |  | E1 revised initial reporter first name as it was entered incorrectly on the initial report that was submitted. g3 date should of reflected (b)(6) 2025 on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. ischemia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3861 | 1220648-2025-47906 | 23412369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412369&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1738105 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller failed to detect a purge disk. no patient involvement was reported. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | Investigation summary: data review: the console logs were consistent with what was reported in the complaint. purge disc not detected alarms were observed in the logs. device analysis: the console was tested. the reported purge disc detection issue was able to be reproduced. further inspection of the console revealed that the purge disc flag was broken (missing at blue retainer). conclusion: the root cause of the reported purge disc detection issue was a broken purge disc flag. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of the console.
3862 | 1220648-2025-47907 | 23412455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412455&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705339 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that on the first day of support the patient experienced signs of hemolysis with red colored urine and blood samples hemolyzed. it was reported that the impella was found to be too deep within the ventricle and the impella was repositioned. the following day, the hemolysis symptoms were reported to have been resolved. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary ==> no product was returned. hematuria: the cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: suspected hemolysis; red urine and blood sample hemolytic; 4pm pump repositioned. no hemolysis signs any more. no logs are available to review. the cause of hemolysis was due to position related but the cause of position issue is unknown. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1949065. device history batch ==> subcomponent lot: n/a. device history review ==> device (sn: (b)(6)) passed all post sterile inspection checks.
3863 | 1220648-2025-47910 | 23412788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23412788&pc=OZD | 10/14/2025 |  |  | 10/29/2025 |  |  | 10/14/2025 | 10/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026752400 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a patient encountered complications during impella cp support. the medical doctor (md) reported about a complication during impella implantation, which required the groin to be surgically opened. the cardiologists subsequently sutured the impella, so it could no longer be repositioned. additionally, the impella had a red alarm "impella in aorta." transthoracic echocardiography was not possible, only transesophageal echocardiography. however, there was a concern for continuing support to a generalized bleeding problem (not impella-related). the situation was discussed in detail with the md, and they were advised to verify the impella position and correct it if necessary. the md was instructed on how to mute the position alarm and wanted to discuss further the situation with the senior physician. the decision was made to escalate the patient's therapy from an impella cp to an impella 5.5 due to complications with the impella cp and need for higher support. however, after a few hours on the impella 5.5, the patient passed away. | [5 empty] | Additional information was obtained from the physician. the patient was in very critical condition and the prognosis was bad. the patient expired due to multi organ failure. | No product was returned for investigation. the cause of the bleed was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | Patient information was not provided. therefore, section a was left blank. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety/clinical reviewed the event. a patient of unknown age and sex presented in cardiogenic shock (scai stage unknown) secondary to an acute myocardial infarction and was implanted with an impella cp device for mechanical circulatory support. on post-implant day 1, a complication involving the impella was reported. the groin access site required surgical reopening, and the physician sutured the impella in place, after which the device could no longer be repositioned. the device subsequently displayed a red alarm indicating ¿impella in aorta.¿ transthoracic echocardiography (tte) was not feasible; only transesophageal echocardiography (tee) could be performed. however, due to a generalized bleeding condition (reported as not related to the impella), the physician preferred to avoid tee. the physician was advised to verify device position and correct it if necessary. the physician inquired about muting the position alarm and indicated intent to discuss the case with a senior physician. a subsequent update indicated that, due to complications with the impella cp and the need for a higher level of support, the patient was escalated to an impella 5.5 device. approximately seven hours later, the patient expired. the physician reported that the patient had been in critical condition with a poor prognosis. the cause of death was reported as multi-organ failure and was not attributed to the impella device. note: h6. component code and investigation type/findings/conclusion coding remain unchanged as previously reported.
3864 | 1220648-2025-47913 | 23413007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413007&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705974 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon insertion of the impella 5.5, an impella position unknown alarm appeared on the automated impella controller. the physician opted to prime and place an additional impella 5.5. the patient later expired on the other impella 5.5 pump which was reported on another report. refer to section h.10 for the related report number. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the delivery issue was unable to be determined as insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3865 | 1220648-2025-47903 | 23413059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413059&pc=OZD | 10/01/2025 |  |  | 10/29/2025 |  |  | 10/01/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683567 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient had triple cardiac bypass grafting procedure, and when separating from the cardiac bypass, the patient¿s requirements for inotrope and vasopressor significantly increased. to hemodynamically support the patient, an impella 5.5 was implanted. on the second day of impella support, it was reported that the pump had alarms for high purge pressure and purge system blocked. the medical team decided to switch the purge solution to tissue plasminogen activator, which resolved the alarms. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: the product was not returned. high purge pressure/purge system blocked: as per clinical purge pressure high alarms reported on 10/01/25. all other troubleshooting info are unknown. the data logs were reviewed and several occurrences of high purge pressure/purge system blocked alarms on 10/1 & 10/02 were found. logs indicated gradual rise of purge pressure for around 28 mins from 600 to 1100 mmhg. no mc spikes observed, but mc variation with rise in pump flows seen. pp gradually decrease from 1100 to 800 mmhg for over ~12 hours and continuously purging at 800 mmhg pressure and purge flows ~3 m/hr (from < 1 ml/hr) for next ~31hours. this trend seems to be biomaterial buildup trend. the cause of high purge pressure/purge system blocked issue was due to biomaterial buildup based on log analysis. device history lot: device lot: 1929996 device history review: the pump sn (b)(6) passed all post sterile inspection checks.
3866 | 1220648-2025-47911 | 23413116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413116&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720048 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 decompensated and expired. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the device passed all post sterile inspection checks.
3867 | 1220648-2025-47912 | 23413149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413149&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683589 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 81 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp a patient experienced bleeding at the access site. perclose stitches were placed and angioseal was used to achieve hemostasis. following insertion, a mobile left ventricular apical thrombosis was observed under fluoroscopy. pump implantation was aborted and the patient survived. | [5 empty] | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3868 | 1220648-2025-47909 | 23413163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413163&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720798 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female | 56 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was admitted for drug induced ventricular fibrillation arrest and received cardiopulmonary resuscitation in the field. they became increasingly hypotensive in the intensive care unit, with increasing pressor and inotropic requirements. patient was placed on impella cp for mechanical circulatory support. less than 3 hours after impella implant, there were no pulses noted via doppler signal in either lower extremity. patient was intermittently going into ventricular tachycardia. amiodarone and bicarbonate pushes. left ventricular diastolic becoming increasingly negative. care was withdrawn and the patient expired. | [5 empty] | H6 added b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation was completed and no product was returned for investigation. ischemia/ tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3869 | 1220648-2025-47914 | 23413168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413168&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was prepared for impella support due to acute myocardial infarction/cardiogenic shock. while attempting to implant the impella, the physician was unable to gain access to either femoral site due to both being occluded. the impella cp placement was aborted due to vascular anatomy. | [5 empty] | H6 revised health effect - impact code since the initial report was submitted. investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and both femoral sites were occluded preventing successful delivery of impella. device history review: the product history review was not possible since the serial number of the device is unknown. | D.4. lot number, serial number, expiration date and completed udi are unknown. h.4. device manufacture date is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
3870 | 1220648-2025-47915 | 23413385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413385&pc=OZD | 09/30/2025 |  |  | 10/29/2025 |  |  | 09/30/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693813 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the access site. pressure was held and the site was redressed. additionally, hematuria was noted so volume was given. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the root cause of the bleeding was determined to be use issue due to improper anticoagulation management. the device passed all post sterile inspection checks.
3871 | 1220648-2025-47916 | 23413409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413409&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698975 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the pump support was initiated, there were no wave forms visible on the automated impella controller for the pump. the pump¿s position was verified and it was noted that the impella flows were 0.5 liters per minute and suction alarms were present at all p levels. as a result of the issues, the pump was explanted and replaced with a new impella cp and cannulated for extracorporeal membrane oxygenation. no patient harm was reported in relation to this event. follow-up efforts were unsuccessful in obtaining additional information, and the final patient outcome is unknown. | [5 empty] | Investigation summary impella cp sn (b)(6) returned pump suction/low pump flow: clinical reported that upon starting the pump suction was triggering at all p-levels. there were also no waveforms observed on the aic, and the pump was getting low flows. the product was returned for investigation and biomaterial was observed inside the cannula. data logs were not returned, so it cannot be confirmed whether suction alarms were accurately triggering during the case. the pump was run in water, and suction alarms did not reproduce. flows were within spec, achieving 3.6l/min at p-9. since data logs were not returned and flow issues did not reproduce during testing, the cause of the suction and reduced flows could not be determined. optical sensor issue: clinical reported that waveforms were not showing up on the console upon starting the pump. data logs were not returned so the lack of waveforms during the case cannot be confirmed. the product was returned for investigation and biomaterial was observed inside the cannula. the pumps optical bench parameters where within specifications and it passed light continuity testing. since data logs were not returned and the pump passed optical signal testing, the cause of the optical signal issue could not be determined. device history lot device lot number: 1943294 device history batch subcomponent lot: n/a device history review pump sn (b)(6) passed all post sterile inspection criteria. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D9 has been updated to reflect product was returned for investigation. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
3872 | 1220648-2025-47917 | 23413444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413444&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1237409 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During training, the console presented a "battery failure" alarm. there was no patient involvement. | [5 empty] | E.1 initial reporter first and last name are unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3873 | 1220648-2025-47918 | 23413452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413452&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740190 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 80 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia. cardiopulmonary resuscitation was provided until return of spontaneous circulation was achieved. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the cardiac arrest was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3874 | 1220648-2025-47919 | 23413669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413669&pc=OZD | 10/03/2025 |  |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736324 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 78 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support had a positive blood culture and was started on antibiotics. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the infection was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3875 | 1220648-2025-47920 | 23413778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413778&pc=OZD | 10/03/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728551 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 63 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had placement signal not reliable alarms. placement signal monitoring was disabled. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the optical sensor issue was not established as no product or logs were returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3876 | 1220648-2025-47921 | 23413798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413798&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733794 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 68 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was not able to be implanted due to the patient's vascular anatomy. the impella 5.5 placement was aborted. | [5 empty] | The impella device was discarded by the operating room staff and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and resistance at innominate artery preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
3877 | 1220648-2025-47922 | 23413878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23413878&pc=OZD | 10/03/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026733480 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Flutter | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia for which the patient was cardioverted and amiodarone was given. the patient also experienced hypotension. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the atrial flutter was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3878 | 1220648-2025-47923 | 23414230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23414230&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589758 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 53 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Low Blood Pressure/ Hypotension | Pulmonary Edema | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after a bypass, the impella 5.5 was placed via direct approach. in the intensive care unit, the patient was defibrillated three times. the patient also had gastrointestinal bleed requiring bivalirudin to be paused. the patient was cannulated for right ventricular assist device with protek duo, but impella flows were kept at p-3 so the patient experienced flash pulmonary edema with severe hypotension with lactic climber to 15. since then, impella flows were increased to p-8. the patient experienced atrial fibrillation with rapid ventricular response. the patient was cardioverted and lidocaine was given before gaining normal sinus rhythm back. it was further reported that the family decided to withdraw care and the patient expired. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned for investigation. paroxysmal atrial fibrillation: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pulmonary edema/ hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3879 | 1220648-2025-31892 | 23414657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23414657&pc=OZD | 07/23/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 07/23/2025 | 10/29/2025 |  |  | 12/31/2024 |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024291833 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the device exhibited placement signal issues. it was also reported that care was withdrawn and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no product was returned for evaluation. data logs were returned and impella position in aorta were observed while pump was running on p2. there was no variation or impact to optical signal 5s parameters. the cause of the optical signal issue was most likely patient condition related based on the data log review. all pertinent information available to abiomed has been submitted at this time.
3880 | 1220648-2025-31893 | 23414754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23414754&pc=OZD | 03/14/2023 | 12/01/2025 |  | 10/29/2025 |  |  | 03/14/2023 | 10/29/2025 |  |  | 12/31/2024 |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024307509 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported loss of placement signal during use of the device. use of the device continued via monitoring motor current and flows. | [5 empty] | H6: investigation findings, corrected the code based on the results of the investigation. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the 5s pattern in the logs during the time that the ps spikes to 3000 is consistent with a damaged fiber line. a visual inspection of the pump revealed a bond issue at the catheter to kink protector bond. the cause of the optical signal issue was a damaged optical fiber line in the red handle caused by the bond issue at the catheter to kink protector bond. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3881 | 1220648-2025-47924 | 23414829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23414829&pc=PYX | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  | 10/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413695 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during bipella support of an impella cp and impella rp flex the patient had hemolysis/hematuria. a chest x-ray revealed the impella rp flex was in too deep, and recommended repositioning (pull-back). the physician agreed and repositioned the impella rp flex under fluoroscopy and continued with bipella support. later, the impella cp and impella rp flex were removed and an impella 5.5 was used. this report is for the impella rp flex and another report will be sent for the impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6. medical device problem code a01 was not included in the initial mdr and is now correctly added. | The investigation of the reported failure mode has been completed. no product was received for evaluation. device in wrong position (positioning issue): the cause of positioning issue cannot be determined since insufficient clinical details were provided. hemolysis: after review of logs, aside from couple of suction alarms at the start of case, no other were observed through out the runtime. no motor current spike or trend in placement signal were observed. the cause of hemolysis was most likely use issue related due to improper positioning since it was successfully managed by repositioning. device history review: the device passed all post sterile inspection checks.
3882 | 1220648-2025-31894 | 23414910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23414910&pc=OZD | 02/14/2023 | 10/29/2025 |  | 10/29/2025 |  |  | 02/14/2023 | 10/29/2025 |  |  | 11/30/2024 |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023275784 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported console emitting an alarm during setup of the device. a new controller was used, however, the alarm persisted. a new pump was opened and used without issue. no harm to any patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs were not returned and the pump was. the root cause of the optical signal issue is due to a damaged fiber line. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3883 | 1220648-2025-47925 | 23415783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23415783&pc=OZD | 09/29/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691229 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  | White | Death | Yes | Yes | Computer Operating System Problem | [6 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was placed onto impella support for acute myocardial infarction/cardiogenic shock. while on support, the automated impella controller (aic) presented an "input/output error" upon startup. the aic was changed before the pump was implanted. the patient was on support for less than an hour. care was withdrawn and the patient expired. there were reportedly no product problems with the impella pump nor adverse events that occurred during support. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The product with event logs were received and the investigation was completed. product history review summary: there were no issues related to the complaint discovered when reviewing the product history of console. console logs were consistent with what was reported. device analysis: the console was tested after arriving in field service. the reported input/output error was able to be reproduced. troubleshooting by swapping out the compact flash cards on the consoles 4-in-1 resolved the issue. as a precaution, field service was asked to power cycle the console 10 times after troubleshooting with known good compact flash cards. no issues observed after swapping out the compact flash cards with known good ones and power cycling the console 10 times. conclusion: the root cause of the reported start up issue was able to be isolated to the compact flash cards during troubleshooting. likely due to corruption of the compact flash cards. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3884 | 1220648-2025-47926 | 23415788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23415788&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707223 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella 5.5 the patient remains febrile. also at time of the explanation for impella 5.5, while pulling pump back there was noted ¿massive amount¿ of blood coming back into the graft. at that time the patient began to decompensate, and the decision was made to emergently bring patient to operating room (or) for chest exploration. once in or chest was reopened, and bleeding was noted at the graft site attached to aorta. hemostasis was achieved with surgical intervention. the patient was stabilized and taken back to intensive care unit. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the access site bleeding issue was not determined without returned product and sufficient clinical details. the cause of the removal issue was not determined without returned product and sufficient clinical information. the cause of the fever was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3885 | 1220648-2025-47929 | 23416098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416098&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728948 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 111 KG |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Anemia | Arrhythmia | Pain | Pericardial Effusion | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk biobental hemiarch procedure was implanted with an impella 5.5 device for mechanical circulatory support. on support day 13, the purge flow alarmed for purge pressure high and purge system blocked. it was noted that the purge flow had decreased over the last two days and was now <2 milliliters per hour. tissue plasminogen activator was added to the purge solution to resolve the issue. on support day 16, the patient experienced an episode of atrial fibrillation with rapid ventricular response. no intervention for the arrhythmia was documented or reported. on support day 20, the patient had noted pericardial effusion, and the pericardial chest tube which was placed previously was pulled slightly back to address the effusion. additionally, the patient was transfused with one unit of packed red blood cells for a low hemoglobin of 7.6 grams per deciliter. the following day it was noted that the patient¿s pain had improved and the patient no longer required oxycodone. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complications. the patient's outcome at the time of explant was survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3886 | 1220648-2025-47930 | 23416139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416139&pc=OZD | 10/02/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761522 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp the patient had lost pulses at the limb post procedure, no flow to right leg. decision was made to take the operating room (or) for 5.5 escalation. cp explanted and 5.5 implanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d3, e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3887 | 1220648-2025-47931 | 23416172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416172&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749920 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support there was a plan of care to gain hemostasis and resolve a hematoma in the groin and scrotum. a mattress suture was placed and the team is aware of the situation and was considering impella 5.5 escalation. blood products were given to the patient. the day after the patient was doing well. the pump was explanted four days later. the patient survived. | [5 empty] | Investigation summary: the investigation has been completed. hematoma: the cause of the injury was most likely use related since the surgeon had placed a mattress suture and pressure to achieve hemostasis. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3888 | 1220648-2025-31895 | 23416283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416283&pc=PYX | 05/16/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 06/16/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625168 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex experienced a drop in purge flow and an increase in purge pressure. correct positioning of the pump was confirmed, and the patient was treated with alteplase. later patient expired. | [5 empty] | A4 is unknown. 'high purge pressure' failure mode added due to purge issue complaint by customer. no product or images were returned for analysis. no data logs were returned for analysis. packaging issue: no product or images were returned for analysis. clinical mention 0.027" guidewire was missing from packaging. the root cause of the packaging issue was not determined as no product or images were returned for analysis. high purge pressure: no product or data logs were returned for analysis. the root cause of the high purge pressure was not determined as no product or data logs were returned for analysis.
3889 | 1220648-2025-47932 | 23416303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416303&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710588 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 73 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the pump was aborted due to patient vascular anatomy. it was unable to pass through the vasculature. there was no patient harm. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and resistance at axillary artery preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3890 | 1220648-2025-47927 | 23416348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416348&pc=OZD | 10/02/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026756624 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient had a right groin hematoma. it was also said that hemolysis-pink/red urine resolved when the impella cp changed to an impella 5.5. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. access site bleeding/hematoma: the cause of access site bleeding/hematoma was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3891 | 1220648-2025-47933 | 23416350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416350&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750271 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during bipella support of impella cp the patient had hemolysis/hematuria noted. impella cp repositioned. this morning, the patient became hypotensive, heparin gtt changed to bival gtt due to concern of possible heparin-induced thrombocytopenia. both the rp and the cp was removed and the patient was put on impella 5.5 support. this report is for the cp and another report will be sent for the rp (serial number: (b)(6). | Additional information was received. the plan was to continue with current therapies. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 additional information was received. h6 added code e0303 and b13 as they were omitted from the initial manufacturer device report number. h10 added related report number. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ investigation summary: the investigation has been completed. clinical symptoms (thrombocytopenia, hypotension): the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. mechanical interaction with blood(hemolysis): the pump was not returned for investigation. data log was only provided until (b)(6) 2025, 1:33 am. there was a sudden drop in motor current, pump flow, and placement signal noted on (b)(6) 2025 at 12:20 am, with an active suction alarm. there was no motor current spike or positioning alarm recorded on the data log. according to the clinical details (b)(6) 2025, 10:54 am, both the impella cp and rp pumps were repositioned, and the doctor noted resolution after that. according to the impella rp investigation, no device-related abnormalities were confirmed during the time of hemolysis. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical information. device history review: the pump passed all post sterile inspection checks.
3892 | 1220648-2025-47935 | 23416352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416352&pc=OZD | 10/02/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734935 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 94 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported post impella 5.5 placement the patient continued to decompensate with runs of ventricular tachycardia and ventricular fibrillation and multiple shocks. a decision was made to have va ecmo cannulation. the impella was explanted as planned. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia/ ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3893 | 1220648-2025-47939 | 23416569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416569&pc=OZD | 10/01/2025 |  |  | 10/29/2025 |  |  | 10/01/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721358 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, it was recommended to have an echocardiogram done on the patient to check placement considering one had not been done for a few days. the staff agreed and an echocardiogram was ordered. white blood count was elevated so the patient was started on antibiotics and cultures were sent. the day after the patient had ventricular tachycardia episode early that morning. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Investigation summary: the investigation has been completed. no product was returned for investigation. tachycardia/infection: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3894 | 1220648-2025-47938 | 23416639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416639&pc=OZD | 10/02/2025 |  |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1708923 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 70 KG | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a purge disc not recognized error when changing purge cassette. the aic console was changed. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report number.
3895 | 1220648-2025-31896 | 23416734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416734&pc=OZD | 04/28/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 09/05/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611450 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Priming Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 experienced ventricular tachycardia (vt) overnight and rosc obtained after 1 round of cpr and 1 lido administered. the next morning the patient went into vt again and 1 lido was given then subsequently started on a lido infusion. the impella p-level was raised post vt. several days later the patient was successfully transplanted. | [5 empty] | Purge cassette failure: data logs were reviewed and the "purge system failure" alarm was confirmed. during case start, the purge cassette was inserted and successfully recognized by the console. then, when attempting to prime the purge cassette, the piston strokes were not able to reach full length. for this reason, "purge system failure" triggered, indicating that the piston was blocked with only 18% of its range remaining, per imc logs. returned product was investigated, and again, there was no issue with console recognition of the purge cassette being inserted. based on data log review and investigation of returned product showing no issue with purge cassette recognition, it was determined that there was no problem found. priming issue: as mentioned above, data log were review confirmed the alarm triggered due to piston blockage. when the purge cassette was replaced with a new one, it successfully primed on the first attempt. the returned purge cassette was successfully de-aired and primed a pump. however, shortly after, the alarm reproduced. in reviewing these data logs, the strokes of the piston gradually became shorter until the range was small enough to reproduce the alarm. the purge cassette was torn down, and the teeth of the driveshaft showed clear damage and/or wear, which resulted in the driveshaft/pinion not being able to advance the rack/cylinder. based on data log review and the wear noted on the returned cassette's driveshaft teeth, the root cause of the priming issue was determined to be a piston failure. vt/vf: it was reported that the patient had several rounds of vt on (b)(6) 2025. it was treated with cpr and a lido gtt. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
3896 | 1220648-2025-47934 | 23416772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416772&pc=OZD | 10/02/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725945 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 80 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, the patient had weak gag, absent pedal pulses, and no impella issues. the patient was given albumin and weaning was tried. no other information was available. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3897 | 1220648-2025-47936 | 23416842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416842&pc=OZD | 10/01/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/02/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1470310 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella optical sensor was not recognized even when the pump was connected to the automated impella controller (aic). the aic was replaced and the issue was resolved. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
3898 | 1220648-2025-47937 | 23416998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23416998&pc=OZD | 10/01/2025 | 12/16/2025 |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574195 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was escalated from an impella cp to an impella 5.5. nine days after escalation to an impella 5.5, a head computed tomography scan revealed evidence of a cerebral infarction. systemic heparin was discontinued for a period due to a bleeding tendency. after consultation with the patient's family, care policy (do not attempt resuscitation) was adopted. | [5 empty] | The investiigation for the reported stroke has been complete. the device was not returned for ecvaluation. the root cause of the stroke could not be deteremined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
3899 | 1220648-2025-31897 | 23417017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23417017&pc=OZD | 04/09/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 04/09/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2025595402 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 72 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  | No | Initial submission | [9 empty] | The complainant reported that there was a potential pump stop issue. the pump was returned for investigation along with the automated impella controller (aic) for an issue concerning "verification failed, exchange the console" where the aic was unable to be switched on again. | [5 empty] | Md reported about an aic failure "verification failed". no pump was returned. no data logs returned. the cause of the issue was not determined as no product or logs were returned for analysis. for an investigation into the console please refer to (b)(4).
3900 | 1220648-2025-47940 | 23420134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23420134&pc=OZD | 09/29/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026735093 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Cardiogenic Shock | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, patient was placed onto veno-arterial extracorporeal membrane oxygenation due to worsening shock. patient experienced atrial tachycardia; amio bolus given and drip started. patient was oliguric and placeed onto continuous renal replacement therapy anticoagulation was withheld due to bleeding concerns (also related to menstrual cycle). oozing noted at groin site. 1 uncross-matched packed red blood cells going to be given which is a total of ten units since admission. occasional ectopy appreciated on monitor; impella pulled back 1cm. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleed cannot be determined since insufficient clinical details were provided. arrhythmia, tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiogenic shock, renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3901 | 1220648-2025-47942 | 23421076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23421076&pc=OZD | 09/29/2025 |  |  | 10/29/2025 |  |  | 09/29/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686699 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 106 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support after experiencing a cardiac arrest. the patient was supported by impella cp during a percutaneous coronary intervention to the right coronary and left anterior descending coronary arteries. on support day 4, the patient experienced a gastrointestinal bleed which heparin was withheld to manage the bleeding. the following day the patient was transfused with one unit of red blood cells, and it was noted that the patient¿s stool remained black and tarry. it was noted that esophagogastroduodenoscopy was performed, and a gastric ulcer was confirmed as the source of the bleed. additionally, the patient experienced premature ventricular contractions and received a bolus and continuous intravenous infusion of amiodarone to resolve the issue. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. | [5 empty] | Investigation summary: major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot: device lot: 1931506. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3902 | 1220648-2025-47943 | 23421077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23421077&pc=OZD | 10/03/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026730253 |  | 000000000001000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 65 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after impella 5.5 was placed and the patient was taken to intensive care unit, the impella had sensor drift and erratic placement signal waveforms but motor current was stable. impella placement was confirmed. daily echoes are being conducted to confirm placement due to the placement signal unreliable alarms. the patient is still on impella support waiting on a heart transplant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3903 | 1220648-2025-47944 | 23421078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23421078&pc=OZD | 10/03/2025 | 10/29/2025 |  | 10/29/2025 |  |  | 10/03/2025 | 10/29/2025 |  |  |  |  | 10/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727178 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 39 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp the patient had hematoma at the access site. during transport from chronic lymphocytic leukemia (ccl) to the cardiac intensive care unit he developed a hematoma. aware of the hematoma, will pull the pump if it does not improve. while the registered nurse was pressing on the site, he vagaled and his pressure dropped. fluids were given and his hr improved from 45 to 76. due to the hematoma, it may lead to the pump being removed soon. the patient was taken down to the ccl to remove the impella. the risk of the hematoma growing was greater than keeping him on proactive support overnight for continued lv unloading. dr. sadeghi used one perclose to close the access site. | [5 empty] | Correction e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks. device history batch ==> sub-component lot : n/a device history review ==> the pump s/n: 616756 passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3904 | 1220648-2025-47865 | 23401243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23401243&pc=OZD | 09/28/2025 |  |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716193 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support after being admitted to the hospital for acute on chronic heart failure and diuresis as a bridge to transplant. on the twenty-fourth day of support, it was reported that the patient tested positive for a heparin antibody, and as a result heparin was discontinued and the patient was switched to angiomax for anticoagulation therapy. on the twenty-eighth day of support, it was reported that the patient experienced ectopy. the impella 5.5 was imaged with an echocardiogram and confirmed to be well positioned. the patient¿s bedside nurse noted that the patient had an electrolyte imbalance which likely contributed to the patient¿s ectopy. the patient¿s electrolytes were repleted and the ectopy resolved. the patient remains on impella 5.5 support and is awaiting a heart/kidney transplant. | Additional information was received indicating that the impella device was explanted. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Correction: b1 the investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue - no abnormalities were found to the 5s parameters on the data logs that were recovered. the cause of the optical issue was not determined as no product was returned, insufficient data logs were recovered, and insufficient clinical details. thrombocytopenia: the cause of the was not determined due to insufficient clinical details. arrhythmia: the cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | Additional information received for d6 explant. section d catalog number was corrected.
3905 | 1220648-2025-47853 | 23401353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23401353&pc=OZD | 09/29/2025 | 12/12/2025 |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738764 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was febrile during impella cp support. | [5 empty] | The investigation for the reported fever has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the fever could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3906 | 1220648-2025-47866 | 23401445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23401445&pc=OZD | 09/28/2025 |  |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721163 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced a gastrointestinal bleed. anticoagulant was held. the patient is still on impella support awaiting to be bridged to a left ventricular assist device. | [5 empty] | D6b explant date was added. h4 revised device manufacture date as it was entered incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3907 | 1220648-2025-47867 | 23401685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23401685&pc=OZD | 10/10/2025 |  |  | 10/28/2025 |  |  | 10/10/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749245 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the support they encountered positioning issues. the impella was in the aorta. the doctor planned on discussing with the senior physician whether the impella should be removed or replaced in left ventricle again. p-level reduction to p2, disconnected & console turned off. the impella was removed shortly afterwards patient is stable post explant. | [5 empty] | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3908 | 1220648-2025-47868 | 23401703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23401703&pc=OZD | 09/26/2025 |  |  | 10/28/2025 |  |  | 10/17/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720042 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 tested positive for drug-resistant bacteria, resulting in a total device replacement (intra-aortic balloon pump). no patient harm was reported. | [5 empty] | A2, a4, and a5 are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The cause of the infection was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | Health effect impact code added.
3909 | 1220648-2025-47869 | 23402070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23402070&pc=OZD | 10/14/2025 | 01/29/2026 |  | 10/28/2025 |  |  | 10/15/2025 | 10/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299285 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [36 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon implant of an impella cp no position waveform was displayed and an automated impella controller (aic) alarm occurred. the aic was replaced to continue patient support. the patient expired on impella support. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety reviewed the event and noted the following: upon initiation of impella cp support, ultimate controller errors occurred. an aic transfer was successfully completed without interruption in support. the patient expired four days later; however, the device malfunction was not a contributing factor to the death, which was likely due to the patient¿s clinical condition. the event will be conservatively reported as a device malfunction because of the nature of the alarm. change in reportability. after review of the case by medical safety, it was determined the event is a malfunction type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical/impact codes, medical device problem code and component code) have been updated/corrected fully repopulated accordingly. device status. further additional information noted the automated impella controller device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly.
3910 | 1220648-2025-47870 | 23402666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23402666&pc=OZD | 09/30/2025 |  |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730438 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 116 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Cardiac Tamponade | Renal Impairment | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to postcardiotomy cardiogenic shock / low cardiac output syndrome. while on support, chest xray showed suspected bleeding from lungs, so systemic anticoagulation was withheld. profound hypotension was unresponsive to methylene blue, calcium chloride, intravenous pyelogram, fluid resuscitation in addition to 5 high dose vasoactive drops. patient was placed onto continuous renal replacement therapy for acute kidney injury. a chest washout was performed for tamponade. patient remains off heparin due to open chest and surgical bleeding risk. patient received 6 units packed red blood cells and platelets for gastrointestinal (gi) bleeding and eventually had gi intervention. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. hypotension / cardiac tamponade/ acute kidney failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3911 | 1220648-2025-47871 | 23403143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23403143&pc=OZD | 09/28/2025 |  |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711550 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 84 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated from an impella cp to an impella 5.5. while on impella 5.5 support, the patient experienced gastrointestinal bleeding and oozing from the new dialysis line. bivalirudin was stopped. ten units of packed red blood cells were administered and a bronchoscopy was performed. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The impella device wis not available for return. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks.
3912 | 1220648-2025-47872 | 23403178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23403178&pc=OZD | 09/29/2025 |  |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680052 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 58 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Cardiac Perforation | Pericardial Effusion | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented to the emergency department with complaints of chest pains that had been occurring for two weeks and was in st elevation myocardial infarction. an impella cp was successfully placed, and the patient had a percutaneous coronary intervention that addressed the left main and left anterior descending coronary arteries with two drug-eluting stents placed. the patient was then transferred to another hospital for advanced care. on the third day of impella support, the patient quickly decompensated and coded twice. return of spontaneous circulation was achieved and ultrasound imaging revealed a posterior pericardial effusion. it was also noted that the patient had a previous anterior apex aneurysm. the physician suspected that the patient had a posterior left ventricular perforation. it was noted that the impella cp was positioned towards the posterior wall. the patient was noted to be stable, and received four units of red blood cells, three units of platelets, and four units of fresh frozen plasma. the patient¿s family decided to withdraw care. the patient¿s care was withdrawn, and the patient expired. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. no product was returned for investigation. cardiac arrest: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac perforation/ pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | . although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3913 | 1220648-2025-47873 | 23403406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23403406&pc=OZD | 09/30/2025 |  |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722083 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 109 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. while on support, the patient experienced internal bleeding suspected to be gastrointestinal. the patient coded and expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. cardiac arrest: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3914 | 1220648-2025-47856 | 23403659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23403659&pc=OZD | 10/09/2025 | 10/27/2025 |  | 10/28/2025 |  |  | 10/09/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655088 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male | 57 KG | Non Hispanic | Asian | Malfunction | No | Yes | Difficult or Delayed Positioning | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while adjusting the position of an implanted impella 5.5, the pump was completely removed into the aorta. the pump was exchanged for a new impella 5.5 to continue patient support. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the user as resistance was felt causing the 5.5 pump to be completely removed into the aorta. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3915 | 1220648-2025-47874 | 23403678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23403678&pc=OZD | 10/09/2025 |  |  | 10/28/2025 |  |  | 10/09/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1411749 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during preparing the impella purge system for implant, the automated impella controller (aic) showed a "purge not detected¿ alarm. the staff tried to reposition the purge system but the alarm was still there. they decided to change the purge system and still had the issues. then they changed the aic and everything worked without any issues. | [5 empty] | The device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. console logs were not consistent with what was reported in the complaint. no evident purge detection issues were observed in the imc logs. the reported purge detection issue could not be reproduced. however, it was observed that the purge stepper motor was stuck. no problem found in relation to the purge detection issue. d4 product information was updated to correct product. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Additional information received. d9 device was received for investigation. once investigation results are received a supplemental report will be submitted.
3916 | 1220648-2025-47875 | 23404219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23404219&pc=OZD | 09/30/2025 |  |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707237 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to post-cardiotomy cardiogenic shock. while on support, patient experienced hematuria. the next day, patient presented with an unspecified arrhythmia as well as ventricular premature complexes, for which bolus of lidocaine, amio, and drip were administered. the following day, patient had a 250ml chest dump of old blood when he stood up. heparin was stopped and he was transfused with one packed red blood cells and one platelet. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. arrhythmia : the cause of the injury was unable to be determined due to insufficient clinical details. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3917 | 1220648-2025-47876 | 23404763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23404763&pc=OZD | 09/30/2025 | 10/28/2025 |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654167 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 51 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hypovolemia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, drainage was noted and patient was taken to operating room for washout. the patient was hypovolemic and received two units of packed red blood cells, 1 platelet, 1 fresh frozen plasma and 2cryo as well as albumin. bival has been off use due to bleeding issues. the patient had hemorrhagic stroke on computed tomography scan, decision made to remove impella to eliminate need for anticoagulants. during removal, impella was noted to have had a clot and it appears that clots showered throughout the body during the removal. patient now showing multiple strokes, ischemic mesenteric, splenic and renal infarcts. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary the device was not returned for evaluation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. thrombosis / stroke/ ischemia: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypovolaemia: the cause of the injury was not determined due to insufficient clinical details. device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
3918 | 1220648-2025-47877 | 23405868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23405868&pc=OZD | 09/28/2025 |  |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026732548 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 support developed thrombocytopenia and there was concern of a gastrointestinal bleed. systemic anticoagulation was suspended. impella support continues. | [5 empty] | The cause of the thrombocytopenia was not determined due to insufficient clinical details. the cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
3919 | 1220648-2025-47878 | 23405935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23405935&pc=OZD | 09/28/2025 |  |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684255 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed hemolysis and hematuria, so the p-level was lowered. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The pump was not returned for review. data logs confirmed the complaint. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the pump passed all post sterile inspection checks.
3920 | 1220648-2025-47879 | 23406100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23406100&pc=OZD | 09/28/2025 | 10/28/2025 |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026730434 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 67 KG | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella cp experienced multiple pump suction alarms. blood volume was given to the patient. additionally, levophed medication along with further blood volume was given due to hemodynamic instability. ultimately, the impella cp was explanted and was replaced with another impella cp. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. pump suction: data log review confirms suction alarms occurring towards the beginning of the case. the cause of the suction was most likely related to patient condition, as placement signal was trending down during the suction events, and resolved after volume was given to the patient. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. these one of two related reports. this report represents the pump. another report will be submitted for the automated impella controller. refer to section h.10 for the related report number.
3921 | 1220648-2025-47880 | 23406163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23406163&pc=OZD | 09/28/2025 | 10/28/2025 |  | 10/28/2025 |  |  | 09/28/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108710 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 67 KG | Non Hispanic | Asian | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) experienced an on going suction alarm. | [5 empty] | D9 the device was returned and the date received was added. h6 codes were updated to reflect the device being returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report number.
3922 | 1220648-2025-47882 | 23407259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407259&pc=OZD | 09/29/2025 | 10/28/2025 |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713900 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Atrial Fibrillation | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient coronary artery bypass graft and then impella 5.5 was placed via direct aorta. in intensive care unit, there was bleeding from the mediastinal chest tube, 400ml out in first 40 minutes. blood products were required, and heparin was held. additionally, the impella had suction and albumin was given. on day three of 5.5 support transthoracic echocardiography showed impella close to septum. the physician attempted to reposition the impella but was unsuccessful. the patient also experienced atrial fibrillation with rapid ventricular response and amio was given. it was further reported that the patient ldh was elevated to 1438. the physician elected to take patient for explant in setting of likely hemolysis and heart recovery. the patient survived. | [5 empty] | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary the device was returned for evaluation. data was not returned for review. pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. pump suction: the cause of suction issue was unable to be determined as limited clinical details were provided and no problem was found on the pump. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no problem was found on the pump. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely patient condition related since there was no change on urine color & the clinical team confirmed ldh increasing was due to volume overloaded and liver congestion. major bleed: the cause of major bleed was unable to be determined as limited clinical details were provided and no problem was found on the pump. paroxysmal atrial fibrillation: in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review ==> pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3923 | 1220648-2025-47884 | 23407271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407271&pc=OZD | 09/29/2025 | 12/16/2025 |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727505 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 65 KG | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Multiple Organ Dysfunction Syndrome | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that within two hours of impella cp implant, the patient developed junctional rhythm rate of 60 with frequent paroxysmal supraventricular tachycardia (psvt) at a rate of 140. additional medication was given and adjustments were made. the pump's position became slightly shallow but cardiology declined repositioning as there were no issues. suspect cardiorenal syndrome with alanine aminotransferase and aspartate aminotransferase greatly increased overnight. developed ileus and right groin ooze requiring redressing. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported renal failure/ organ failure/ minor bleed, positioning issue, and tachycardia has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of the renal failure/ organ failure/ minor bleed, positioning issue, and tachycardia could not be determined.
3924 | 1220648-2025-47885 | 23407272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407272&pc=OZD | 09/29/2025 | 02/18/2026 |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657845 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 112 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported additional information: "no pump is available for return unfortunately." | The user facility reported that the decision for impella cp was made when a patient was requiring multiple pressors post operation transesophageal echocardiogram and cardioversion. the cp was placed without difficulty. during support it the patient had marginal flows on the impella. it was noted that the pump migrated into left ventricular outflow tract. the patient was without post tibial pulses despite reperfusion catheter. the reperfusion catheter was clotted off. the physician re-flushed, and the patient had post tib pulses. repositioning was attempted under transesophageal echocardiogram, pigtail was entrapped in cordae. plan was to reposition under fluoro. it was reported that there were no pulses to right lower extremity despite reperfusion cannula. unable to obtain adequate flows due to malposition of the pump, repositioned under fluro. the right lower extremity vessels were angio¿d in and patient was with adequate flow and pulses dopplered. the patient significantly improved and the cp was successfully weaned and explanted. |  |  |  |  | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia, thrombosis: the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. low pump flow: after review of logs, multiple suction alarms were observed. no motor current spike or trend in placement signal was observed. the cause of low pump flow was most likely use issue related due to incorrect positioning of the pump per clinical but unknown why it was out of position. device history review: the device passed all post sterile inspection checks. | B5 updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3925 | 1220648-2025-47886 | 23407293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407293&pc=OZD | 09/29/2025 |  |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost distal flows in the left lower extremity, which was confirmed using doppler. additionally, the patient developed ectopy in response to milrinone so dobutamine was given instead along with vasopressers. the patient also exhibited hemolysis. the patient elected to receive comfort care so care was withdrawn and the patient expired. | [5 empty] | The cause of the cardiac arrhythmia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3926 | 1220648-2025-47887 | 23407295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407295&pc=OZD | 09/29/2025 |  |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700164 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 78 KG |  |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced bleeding. multiple blood products were transfused without resolution. later, care was withdrawn and the patient expired. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3927 | 1220648-2025-47888 | 23407415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407415&pc=OZD | 09/29/2025 |  |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743331 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced purge pressure issues. the patient requested care to be withdrawn and the patient expired. | [5 empty] | The cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3928 | 1220648-2025-47890 | 23407508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407508&pc=OZD | 09/30/2025 | 02/18/2026 |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674001 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experience bleeding at the insertion site. the dressing was changed, sandbags were used, and heparin was withheld.later, the patient was successfully weaned from the pump and survived. | The complainant reported additional information upon request: "no pump has been thrown away." |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
3929 | 1220648-2025-47883 | 23407513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407513&pc=OZD | 09/29/2025 | 10/28/2025 |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025577878 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient whom experienced an arrhythmia going into atrial fibrillation with rapid ventricular response while on impella 5.5 with smartassist for left ventricle unloading. treated with amiodarone and converted back to normal sinus rhythm. developed heparin induced thrombocytopenia, as platelets dropped from 158 to 81. heparin switched to argatroban for anticoagulation. hemodynamics were stable but patient became symptomatic during weaning trial. pump explanted and hemostasis was achieved by vascular stapling a graft and oversewn, suturing the wound shut. the patient curvived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia / paroxysmal atrial fibrillation: the cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3930 | 1220648-2025-47891 | 23407529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407529&pc=PYX | 09/30/2025 |  |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026699658 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 60 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a drop in hemoglobin. an x-ray revealed a site of suspected bleeding. the patient was transfused four units of packed red blood cells and was taken back to surgery for re-exploration of a thoracotomy. also, heparin was withheld. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3931 | 1220648-2025-47892 | 23407533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407533&pc=OZD | 09/30/2025 |  |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713888 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 80 KG | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was low purge pressure during impella cp patient support. during the coronary artery bypass grafting procedure, the purge flow began to go to 30 and the purge pressure dropped below 200. the purge cassette was replaced and the controller was changed. neither resolved the issue. the doctor replaced the cp in the operating room. the new cp was placed and was working properly. the patient survived. | [5 empty] | The product with event logs were received and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. purge pressure low: clinical details note that during a coronary artery bypass graft (cabg) procedure on the second day of impella support, the purge flow began to increase to 30 ml/hr. and the purge pressure dropped below 200. replacing the purge cassette and controller did not resolve the issue. it was stated there was no leak observed and no excessive force used. data log review showed an increase in purge flow from ~7 to 30 ml/hr. within an hour. surgical mode was enabled approximately two hours after the purge flow got to 30 ml/hr. the returned product was evaluated and there was a crack found on the back of the purge sidearm filter and the low purge pressure reproduced with a leak observed from the filter crack. the cause of the low purge pressure was a leak from the damaged sidearm filter, which was use-related based on the crack observed on the back of the returned purge sidearm filter. correspondingly, sections d9: device received by manufacturer with receipt date and h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3932 | 1220648-2025-47893 | 23407548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407548&pc=OZD | 09/30/2025 |  |  | 10/28/2025 |  |  | 09/30/2025 | 10/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709867 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a pump stop and suspected thrombus during impella cp patient support. on the 14th day of impella support, the pump suddenly stopped. multiple attempts were made at performance levels 0-9 without success. there was a suspected clot. the impella was removed and the decision was made to not implant another impella. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product was received and the investigation was completed. correspondingly, fields d9: device received by manufacturer with receipt date were updated accordingly. device history record review was completed, and pump passed all post-sterile inspection checks. thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. pump stop: the pump was returned cut so serial number could not be verified. the returned pump was inspected, and an abnormally large purge gap was observed. clinical reported pump stopped after 14 days of support and could not be restarted. support was from 16-sep to 30-sep and logs returned were from 16-sep to 25-sep. in the logs returned there was no pump stop. the pump was used for 14 days, which is beyond the 8-day design life of impella cpc9. the cause of the pump stop was determined to be due to bearing wear as evidenced by abnormally large purge gap on returned product aligning with clinical information of use beyond indication. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3933 | 1220648-2025-47894 | 23407560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407560&pc=OZD | 09/29/2025 | 12/16/2025 |  | 10/28/2025 |  |  | 09/29/2025 | 10/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698773 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Thrombosis/Thrombus | Heart Block | Heart Failure/Congestive Heart Failure | [31 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella cp experienced ectopy and ventricular tachycardia followed by a heart block after explantation of the pump. an echocardiogram showed mild tricuspid and mitral regurgitation, moderately reduced right ventricle failure, an akinetic anterior wall, and reduced ejection fraction. the pump's p level was reduced and the patient was transfused packed red blood cells. cannulas were placed in the right lower extremity and the left femoral artery due to ischemia. the patient was escalated to an impella 5.5 for increased support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was returned for analysis. the cause of the suction could not be determined. the cause of the access site adverse event was determined to be an unintended use issue. the root cause of the arrhythmia, tachycardia, ischemia, and thrombus were unable to be determined due to insufficient clinical details. the cause of the heart block and cardiac failure cannot be determined as insufficient clinical details were provided. the device passed all post-sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3934 | 1220648-2025-47895 | 23407561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407561&pc=OZD | 09/30/2025 | 12/30/2025 |  | 10/28/2025 |  |  | 10/01/2025 | 10/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607631 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was anemia, a minor bleed and pump suction during impella cp patient support. while on support, there was minor bleeding noted around the impella catheter at insertion site. it stopped after the use of a pressure bandage. it was noted that the insertion of the device in the cathlab was hasty. therapy continued as planned. the patient received two units of packed red blood cells at a hemoglobin of five. the ¿no suction¿ alarm was discussed but the constant negative diastolic left ventricle pressure was below -30. the suction was resolved through lowering the performance level and fluids. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: minor bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. anemia : the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot : 1878514. device history batch: sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
3935 | 1220648-2025-47896 | 23407615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407615&pc=OZD | 10/14/2025 |  |  | 10/28/2025 |  |  | 10/14/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706865 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 91 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced sepsis. the patient was escalated to an impella 5.5. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3936 | 1220648-2025-31887 | 23407628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407628&pc=OZD | 08/05/2021 | 10/28/2025 |  | 10/28/2025 |  |  | 10/27/2021 | 10/28/2025 |  |  | 04/30/2023 |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070465 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Leak/Splash | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, there was a leak at the yellow luer connection. the cassette was changed twice followed by the de-air procedure, however the leak continued. the nurse was instructed to perform a purge system bypass to resolve the issue. additionally, it was reported that the patient experienced bleeding at the impella access site, and as a result heparin was withheld until hemostasis was achieved. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
3937 | 1220648-2025-47843 | 23407644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407644&pc=OZD | 10/08/2025 |  |  | 10/28/2025 |  |  | 10/08/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734054 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 had high purge pressure and purge system blocked alarms. lysis was administered. it was further reported that the patient¿s therapy was discontinued, the reason was not provided due to data protection of the patient. this decision was not impella related. the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction b1. the investigation of the reported failure mode has been completed. no product was received for evaluation. purge pressure high: data log review confirmed the reported issue. the cause of the high purge pressure was biomaterial buildup since data log review showed gradual purge pressure buildup and no motor current (mc) spikes. device history review: the device passed all post sterile inspection checks.
3938 | 1220648-2025-31888 | 23407816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407816&pc=OZD | 05/12/2022 | 10/28/2025 |  | 10/28/2025 |  |  | 05/12/2022 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUM |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1488594 |  | 0048-0032 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with known coronary artery disease, mitral valve disease, an ejection fraction of 20%, and noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. it was reported that the staff noticed a purge system open alarm which they were unable to resolve. as a result, the impella cp was removed and the hrpci was postponed until the following morning. there were no adverse events reported as a result of the purge alarm. | [5 empty] | D4: unable to locate an expiration date, manufacture date, or udi# for this product returned purge cassette and pump visually inspected and no abnormalities were observed. complaint purge cassette and pump connected, and no persistent purge system open alarm reproduced. purge value was with specification. cut open purge cassette and no leaks were found.
3939 | 1220648-2025-31889 | 23407830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23407830&pc=OZD | 07/20/2025 | 12/04/2025 |  | 10/28/2025 |  |  | 09/03/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314855 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 showed an erratic placement signal with position alarms on the automated impella controller (aic). the white cable was pushed in all the way. staff had to disable placement signal monitoring. motor current and maximum/minimum flows were reported to be good and appropriate. care was eventually withdrawn, and the patient expired. | [5 empty] | Data analysis: the pump investigation reviewed the aic log data and concluded: "data logs were reviewed and the position in aorta alarms were confirmed. the alarms aligned with significant drops in the snr and contrast. leading up to the alarms, an rt log showed that raw optical signal had a significant amount of noise that correlated with the initial drop in snr. the ps began to jump back and forth between values around 200 mmhg and 70 mmhg. however, the algorithm initially was able to filter out the jumps up to 200 mmhg. later however, they were not being filtered out, so the position in aorta alarm was falsely triggering. this is because jumps up to a higher psmax made the mc and ps waveforms go out of phase, so the algorithm calculation believed there was insufficient mc modulation. later in the case, the ps returned to showing steady values around 70 mmhg and improved snr/contrast without any clear change in any other pump metrics. that being said, the waveform did not appear to be aortic at all. there was also no clear pressure value range that would cause snr/contrast to drop. the performance of 2 pumps used on the same console prior to the complaint pump were checked, and there were no indications that this same complication occurred on either pump. the pump was not returned for investigation. " device analysis: given the console ic5647 was not returned for this investigation, the cause of the optical signal issue could not be determined. d6a / d6b should have been left blank. | The impella device has not been returned by the customer. a supplemental report will be filed upon device return or receipt of any additional information.
3940 | 1220648-2025-47897 | 23408482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23408482&pc=OZD | 10/01/2025 |  |  | 10/28/2025 |  |  | 10/01/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707312 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after cabg x 5 was performed, impella 5.5 was placed via direct aortic route. in icu, the patient had sustained ventricular tachycardia followed by sinus tachycardia. amiodarone was added. the patient continued impella support and was successfully weaned and explanted as planned | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3941 | 1220648-2025-47898 | 23408833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23408833&pc=OZD | 10/02/2025 |  |  | 10/28/2025 |  |  | 10/02/2025 | 10/28/2025 |  |  |  |  | 10/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683446 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male | 124 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a thrombus on the impella catheter and a placement signal not reliable alarm. the patient was escalated to an impella 5.5 and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product returned. data logs received were missing the last few hours of support in comparison to the reported explant date/time. the cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the hemolysis was unable to be determined due to lack of clinical data, inconclusive data log review, and no product returned for investigation. the cause of the optical signal issue was determined to be console related per data log review. the device passed all post sterile inspection checks.
3942 | 1220648-2025-47730 | 23389178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389178&pc=OZD | 09/10/2025 |  |  | 10/27/2025 |  |  | 09/10/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703324 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Fever | Hemorrhage/Blood Loss/Bleeding | Ischemia | Great Vessel Perforation | Hematuria | Thrombosis/Thrombus | Hemodynamic instability | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was placed on an impella cp for mechanical circulatory support. the case noted the patient had a left anterior descending (lad) stent placed. decompensated stent thrombosis impella, heart non-pulsatile, pink-red urine output, and medication requirements were noted. suction and volume issues, gastrointestinal bleeding, and groin drainage were also noted. the pump was repositioned after suction alarms were noted. pedal pulses were lost after turning and repositioning. the left foot was mottled with absent distal pulses. it was determined to remove the pump the following day. vascular injury requiring three units of red blood cells and vascular team was consulted. significant inotropic support was needed. seizure like activity and inability to escalate care/perform further diagnostics due to hemodynamic instability with movement. the patient became febrile, hypothermic, anuric requiring continuous veno-venous hemodialysis. impella site ooze was noted and an echocardiogram showed severe anterior hypokinesis. the impella cp was removed and an intra-aortic balloon pump was placed. the impella was successful explanted due to complication and closure was successful. | [5 empty] | No product was returned. the cause of the pump suction was not determined due to no product returned, no data logs recovered, and insufficient clinical details. the root cause of the groin damage, hemolysis, hematuria, bleeding, hemodynamic instability, thrombosis, fever, and ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3943 | 1220648-2025-47798 | 23389245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389245&pc=OZD | 10/01/2025 |  |  | 10/27/2025 |  |  | 10/01/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026750134 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. upon removal of the impella pump, a clot/tissue was observed on inlet cage after removal. the medical doctor elected to send to pathology for further evaluation. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number.
3944 | 1220648-2025-47797 | 23389269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389269&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122975 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the automated impella controller (aic) experienced a controller error. the clinical team was advised to ensure that the cable was fully seated in the aic. the alarm resolved immediately after the cable was assessed. the aic console was not replaced and support continued. no patient harm was reported due to the issue. | [5 empty] | The investigation of the reported failure mode has been completed. data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench are represented with negative values due to the crc error. s device analysis: the console was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the console operated as designed. root cause: the root cause of the ¿controller error and optical signal issue was due to a firmware issue (optical bench fw anomaly). device history review: the device passed all the post sterile inspection checks. | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to section h.10 for the related report number.
3945 | 1220648-2025-47799 | 23389359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389359&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639740 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Pulmonary Embolism | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support after an unsuccessful percutaneous coronary intervention to rest and unload the heart. on the third day of impella support, it was reported that there were intermittent suction alarms. albumin and intravenous fluids were administered, and during administration of fluids the patient experienced significant chest pain, and the patient received a nitroglycerin drip. the patient¿s nurse reported that the patient had a small bilateral pulmonary emboli. echocardiography was performed and revealed that the impella was sitting very shallow. the impella was repositioned under fluoroscopy and verified with an echocardiogram. the impella was advanced into the left ventricle. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary pain/pulmonary embolism: the cause of the injury was not determined due to insufficient clinical details. pump suction: clinical reported the patient was having intermittent suction. the product was not returned for investigation. data log review confirms suction alarms occurring during the case. placement signal was seen trending downward leading up to the suction alarms, and clinical stated that albumin was given, meaning the patient was likely experiencing volume issues. the positioning issues reported did not occur until after the suction was triggered and resolved. the cause of the suction was most likely related to patient condition, and placement signal was trending down at the time and albumin was given to address the suction conditions. positioning issue: clinical reported after repositioning there was concern that the pump was sitting shallow on echo. the product was not returned. the cause of the positioning issue could not be determined due to insufficient clinical details and no return product. device history lot device lot number: 1897520 device history review pump sn (b)(6) passed all post sterile inspection criteria. qn #(b)(4) was opened because particle was found inside the tray and package for pump sn (b)(6). the pump was reworked and approved. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
3946 | 1220648-2025-47800 | 23389400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389400&pc=PYX | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026668123 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute heart failure was implanted with an impella rp flex for mechanical circulatory support following a heart transplant and cannulation with venoarterial extracorporeal membrane oxygenation (va ecmo). the patient originally was supported by an impella 5.5 as a bridge to transplant, successfully received a heart transplant, and during surgery for the heart transplant the impella 5.5 was explanted. following the transplant, the patient was placed on va ecmo. after four days of va emco support, the decision was made to wean the patient from va ecmo, but the patient required right heart support. the decision was made to implant the patient with an impella rp flex to wean the patient off of va ecmo. the patient was successfully weaned and decannulated from va emco the following day. on the third day of impella rp flex support, it was reported there was an ongoing placement signal unreliable alarm (psnr). as pump functionality was not impeded, the alarm was disabled, and the medical team was instructed by a company representative to monitor motor current and pump flows. two days later, the patient¿s nurse reported that changing the purge cassette resolved the psnr. a company representative advised not using the displayed metrics for documentation and discussed that there is no way to know if the sensor would fail again. the following day, it was reported that there was a drop in pump flows that was preceded by an increase in motor current and placement signal. it was noted that the motor current flattened out and then was followed by a spike and the motor current doubled. the motor current never returned to baseline, and the pump flows remained reduced. the patient remained stable throughout the pump issues. the company representative advised there was a potential clot entrainment causing the pump¿s decreased flows and high motor current. the pump was subsequently explanted due to the low pump flows; however, the patient¿s right ventricular function had improved enough that the medical team determined no other temporary support was needed to replace the impella rp flex. the pump was successfully explanted; however, the overall patient outcome is unknown. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product returned. thrombosis: the cause of the injury was unable to be determined due to insufficient clinical details. low flow: clinical reports low flows after motor current (mc) spike that does not resolve. activated clotting time (act) was in range. blood volume and pump repositioning not attempted as troubleshooting. field data logs confirm the mc spike with speed deviation and upward shift in placement signal (ps) on sep-30. flows drop and remain low for remainder of case. the cause of the low flow was biomaterial ingestion based on data log review. placement signal issue: placement signal not reliable (psnr) reported on sep-27. data logs confirm the psnr alarms (opm bench invalid) and crc controller error alarms. snr drops to 0, gain & light drop, lamp steps up, and contrast oscillates between about -/+ 3200. the issue lasts ~46 hours and then resolves on sep-29. the cause of the placement signal issue was determined to be console related per data log review. device history lot: device lot: 1912982. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
3947 | 1220648-2025-47801 | 23389511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389511&pc=PYX | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025507455 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 92 KG |  | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Fever | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | Hemodynamic instability | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella rp flex experienced bleeding at the access site. an additional modified purse string suture was required to achieve hemostasis. pump alarms occurred for suction so fluid and blood products were given. the patient experienced atrial fibrillation along with a drop in impella flows and hemodynamic instability requiring an increase in pressors given. cardioversion was attempted but was unsuccessful. the patient also experienced increasing hypoxia and fevers. the patient's family decided to give comfort care and the patient expired. | [5 empty] | A5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | Data logs were reviewed and confirm the complaint. the device was not returned for investigation. the cause of the low or blocked pump flow/pump suction was most likely patient condition related since the patient had severe heart failure which not tolerate with treatment. the cause of major bleed was most likely due to use related since the additional modified purse string sutured required for hemostasis. the root cause of the fever and arrhythmia was unable to be determined due to insufficient clinical details. the cause of the respiratory failure and hemodynamic instability was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A review of the data logs confirmed the complaint. the pump was not returned for review. the cause of the low or blocked pump flow/pump suction was most likely patient condition related since the patient had severe heart failure which not tolerate with treatment. the cause of major bleed was most likely due to use related since the additional modified purse string sutured required for hemostasis. the root cause of the arrhythmia and fever was unable to be determined due to insufficient clinical details. the cause of the respiratory failure and hemodynamic instability was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3948 | 1220648-2025-47802 | 23389587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389587&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727022 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 117 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support following an episode of ventricular arrest which required multiple rounds of defibrillations to resolve. the patient was presented for a percutaneous coronary intervention following impella cp support to address a chronic total occlusion of the left anterior descending coronary artery and lesion of the right coronary artery. the physician was unsuccessful at opening the right coronary artery, and the patient was transferred to another hospital for a possible coronary artery bypass graft. upon arrival to the hospital, there were alarms for impella in aorta. the decision was made to remove the impella cp due to the positioning issue. the outcome of patient is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Investigation summary: the investigation has been completed. device in wrong position: the pump was not returned for investigation. data log shows an active "pump position in aorta" alarm, with trends in motor current and flow confirming the pump was out of position. the cause of the positioning issue was most likely unintentional use-related, as the issue occurred during patient transfer. device history review: the pump passed all post sterile inspection checks.
3949 | 1220648-2025-47803 | 23389711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389711&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690152 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female | 55 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a gastrointestinal bleed and suspected heparin-induced thrombocytopenia, with low hemoglobin and platelet values. the patient was transfused packed red blood cells and platelets. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | No product was returned for investigation. the cause of the bleeding is determined to be use issue. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3950 | 1220648-2025-47804 | 23389875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389875&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107436 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a yellow controller error alarm on the supporting automated impella controller. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3951 | 1220648-2025-47790 | 23389956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23389956&pc=OZD | 09/27/2025 | 10/24/2025 |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725985 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock and was escalated to an impella 5.5. during impella 5.5 support, the patient experienced a gastrointestinal (gi) bleed. later, suction events occurred along with increased chemical support requirements. the patient had low hemoglobin (hgb) and received albumin and red blood cells. the patient was given "a lot of products overnight." the patient continued with low hgb and was given two units of packed red blood cells (prbcs). furthermore, the patient reported having large black tarry stools with a drop in hgb with one unit of prbcs given. additional blood transfusion was received due to gi bleeding. the patient was escalated to another impella 5.5 and remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 added code b13 as it was omitted from the initial report that was submitted. h10 added related report number. this is one of two related reports. this report is for the second pump used and another report was submitted for the third pump used. investigation summary: the investigation has been completed. major bleed/anemia: the cause of the injury was not determined due to insufficient clinical details. pump suction: clinical details reported suction events and albumin was given. the product was not returned for investigation. data log review confirmed suction occurred. at the time suction started to trigger, placement signal started to trend down from 110mmhg to approximately 70mmhg. clinical details also reported giving albumin to address the suction. suction did not occur for the remainder of the case and placement signal increased again, indicating that the patient was having volume issues but after getting albumin, they resolved. the cause of the suction was most likely related to the patient's condition based on clinical details and the downward trend in placement signal during suction. device history review: the pump passed all post sterile inspection criteria.
3952 | 1220648-2025-47806 | 23390076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390076&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Backflow | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  | Stroke/CVA | Hemolysis | Multiple Organ Dysfunction Syndrome | Embolism/Embolus | [9 empty] | Required Intervention | Death |  |  |  |  | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support following a high-risk surgery for a reoperation on their mitral valve. the mitral valve reoperation required over eight hours of support from a cardiopulmonary bypass machine and the patient was cannulated for veno-arterial extracorporeal membrane oxygenation (va ecmo) and implanted with an intra-aortic balloon pump (iabp) before being taken to the cardiac catheterization laboratory for impella cp placement for left ventricular unloading. after the impella cp was placed, the iabp was removed, but the patient remained on va ecmo support with the impella cp. the following day, it was reported that there had been suction alarms from the impella overnight. the performance level was lowered to p2 and the patient was given albumin. the patient¿s laboratory results were monitored for signs of hemolysis. the next day, the impella measured deep then ideal at 5.8-6 centimeters (cm). it was also noted that the plasma free hemoglobin (pfhgb) was down slightly at 40 milligrams per deciliter (mg/dl) and the lactate dehydrogenase (ldh) results were stable but elevated. the impella was repositioned and pulled back to 3.7 cm, however there was concern it may be slightly orientated towards the mitral valve so a full beside echocardiogram was ordered. after pulling the impella back to 3.7 cm, the medical team reported alarms for impella in aorta. the impella was advanced across the aortic valve to resolve the alarm and correct the impella position. the next day, on the fourth day of support, it was reported that the patient was in multi-organ failure, and aggressive hemolysis with a pfhgb result of 150 mg/dl and an ldh of 1646. the patient was placed on dialysis and due to the hemolysis, the patient was transfused with one unit of packed red blood cells. on the fifth day of support, it was noted that the patient had a poor neurological prognosis after being on cardiopulmonary bypass for 12 hours. the patient continued to hemolyze however the pfhgb results improved from 150 to 100 mg/dl. it was noted that both the continuous renal replacement therapy (crrt) and va ecmo circuits clotted and required replacement. the medical team had concerns for potential micro emboli causing hemolysis. of note, there were no impella concerns, and echocardiography showed the impella in optimal positioning. it was noted that there were occasional preventing retrograde alarms due to the patient¿s high afterload. on the sixth day of support, it was noted that the patient¿s hemolysis had decreased, although the pfhbg was still elevated at 60 mg/dl. on the eighth day of support, computed tomography was performed and revealed a large ischemic stroke affecting the left middle cerebral artery. it was also noted that the patient was unable to be weaned and decannulated from va ecmo the previous day. the patient¿s family decided to withdraw the patient¿s care. the patient¿s care was withdrawn, and the patient expired. | [5 empty] | Based on the information available, it is unlikely that the stroke was related to the use of the impella device; however, a device-related contribution cannot be definitively ruled out. although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | This report is being submitted to correct h6 medical device problem codes which were previously incorrectly reported. the codes have now been updated to reflect the correct codes
3953 | 1220648-2025-47807 | 23390163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390163&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688128 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support following a hypotensive episode and increased lactic acid results that occurred post percutaneous coronary intervention to the left anterior descending and obtuse marginal on (b)(6) 2025. on the second day of support, it was reported that the patient experienced ventricular fibrillation arrest. the impella was found to be too deep at 6.3 centimeters (cm) and potentially causing arrhythmia. the impella was repositioned successfully and pulled back to 4.88 cm. antiarrhythmics were administered and the patient was ultimately shocked four times to resolve the arrythmia. no additional arrythmias were noted following the repositioning, and it was noted that the patient remained in critical condition. on the seventh day of support, it was reported that the patient continued to have occasional hypotension with arrhythmias on and off bilevel positive airway pressure. no interventions were performed to resolve the occasional episodes. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary ==> no product was returned. ventricular fibrillation, arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history lot- device lot: 1932171. device history review- device (sn: (b)(6) passed all post sterile inspection checks.
3954 | 1220648-2025-47809 | 23390343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390343&pc=OZD | 09/28/2025 | 10/23/2025 |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709958 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | Valvular Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered placement signal issues and that the patient had stroke, thrombus and mitral regurgitation during impella 5.5 support. it was noted that placement signal not reliable alarm (psnr) was on and that ps (placement signal) drift was noted prior to psnr. additionally, patient encountered stroke. cat scan revealed that patient had a clot to right mca (middle cerebral artery). patient was not a candidate for mechanical thrombectomy. patient experienced hemorrhagic conversion of prior ischemic stroke and anticoagulant was discontinued. moreover, patient had severe mitral regurgitation for which the medical team planned on increasing medical management for patient also encountered hematoma. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted to correct a duplicate mdr. it has been determined that mdr 1220648-2025-47809 and mdr 1220648-2025-46531 report the same event. this report (1220648-2025-47809) is being closed as a duplicate of the initial, valid report (1220648-2025-46531).
3955 | 1220648-2025-47810 | 23390362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390362&pc=OZD | 09/28/2025 | 12/09/2025 |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724592 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 114 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had access site bleeding during impella cp support. it was noted that patient had oozing from site as well as oozing from the suture puncture marks where the needle went into the skin. after repositioning, physician decided to place another forward tension suture. the medical team was talked through angle matching, and manual pressure was also applied by staff for 20 mins. after 20 mins, hemostasis was achieved with no visible oozing and dressing was applied. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed has been completed. no product was returned for review. the root cause of the major bleed is most likely patient condition related since patient has oozing from site as well as oozing from the suture puncture marks where the needle went into the skin.
3956 | 1220648-2025-47811 | 23390385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390385&pc=OZD | 09/28/2025 |  |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727027 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that there was arrhythmia, positioning issue, hematuria, and suction events during impella cp patient support. the patient presented in bradycardia. initial positioning of the pump was close to the left ventricle septum and was beginning to show signs of possible suction events so the p level was lowered and volume was added. the patient also exhibited hematuria. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned for investigation. data log were reviewed. the cause of the hemolysis was not determined without sufficient clinical information and returned product investigation. the cause of the positioning issue was not determined without sufficient clinical information and returned product investigation. there was no suction alarm reported during the case run. no suction defect was found in this case. the issue could not be confirmed. the device passed all post sterile inspection checks.
3957 | 1220648-2025-47813 | 23390409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390409&pc=OZD | 09/28/2025 |  |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026730256 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed frank hematuria. transthoracic echocardiography was performed and impella was retracted and redirected with clockwise torque to mid-ventricular cavity. impella measured at 4.3. hematuria was beginning to improve. intensive care unit team was also working on reducing the afterload to also help with the hematuria. heparin was held after noticing hematuria. impella was repositioned again, more anteriorly directing towards the mitral apparatus as hematuria continued. the impella was now in the midventricular space measuring 4.5 cm. the next day, hemolysis was noted that increased when the 5.5 flow was increased from p4 to p6 upon extracorporeal membrane oxygenation (ecmo) decannulation. the 5.5 was repositioned during ecmo decannulation. on day six of impella support, impella in aorta alarm occurred. position was evaluated and repositioned was attempted but unsuccessful. the catheter was kinked. impella was removed. the patient was stable and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3958 | 1220648-2025-47814 | 23390558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390558&pc=OZD | 09/28/2025 |  |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026734274 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced blood in bowel movements requiring removal of an abdominal abscess. additionally, the patient was on vasoactive medications and continuous renal replacement therapy for fluid removal. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the bleeding was determined to be patient condition because of the bleeding from the non-impella site. the cause of the renal failure was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3959 | 1220648-2025-47815 | 23390565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390565&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723590 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Dyspnea | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was unable to be inserted due to the patient calcified iliac. the cp was removed and iliac was ballooned but insertion was unsuccessful. the third attempt, the iliac was stent but still unsuccessful cp placement. a decision made to abort impella insertion. the patient on the operation room table was complaining of chest pain and shortness of breath. the impella 5.5 was implanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and noted the following: since the impella cp placement via groin was unsuccessful the physician strategy changed. the plan was to place impella 5.5 via axillary artery. however, the impella 5.5 was never implanted. as axillary cut down was being performed the physician placed cp via axillary instead due to the patient quickly decompensating. the outcome of the patient following the second implant is unknown at the time of this mdr writing. h6 health effect - impact code required correction to reflect the impella pump was replaced. consequently, h6 health effect - clinical code/impact code and medical device product code were updated/fully repopulated. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the dyspnea, the cause was not determined due to insufficient clinical details. regarding the unable to deliver or advance (delivery issue), the cause of the deliver issue was determined to be patient condition as the patient had poor calcified iliac/diseased vessel preventing successful delivery of impella. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3960 | 1220648-2025-47816 | 23390579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390579&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734905 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient had a hematoma noted upon arrival at intensive care unit. site oozed for four hours with manual pressure held above repositioning sheath. noted bleeding from needle stick for forward tension sutures as well. activated clotting time decreased but oozing continued. the physician returned and placed a figure of 8 suture to the groin with no relief of oozing. decision made to remove the pump due to ongoing bleeding issues and changes in labs. the physician removed impella at bedside and tightened percloses down for vessel control. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed/ hematoma: the root cause of the bleeding/ hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
3961 | 1220648-2025-47808 | 23390585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390585&pc=OZD | 09/28/2025 |  |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694715 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support had become agitated despite sedation and removed the impella catheter. the patient was taken back to the catheterization lab where the impella cp was advanced back into position. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient. the device passed all post sterile inspection checks.
3962 | 1220648-2025-47818 | 23390590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390590&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685148 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 99 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient was repositioned 30 minutes prior to event. after repositioning, fluid bolus and p level reduction suction remained. ¿the suction won't break let¿s remove the impella. the patient doesn¿t need the pump anymore." there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B7 added information that was omitted from the initial report that was submitted. d4 revised primary udi number as it was entered incorrectly from the initial report that was submitted. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. pump suction: clinical details reported suction events started after the patient was repositioned, and troubleshooting methods did not resolve the suction. the product was not returned for investigation. data log review confirmed suction alarms occurring towards the end of the case. at the time suction starts, placement signal was stable, and there were no abnormalities in 5s parameters, but motor current pulsatility decreases which could suggest obstructed inflow or improper positioning, but no information was provided regarding pump positioning at this time. the cause of the suction could not be determined due to insufficient clinical details on pump positioning. device history review: the pump passed all post sterile inspection criteria.
3963 | 1220648-2025-47819 | 23390634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390634&pc=OZD | 09/28/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717783 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella cp had a leak on the purge filter. no changes occurred in motor current, purge flow, or purge pressure. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The purge cassette and pump were returned. the returned pump was visually inspected and a filter vent hole was circled with marker indicating the location of the field leak. no cracks or other abnormalities were observed on the filter. reproduction leak testing was performed on the returned pump and the filter vent hole leak reproduced from the circled hole. the filter was torn down and the back of the failing membrane was inspected and found to have the ptfe fibers exposed. the other membrane appeared to be intact with a smooth non-fibrous membrane surface. the cause of the pump leak was a damaged filter membrane resulting in a leak from the vent hole. the pump passed all post-sterile inspection criteria. | D.9 added device was returned and date. h.6 updated codes as the device was returned. h.11 updated as the device was returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3964 | 1220648-2025-47820 | 23390645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390645&pc=OZD | 10/01/2025 |  |  | 10/27/2025 |  |  | 10/01/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1315810 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced yellow "controller error" alarm associated with a power or battery issue. the aic was in the storage room. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
3965 | 1220648-2025-47821 | 23390665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390665&pc=OZD | 10/03/2025 |  |  | 10/27/2025 |  |  | 10/03/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025497117 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon explantation of an impella cp the ostium secundum was found to be dissected. an endovascular thrombectomy was performed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the vascular dissection was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
3966 | 1220648-2025-47822 | 23390714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390714&pc=OZD | 09/28/2025 | 11/06/2025 |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288833 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 64 KG | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. on the fourth day of impella cp support, it was reported that the automated impella controller (aic) stopped responding when buttons were pressed following a purge cassette and tubing change. the nurse attempted multiple times to press all buttons without any response received. the nurse observed that a small amount of purge fluid encountered the screen when spiking the new bag of purge fluid. as a result of the issue, the aic was exchanged successfully and the purge cassette was successfully primed. there was no patient harm reported. the impella cp was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: the first log entries indicating ¿imc-kbd error: 0xa1 0x01¿ appeared, suggesting that the issue occurred when a touch was detected outside the soft-key area immediately after the purge change was initiated¿approximately three days after the pump was started. subsequently, the complete the procedure (complete the purge wizard) alarm was triggered. log data showed purge pressure and purge flow dropping to zero and never recovered. device analysis: upon arrival, the reported issue could not be reproduced, even after performing multiple power cycles. the console was tested with both the pump and purge cassette for several hours, and no abnormalities were observed. an inspection of the internal circuitry and cables revealed everything to be in good condition. root cause: the root cause for the reported issue ¿the buttons and wheel on the console did not respond to any touch.¿ could not be determined due to the inability to reproduce. phr summary: there were no issues related to complaint discovered when reviewing the product history of console ic5072.
3967 | 1220648-2025-47823 | 23390820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23390820&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731100 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 71 KG | Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient had a hematoma noted at the access site. manual pressure was held and pressure dressing was applied. hemoglobin was unchanged. duration of use was longer than food and drug administration indication. leak in purge side arm required pump exchange. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The pump was not returned for investigation. the cause of the hematoma was not determined without returned product and sufficient clinical information. the cause of the purge leak issue was not determined without returned product and sufficient clinical details. the device passed all post sterile inspection checks.
3968 | 1220648-2025-31882 | 23391307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391307&pc=OZD | 02/27/2023 | 10/27/2025 |  | 10/27/2025 |  |  | 02/27/2023 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1395260 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that, after pump set up and insertion/activation, a placement signal unreliable alarm appeared on their impella cp / automated impella controller (aic), and no aorta or left ventricle placement signals were visible. the pump was eventually weaned and explanted successfully, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event (although having incorrect timestamps) are consistent with complaint and show that placement signal was unavailable since the moment pump 425490 was connected to the console. as a result optical sensor calibration could not be executed, and operators were prompted to disconnect and reconnect the pump (twice), which did not resolve the issue. the issue of placement signal loss had been determined by the pump investigation 23-13701-2 to have been caused by fiber break within the pump red handle. during log analysis and testing of the console, it¿s been observed that the real-time-clock did not maintain the time, which later troubleshooting confirmed to have been caused by depleted rtc coin-cell battery. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3969 | 1220648-2025-31883 | 23391329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391329&pc=OZD | 02/27/2023 | 10/27/2025 |  | 10/27/2025 |  |  | 02/27/2023 | 10/27/2025 |  |  | 12/31/2024 |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295328 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. the pump triggered an "unreliable placement signal" alarm, and no aortic or left ventricular placement signals were visible. the pump remained in use for continued support. no harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the product was returned and the 5s parameters suggested a hard optical failure. the laser light test was performed, and light was found to be leaking from the red handle, indicating a break in the optical fiber. the red handle was opened and the fiber was confirmed to be broken around the post of the red handle. the root cause of the optical signal issue was determined to be a damaged optical fiber. this is likely due to their not being enough slack for the fibers to translate when the forces pump insertion were applied, kinking the fiber around the post in the red handle. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
3970 | 1220648-2025-47827 | 23391456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391456&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729352 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 158 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [5 empty] | VA ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 and cannulated for veno-arterial extracorporeal membrane oxygenation (va ecmo) for mechanical circulatory support after experiencing poor left ventricular function following a complex cardiac surgery that included coronary unroofing, coronary artery bypass grafting, mitral valve replacement and left atrial appendage ligation. on the twelfth day of va ecmo and impella 5.5 support, it was reported that the patient had an episode of atrial fibrillation during the night. the patient was administered with a bolus of amiodarone followed by a continuous intravenous drop of amiodarone to treat the arrhythmia. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | Additional information was provided that the patient had pre-existing atrial fibrillation and the medical team did not believe that the impella caused the arrythmia. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5: additional information was provided by the medical team regarding the patient's pre-existing conditions. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary paroxysmal atrial fibrillation : the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot : 1962574. device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
3971 | 1220648-2025-47829 | 23391671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391671&pc=PYX | 09/16/2025 |  |  | 10/27/2025 |  |  | 09/16/2025 | 10/27/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026630673 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp had suction which was resolved at p6 level and albumin was given. on day seven of impella rp support the impella was removed due to a clot that formed on the catheter in the middle of the night and decision made to explant impella. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | D2a and d2b was erroneously entered on the initial manufacturer device report number and updated correctly on this report. investigation was completed. low pump flow: clinical reported that placement signal suddenly increased and flows began to drop. a motor current spike also occurred before the change. the product was not returned for investigation and data logs were not provided. a motor current spike could be an indication of biomaterial ingestion, but without returned product or data logs, it cannot be confirmed as the cause of low flow. the cause of the low pump flow could not be determined due to lack of data logs and no product returned. device history review was completed and passed all post sterile inspection criteria.
3972 | 1220648-2025-47830 | 23391759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391759&pc=PYX | 09/16/2025 |  |  | 10/27/2025 |  |  | 09/16/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026683432 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 106 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered access site bleeding during impella rp flex support. during insertion of the pump, after the repositioning sheath was secured, some oozing was noted at the access site. this was treated with placing additional sutures around the access site. the patient was placed on extracorporeal membrane oxygenation and the impella was removed, patient survived. | [5 empty] | B3 revised date of event as it was reported incorrectly on the initial report that was submitted. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. g4 revised pma/ 510(k) as it was entered incorrectly on the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. revised b17 to b18 as the device was discarded. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ investigation summary: the investigation has been completed. the device was discarded. major bleed: the cause of the injury is most likely use related since placing additional sutures around the access site helped reduce bleeding. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3973 | 1220648-2025-47831 | 23391853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391853&pc=OZD | 09/27/2025 |  |  | 10/27/2025 |  |  | 09/27/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026736379 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 70 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was escalated from an impella cp with an impella 5.5 for mechanical circulatory support. on the morning of (b)(6) 2025 the decision was made by team to escalate the patient to an impella 5.5. at the time, the patient was on an impella cp. impella 5.5 support was started and the impella cp was pulled into the descending and explanted with no issues. at implant, time waveforms on the automated impella controller (aic) were able to be increased to p8-p9 flows. after about 30 minutes since the implant, the left ventricle (lv) waveforms went negative systolic/diastolic and were not able to flow above p5 flows. the right ventricle function was assessed with good function. transesophageal echocardiogram (tee) measurement confirmed the impella position was still measuring 5.5 centimeters (cm) distance from the aortic annulus direction towards the apex. the decision was made to transfer to the cardiovascular intensive care unit (cuicu) and stabilizing continues to reassess flows. the patient was given 14 mmhg albumin and was started on an epinephrine drip which the patient responded too and was able to increase to p7 flows 2.7-3 liters for a short time but did continue to have suction alarms. tte was performed at bedside and confirmed the position was still adequate at 5.3 cm from aortic annulus. a discussion with the team was had even though it was in good position. the patient was having suction alarms and could not increase flow to attempt to reposition. the team stated at this time they were attempting to wean the ventilator. over the next few hours, the patient continued to decompensate and required more vasoactive support. the patient code blue and cardiopulmonary resuscitation was performed and on automated impella controller (aic) waveforms appeared to have flow saturation min 6.2 max 6.2 computed tomography surgeon placed the patient on peripheral veno-arterial extracorporeal membrane oxygenation and the patient was taken emergently back to cardiovascular operation room to exchange the impella 5.5 and successful placement of the new impella 5.5 was noted. the patient's care was withdrawn and the patient expired after 12 days on support on the new impella 5.5. this report specifically addresses the first impella 5.5 and the second impella 5.5 (serial number (b)(6)) will be reported under a separate report. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. pump suction/low flow/saturated flow. the data logs show that pump flows were within expected range at the start of the case but after about 12 minutes into support, there was a motor current (mc) spike and speed deviation followed by lower flows and intermittent suction alarms. flows were lower than the expected range for approximately 3 hours and then the map can been seen to be trending down from about 75mmhg to 40mmhg over the span of about 2 hours. shortly after this, the motor current and flows lose pulsatility and can be seen to be trending up and then the flow is saturated for the duration of the case. the pump was explanted shortly after. the returned product had biomaterial all over the outflow cage and inflow case. the product was soaked in warm water with tergazyme but the biomaterial remained. the pump was unable to be run in a bucket due to the amount of biomaterial present in both cages. based on the data logs and returned product analysis, the root cause of the pump suction, low flows and saturated flow is due to biological material ingestion that initially caused lower flows due to the position of the biomaterial and then contributed to saturated flows after cpr as the clot most likely migrated back to the impeller. ventricular tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
3974 | 1220648-2025-47826 | 23391876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23391876&pc=OZD | 08/21/2025 | 12/01/2025 |  | 10/27/2025 |  |  | 10/08/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1460542 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 72 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 the placement signal was not reliable with absent ao and left ventricle placement signals. the rep called account to help with persistent peak signal-to-noise ratio alarm with absent ps waveforms. rn reports us ordered to verify position. he denied kinks and confirmed white plug was secured in aic. reviewed alarm and explained it does not impact performance of device. instructions provided to disable alarm audio. no further needs currently. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
3975 | 1220648-2025-47825 | 23392235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392235&pc=OZD | 07/24/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 10/06/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1379919 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the malfunction of optical signal was lost and multiple placement signal not reliable alarms, if it were to recur, is likely to cause or contribute to a serious injury or death. the patient was unable to be transitioned to a left ventricular assist device (lvad) and therefore, was continued on the impella 5.5 even after the expected usage period has expired. investigation of the impella pump determined the cause may be related to the automated impella controller (aic). the aic has been requested for further evaluation. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
3976 | 1220648-2025-47833 | 23392263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392263&pc=OZD | 09/16/2025 |  |  | 10/27/2025 |  |  | 09/16/2025 | 10/27/2025 |  |  | 10/31/2025 |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411975 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 113 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue during insertion attempt for impella 5.5 (second pump for patient case). physician had a tough time getting past aortic cannula but eventually crossed the valve. however, wire popped out when impella catheter was right at aortic valve. physician attempted to cross by manually making valve incompetent but could not cross. started over and again crossed with wire. same occurrence with wire into lv (left ventricle), it popped out when impella catheter was right at av (aortic valve). attempted to cross again unsuccessfully. thus, decision was made to centrally cannulate for ecmo and place cp in groin for vent. (b)(4) is for first pump (cp). (b)(4) is for second pump (5.5). (b)(4) is for third pump (cp). (b)(4) is for fourth pump (5.5). | [5 empty] | The investigation was completed and no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy and aorta preventing successful delivery of impella. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3977 | 1220648-2025-47835 | 23392284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392284&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661765 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered ventricular tachycardia (vt) during impella 5.5 support. care team mentioned that vt was not caused by impella 5.5. however, after a few days of support, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no product was returned. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
3978 | 1220648-2025-47834 | 23392318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392318&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719235 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 74 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Renal Impairment | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support to manage multiple episodes of ventricular tachycardia that developed after quadruple bypass surgery. the impella 5.5 support was initiated, and support was slowly advanced up to p5 with flows to 2.7 liters per minute and the medical team decided to not increase further due to the patient¿s poor right ventricular function. on the fourth day of impella 5.5 support it was noted that the patient¿s platelet count was dropping with a platelet count of 38,000 platelets per microliter of blood. heparin was withheld and it was additionally noted that the patient was transfused with two units of blood during the previous night. it was also noted that the patient¿s urine production was decreased. on the fifth day of support, it was noted that heparin was continued to be held due to a possible gastrointestinal bleed. the following day, the patient was placed on continuous renal replacement therapy, and it was noted that the patient¿s platelet count had slightly increased to 48,000 platelets per microliter of blood. on the seventh day of support, it was reported that the patient was not a candidate for advanced options or recovery. the patient¿s family decided to withdraw care of the patient. the patient¿s care was withdrawn, and the patient expired. | [5 empty] | Although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Investigation summary : major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. acute kidney failure/ thrombocytopenia : the cause of the injury was not determined due to insufficient clinical details. device history lot :device lot : 1955743. device history review : the pump s/n: (b)(6) passed all the post sterile inspection checks.
3979 | 1220648-2025-47837 | 23392364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392364&pc=OZD | 09/30/2025 |  |  | 10/27/2025 |  |  | 09/30/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729913 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support was bleeding from the mouth and nose requiring heparin to be on hold. additionally, the patient appeared to have sepsis. the impella was removed, and sepsis was treated. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the injury is most likely patient condition related since the patient was bleeding from mouth and nose. sepsis: the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3980 | 1220648-2025-47838 | 23392413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392413&pc=OZD | 09/29/2025 |  |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729650 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient's right femoral impella site had substantial bleeding. additional stitches were placed. several days later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Investigation was completed. major bleed: the cause of the injury is most likely use related since additional stitches placed and hemostasis was achieved. device history review was completed and passed all the post sterile inspection checks.
3981 | 1220648-2025-47839 | 23392427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392427&pc=OZD | 09/16/2025 |  |  | 10/27/2025 |  |  | 09/16/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729354 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female | 53 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp that there was bleeding noted from impella site, successfully managed with blood products and fem stop. fem stop removed this morning, and no further bleeding from site reported. anticoagulation on hold due to bleeding overnight, possibly will start later today. bleeding resolved. the patient was successfully weaned and survived. | [5 empty] | The investigation was completed and no product was returned for investigation. major bleed: bleeding around the impella site. hgb dropped from 9.6 to 7.1. successfully managed with blood product administration for volume loss, and fem-stop placed at groin site. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
3982 | 1220648-2025-47832 | 23392433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392433&pc=OZD | 09/28/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026745517 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pain | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support was presented with thrombocytopenia which required heparin to be withheld. additionally, the patient experienced severe chest pain and was taken to the emergent catheterization lab for two stents. amiodarone and lidocaine were infused due to ventricular tachycardia runs. the patient was eventually successfully weaned from support and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombocytopenia/ pain : the cause of the injury was not determined due to insufficient clinical details. tachycardia : the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3983 | 1220648-2025-47840 | 23392534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392534&pc=OZD | 09/16/2025 |  |  | 10/27/2025 |  |  | 09/16/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653706 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had plasma free hemoglobin elevated at 16.3 but is most likely due to drawing labs from piv. no other signs of hemolysis were noted. the patient continued support and was bridged to heart transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. mechanical interaction with blood (hemolysis): the pump was not returned for investigation. data log shows no major abnormalities related to suction, motor current spikes, or positioning issues during the time of the hemolysis event. the fluctuating placement signal trend was also inconclusive. according to the clinical details, hemolysis was diagnosed based on elevated plasma free hemoglobin (pfhb). the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. device history review was completed and passed all post sterile inspection checks.
3984 | 1220648-2025-47828 | 23392545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392545&pc=OZD | 09/16/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 09/16/2025 | 10/27/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739494 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Ischemia | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was escalated from a intra-aortic balloon pump to a impella 5.5 to continue with heart failure work up. however the first attempt to place a impella 5.5 (serial number (b)(6) was aborted due to suspected calcium within axillary artery. impella cp was placed. during cp support it was reported that the patient became hypotensive with low central venous oxygen saturation (sv02), started dobutamine and increased dose when no improvement was shown, however precipitated arrhythmia. the cp was placed in the right axillary and patient's right hand was dusky with low flow. shock x5 and patient was placed on another impella 5.5 (serial number (b)(6). this report is for the second pump (serial number (b)(6). additional reports will be sent for the 1st and 3rd pumps. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 added health effect - impact code and added code b13 to type of investigation codes as they were omitted from the initial report that was submitted. h10 related report numbers were added. investigation summary: the investigation has been completed. no product was returned for investigation. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. ischemia/ arrhythmia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review. medical safety/clinical reviewed the event. the patient was admitted for congestive heart failure exacerbation and supraventricular tachycardia, underwent cardioversion, and was initially transferred to telemetry. due to worsening cardiogenic shock, the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va-ecmo) on (b)(6) 2025 (4 days prior to impella support), subsequently weaned from vasopressor and inotropic support, had diuretics discontinued, and was supported with a femoral intra-aortic balloon pump (iabp). four days later, escalation to impella support was planned. on (b)(6) 2025, the patient was taken to the cardiovascular operating room for implantation of an impella 5.5 (pump 1). a 10-mm graft was anastomosed to the right axillary artery, and impella 5.5 advancement was attempted under fluoroscopic. the procedure was aborted due to suspected calcification of the axillary artery, and the impella 5.5 was not implanted. an impella cp (pump 2) was subsequently implanted and following this implant, the patient became hypotensive with low mixed venous oxygen saturation (svo2). dobutamine was initiated and escalated without improvement and was associated with arrhythmia, requiring multiple cardioversions. the right hand was noted to be dusky with low flow. the patient required multiple defibrillation shocks. plans were made to escalate support with a left axillary impella 5.5, and the impella cp was explanted on (b)(6) 2025. an impella 5.5 (pump 3) was implanted and maintained through (b)(6) 2025. on 9/17/2025, nursing staff reported increasing vasopressor requirements, administration of multiple blood products and fluids, and concern for possible sepsis. nephrology adjusted continuous renal replacement therapy (crrt). the following day on (b)(6) 2025, interrogation of the patient¿s implanted cardioverter-defibrillator revealed frequent episodes of atrial fibrillation with rapid ventricular response, requiring cardioversion with an external defibrillator. on (b)(6) 2025, the patient developed sustained ventricular tachycardia overnight, receiving one icd shock and one external defibrillation at 100 joules, resulting in conversion to a paced rhythm. the patient received an amiodarone bolus and remained on continuous infusion. labs revealed an elevated plasma-free hemoglobin of 189, increased from the prior day. central venous pressure was reported as 10 mmhg, and the left ventricular pressure on the automated impella controller was reported as ¿31 mmhg, raising concern for possible impella positioning. echocardiography was pending and possible impella pump repositioning (echo results and if the impella was repositioned are unknown; no further details were noted on this). impella 5.5 and va-ecmo support were continued. subsequently, impella support was withdrawn overnight. no additional details regarding the exact timing or circumstances of withdrawal or patient outcome were noted. based on the information at hand, the aborted implantation of impella 5.5 (pump 1) was attributed to patient anatomy, specifically suspected axillary artery calcification. this a malfunction type of reportable event. the adverse events observed during impella cp (pump 2) with no device malfunction is a serious injury. regarding impella 5.5 (pump 3) support, there were recurrent atrial and ventricular arrhythmias, rising plasma-free hemoglobin, and concern for possible device positioning were noted. there is not enough evidence to exclude the impella 5.5 pump 3 as an association factor to the patient's outcome. however, patient's underlying condition contributed most to the outcome (advanced heart failure progression). the event story involves three product complaints mpella 5.5 - mdr 1220648-2025-47536: malfunction impella cp pump 2 - mdr 1220648-2025-47828: serious injury impella 5.5 pump 3 - mdr 1220648-2025-48483: death.
3985 | 1220648-2025-47841 | 23392617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392617&pc=OZD | 09/30/2025 |  |  | 10/27/2025 |  |  | 09/30/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026761449 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 79 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Ventricular Fibrillation | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was ventricular fibrillation (vf), arrhythmia, and cardiac arrest during impella cp patient support. during the procedure, the patient had multiple episodes of vf, which required defibrillation. post procedure, the patient was bradycardic, which required the placement of a temporary pacer. the patient coded overnight and was given two rounds of cardiopulmonary resuscitation and was shocked. the impella was repositioned by the doctor post arrest. the echocardiogram showed an acceptable position around 2.7 centimeters from the aortic valve. the patient had ventricular tachycardia with a pulse and was defibrillated once. the patient was being paced but underlying rhythm was bigeminy which caused the suction alarms. the patient was converted back to a normal sinus rhythm. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Additional information was received and noted the following: the pump was placed because the patient experienced a ventricular fibrillation arrest following a percutaneous coronary intervention. it was further clarified that the impella device was placed after the ventricular fibrillation arrest. care was withdrawn. based on the additional information complaint file clinical code ventricular fibrillation was removed. therefore, h6 health - clinical codes/impact codes were repopulated to represent the reported event. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record confirmed the device passed all the post sterile inspection checks. regarding the cardiac arrest/ arrhythmia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3986 | 1220648-2025-47842 | 23392821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23392821&pc=OZD | 09/14/2025 |  |  | 10/27/2025 |  |  | 09/14/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697810 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 141 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on impella cp for heart recovery. during insertion a hematoma was noted at insertion site. it was successfully managed by dressing repositioned for angle matching. the patient continued on support until successful explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B7 added information that was omitted from the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. h10 this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. investigation summary: the investigation has been completed. no product was returned for investigation. hematoma: hematoma at insertion site right femoral artery was noted. a swan was in place at the right femoral vein. the cause of the hematoma is determined to be use issue because hematoma is resolved by dressing changed and angle matching performed. device history review: the device passed all post sterile inspection checks.
3987 | 1220648-2025-47844 | 23393036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393036&pc=PYX | 09/28/2025 | 12/26/2025 |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  | 11/30/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024434396 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the impella rp flex had high purge pressure and purge system blocked alarm. tissue plasminogen activator was given but the purge pressure and flow remained the same. the impella rp flex was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the high purge pressure and placement signal issue has been completed. the data logs confirm high purge pressure and purge system blocked alarms and show motor current spikes followed by a gradual rise of purge pressure for about 17 hours before alarms trigger. bag and cassette changes prior to alarm were within standard time. logs show several motor current spikes with placement signal shifts and effects on pump flow. the logs also show noisy motor current and several motor current spikes outside p-level changes. cvp was negative before motor current spikes which suggest patient in continuous though biomaterial could have been a contributor. the root cause of the purge pressure high and purge system blocked issues were determined to be due to biomaterial though not classified as ingestion or build up as log analysis showed motor current spikes as well as a gradual rise of purge pressure. returned product also showed no abnormalities. the root cause of the placement signal issue was determined to be patient condition related as evidenced by log analysis showing continuous suction with stable 5s parameters and returned product analysis confirming no defects | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. h6 type of investigation code b19 added.
3988 | 1220648-2025-47847 | 23393316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393316&pc=OZD | 09/28/2025 | 12/12/2025 |  | 10/27/2025 |  |  | 09/28/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735026 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp support developed hematuria. there was also concern about possible hemolysis. the impella was in good position, but it was repositioned, and urine was clearing up after adjustment. it was further reported that the patient did not have hemolysis, and the hematuria was due to a traumatic foley inserted. the patient continued with support and the impella was explanted as planned. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported hematuria has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematuria could not be determined.
3989 | 1220648-2025-47846 | 23393551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393551&pc=OZD | 09/11/2025 |  |  | 10/27/2025 |  |  | 09/11/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026697610 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 91 KG |  | Asian | Injury | Yes | No | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp insertion, the device became entangled in the papillary muscle of the aortic valve, which caused the pump to be unable to be fully delivered into the left ventricle. the pump was pulled back and repositioning was attempted several times without success. the physician removed the pump to rewire to access the left ventricle, however upon removal of the pump the decision was made to replace the pump and deploy a new device. the new device was successfully implanted, and a successful percutaneous coronary intervention was performed. prior to leaving the cardiac catheterization laboratory, the medical team reported that there was bleeding around the repositioning sheath. an additional suture was added, and the site was dressed with angle matching. the patient was then transported to the intensive care unit, and it was noted that the site continued to bleed. a femstop was deployed and hemostasis was achieved. additionally, it was reported that there was blood leaking from the red impella handle. no intervention was performed for the red handle leak, and the leaking stopped the following day. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. this complaint involves two impella devices: pump 1: impella cp with serial number: (b)(6) pump 2: impella cp with serial number: (b)(6). this mdr specifically addresses the second pump. impella cp with serial number, which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Investigation summary: no product was returned for evaluation. major bleed: clinical details noted bleeding at insertion site, additional skin suture was applied but did not resolve issue. when patient arrived in unit, bleeding was still present and a femstop was applied and hemostasis was achieved. the cause of the major bleed could not be determined as limited clinical details were provided. red handle leak: clinical details noted that blood was leaking from red impella handle. no troubleshooting was noted. the leak resolved the following day. the cause of the red handle leak could not be determined as no product was returned for evaluation. device history lot: device lot: 1949377. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | Product status: the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
3990 | 1220648-2025-47849 | 23393641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393641&pc=OZD | 09/29/2025 | 11/05/2025 |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705303 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 101 KG | Non Hispanic | White | Death | Yes | No | Connection Problem | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiac Arrest | Hematoma | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported after the impella cp was placed, placement signal (ps) had appropriate arterial waveform. once the impella was started, ps became highly erratic with massive ps spikes. motor current was appropriate, although there was a large differential between systolic/diastolic. flows were appropriate for the p level once the position was optimized. it was discovered that the white cable was not fully seated into the automated impella controller which once the staff pushed in, the placement signal unreliable alarm resolved. additionally, towards the end of the impella insertion on the right femoral artery and removal of the sterile drape, site of the left arterial access, a large hematoma was noted. upon assessment of the right femoral site, another hematoma was noted but smaller. it is usure if the right hematoma was related to the pulmonary artery catheter or impella repositioning sheath. no external bleeding noted. manual pressure was held. cangrelor was held. the patient previously already unstable so a decision was made by the team to transfer the patient to intensive care unit and hematomas were measured further there. during transfer, the patient pulseless electrical activity arrested. two units of platelets, and two units packed red blood cells were administered. the patient expired and the impella was turned off and the device was left in patient. | [5 empty] | H6: code updated per information in the complaint file. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. h6 type of investigation code b19 added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
3991 | 1220648-2025-47850 | 23393657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393657&pc=OZD | 09/11/2025 |  |  | 10/27/2025 |  |  | 09/11/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725085 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 115 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Respiratory Failure | Hematuria | [9 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient was escalted from impella cp to an impella 5.5 support. during support of the impella 5.5 the patient was found to have blood-tinged urine, which resolved on its own and cleared up. additionally, the patient had runs of ventricular fibrillation/ventricular tachycardia overnight. moreover, patient encountered hypoxia and was changed to bilevel on ventilator. however, ultimately care was withdrawn and the patient expired. the death is currently associated with the first pump (cp) and reported on (b)(4). | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria, respiratory failure: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. a 71-year-old male presented to the emergency department with st-elevation myocardial infarction (stemi) and was emergently transferred to the cardiac catheterization laboratory. during the initial coronary angiogram, the patient decompensated and required one round of cardiopulmonary resuscitation (cpr), with return of spontaneous circulation (rosc). an impella cp was successfully implanted for mechanical circulatory support, resulting in hemodynamic stabilization. percutaneous coronary intervention of the left main into the left anterior descending artery was performed. the patient was subsequently escalated to an impella 5.5 for continued support. on (b)(6) 2025, the patient underwent a bedside left lower extremity fasciotomy for reported compartment syndrome related to an intraosseous line with wound vacuum placement. the patient later experienced ventricular arrhythmias, atrial flutter, fever (antibiotics regimen adjusted), hypoxia, and required mechanical ventilation and continuous renal replacement therapy. despite continued mechanical circulatory and supportive care, the patient¿s condition deteriorated. a decision was made to withdraw care, and the patient expired. the impella 5.5 will be reported as death type of reportable event. however, based on available information, the patient's underlying condition (including progression of cardiogenic shock [scai stage d], acute myocardial infarction) contributed most to the patient¿s outcome. correction: after review of the case by medical safety, it was determined the impella 5.5 is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. other corrections include patient-specific information for a4. weight and b7. medical history details added. note: the exact date of death is unknown but has been captured as the date of explant. if actual information becomes known, a supplemental report will be submitted.
3992 | 1220648-2025-47852 | 23393699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393699&pc=OZD | 10/08/2025 |  |  | 10/27/2025 |  |  | 10/08/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026749622 |  | 000000000001000851 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male | 87 KG |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered positioning issues during impella cp support. it was noted that the patient moved, so the impella migrated into the aorta and had to be replaced with a new one. the patient survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient when he moved. the device passed all post sterile inspection checks.
3993 | 1220648-2025-47851 | 23393778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393778&pc=OZD | 09/29/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701052 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that in the intensive care unit, the patient on impella cp support had bloody urine and some labs were hemolyzing. the next day, labs were steadily improving. the device was in adequate position per bedside transthoracic echocardiography. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood (hemolysis): data log shows some step-up in placement signal after case start, with flow and motor current dropping for about 2 hours. later, the pump was continued running at a higher p-level, with some suction alarms and minor fluctuations in pump flow and motor current. the step up placement signal gradually decreased and returned to 50 mmhg pump run was stable without abnormalities later in the case. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
3994 | 1220648-2025-47854 | 23393897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393897&pc=OZD | 10/08/2025 |  |  | 10/27/2025 |  |  | 10/08/2025 | 10/27/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1848957 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A japan complainant reported that during impella cp support the automated impella controller (aic) screen suddenly froze. the aic was exchanged. no harm to the patient was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b, date of explant, has been added.
3995 | 1220648-2025-47857 | 23393906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23393906&pc=OZD | 10/09/2025 |  |  | 10/27/2025 |  |  | 10/09/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026669544 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was noted that a few minutes after the impella was correctly positioned, patient's blood pressure dropped dramatically. the doctors performed an echo and detected a pericardial effusion. as a result, the impella was removed and team performed a pericardiocentesis. patient is stable post explant. | [5 empty] | Section a: patient age and weight is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
3996 | 1220648-2025-47858 | 23394337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23394337&pc=OZD | 08/17/2025 |  |  | 10/27/2025 |  |  | 08/17/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720651 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. it was reported that during the hrpci and while on impella cp support, the patient became agitated, sat up, and had to be restrained. the patient then reached up and grabbed the repositioning unit of the impella cp and forcibly pulled on it. an angiogram was performed afterwards and a dissection in the right iliac was revealed. the decision was made to remove the impella cp and hold manual pressure. manual pressure was held for sixty minutes and then a femstop was applied to the site. distal pulses were confirmed, and the patent was released to the cardiac care unit in stable condition. | [5 empty] | This mdr is being submitted under the correct device to replace mdr 1220648-2025-46459, which was previously submitted under an incorrect device (introducer) due to an administrative error. the initial mdr was filed within the required reporting timeframe, therefore this replacement mdr is not considered late. a.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Complaint product experience coding was updated to better describe the reported event. therefore, h6 medical device problem coding was revised accordingly. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the vascular dissection, according to the clinical details, right iliac artery dissection was noted after the patient forcibly pulled on the pump. the complication was successfully managed. the cause of the vascular damage issue is most likely user-related, due to applied excessive force, as per the clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
3997 | 1220648-2025-47859 | 23395004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23395004&pc=OZD | 09/15/2025 |  |  | 10/27/2025 |  |  | 09/15/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654265 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 105 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Anemia | Arrhythmia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced arrhythmia, anemia, and suction events during impella 5.5 patient support. while on support, there was a new onset of atrial fibrillation with rapid ventricular response. the patient was given a combination of digoxin and amiodarone. there were suction alarms and 250ml of albumin were given. there were occasional suction events after that. hemoglobin was low so one unit of packed red blood cells was administered. the care was withdrawn and patient expired. the hospital care team do not believe the impella use was related to patients' death as family withdrew care. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. anemia: the cause of the injury was not determined due to insufficient clinical details. pump suction: data log review confirms suction events occurring throughout the case. placement signal was generally low throughout the case, and it's also seen trending down over time. placement signal would increase following suction events, and clinical reported administering volume, which would resolve the suction. however, the placement signal was eventually decrease again, triggering suction multiple times during care. there were no motor current spikes observed. the cause of the suction was most likely due to the patients condition, as the downward trend in placement signal indicates that the patient could have been decompensating. the low placement signal values throughout the case could also mean the patient was experiencing volume issues, which is confirmed by suction resolving after volume administration, as stated in the clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
3998 | 1220648-2025-47860 | 23398177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23398177&pc=OZD | 09/29/2025 | 11/05/2025 |  | 10/27/2025 |  |  | 09/29/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026730823 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that there was a delivery issue during impella 5.5 implant. the impella could not be progressed forward due to the patient¿s anatomy. the implant was aborted. | [5 empty] | The device was returned. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the issue of unable to deliver or advance. only the pump was returned for investigation. the cannula and catheter were found kinked; no other damage was noted on the returned pump. according to the clinical details, the pump could not be advanced due to patient anatomy. the cause of the delivery issue was most likely related to the patient's anatomical condition, which also resulted in kinking of the catheter and cannula during the procedure. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. h6 type of investigation code b19 added.
3999 | 1220648-2025-47861 | 23398353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23398353&pc=OZD | 09/30/2025 |  |  | 10/27/2025 |  |  | 09/30/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681798 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 63 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Sepsis | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was arrhythmia, sepsis, stroke and a major bleed during impella cp patient support. while on support, the patient was cardioverted in the sticu after atrial tachycardia and converted on the second cardioversion attempt. additionally, there was a concern of septic shock. the doctor was concerned that the patient had disseminated intravascular coagulation vs heparin induced thrombocytopenia due to bleeding in the endotracheal tube, at foley insertion and peripheral iv sites. heparin was stopped. the patient became unresponsive. elevated sodium levels were revealed and attempted to be corrected. there were concerns made by care team that the patient had developed diabetes insipidus with 1300 output yesterday. the electroencephalogram showed minimal to no brain activity. care was withdrawn and the patient expired. | [5 empty] | B.2. added selection as it was omitted from the initial report that was submitted. h.6. added code b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. arrhythmia/sepsis/stroke: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the patient was bleeding in the endotracheal tube and at foley insertion and peripheral intravenous sites. the root cause of bleeding was determined to be patient condition because of the bleeding from multiple sites. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4000 | 1220648-2025-31884 | 23398638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23398638&pc=PYX | 04/06/2021 | 10/27/2025 |  | 10/27/2025 |  |  | 04/06/2021 | 10/27/2025 |  |  | 01/31/2023 |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC |  |  | US |  | 2022030392 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The pump had a high motor current causing the pump to stop, another device was used, and there was no reported harm to the patient. | [5 empty] | The returned product was visually inspected, and blue fibers were observed entraining the impeller. blue fibers likely came from impella contact with a towel or other fibrous material prior to insertion.the cause of the failure mode was most likely use related since blue fibers were observed on the impeller.
4001 | 1220648-2025-31885 | 23398741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23398741&pc=OZD | 05/31/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 07/08/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664159 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Renal Failure | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaint team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured 2nd degree type 2 block with a "possible" relationship to the impella device used: ¿patient started having long pauses. electrocardiogram (ekg) obtained showing 2nd degree type 2 block. amiodarone was stopped and dobutamine was increased¿. the patient recovered with no sequelae. in addition, the patient endured renal failure with a "possible" relationship to the impella device used: ¿baseline creatinine on was 1.07 mg/dl. patient noted to have acute kidney injury in the ¿setting of cardiogenic shock and impella.¿ continuous renal replacement therapy (crrt) initiated. crrt used intermittently. hemodialysis started. patient was discharged from hospital on hemodialysis without hope for renal recovery. during support it was noted that there were issues with placement requiring repositioning 3 times. an echocardiogram showed it had pulled back. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. renal failure and 2nd degree type 2 block with a "possible" relationship to the impella device used was reported. there were issues with pump positioning requiring repositioning of the device multiple times on support. the pump was not returned for investigation, and insufficient clinical details were provided. the cause of the renal failure, 2nd degree type 2 block, and the positioning issues could not be determined.
4002 | 1220648-2025-31886 | 23398760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23398760&pc=OZD | 07/09/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 07/09/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671806 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a impella 5.5 on high-risk surgical patient. over 2 months after first allegation, the impact study reported upon multiple adverse events all with relationship listed as "definite" to the pump. they are coagulopathy treated by blood products and chest tube, the next adverse event is acute right ventricle dysfunction due to either pump use or bypass, non-sustained ventricular tachycardia was also reported, and finally the adverse event of acute kidney injury which prompted the diuresis plans to be changed/modified. patient weaned and explanted. | [5 empty] | The investigation has been completed for tachycardia (vt), renal failure, and right heart insufficiency. the pump was not returned for evaluation. additionally, clinical details were insufficient to determine the root cause. therefore, the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the root cause of the vt was unable to be determined due to insufficient clinical details. the cause of the renal failure was not determined due to insufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ a.5 unknown.
4003 | 1220648-2025-47862 | 23398765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23398765&pc=OZD | 10/09/2025 | 10/27/2025 |  | 10/27/2025 |  |  | 10/09/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676611 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male | 83 KG | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 insertion was aborted due to the patient having congenital pathology (tgv) and the aortic arch was too narrow to rotate the 5.5 properly. | [5 empty] | The investigation for the reported event, unable to deliver or advance (delivery issue), has been completed. no product was returned. the cause of the deliver issue was determined to be patient condition as the patient had narrow aortic arch and resistance there preventing successful delivery of impella. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4004 | 1220648-2025-47864 | 23399088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23399088&pc=OZD | 10/12/2025 |  |  | 10/27/2025 |  |  | 10/12/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697030 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop. a motor current increase alarm occurred but no thrombus could be found in the pump. the percutaneous coronary intervention procedure was performed without the use of the impella. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4005 | 1220648-2025-47863 | 23399091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23399091&pc=OZD | 10/10/2025 | 01/07/2026 |  | 10/27/2025 |  |  | 10/10/2025 | 10/27/2025 |  |  |  |  | 10/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594706 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female | 58 KG |  | Asian | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a false alarm during impella cp support. during support, impella in ventricle alarm was displayed. on fluro, impella was clearly across aortic valve. aortic placement signal systolic 98 diastolic 5. motor current signal systolic 425 diastolic 370. physician advised false alarm due to metrics as fluro clearly demonstrated correct position. | [5 empty] | Investigation summary: no product returned. false alarm / buzzer - after less than 4 hours on support, impella in ventricle alarm displayed despite fluro showing adequate positioning. a/o ps systolic 98 diastolic 5. pt noted to have small lv and low diastolic bp. the cause of the false alarm / buzzer could not be determined due to no product returned, no data logs recovered, and insufficient clinical details. device history lot: device lot: 1821990. device history batch: subcomponent lot: n/a. device history review: pump s/n (b)(6) passed all post sterile inspection checks. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4006 | 1220648-2025-47762 | 23378034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23378034&pc=OZD | 09/12/2025 | 10/23/2025 |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717693 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | Embolism/Embolus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient had shock with worsening ventilation and organ perfusion labs. as a result, va ecmo was cannulated which improved the situation. additionally, patient had active pulmonary embolisms resulting in the switching of purge over to heparin. thirdly, the patient had retroperitoneal bleed and the team was deciding on maybe going to interventional radiology for bleed management. however, no further intervention was mentioned. patient was successfully weaned off the impella and is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the retroperitoneal hemorrhage: purge switched over to heparin d/t active pes per pharmacy. the cause of the hemorrhage was not determined due to insufficient clinical details. cardiogenic shock/ embolism: the cause of the clinical symptom was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
4007 | 1220648-2025-47768 | 23378098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23378098&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 95 KG |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was positioning issues and ischemia during impella cp patient support. while on support, it was noted the impella came back into the aorta. at that time, the doctors were concerned about the patient¿s hypercoagulable state and the perfusion of his leg with impella in place, therefore, impella was removed. | [5 empty] | The investigation was completed and no product was returned. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue cannot be determined since insufficient clinical details were provided. device history reviewed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4008 | 1220648-2025-47769 | 23378185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23378185&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 115 KG | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there was an inaudible or missed buzzer on the automated impella controller (aic) during impella 5.5 support. the pump alarmed because it was in the left ventricle, but there was no accompanying aic alarm. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4009 | 1220648-2025-47770 | 23378572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23378572&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683681 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the medical team pulled the impella to descending aorta to be explanted due to continues suction. during impella repositioning suction alarm was noted even at p2. so, since the patient hemodynamics was stable, physician decided to pull the impella to the descending aorta and explant when act was below 150. patient hemodynamics were stable with stable condition. placement signal monitoring was also disabled. additionally, the same day, the patient started having hematoma. pump was explanted the following day. patient survived the procedure and is stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the injury was not determined due to insufficient clinical details. pump suction: data log review confirms suction alarms triggering around the time it was reported. the cause of the suction was most likely biomaterial ingestion, as seen by the motor current (mc) spike right before the suction alarms started, which led to altered waveforms and decreased flow rates. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4010 | 1220648-2025-47771 | 23378671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23378671&pc=PYX | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529640 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support. the patient encountered placement signal issue and bipella patient had pulmonary hemorrhage during impella rp flex support. placement signal (ps) not reliable alarms occurred shortly after proactive calls. the ps was significantly reduced and ultimately failed. additionally, a few days later patient had pulmonary hemorrhage and as a result heparin was put on hold. patient remained off heparin due to bleeding in the lungs. patient remained stable and pump was explanted shortly afterwards. this complaint is for the impella rp flex. patient had impella cp (sn # (b)(6)) which is reported under separate report. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. a cut motor end catheter was returned for investigation. no investigation could be performed on it. data log shows a temporary loss in placement signal on the implant day, with a decreasing trend in placement signal that correlated with an upward trend in pump flow. both placement signal and pump flow returned to normal after eight hours of the case run. no abnormalities were related to optical signal issues during the case. placement signal issue: placement signal issue does not match any known failure trend, and the product returned was insufficient for investigation. therefore, the cause of the placement signal issue was not determined. clinical symptoms (hemorrhage/bleeding): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4011 | 1220648-2025-47772 | 23378792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23378792&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712064 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 101 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had high purge pressure alarms requiring tissue plasminogen activator (tpa) infusion through the purge system. there were some increase in purge flows so tpa was continued. some suction that resolved with fluid administration. placement of pump was confirmed via echocardiogram. the patient remained on support until successful weaning and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. purge pressure high - data log review showed purge pressure trend up from case start through 11sep25 when it began to trend down. the cause of the purge pressure high was patient condition related to biomaterial build-up that resolved with the use of lysis solution. pump suction - data log review did not show suction alarms. the cause of the pump suction could not be determined as the incomplete data logs recovered did not show suction alarms, no product was returned, and insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4012 | 1220648-2025-47773 | 23379039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23379039&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026721574 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after implantation of an impella cp a vessel perforation and bleeding was noted in the left common femoral artery. the perclose did not deploy correctly and the angio seal had failed. as a result, the impella was removed from the right side in order to get balloon tamponade and ultimately a covered stent to the left side. the patient was treated with two units of packed red blood cells and pressors. after the covered stent was placed and team had controlled the bleeding on the left side, the decision was made to place another impella as the patient still had escalating pressor needs and low native cardiac function. the patient ultimately expired on the second impella cp. | [5 empty] | No product was returned for investigation. the cause of the bleed was most likely use issue related since the perclose did not deploy correctly and failed. the cause of the vessel perforation was most likely use issue related since the perclose did not deploy correctly and failed. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4013 | 1220648-2025-47774 | 23379070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23379070&pc=OZD | 09/11/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694748 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Heart Block | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support placed as care escalation from the impella cp. extracorporeal membrane oxygenation was additionally placed for the patient with acute myocardial infarction/cardiogenic shock. the pump was supporting when there was complete heart block and ventricular tachycardia. the pump remained on for 9 days, and the pump was then weaned and explanted. the patient survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b11 and b13 h11: added the results of the investigation. investigation summary the device was not returned for evaluation. heart block/ ventricular tachycardia: in order to make a cause determination, e clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device history lot ==> device lot: 1934685 device history review ==> the pump s/n: (b)(6) passed all the post sterile inspection checks.
4014 | 1220648-2025-47775 | 23379154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23379154&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026699977 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 116 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the patient had known cardiac comorbidities and had the pump for the pci and was then explanted. 2 perclose were deployed pre impella. cinched down post impella removal with significant oozing. 8f angioseal attempted to be deployed and failed. manual compression applied until hemostasis achieved. patient is stable post explant. | [5 empty] | B.7 added information that was omitted from the initial report that was submitted. g.1 manufacturing site name was reported incorrectly on the initial report that was submitted and should of been left blank. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. no product was returned for investigation. major bleed: manual compression was applied until hemostasis was achieved. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4015 | 1220648-2025-47776 | 23379244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23379244&pc=OZD | 09/11/2025 |  |  | 10/24/2025 |  |  | 09/11/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713907 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Unknown | 118 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Atrial Fibrillation | High Blood Pressure/ Hypertension | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hypertension, atrial fibrillation, and tachycardia. an echocardiogram confirmed the pump was positioned deep. the pump was pulled back but met resistance. after multiple unsuccessful attempts the pump was explanted and exchanged for a new impella 5.5 pump. | [5 empty] | E.1 added initial reporter phone number as it was omitted from the initial report that was submitted. h.6 code e0601 was reported incorrectly on the initial report that was submitted. a new code was added. investigation summary: no product was returned. tachycardia, paroxysmal atrial fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypertension: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4016 | 1220648-2025-47777 | 23379403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23379403&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026670064 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the pump had been placed via the left axillary and had supported for 15 days when explanted. at the time of explant there was a drop in the mean arterial pressure (map). the map dropped to 28. the drop in map prompted the team to enlist vascular surgery. the surgeon had to perform balloon embolectomy with the fogarty. patient survived the removal issues, thrombosis, and drop in map. patient was escalated to 5.5 | [5 empty] | E1 added initial reporter email address as it was omitted from the initial report that was submitted. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
4017 | 1220648-2025-47778 | 23379440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23379440&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 101 KG | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that there was difficulty implanting an impella cp due to the patient's calcified arteries. the guidewire became caught in the papillary muscle and met resistance when attempting to move the wire forward. the left ventricle was re-accessed with better wire placement, and the impella was successfully placed. a small hematoma was noted at the left femoral artery, which dried and flattened by the end of the procedure. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.1 adverse event was incorrectly selected on the initial report that was submitted. b.2 required intervention was selected incorrectly on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name should of been left blank on the initial report that was submitted. h.4 codes e0506 and a24 were reported incorrectly on the initial report that was submitted. investigation summary: no product was returned. patient anatomy or condition prior to therapy: the cause of patient anatomy or condition prior to therapy was most likely patient condition related since patient had calcified arteries. difficult to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had calcified arteries preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
4018 | 1220648-2025-47779 | 23380474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380474&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739177 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 103 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding when the perclose failed at the time of explantation. a balloon tamponade and covered stent were used to successfully support the patient instead. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleed is most likely use related since post explant perclose failed requiring balloon tamponade and covered stent placement to rt axillary artery. the device passed all the post sterile inspection checks.
4019 | 1220648-2025-47780 | 23380588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380588&pc=OZD | 09/24/2025 | 02/13/2026 |  | 10/24/2025 |  |  | 09/24/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711387 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male | 95 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient on impella 5.5 was presented with ventricular tachycardia and required administration of amiodarone. additionally, during cranial cat scan the patient was shown to have had multiple cerebral infarction while on support. the patient was successfully weaned and survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4, g1 the investigation of the reported failure mode has been completed. no product was received for evaluation. stroke/ arrhythmia/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4020 | 1220648-2025-47781 | 23380602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380602&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676292 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 111 KG | Non Hispanic |  | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular fibrillation requiring defibrillation. the patient also developed a hematoma at the impella access site and oozing from the right axillary site. heparin was withheld and the dressing was changed. additionally, the placement signal was low and alarmed with ambulation. the alarm appeared to be due to poor cardiac function; not position related due to low diastolic pressure. the patient continues on impella support. | [5 empty] | The root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of the bleeding issue was not determined without returned product investigation and sufficient clinical details. the cause of the optical signal issue was not determined without returned product investigation. the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4021 | 1220648-2025-47782 | 23380605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380605&pc=OZD | 09/25/2025 | 12/09/2025 |  | 10/24/2025 |  |  | 09/25/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026743370 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 111 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella c the patient had been on p4 for several hours with vitals and labs remaining stable. medical doctor(md) assessed placement of impella on echocardiogram. he pulled back device a few centimeters, but waveforms remained about the same. weaning screen showed no changes in native function which remained stable. the patient was oozing at the site, no swan present. the patient on p4 overnight so decision made to explant the pump. md said patient was very oozy even before impella placement. patient survived. | [5 empty] | The investigation for the reported major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4022 | 1220648-2025-47783 | 23380611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380611&pc=OZD | 09/25/2025 | 10/23/2025 |  | 10/24/2025 |  |  | 09/25/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725090 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella 5.5 the patient had bleeding. slight drop in hand from unknown origin. 1-unit of packed red blood cells (prbc). no additional information was given. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
4023 | 1220648-2025-47784 | 23380633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380633&pc=OZD | 09/25/2025 | 12/09/2025 |  | 10/24/2025 |  |  | 09/25/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672859 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 100 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [27 empty] | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella cp the patient arrived to 9914 on impella support. hematoma noted at the impella site on arrival. fem stop was applied. distal pulses present. pending repeat labs. echo done measuring at 3.4. the patient was started on nitric oxide. plan to insert swan. will continue to optimize care. care was withdrawn and the patient expired the same day. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma could not be determined.
4024 | 1220648-2025-47786 | 23380656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380656&pc=OZD | 09/25/2025 |  |  | 10/24/2025 |  |  | 09/25/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543830 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the insertion was aborted due to the patient¿s anatomy. there was no patient harm reported. patient outcome is unknown. | [5 empty] | . the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
4025 | 1220648-2025-47787 | 23380701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380701&pc=OZD | 09/26/2025 |  |  | 10/24/2025 |  |  | 09/26/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1155706 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the automated impella console (aic) failed to recognize impella catheter rendering impella catheter inoperative. three aic¿s could not recognize impella catheter and required md to replace inoperative impella. this mdr specifically addresses automated impella controller, which is the subject of this report. a separate mdr was submitted for the impella cp associated with this complaint. | [5 empty] | Section a4 and a6 : the patient race and weight is unknown . the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4026 | 1220648-2025-47788 | 23380709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380709&pc=OZD | 09/27/2025 | 12/09/2025 |  | 10/24/2025 |  |  | 09/27/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722238 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient came to the emergency room in cardiac arrest. the patient was taken to cath lab for heart catheterization and impella cp placement. in intensive care unit, platelets were dropping, and the team was discussing stopping heparin. it was further reported that the family decided to withdraw care as the patient had anoxic brain injury. the patient expired | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined.
4027 | 1220648-2025-47785 | 23380879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23380879&pc=OZD | 09/12/2025 |  |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725978 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding so heparin was withheld. additionally, the patient went into a torsade¿s rhythm and lost pulsatility, but remained awake and alert with flows of 4.2 on the pump. the patient was given versed and lidocaine, and was shocked out of acute rhythm. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the bleeding was not determined since product was not returned and insufficient clinical details provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4028 | 1220648-2025-47789 | 23381080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23381080&pc=OZD | 09/12/2025 | 10/16/2025 |  | 10/24/2025 |  |  | 09/12/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722148 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 94 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision was made for a escalation of intra-aortic balloon pump (iabp) to impella 5.5 in the continued setting of cardiogenic shock. the 5.5 was inserted in the right axillary/subclavian artery without incident. it was reported that after iabp removal the patient no longer had pedal pulses on that leg. in addition to having suction alarms however resolved after 500ml of albumin. it was noted that vascular surgeon at bedside to perform fasciotomy of right lower extremity due to no distal flow post iabp. the patient also went into ventricular tachycardia and was shocked x1. the patient was also febrile. it was also noted that gastrointestinal bleed persisted but h&h stabilized, not requiring transfusion. however, heparin was turned off due to bloody stools and bloody mucosal drainage. 1unit packed red blood cells were transfused. the patient remains on support awaiting left ventricle assisted device workup. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved quickly. pump suction: the cause of suction issue was most likely patient condition related since the suction was resolved after albumin given. fever/ventricular fibrillation/ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. gi bleed/epistaxis: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4029 | 1220648-2025-47791 | 23381224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23381224&pc=OZD | 09/13/2025 |  |  | 10/24/2025 |  |  | 09/13/2025 | 10/24/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720782 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Difficult to Open or Close | [5 empty] | Bradycardia | Stroke/CVA | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support following a percutaneous coronary intervention (pci) and subsequent ventricular tachycardia arrest. the impella cp was placed emergently for support to re-explore the stents placed during the pci. following implantation of the impella cp, the decision was made to transfer the patient to another hospital for a higher level of care. the patient was transferred to the intensive care unit (icu) while waiting transport to the outside hospital and it was reported that just after arriving to the icu, there were alarms for impella in aorta. the touhy lock was found to be unsecured. the pump was repositioned, and then the groin was re-sutured and redressed. additionally, the patient experienced an episode of bradycardia and it was noted the patient¿s heartrate dropped to the 50¿s. volume, a dose of epinephrine, and sodium bicarbonate were administered to resolve the bradycardia. the patient was successfully transferred to the outside hospital, and the patient had a revascularization of the affected coronary arteries. post re-vascularization, the impella cp was successfully weaned and explanted. the day following explant, it was reported that the patient suffered a hemorrhagic stroke which affected the patient¿s left side. no further information was provided on interventions or patient outcome. | [5 empty] | H6. medical device problem code malposition of device was not included in the initial mdr and is now correctly added. | A5, a6 are unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4030 | 1220648-2025-47792 | 23381409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23381409&pc=OZD | 09/14/2025 | 11/28/2025 |  | 10/24/2025 |  |  | 09/14/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186705 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 90 KG |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility had a console and 5.5 supporting the patient admitted in with a high risk surgery. the console failed and was replaced during the support. there was no harm or injury from the malfunction, the support proceeded until explant after 80 hours of support. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation for the reported unexpected reboot has been completed. : the cause of the unexpected reboot could not be determined since the failure was not reproduced.
4031 | 1220648-2025-47793 | 23381445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23381445&pc=OZD | 09/13/2025 | 01/06/2026 |  | 10/24/2025 |  |  | 09/14/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711230 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female | 60 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a 5.5 pump supporting the patient admitted in with cgs and cardiomyopathy. the pump was supporting while left ventricular assist device being considered. the team did observe high pp and gave lytic to resolve, and additionally there was optical signal loss. the malfunctions did not cause any harm or injury. pump remained on until wean and explant after 32 days. | [5 empty] | Investigation summary: no product or logs returned for evaluation. high purge pressure: clinical details noted that pump underwent lysis protocol due to high purge pressure. no additional details were provided. the cause of the high purge pressure could not be determined as no products or logs were returned for evaluation and clinical details were provided. optical signal issue: clinical details noted that aic was not displaying ps or lv curves or values. pump position was confirmed, pump was not replaced. no additional details were provided. the cause of the optical signal issue could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. device history lot: device lot: 1945136. device history batch: subcomponent lot: n/a. device history review pump sn (b)(6) passed all post-sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4032 | 1220648-2025-47795 | 23383425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23383425&pc=OZD | 09/15/2025 |  |  | 10/24/2025 |  |  | 09/15/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711423 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 87 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient underwent aortic root/valve surgery. noted was successful insertion of impella 5.5 post aortic valve/arch replacement. intra-aortic balloon pump removed. extracorporeal membrane oxygenation initiated using bypass cannulas. sternotomy left open post procedure for return to further assess bleeding and close. during support it was reported that the patient was returning from cath lab for an attempt to removal of atrial thrombus however failed, cardiologist confirmed that plan was to bring the patient to operating room for open removal of epicardial thrombus that was partially compressing right atrium. it was reported that after returning from the operating room the patient was requiring increasing pharmacological support. decision was made to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no product was returned. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
4033 | 1220648-2025-47794 | 23383488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23383488&pc=OZD | 09/14/2025 | 10/24/2025 |  | 10/24/2025 |  |  | 09/14/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725399 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 113 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered continuous pump suction during impella cp support (first pump for patient case). suction was encountered above p5 and a few suction alarms were noted. the patient remained on cp pump support for total of 3 days before escalating to impella 5.5. this report is for first pump (cp). (b)(4) is for second pump (5.5). (b)(4) is for third pump (cp). (b)(4) is for fourth pump (5.5). | [5 empty] | The investigation was completed. pump suction: clinical reported suction events on the second and third day of care. the product was not returned for investigation. placement signal was low throughout the case, staying under 100mmhg, and started to gradually trend down starting about a day after case start. around this time, suction starting triggering. clinical reported plans of volume resuscitation around this time as well. the patient was reportedly on va ecmo, which can cause potential suction events by mechanically drawing out volume from the heart. the cause of the suction was most likely related to the patients condition, as low pressures in the heart along with the trend down overtime could indicate that the patient was experiencing volume issues. device history review was completed and passed all post sterile inspection criteria. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4034 | 1220648-2025-47796 | 23383758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23383758&pc=OZD | 09/15/2025 | 10/24/2025 |  | 10/24/2025 |  |  | 09/15/2025 | 10/24/2025 |  |  |  |  | 10/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716180 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Renal Failure | Cardiogenic Shock | Hematuria | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was renal failure, shock, ischemia, arrhythmia, and hematuria during impella 5.5 support. while on support the patient was experiencing renal failure, low urine output and red urine. the clinical staff were actively replacing the continuous renal replacement therapy circuit during rounding. the pump was repositioned, and the patient had some ongoing ectopy. the patient was experiencing shock. left lower extremity pulses were lost. the impella was successfully weaned and the patient survived. | [5 empty] | E1 added initial reporter address line 1 and initial reporter phone number as they were omitted from the initial report that was submitted. h6 added b13 code as it was omitted from the initial report that was submitted. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. renal failure/ hematuria: the cause of the injury was not determined due to insufficient clinical details. cardiogenic shock/ arrhythmia/ ischemia : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.